## bar1
Biobased content | Biobased content error | Biobased content multiplier | Blank rts assigned | Blank rts err assigned | Blanks fm av err Calc | Blanks fm average Calc | C12 | C12C13 ratio | C14C13 ratio | Category Field | Category In Calculation | Collection Date | Collection Date Saved | Colour Description | Comment | CRA | CRA Report | CRA Report Saved | CRA_Error | CRA_Error Report | CRA_Error Report Saved | CRA_OneSigmaMinus | CRA_OneSigmaPlus | Creation User | D14C | D14C_Error | Date Run | DateLastRun | DaysAsGraphite | DaysFromCurrentRun | DCC | DCC_error | DCCstd | DCCstd_error | delta13C_AMS | delta13C_AMS Saved | delta13C_AMS_Error | delta13C_AMS_Error Stored | delta13C_In_Calculation | delta13C_In_Calculation_Error | delta13C_In_Calculation_Source | delta13C_IRMS | delta13C_IRMS_Error | delta13C_IRMS_Source | delta13C_std_av_Match | delta13C_stds_av | delta13C_stds_av_Calculation | delta13C_stds_av_error | delta13C_stds_av_error_Calculation | DELTA14C | DELTA14C Saved | DELTA14C_Error | DELTA14C_Error Saved | deltalower14C | deltalower14C_Error | dlower14C | dlower14C_Error | Email Body | Email sender | End Date | End Operator | End Time | F_corrected_normed | F_corrected_normed Saved | F_corrected_normed_error | F_corrected_normed_error Saved | F_corrected_normed_error_no_WTW_Error | F_corrected_normed_error_no_WTW_noSYS_Error | FeC | Find_TW Global | Flake Description | Fm_normalized | Fm_normalized_error | ID | Import XCAMS Results Date | Import XCAMS Results Time | Import XCAMS Results User | Job | MCC | MCC Saved | MCC_error | MCC_error Saved | MCC_fractional_error | Mode | Modification Date | Modification Time | Modification User | Normalization_exp_factor | Packing Comments | pM | pM_Error | Position | Process Group | Processes Order | Quality Flag | Ratio to standard | Ratio to standard error | Ratio_13C_standard | Ratio13C_denominator | Ratio13C_numerator | Ratio13C_standard_error | Ratio13C_std_av | Ratio13C_std_av Calc | Ratio13C_std_av_Match | Remove System Error | Result Comment for Submitter | Result Comment Internal | rts_bl_av_Calculation | rts_bl_av_Calculation Old | rts_bl_av_error_Calculation | RTS_corrected | RTS_corrected Saved | RTS_corrected_error | RTS_corrected_error Saved | rts_std_av_Match | rts_stds_av | rts_stds_av_Calculation | rts_stds_av_error | rts_stds_av_error_Calculation | Silvia_F_corrected_normed | Silvia_F_corrected_normed_Check | Silvia_F_corrected_normed_error | Silvia_F_corrected_normed_error Check | Silvia_F_corrected_normed_error_no_WTW_Error Check | Silvia_MCC | Silvia_MCC Check | Silvia_MCC_error | Silvia_MCC_error Check | Silvia_RTS_corrected | Silvia_RTS_corrected Check | Silvia_RTS_corrected_error | Silvia_RTS_corrected_error Check | Size Category | SmallSamples values calculated via Excel | Standard 13C Value Constant | Standard Specific Activity Constant | Standard Type Number | Start Date | Start Operator | Start Time | Static Description | Status | tmp_jct_F_new | tmp_jct_F_old | tmp_jct_rts_corr_new | tmp_jct_rts_corr_old | TP | TP Graphite Submitted Minimum Value | TPs Blanks | TPs Blanks Report | TPs Standards | TPs Standards Report | TW | TWs Blanks | TWs Standards | Update Blanks Average Date | Update Blanks Average Time | Update Blanks Average User | Use Assigned Blank | Weight Initial | Wheel To Wheel Error | wtgraph | Date Run | Ratio to standard | TP | TW | wtgraph | Ratio to standard | Ratio to standard error | TP | Standard Prep Type | AMS System | Standard Prep Type | System Error | Wheel Result Comment For Submitter | Wheel Result Comment Internal | delta13C | Surname | delta13C | Position Number | Run Number | delta13C | Run Number | AMS Category | delta13C Estimate | External Due Date | Fraction to date | GlobalStatus | Job | Job Invoicing Notes | Job notes | NZA | Precision | R | Virtual TW | Contaminant observed suspected | Process End Date | Process ID | Process Name | Process Status | R Composed | Sample Description | Comments From Submitter | Comments On Sample Lab Internal | Contaminants | delta13C Submitted | Expected age | Pretreatment by Submitter | Sample Description | Sample ID | Submitter Description | Submitter Requested Pretreatment
[7 empty] | 118.3 | 89.5209 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 50994 | 50993 | 50995 | 467 | 467 | 459 |  |  | Silvia Canessa | -998.25 | 0.10 ‰ | 2012-06-25 | 2012-06-25 | [10 empty] | 2.46 ‰ | 0.20 ‰ | C13 | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.2817 | 0.04 | 0.07071 | -998.25 ‰ |  | 0.1 |  | -998.23 | 0.10 ‰ | -998.23 | 0.10 ‰ |  | Silvia Canessa | 2012-06-28 | Silvia Canessa | 13:28 | 0.0017502 | 0.00175 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.3 |  |  | 0.00175 | 0.0001 | 160 |  |  |  | 111928 | [5 empty] | 4 | 5/07/22 | 15:24:43 | Albert Zondervan | 1 |  | 0.18 | 0.01% | 31 | 5 | 6 | ... | 0.00172 | 0.0001 |  |  |  |  | 1.0092 | 1.0092 | Yes | [6 empty] | 0.00172 |  | 0.0001 |  | No | 1 | 1 | 0.001 | 0.001 | 0.00175 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00172 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-06-22 | Silvia Canessa | 15:00 |  | Completed |  |  |  |  | 48827 | 2000000 |  |  | 48973 48974 48975 48976 48977 48980 48985 48988 | 48973~48974~48975~48976~48977~48980~48985~48988 | 2761 |  | 2761 2761 2761 2761 2761 2761 2761 2761 |  |  |  |  | 1.30 mg |  | 1.3 | [5 empty] | 1.0001 | 0.00286 | 48973 |  | XCAMS | Both |  |  |  | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-08-17 | Carbonate |  | JOB-111928 |  | Job 111928 was created in Access and entered into RLIMS with delta | 50112 | Normal | 14047/1 |  |  | 2012-06-22 | 464 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0052 | 0.00287 | 48974 | [26 empty] | 2012-06-22 | 62 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00284 | 48975 | [26 empty] | 2012-06-22 | 140 | Graphite | Cancelled
[168 empty] | 0.99956 | 0.00283 | 48976 | [26 empty] | 2012-06-22 | 137 | Graphite LG1 | Completed
[168 empty] | 1.0015 | 0.00284 | 48977 | [26 empty] | 2012-06-22 | 99 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00283 | 48980 | [26 empty] | 2012-06-28 | 160 | AMS Submission Results | Completed
[168 empty] | 0.99648 | 0.00281 | 48985 | [26 empty] | 2012-06-29 | 108 | Review Final Report | Completed
[168 empty] | 0.99629 | 0.00284 | 48988 | [26 empty] | 2012-07-20 | 128 | Report Results | Completed
[197 empty] | 2012-07-20 | 403 | Close Job | Completed
[7 empty] | 104.7 | 89.3484 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 50635 | 50635 | 50636 | 456 | 456 | 439 |  |  | Dawn Chambers | -998.17 | 0.10 ‰ | 2012-07-02 | 2012-07-02 | [6 empty] | 5.19‰ | 5.19‰ | 2.00‰ | 2.00‰ | 5.19 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.0357 | 0.04 | 0.70711 | -998.17 ‰ |  | 0.1 |  | -998.15 | 0.11 ‰ | -998.15 | 0.11 ‰ |  | Dawn Chambers | 2012-07-06 | Dawn Chambers | 14:24 | 0.00183 | 0.00183 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.9 |  |  | 0.00183 | 0.0001 | 238 |  |  |  | 112005 | [5 empty] | 5 | 5/07/22 | 15:24:43 | Albert Zondervan | 0 |  | 0.18 | 0.01% | 12 | 5 | 5 | ... | 0.00176 | 0.0001 |  |  |  |  | 1.0092 | 1.0092 | Yes | [6 empty] | 0.00176 |  | 0.0001 |  | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00183 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00176 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-06-29 | Dawn Chambers | 15:22 |  | Completed |  |  |  |  | 48889 | 2000000 |  |  | 48978 49002 49005 50081 50082 50088 50091 50103 | 48978~49002~49005~50081~50082~50088~50091~50103 | 2763 |  | 2763 2763 2763 2763 2763 2763 2763 2763 |  |  |  |  | 1.96 mg |  | 1.96 | [5 empty] | 0.9986 | 0.00168 | 48978 |  | XCAMS | Both |  |  |  | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-08-17 | Carbonate |  | JOB-112005 |  | Job was done in Access, completed in RLIMS w/o delta 13C information | 50193 | Normal | 14047/1 |  |  | 2012-06-22 | 468 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99958 | 0.00176 | 49002 | [26 empty] | 2012-06-22 | 63 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99862 | 0.0017 | 49005 | [26 empty] | 2012-06-22 | 138 | Graphite LG1 | Completed
[168 empty] | 1.0019 | 0.00166 | 50081 | [26 empty] | 2012-06-22 | 96 | C13 Measurement | Completed
[168 empty] | 1.0033 | 0.00226 | 50082 | [26 empty] | 2012-07-06 | 238 | AMS Submission Results | Completed
[168 empty] | 0.99918 | 0.00216 | 50088 | [26 empty] | 2012-08-21 | 199 | Review Final Report | Completed
[168 empty] | 1.0013 | 0.00165 | 50091 | [26 empty] | 2012-08-21 | 671 | Report Results | Completed
[168 empty] | 0.99844 | 0.00168 | 50103 | [26 empty] | 2012-08-23 | 897 | Close Job | Completed
[7 empty] | 104.9 | 89.5346 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 49771 | 49770 | 49763 | 410 | 409 | 394 |  |  | Dawn Chambers | -997.96 | 0.10 ‰ | 2012-07-06 | 2012-07-06 | 22 | [5 empty] | 2.41‰ | 2.41‰ | 2.00‰ | 2.00‰ | 2.41 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0.04 | 0.70711 | -997.96 ‰ |  | 0.1 |  | -997.94 | 0.11 ‰ | -997.94 | 0.11 ‰ |  | Dawn Chambers | 2012-07-10 | Dawn Chambers | 14:02 | 0.002038 | 0.00204 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 3.4 |  |  | 0.00204 | 0.0001 | 298 |  |  |  | 190334 | [5 empty] | 5 | 5/07/22 | 15:24:43 | Albert Zondervan | 0 |  | 0.2 | 0.01% | 6 | 5 | 4 | ... | 0.00196 | 0.0001 |  |  |  |  | 1.0086 | 1.0085 | Yes | [6 empty] | 0.00196 |  | 0.0001 |  | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00204 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00196 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-03 | Dawn Chambers | 17:39 |  | Completed |  |  |  |  | 60101 | 2000000 |  |  | 60070 60131 60132 60133 60134 60159 60160 60178 | 60070~60131~60132~60133~60134~60159~60160~60178 | 2765 |  | 2765 2765 2765 2765 2765 2765 2765 2765 |  |  |  |  | 0.57 mg |  | 0.57 | [5 empty] | 0.99893 | 0.00171 | 60070 |  | XCAMS | Both | 0% |  | - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Inorganic |  | 2012-08-08 | Carbonate |  | JOB-190334 |  | marble std for RG5 graphite | 50249 | Normal | 14047/1 |  |  | 2012-06-13 | 422 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99918 | 0.0017 | 60131 | [26 empty] | 2012-06-13 | 45 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99926 | 0.00168 | 60132 | [26 empty] | 2012-06-14 | 111 | Graphite | Completed
[168 empty] | 0.99937 | 0.00171 | 60133 | [26 empty] | 2012-07-10 | 298 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.0017 | 60134 | [26 empty] | 2012-08-02 | 255 | Review Final Report | Completed
[168 empty] | 1.0003 | 0.00177 | 60159 | [26 empty] | 2012-08-20 | 402 | Report Results | Completed
[168 empty] | 1.0006 | 0.00176 | 60160 | [26 empty] | 2012-08-20 | 754 | Close Job | Completed
[168 empty] | 1.0009 | 0.00174 | 60178
[7 empty] | 113.2 | 89.3833 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 54558 | 54558 | 54580 | 669 | 669 | 646 |  |  | Jenny Dahl | -998.88 | 0.09 ‰ | 2012-07-17 | 2012-07-17 | [6 empty] | 2.12‰ | 2.12‰ | 2.00‰ | 2.00‰ | 2.12 ‰ | 2.00 ‰ | AMS | 2.15‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.04 | 0.70711 | -998.88 ‰ |  | 0.09 |  | -998.86 | 0.09 ‰ | -998.86 | 0.09 ‰ |  | Jenny Dahl | 2012-07-23 | Jenny Dahl | 8:56 | 0.001123 | 0.00112 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.4 |  |  | 0.00112 | 9e-05 | 446 |  |  |  | 190429 | [5 empty] | 5 | 5/07/22 | 15:24:43 | Albert Zondervan | 0 |  | 0.11 | 0.01% | 2 | 5 | 5 | ... | 0.00108 | 9e-05 |  |  |  |  | 1.0066 | 1.0065 | Yes | [6 empty] | 0.00108 |  | 9e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00112 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00108 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-16 | Jenny Dahl | 10:37 |  | Completed |  |  |  |  | 50122 | 2000000 |  |  | 50105 50106 50107 50108 50181 50273 50276 50279 | 50105~50106~50107~50108~50181~50273~50276~50279 | 2769 |  | 2769 2769 2769 2769 2769 2769 2769 2769 |  |  |  |  | 1.29 mg |  | 1.29 | [5 empty] | 1.0023 | 0.00188 | 50105 |  | XCAMS | Both |  |  | Results signed off | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2012-08-14 | Carbonate |  | JOB-190429 |  |  | 50383 | Normal | 14047/1 |  |  | 2012-06-19 | 439 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9977 | 0.00185 | 50106 | [26 empty] | 2012-06-19 | 50 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00191 | 50107 | [26 empty] | 2012-06-21 | 122 | Graphite LG1 | Completed
[168 empty] | 1.0009 | 0.00187 | 50108 | [26 empty] | 2012-06-27 | 94 | C13 Measurement | Completed
[168 empty] | 0.99627 | 0.00187 | 50181 | [26 empty] | 2012-07-23 | 446 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.0019 | 50273 | [26 empty] | 2012-07-25 | 389 | Review Final Report | Completed
[168 empty] | 0.99962 | 0.00188 | 50276 | [26 empty] | 2012-08-20 | 248 | Report Results | Completed
[168 empty] | 1.0012 | 0.00192 | 50279 | [26 empty] | 2012-08-20 | 789 | Close Job | Completed
[7 empty] | 114.2 | 89.3022 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 52027 | 52027 | 52021 | 488 | 488 | 469 |  |  | Jenny Dahl | -998.46 | 0.09 ‰ | 2012-07-17 | 2012-07-17 | [6 empty] | 3.03‰ | 3.03‰ | 2.00‰ | 2.00‰ | 3.03 ‰ | 2.00 ‰ | AMS | 2.12‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.04 | 0.70711 | -998.46 ‰ |  | 0.09 |  | -998.44 | 0.09 ‰ | -998.44 | 0.09 ‰ |  | Jenny Dahl | 2012-07-23 | Jenny Dahl | 8:56 | 0.0015389 | 0.00154 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.4 |  |  | 0.00154 | 9e-05 | 447 |  |  |  | 190430 | [5 empty] | 5 | 5/07/22 | 15:24:43 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 17 | 5 | 5 | ... | 0.00148 | 9e-05 |  |  |  |  | 1.0066 | 1.0065 | Yes | [6 empty] | 0.00148 |  | 9e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00154 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00148 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-16 | Jenny Dahl | 10:37 |  | Completed |  |  |  |  | 50123 | 2000000 |  |  | 50105 50106 50107 50108 50181 50273 50276 50279 | 50105~50106~50107~50108~50181~50273~50276~50279 | 2769 |  | 2769 2769 2769 2769 2769 2769 2769 2769 |  |  |  |  | 1.39 mg |  | 1.39 | [5 empty] | 1.0023 | 0.00188 | 50105 |  | XCAMS | Both |  |  | Results signed off | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2012-08-14 | Carbonate |  | JOB-190430 |  |  | 50384 | Normal | 14047/1 |  |  | 2012-06-19 | 440 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9977 | 0.00185 | 50106 | [26 empty] | 2012-06-19 | 51 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00191 | 50107 | [26 empty] | 2012-06-21 | 123 | Graphite LG1 | Completed
[168 empty] | 1.0009 | 0.00187 | 50108 | [26 empty] | 2012-06-27 | 95 | C13 Measurement | Completed
[168 empty] | 0.99627 | 0.00187 | 50181 | [26 empty] | 2012-07-23 | 447 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.0019 | 50273 | [26 empty] | 2012-07-25 | 390 | Review Final Report | Completed
[168 empty] | 0.99962 | 0.00188 | 50276 | [26 empty] | 2012-08-20 | 249 | Report Results | Completed
[168 empty] | 1.0012 | 0.00192 | 50279 | [26 empty] | 2012-08-20 | 790 | Close Job | Completed
[7 empty] | 104.1 | 89.3367 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 51866 | 51866 | 51867 | 585 | 585 | 563 |  |  | Dawn Chambers | -998.43 | 0.11 ‰ | 2012-07-23 | 2012-07-23 | [6 empty] | 4.55‰ | 4.55‰ | 2.00‰ | 2.00‰ | 4.55 ‰ | 2.00 ‰ | AMS | 2.27‰ | 0.2‰ | C13 | No | -19 | -19.0065 | 0.04 | 0.70711 | -998.43 ‰ |  | 0.11 |  | -998.41 | 0.12 ‰ | -998.41 | 0.12 ‰ |  | Dawn Chambers | 2012-07-31 | Dawn Chambers | 9:34 | 0.0015701 | 0.00157 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.1 |  |  | 0.00157 | 0.00011 | 592 |  |  |  | 190648 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.16 | 0.01% | 2 | 5 | 5 | ... | 0.00151 | 0.00011 |  |  |  |  | 1.0085 | 1.0085 | Yes | [6 empty] | 0.00151 |  | 0.00011 |  | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00157 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00151 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-20 | Dawn Chambers | 10:47 |  | Completed |  |  |  |  | 50179 | 2000000 |  |  | 50214 50240 50241 50253 50254 50256 50286 50301 | 50214~50240~50241~50253~50254~50256~50286~50301 | 2772 |  | 2772 2772 2772 2772 2772 2772 2772 2772 |  |  |  |  | 1.21 mg |  | 1.21 | [5 empty] | 0.99868 | 0.00195 | 50214 |  | XCAMS | Both |  |  | Signoff with CP/MN treatment | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2012-08-22 | Carbonate |  | JOB-190648 |  |  | 50486 | Normal | 14047/1 |  |  | 2012-06-27 | 483 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99873 | 0.00192 | 50240 | [26 empty] | 2012-06-27 | 67 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99858 | 0.00196 | 50241 | [26 empty] | 2012-06-28 | 179 | Graphite LG1 | Completed
[168 empty] | 1.0027 | 0.00192 | 50253 | [26 empty] | 2012-07-03 | 141 | C13 Measurement | Completed
[168 empty] | 0.99775 | 0.00202 | 50254 | [26 empty] | 2012-07-31 | 592 | AMS Submission Results | Completed
[168 empty] | 1.0035 | 0.00193 | 50256 | [26 empty] | 2012-08-21 | 503 | Review Final Report | Completed
[168 empty] | 0.9999 | 0.00193 | 50286 | [26 empty] | 2012-08-21 | 853 | Close Job | Completed
[168 empty] | 0.99987 | 0.00195 | 50301
[7 empty] | 112.5 | 89.3582 | 0.0001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 59979 | 59978 | 60005 | 1168 | 1168 | 1127 |  |  | Dawn Chambers | -999.43 | 0.08 ‰ | 2012-07-23 | 2012-07-23 | [6 empty] | 4.31‰ | 4.31‰ | 2.00‰ | 2.00‰ | 4.31 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.0065 | 0.04 | 0.70711 | -999.43 ‰ |  | 0.08 |  | -999.42 | 0.08 ‰ | -999.42 | 0.08 ‰ |  | Dawn Chambers | 2012-07-31 | Dawn Chambers | 9:34 | 0.00057189 | 0.00057 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1 |  |  | 0.00057 | 8e-05 | 615 |  |  |  | 190873 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.06 | 0.01% | 12 | 5 | 5 | ... | 0.00055 | 8e-05 |  |  |  |  | 1.0085 | 1.0085 | Yes | [6 empty] | 0.00055 |  | 8e-05 |  | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00057 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00055 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-20 | Dawn Chambers | 10:49 |  | Completed |  |  |  |  | 50263 | 2000000 |  |  | 50214 50240 50241 50253 50254 50256 50286 50301 | 50214~50240~50241~50253~50254~50256~50286~50301 | 2772 |  | 2772 2772 2772 2772 2772 2772 2772 2772 |  |  |  |  | 2.13 mg |  | 2.13 | [5 empty] | 0.99868 | 0.00195 | 50214 |  | XCAMS | Both |  |  | Signoff with CP/MN treatment | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-09-03 | Carbonate |  | JOB-190873 |  | was Job 111233 in access data entered in Rlims. | 50509 | Normal | 14047/1 |  |  | 2012-07-09 | 511 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99873 | 0.00192 | 50240 | [26 empty] | 2012-07-09 | 104 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99858 | 0.00196 | 50241 | [26 empty] | 2012-07-11 | 264 | Graphite LG1 | Completed
[168 empty] | 1.0027 | 0.00192 | 50253 | [26 empty] | 2012-07-12 | 215 | C13 Measurement | Completed
[168 empty] | 0.99775 | 0.00202 | 50254 | [26 empty] | 2012-07-31 | 615 | AMS Submission Results | Completed
[168 empty] | 1.0035 | 0.00193 | 50256 | [26 empty] | 2012-08-21 | 526 | Review Final Report | Completed
[168 empty] | 0.9999 | 0.00193 | 50286 | [26 empty] | 2012-08-21 | 862 | Close Job | Completed
[168 empty] | 0.99987 | 0.00195 | 50301
[7 empty] | 113.2 | 89.3858 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 54411 | 54410 | 54437 | 511 | 511 | 493 |  |  | Jenny Dahl | -998.86 | 0.07 ‰ | 2012-07-26 | 2012-07-26 | [6 empty] | 2.21‰ | 2.21‰ | 2.00‰ | 2.00‰ | 2.21 ‰ | 2.00 ‰ | AMS | 2.24‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0.03 | 0.66667 | -998.86 ‰ |  | 0.07 |  | -998.84 | 0.07 ‰ | -998.84 | 0.07 ‰ |  | Jenny Dahl | 2012-08-01 | Jenny Dahl | 9:48 | 0.0011438 | 0.00114 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.2 |  |  | 0.00114 | 7e-05 | 654 |  |  |  | 190974 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.11 | 0.01% | 33 | 5 | 5 | ... | 0.0011 | 7e-05 |  |  |  |  | 1.0067 | 1.0067 | Yes | [6 empty] | 0.0011 |  | 7e-05 |  | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00114 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0011 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-24 | Jenny Dahl | 9:17 |  | Completed |  |  |  |  | 50308 | 2000000 |  |  | 50132 50133 50182 50183 50212 50213 50218 50247 50252 | 50132~50133~50182~50183~50212~50213~50218~50247~50252 | 2774 |  | 2774 2774 2774 2774 2774 2774 2774 2774 2774 |  |  |  |  | 1.64 mg |  | 1.64 | [5 empty] | 0.99998 | 0.00178 | 50132 |  | XCAMS | Both |  |  | Signoff with MN treatment | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2012-09-10 | Carbonate |  | JOB-190974 |  | Last of the old marble standard. | 50591 | Normal | 14047/1 |  |  | 2012-07-16 | 546 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99804 | 0.00176 | 50133 | [26 empty] | 2012-07-16 | 125 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0048 | 0.00181 | 50182 | [26 empty] | 2012-07-17 | 308 | Graphite LG1 | Completed
[168 empty] | 0.99852 | 0.00177 | 50183 | [26 empty] | 2012-07-19 | 245 | C13 Measurement | Completed
[168 empty] | 0.99885 | 0.00173 | 50212 | [26 empty] | 2012-08-01 | 654 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00172 | 50213 | [26 empty] | 2012-08-21 | 597 | Review Final Report | Completed
[168 empty] | 0.9997 | 0.00174 | 50218 | [26 empty] | 2012-08-21 | 628 | Report Results | Completed
[168 empty] | 0.99956 | 0.00175 | 50247 | [26 empty] | 2012-08-23 | 1085 | Close Job | Completed
[168 empty] | 1.0009 | 0.00179 | 50252
[7 empty] | 119.4 | 89.2981 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 53320 | 53320 | 53321 | 510 | 510 | 491 |  |  | Dawn Chambers | -998.69 | 0.08 ‰ | 2012-07-31 | 2012-07-31 | [6 empty] | 6.81‰ | 6.81‰ | 2.00‰ | 2.00‰ | 6.81 ‰ | 2.00 ‰ | AMS | 2.37‰ | 0.2‰ | C13 | No | -19 | -19.0164 | 0.04 | 0.70711 | -998.69 ‰ |  | 0.08 |  | -998.67 | 0.08 ‰ | -998.67 | 0.08 ‰ |  | Dawn Chambers | 2012-08-06 | Dawn Chambers | 9:14 | 0.0013101 | 0.00131 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1 |  |  | 0.00131 | 8e-05 | 731 |  |  |  | 191015 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.13 | 0.01% | 12 | 5 | 5 | ... | 0.00126 | 8e-05 |  |  |  |  | 1.0103 | 1.0103 | Yes | [6 empty] | 0.00126 |  | 8e-05 |  | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00131 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00126 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-30 | Dawn Chambers | 11:56 |  | Completed |  |  |  |  | 50349 | 2000000 |  |  | 48836 50287 50288 50289 50354 50357 50360 50362 | 48836~50287~50288~50289~50354~50357~50360~50362 | 2776 |  | 2776 2776 2776 2776 2776 2776 2776 2776 |  |  |  |  | 1.89 mg |  | 1.89 | [5 empty] | 0.99924 | 0.00175 | 48836 |  | XCAMS | Both |  |  | Signoff with MN treatment | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-09-11 | Carbonate |  | JOB-191015 |  | acid etched 17/7/2012 | 50667 | Normal | 14047/1 |  |  | 2012-07-17 | 559 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0006 | 0.00169 | 50287 | [26 empty] | 2012-07-17 | 137 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99923 | 0.00174 | 50288 | [26 empty] | 2012-07-19 | 349 | Graphite LG1 | Completed
[168 empty] | 1.0024 | 0.00167 | 50289 | [26 empty] | 2012-07-20 | 271 | C13 Measurement | Completed
[168 empty] | 0.99845 | 0.00168 | 50354 | [26 empty] | 2012-08-06 | 731 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00167 | 50357 | [26 empty] | 2012-08-06 | 674 | Review Final Report | Completed
[168 empty] | 0.99849 | 0.00167 | 50360 | [26 empty] | 2012-08-20 | 491 | Report Results | Completed
[168 empty] | 1.0013 | 0.00169 | 50362 | [26 empty] | 2012-08-20 | 735 | Close Job | Completed
[7 empty] | 119.6 | 89.4649 | 0.0001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 57172 | 57172 | 57183 | 721 | 720 | 694 |  |  | Dawn Chambers | -999.19 | 0.07 ‰ | 2012-07-31 | 2012-07-31 | [6 empty] | 4.93‰ | 4.93‰ | 2.00‰ | 2.00‰ | 4.93 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.0164 | 0.04 | 0.70711 | -999.19 ‰ |  | 0.07 |  | -999.18 | 0.07 ‰ | -999.18 | 0.07 ‰ |  | Dawn Chambers | 2012-08-06 | Dawn Chambers | 9:14 | 0.00081104 | 0.00081 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 0.9 |  |  | 0.00081 | 7e-05 | 732 |  |  |  | 191017 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.08 | 0.01% | 17 | 5 | 5 | ... | 0.00078 | 7e-05 |  |  |  |  | 1.0103 | 1.0103 | Yes | [6 empty] | 0.00078 |  | 7e-05 |  | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00081 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00078 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-07-30 | Dawn Chambers | 11:56 |  | Completed |  |  |  |  | 50374 | 2000000 |  |  | 48836 50287 50288 50289 50354 50357 50360 50362 | 48836~50287~50288~50289~50354~50357~50360~50362 | 2776 |  | 2776 2776 2776 2776 2776 2776 2776 2776 |  |  |  |  | 2.30 mg |  | 2.3 | [5 empty] | 0.99924 | 0.00175 | 48836 |  | XCAMS | Both |  |  | Signoff with MN treatment | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-09-12 | Carbonate |  | JOB-191017 |  | acid etched 17 July 2012 details in J 190973 | 50668 | Normal | 14047/1 |  |  | 2012-07-18 | 566 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0006 | 0.00169 | 50287 | [26 empty] | 2012-07-18 | 143 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99923 | 0.00174 | 50288 | [26 empty] | 2012-07-23 | 374 | Graphite LG1 | Completed
[168 empty] | 1.0024 | 0.00167 | 50289 | [26 empty] | 2012-07-25 | 292 | C13 Measurement | Completed
[168 empty] | 0.99845 | 0.00168 | 50354 | [26 empty] | 2012-08-06 | 732 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00167 | 50357 | [26 empty] | 2012-08-06 | 675 | Review Final Report | Completed
[168 empty] | 0.99849 | 0.00167 | 50360 | [26 empty] | 2012-08-20 | 492 | Report Results | Completed
[168 empty] | 1.0013 | 0.00169 | 50362 | [26 empty] | 2012-08-20 | 736 | Close Job | Completed
[7 empty] | 118.2 | 89.4687 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 52192 | 52191 | 52180 | 499 | 498 | 479 |  |  | Jenny Dahl | -998.49 | 0.09 ‰ | 2012-08-09 | 2012-08-09 | [6 empty] | 3.34‰ | 3.34‰ | 2.00‰ | 2.00‰ | 3.34 ‰ | 2.00 ‰ | AMS | 2.17‰ | 0.2‰ | C13 | No | -19 | -19.016 | 0.04 | 0.70711 | -998.49 ‰ |  | 0.09 |  | -998.47 | 0.09 ‰ | -998.47 | 0.09 ‰ |  | Jenny Dahl | 2012-08-10 | Jenny Dahl | 10:39 | 0.0015077 | 0.00151 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.5 |  |  | 0.00151 | 9e-05 | 819 |  |  |  | 191014 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 17 | 5 | 5 | ... | 0.00145 | 9e-05 |  |  |  |  | 1.0087 | 1.0087 | Yes | [6 empty] | 0.00145 |  | 9e-05 |  | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00151 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00145 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-08-03 | Jenny Dahl | 11:17 |  | Completed |  |  |  |  | 50348 | 2000000 |  |  | 50255 50290 50291 50302 50323 50326 50329 50332 | 50255~50290~50291~50302~50323~50326~50329~50332 | 2779 |  | 2779 2779 2779 2779 2779 2779 2779 2779 |  |  |  |  | 1.24 mg |  | 1.24 | [5 empty] | 0.99884 | 0.00183 | 50255 |  | XCAMS | Both |  |  | Signoff with JT/Lab | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2012-09-11 | Carbonate |  | JOB-191014 |  | Acid etched 17/7/2012 details in J190973 | 50755 | Normal | 14047/1 |  |  | 2012-07-17 | 558 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99862 | 0.00178 | 50290 | [26 empty] | 2012-07-17 | 136 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99966 | 0.00177 | 50291 | [26 empty] | 2012-07-19 | 348 | Graphite LG1 | Completed
[168 empty] | 0.99869 | 0.00176 | 50302 | [26 empty] | 2012-07-20 | 270 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00176 | 50323 | [26 empty] | 2012-08-10 | 819 | AMS Submission Results | Completed
[168 empty] | 1.0017 | 0.00177 | 50326 | [26 empty] | 2012-08-21 | 762 | Review Final Report | Completed
[168 empty] | 0.99873 | 0.00177 | 50329 | [26 empty] | 2012-08-23 | 1086 | Close Job | Completed
[168 empty] | 1.0039 | 0.00184 | 50332
[7 empty] | 115.8 | 89.3744 | 0.0001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 64330 | 64329 | 64396 | 1004 | 1004 | 974 |  |  | Jenny Dahl | -999.67 | 0.04 ‰ | 2012-08-16 | 2012-08-16 | [6 empty] | 3.22‰ | 3.22‰ | 2.00‰ | 2.00‰ | 3.22 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0068 | 0.04 | 0.70711 | -999.67 ‰ |  | 0.04 |  | -999.66 | 0.04 ‰ | -999.66 | 0.04 ‰ |  | Jenny Dahl | 2012-08-21 | Jenny Dahl | 14:56 | 0.00033273 | 0.00033 | 4e-05 | 4e-05 | 4e-05 | 4e-05 | 0.9 |  |  | 0.00033 | 4e-05 | 905 |  |  |  | 191272 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.03 |  | 27 | 5 | 5 | ... | 0.00032 | 4e-05 |  |  |  |  | 1.0075 | 1.0076 | Yes | [6 empty] | 0.00032 |  | 4e-05 |  | No | 1 | 0.99998 | 0.00097 | 0.00097 | 0.00033 | 0 | 4e-05 | 0 | 0 |  |  |  |  | 0.00032 | 0 | 4e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-08-15 | Jenny Dahl | 12:26 |  | Completed |  |  |  |  | 50471 | 2000000 |  |  | 50303 50304 50305 50306 50434 50462 50465 50468 | 50303~50304~50305~50306~50434~50462~50465~50468 | 2781 |  | 2781 2781 2781 2781 2781 2781 2781 2781 |  |  |  |  | 2.55 mg |  | 2.55 | [5 empty] | 0.9991 | 0.00273 | 50303 |  | XCAMS | Both | 0.24% |  | - system error = 0.24% - MB subtraction applied - after the first 13 rotations the system was retuned because of high variance amongst standards. This retune had no effect on this variance. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-10-01 | Carbonate |  | JOB-191272 |  | Acid etched 17/7/2012 | 50847 | Normal | 14047/1 |  |  | 2012-08-06 | 608 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0047 | 0.00274 | 50304 | [26 empty] | 2012-08-06 | 163 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00273 | 50305 | [26 empty] | 2012-08-07 | 471 | Graphite LG1 | Completed
[168 empty] | 1.0015 | 0.00275 | 50306 | [26 empty] | 2012-08-10 | 374 | C13 Measurement | Completed
[168 empty] | 0.99545 | 0.00273 | 50434 | [26 empty] | 2012-08-21 | 905 | AMS Submission Results | Completed
[168 empty] | 1.0007 | 0.00275 | 50462 | [26 empty] | 2012-08-21 | 871 | Review Final Report | Completed
[168 empty] | 0.99764 | 0.00272 | 50465 | [26 empty] | 2012-08-23 | 1042 | Close Job | Completed
[168 empty] | 1.0005 | 0.00273 | 50468
[7 empty] | 133.3 | 88.9423 | 0.0001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 60126 | 60126 | 60148 | 1041 | 1041 | 1004 |  |  | Silvia Canessa | -999.44 | 0.07 ‰ | 2012-09-07 | 2012-09-07 | [6 empty] | 4.83‰ | 4.83‰ | 2.00‰ | 2.00‰ | 4.83 ‰ | 2.00 ‰ | AMS | 2.32‰ | 0.2‰ | C13 | No | -19 | -19.003 | 0.04 | 0.70711 | -999.44 ‰ |  | 0.07 |  | -999.43 | 0.07 ‰ | -999.43 | 0.07 ‰ |  | Silvia Canessa | 2012-09-11 | Silvia Canessa | 14:28 | 0.00056149 | 0.00056 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1 |  |  | 0.00056 | 7e-05 | 1146 |  |  |  | 191273 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.06 | 0.01% | 17 | 5 | 5 | ... | 0.00054 | 7e-05 |  |  |  |  | 1.0043 | 1.0043 | Yes |  |  | JCT - This background is exceptionally good, but the kapuni in the wheel is even better. Last couple of C-1 marbles have also been very good, so used this blank as usual. |  |  |  | 0.00054 |  | 7e-05 |  | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00056 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00054 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-06 | Silvia Canessa | 13:50 |  | Completed |  |  |  |  | 50472 | 2000000 |  |  | 48837 48840 48841 50486 50487 50488 50489 50543 | 48837~48840~48841~50486~50487~50488~50489~50543 | 2787 |  | 2787 2787 2787 2787 2787 2787 2787 2787 |  |  |  |  | 2.19 mg |  | 2.19 | [5 empty] | 0.99724 | 0.00204 | 48837 |  | XCAMS | ST | 0.09% |  | -used two regression regions -needed an additional factor of 1.07 to explain the global run-to-run variance - system error of 0.09% - rejected first rotation - rejected first (high) runs of blanks | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2012-10-01 | Carbonate |  | JOB-191273 |  | Acid etched 17/7/2012 |  | Normal | 14047/1 |  |  | 2012-08-06 | 609 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99939 | 0.00194 | 48840 | [26 empty] | 2012-08-06 | 164 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99703 | 0.00196 | 48841 | [26 empty] | 2012-08-07 | 472 | Graphite LG1 | Completed
[168 empty] | 1.0018 | 0.00197 | 50486 | [26 empty] | 2012-08-10 | 375 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00196 | 50487 | [26 empty] | 2012-09-11 | 1146 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.002 | 50488 | [26 empty] | 2012-10-05 | 1740 | Close Job | Completed
[168 empty] | 1.0019 | 0.00196 | 50489
[168 empty] | 0.99999 | 0.00195 | 50543
[7 empty] | 132.5 | 89.2049 | 0.0003 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 |  |  |  | 49608 | 49608 | 49607 | 321 | 321 | 309 |  |  | jenny dahl | -997.92 | 0.08 ‰ | 2012-09-14 | 2012-09-14 | [6 empty] | 3.90‰ | 3.90‰ | 2.00‰ | 2.00‰ | 3.90 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0.04 | 0.8165 | -997.92 ‰ |  | 0.08 |  | -997.89 | 0.08 ‰ | -997.89 | 0.08 ‰ |  | jenny dahl | 2012-09-18 | jenny dahl | 12:03 | 0.0020796 | 0.00208 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.4 |  |  | 0.00208 | 8e-05 | 1255 |  |  |  | 191277 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.21 | 0.01% | 12 | 5 | 5 | ... | 0.002 | 8e-05 |  |  |  |  | 1.0063 | 1.0063 | Yes | [6 empty] | 0.002 |  | 8e-05 |  | No | 1 | 0.99983 | 0.001 | 0.001 | 0.00208 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.002 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-13 | jenny dahl | 15:32 |  | Completed |  |  |  |  | 50482 | 2000000 |  |  | 50669 50670 50671 50672 50725 50731 | 50669~50670~50671~50672~50725~50731 | 2790 |  | 2790 2790 2790 2790 2790 2790 |  |  |  |  | 0.48 mg |  | 0.48 | [5 empty] | 0.99811 | 0.00243 | 50669 |  | XCAMS | EA | 0.19% |  | - a system error of 0.19( was needed to explain the calibration pos-to-pos variance - a few samples have been sputtered through, but effected runs did not need to be rejected as those runs did not influence the quality of the data set - TP#50657 came out | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-10-01 | Carbonate |  | JOB-191277 |  | Acid etched 17/7/2012. Not recombusted |  | Normal | 14047/1 |  |  | 2012-08-06 | 613 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0028 | 0.00244 | 50670 | [26 empty] | 2012-08-06 | 168 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9954 | 0.00275 | 50671 | [26 empty] | 2012-08-07 | 482 | Graphite LG1 | Completed
[168 empty] | 0.99928 | 0.0024 | 50672 | [26 empty] | 2012-08-10 | 385 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00239 | 50725 | [26 empty] | 2012-09-18 | 1255 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.0024 | 50731 | [26 empty] | 2012-10-26 | 2268 | Close Job | Completed
[7 empty] | 119.3 | 89.3341 | 0.0008 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 |  |  |  | 42410 | 42410 | 42418 | 197 | 196 | 189 |  |  | jenny dahl | -994.91 | 0.12 ‰ | 2012-09-14 | 2012-09-14 | [6 empty] | 2.44‰ | 2.44‰ | 2.00‰ | 2.00‰ | 2.44 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0018 | 0.04 | 0.8165 | -994.91 ‰ |  | 0.12 |  | -994.84 | 0.13 ‰ | -994.84 | 0.13 ‰ |  | jenny dahl | 2012-09-18 | jenny dahl | 12:03 | 0.005095 | 0.00509 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 4.6 |  |  | 0.00509 | 0.00012 | 1260 |  |  |  | 191617 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.51 | 0.01% | 26 | 5 | 6 | ... | 0.0049 | 0.00012 |  |  |  |  | 1.0063 | 1.0063 | Yes |  |  | Recombusted and little - did not use to correct samples. |  |  |  | 0.0049 |  | 0.00012 |  | No | 1 | 0.99983 | 0.001 | 0.001 | 0.00509 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.0049 | 0 | 0.00012 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-13 | jenny dahl | 15:32 |  | Completed |  |  |  |  | 50699 | 2000000 |  |  | 50669 50670 50671 50672 50725 50731 | 50669~50670~50671~50672~50725~50731 | 2790 |  | 2790 2790 2790 2790 2790 2790 |  |  |  |  | 0.25 mg |  | 0.25 | [5 empty] | 0.99811 | 0.00243 | 50669 |  | XCAMS | EA | 0.19% |  | - a system error of 0.19( was needed to explain the calibration pos-to-pos variance - a few samples have been sputtered through, but effected runs did not need to be rejected as those runs did not influence the quality of the data set - TP#50657 came out |  | RRL | [5 empty] | Background Inorganic |  | 2012-10-19 | Carbonate |  | JOB-191617 |  | AZ has confirmed AMS category no <0.3 UI in wheel. Recombusted. SIL gave no result for d13C. This job is for a combusted marble standard load <4mg for evolution then combust gas |  | Normal | 14047/1 |  |  | 2012-08-24 | 700 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0028 | 0.00244 | 50670 | [26 empty] | 2012-08-27 | 199 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9954 | 0.00275 | 50671 | [26 empty] | 2012-09-04 | 318 | Sealed Tube Combustion | Completed
[168 empty] | 0.99928 | 0.0024 | 50672 | [26 empty] | 2012-09-05 | 699 | Graphite LG1 | Completed
[168 empty] | 1.0018 | 0.00239 | 50725 | [26 empty] | 2012-09-11 | 544 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.0024 | 50731 | [26 empty] | 2012-09-18 | 1260 | AMS Submission Results | Completed
[197 empty] | 2012-10-26 | 2251 | Close Job | Completed
[7 empty] | 142 | 89.5455 | 0.0004 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 |  |  |  | 48591 | 48591 | 48592 | 318 | 318 | 306 |  |  | jenny dahl | -997.64 | 0.09 ‰ | 2012-09-24 | 2012-09-24 | [6 empty] | -2.92‰ | -2.92‰ | 2.00‰ | 2.00‰ | -2.92 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0.04 | 0.70711 | -997.64 ‰ |  | 0.09 |  | -997.61 | 0.09 ‰ | -997.61 | 0.09 ‰ |  | Christine Prior | 2012-09-26 | Christine Prior | 16:50 | 0.0023603 | 0.00236 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.1 |  |  | 0.00236 | 9e-05 | 1297 |  |  |  | 191276 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.24 | 0.01% | 13 | 5 | 5 | ... | 0.00227 | 9e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.00227 |  | 9e-05 |  | Yes | 1 | 1 | 0.0009 | 0.0009 | 0.00236 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00227 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-13 | jenny dahl | 15:35 |  | Completed |  |  |  |  | 50481 | 2000000 |  |  | 50451 50459 50501 50502 50505 50507 50682 50686 | 50451~50459~50501~50502~50505~50507~50682~50686 | 2794 |  | 2794 2794 2794 2794 2794 2794 2794 2794 |  |  |  |  | 0.46 mg |  | 0.46 | [5 empty] | 1.0008 | 0.00254 | 50451 |  | XCAMS | EA | 0.18% | Many of the samples in this wheel are close to background. These samples were submitted as CO2 Gas, prepared by the submitter from marine forams. The background materials were prepared in our laboratory, and we used IAEA C-1 carbonate as background material. Since the CO2 evolution was done in different laboratories, with different procedures, the background corrections do not accurately reflect the entire process the samples went through. This should be taken into consideration when interpreting these results, especially for those that are >35,000 BP, where the background correction has the largest effect. We are happy to provide more details and discuss this further as needed. | A machine blank count rate subtraction was performed using cathode #36, which was ex XCAMS stock. Overdispersion required the following three uncertainty adjustments: a global factor x1.04 (i.e. for all runs of all cathodes), a further factor x1.3 for all runs of cathode #28 (TP50705), and a 0.18% system error. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-10-01 | Carbonate |  | JOB-191276 |  | sml std large std large std Acid etched 17/7/2012 |  | Normal | 14047/1 |  |  | 2012-08-06 | 612 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99947 | 0.00255 | 50459 | [26 empty] | 2012-08-06 | 167 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99476 | 0.00254 | 50501 | [26 empty] | 2012-08-07 | 481 | Graphite LG1 | Completed
[168 empty] | 0.99924 | 0.00253 | 50502 | [26 empty] | 2012-08-10 | 378 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00254 | 50505 | [26 empty] | 2012-09-26 | 1297 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00255 | 50507 | [26 empty] | 2012-10-26 | 2299 | Close Job | Completed
[168 empty] | 1.002 | 0.00255 | 50682
[168 empty] | 0.999 | 0.00253 | 50686
[7 empty] | 140.7 | 89.0149 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 50727 | 50727 | 50724 | 369 | 369 | 355 |  |  | jenny dahl | -998.19 | 0.08 ‰ | 2012-09-24 | 2012-09-24 | [6 empty] | 3.02‰ | 3.02‰ | 2.00‰ | 2.00‰ | 3.02 ‰ | 2.00 ‰ | AMS | 1.73‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0.04 | 0.70711 | -998.19 ‰ |  | 0.08 |  | -998.17 | 0.08 ‰ | -998.17 | 0.08 ‰ |  | Christine Prior | 2012-09-26 | Christine Prior | 16:50 | 0.0018092 | 0.00181 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.6 |  |  | 0.00181 | 8e-05 | 1301 |  |  |  | 191622 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.18 | 0.01% | 27 | 5 | 5 | ... | 0.00174 | 8e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.00174 |  | 8e-05 |  | Yes | 1 | 1 | 0.0009 | 0.0009 | 0.00181 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00174 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-13 | jenny dahl | 15:36 |  | Completed |  |  |  |  | 50647 | 2000000 |  |  | 50451 50459 50501 50502 50505 50507 50682 50686 | 50451~50459~50501~50502~50505~50507~50682~50686 | 2794 |  | 2794 2794 2794 2794 2794 2794 2794 2794 |  |  |  |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0008 | 0.00254 | 50451 |  | XCAMS | EA | 0.18% | Many of the samples in this wheel are close to background. These samples were submitted as CO2 Gas, prepared by the submitter from marine forams. The background materials were prepared in our laboratory, and we used IAEA C-1 carbonate as background material. Since the CO2 evolution was done in different laboratories, with different procedures, the background corrections do not accurately reflect the entire process the samples went through. This should be taken into consideration when interpreting these results, especially for those that are >35,000 BP, where the background correction has the largest effect. We are happy to provide more details and discuss this further as needed. | A machine blank count rate subtraction was performed using cathode #36, which was ex XCAMS stock. Overdispersion required the following three uncertainty adjustments: a global factor x1.04 (i.e. for all runs of all cathodes), a further factor x1.3 for all runs of cathode #28 (TP50705), and a 0.18% system error. | 1.7 | RRL | [5 empty] | Background Inorganic |  | 2012-10-19 | Carbonate |  | JOB-191622 |  | lrg std load >15 <25mg marble for evolution |  | Normal | 14047/1 |  |  | 2012-08-24 | 705 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99947 | 0.00255 | 50459 | [26 empty] | 2012-08-27 | 204 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99476 | 0.00254 | 50501 | [26 empty] | 2012-08-30 | 647 | Graphite LG1 | Completed
[168 empty] | 0.99924 | 0.00253 | 50502 | [26 empty] | 2012-09-07 | 515 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00254 | 50505 | [26 empty] | 2012-09-26 | 1301 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00255 | 50507 | [26 empty] | 2012-10-26 | 2303 | Close Job | Completed
[168 empty] | 1.002 | 0.00255 | 50682
[168 empty] | 0.999 | 0.00253 | 50686
[7 empty] | 135.2 | 89.0407 | 0.0007 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 |  |  |  | 44020 | 44020 | 44020 | 280 | 280 | 289 |  |  | jenny dahl | -995.83 | 0.15 ‰ | 2012-09-24 | 2012-09-24 | [6 empty] | 2.73‰ | 2.73‰ | 2.00‰ | 2.00‰ | 2.73 ‰ | 2.00 ‰ | AMS | 2.01‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0.04 | 0.70711 | -995.83 ‰ |  | 0.15 |  | -995.78 | 0.15 ‰ | -995.78 | 0.15 ‰ |  | Christine Prior | 2012-09-26 | Christine Prior | 16:50 | 0.0041696 | 0.00417 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 3.5 |  |  | 0.00417 | 0.00015 | 1302 |  |  |  | 191627 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.42 | 0.01% | 17 | 5 | 5 | ... | 0.00401 | 0.00014 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.00401 |  | 0.00014 |  | Yes | 1 | 1 | 0.0009 | 0.0009 | 0.00417 | 0 | 0.00015 | 0 | 0 |  |  |  |  | 0.00401 | 0 | 0.00014 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-13 | jenny dahl | 15:36 |  | Completed |  |  |  |  | 50644 | 2000000 |  |  | 50451 50459 50501 50502 50505 50507 50682 50686 | 50451~50459~50501~50502~50505~50507~50682~50686 | 2794 |  | 2794 2794 2794 2794 2794 2794 2794 2794 |  |  |  |  | 0.22 mg |  | 0.22 | [5 empty] | 1.0008 | 0.00254 | 50451 |  | XCAMS | EA | 0.18% | Many of the samples in this wheel are close to background. These samples were submitted as CO2 Gas, prepared by the submitter from marine forams. The background materials were prepared in our laboratory, and we used IAEA C-1 carbonate as background material. Since the CO2 evolution was done in different laboratories, with different procedures, the background corrections do not accurately reflect the entire process the samples went through. This should be taken into consideration when interpreting these results, especially for those that are >35,000 BP, where the background correction has the largest effect. We are happy to provide more details and discuss this further as needed. | A machine blank count rate subtraction was performed using cathode #36, which was ex XCAMS stock. Overdispersion required the following three uncertainty adjustments: a global factor x1.04 (i.e. for all runs of all cathodes), a further factor x1.3 for all runs of cathode #28 (TP50705), and a 0.18% system error. | 2 | RRL | [5 empty] | Background Inorganic |  | 2012-10-19 | Carbonate |  | JOB-191627 |  | small std |  | Normal | 14047/1 |  |  | 2012-08-24 | 706 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99947 | 0.00255 | 50459 | [26 empty] | 2012-08-27 | 205 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99476 | 0.00254 | 50501 | [26 empty] | 2012-08-29 | 644 | Graphite LG1 | Completed
[168 empty] | 0.99924 | 0.00253 | 50502 | [26 empty] | 2012-09-07 | 506 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00254 | 50505 | [26 empty] | 2012-09-26 | 1302 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00255 | 50507 | [26 empty] | 2012-10-26 | 2304 | Close Job | Completed
[168 empty] | 1.002 | 0.00255 | 50682
[168 empty] | 0.999 | 0.00253 | 50686
[7 empty] | 109.3 | 89.0692 | 0.0003 | RRL-Test-SM | Background Test | 1950-01-01 |  |  |  | 50148 | 50148 | 50167 | 515 | 515 | 497 | 49648.7 | 50681.1 | jenny dahl | -998.06 | 0.12 ‰ | 2012-09-17 | 2012-09-17 | 7 | [5 empty] | 3.67‰ | 3.67‰ | 2.00‰ | 2.00‰ | 3.67 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -998.06 ‰ |  | 0.12 |  | -998.03 | 0.13 ‰ | -998.03 | 0.13 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.0019444 | 0.00194 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 11.6 |  |  | 0.00194 | 0.00012 | 1324 |  |  |  | 191619 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.19 | 0.01% | 4 | 5 | 4 | ... | 0.00187 | 0.00012 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.00187 |  | 0.00012 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.00194 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.00187 | 0 | 0.00012 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:47 |  | Completed |  |  |  |  | 60316 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.17 mg |  | 0.17 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191619 |  | JCT 5/4/13 Invoice GCT1. small std load 4mg or less of marble for evolution |  | Normal | 14047/1 |  |  | 2012-08-24 | 702 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-28 | 201 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-10 | 326 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1324 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1959 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 104.3 | 89.3108 | 0.0005 | RRL-Test-SM | Background Test | 1950-01-01 |  |  |  | 46906 | 46905 | 46910 | 373 | 372 | 386 | 46541 | 47287.5 | jenny dahl | -997.09 | 0.14 ‰ | 2012-09-17 | 2012-09-17 | 6 | [5 empty] | 0.95‰ | 0.95‰ | 2.00‰ | 2.00‰ | 0.95 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -997.09 ‰ |  | 0.14 |  | -997.05 | 0.14 ‰ | -997.05 | 0.14 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.0029114 | 0.00291 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 11.6 |  |  | 0.00291 | 0.00014 | 1329 |  |  |  | 191628 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.29 | 0.01% | 10 | 5 | 4 | ... | 0.0028 | 0.00013 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.0028 |  | 0.00013 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.00291 | 0 | 0.00014 | 0 | 0 |  |  |  |  | 0.0028 | 0 | 0.00013 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60321 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.17 mg |  | 0.17 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191628 |  | JCT 5/4/13 Invoice GCT1. for tiny ant wheel load 1.72 mg for evolution |  | Normal | 14047/1 |  |  | 2012-08-24 | 707 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-29 | 206 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 331 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1329 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1964 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 63.6 | 90.0761 | 0.001 | RRL-Test-TY | Background Test | 1950-01-01 |  |  |  | 40973 | 40972 | 40977 | 411 | 411 | 409 | 40571.9 | 41395.1 | jenny dahl | -993.91 | 0.31 ‰ | 2012-09-17 | 2012-09-17 | 5 | [5 empty] | -7.55‰ | -7.55‰ | 2.00‰ | 2.00‰ | -7.55 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -993.91 ‰ |  | 0.31 |  | -993.83 | 0.32 ‰ | -993.83 | 0.32 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.0060932 | 0.00609 | 0.00031 | 0.00031 | 0.00031 | 0.00031 | 26.3 |  |  | 0.00609 | 0.00031 | 1330 |  |  |  | 191629 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.61 | 0.03% | 20 | 5 | 4 | ... | 0.00586 | 0.0003 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.00586 |  | 0.0003 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.00609 | 0 | 0.00031 | 0 | 0 |  |  |  |  | 0.00586 | 0 | 0.0003 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60329 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.07 mg |  | 0.07 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191629 |  | JCT 5/4/13 Invoice GCT1. for tiny ant wheel load 0.86 mg for evolution |  | Normal | 14047/1 |  |  | 2012-08-24 | 708 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-09-07 | 207 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 339 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1330 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1965 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 52.9 | 89.4716 | 0.0013 | RRL-Test-TY | Background Test | 1950-01-01 |  |  |  | 39088 | 39087 | 39093 | 488 | 487 | 490 | 38614.2 | 39591.1 | jenny dahl | -992.3 | 0.47 ‰ | 2012-09-17 | 2012-09-17 | 6 | [5 empty] | -0.84‰ | -0.84‰ | 2.00‰ | 2.00‰ | -0.84 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -992.30 ‰ |  | 0.47 |  | -992.2 | 0.47 ‰ | -992.2 | 0.47 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.0077049 | 0.0077 | 0.00047 | 0.00047 | 0.00047 | 0.00047 | 37.9 |  |  | 0.0077 | 0.00047 | 1331 |  |  |  | 191630 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.77 | 0.05% | 18 | 5 | 4 | ... | 0.00741 | 0.00045 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.00741 |  | 0.00045 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.0077 | 0 | 0.00047 | 0 | 0 |  |  |  |  | 0.00741 | 0 | 0.00045 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60322 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.05 mg |  | 0.05 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191630 |  | JCT 5/4/13 Invoice GCT1. for tiny ant wheel load 0.43 mg for evolution |  | Normal | 14047/1 |  |  | 2012-08-24 | 709 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-29 | 208 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 332 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1331 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1966 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 28.2 | 90.0042 | 0.0025 | RRL-Test-TY | Background Test | 1950-01-01 |  |  |  | 33954 | 33954 | 33953 | 486 | 486 | 484 | 33481.9 | 34455.7 | jenny dahl | -985.4 | 0.88 ‰ | 2012-09-17 | 2012-09-17 | 6 | [5 empty] | -6.76‰ | -6.76‰ | 2.00‰ | 2.00‰ | -6.76 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -985.40 ‰ |  | 0.88 |  | -985.22 | 0.89 ‰ | -985.22 | 0.89 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.014599 | 0.0146 | 0.00088 | 0.00088 | 0.00088 | 0.00088 | 68.8 |  |  | 0.0146 | 0.00088 | 1332 |  |  |  | 191637 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 1.46 | 0.09% | 23 | 5 | 4 | ... | 0.01404 | 0.00085 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.01404 |  | 0.00085 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.0146 | 0 | 0.00088 | 0 | 0 |  |  |  |  | 0.01404 | 0 | 0.00085 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60323 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.03 mg |  | 0.03 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191637 |  | JCT 5/4/13 Invoice GCT1. for tiny ant wheel load approx 0.25mg marble for evolution this will need to be weighed on 6 place balance |  | Normal | 14047/1 |  |  | 2012-08-24 | 710 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-29 | 209 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 333 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1332 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1967 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 10.4 | 90.2882 | 0.0064 | RRL-Test-TY | Background Test | 1950-01-01 |  |  |  | 26069 | 26068 | 26069 | 491 | 490 | 491 | 25592.1 | 26575.2 | jenny dahl | -961.04 | 2.38 ‰ | 2012-09-17 | 2012-09-17 | 6 | [5 empty] | -9.88‰ | -9.88‰ | 2.00‰ | 2.00‰ | -9.88 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -961.04 ‰ |  | 2.38 |  | -960.56 | 2.41 ‰ | -960.56 | 2.41 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.038961 | 0.03896 | 0.00238 | 0.00238 | 0.00238 | 0.00238 | 207.4 |  |  | 0.03896 | 0.00238 | 1333 |  |  |  | 191638 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 3.9 | 0.24% | 31 | 5 | 4 | ... | 0.03747 | 0.00229 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.03747 |  | 0.00229 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.03896 | 0 | 0.00238 | 0 | 0 |  |  |  |  | 0.03747 | 0 | 0.00229 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60324 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.01 mg |  | 0.01 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Test |  | 2012-10-19 | Carbonate |  | JOB-191638 |  | JCT 5/4/13 Invoice GCT1. for tiny ant wheel load approx 0.08mg marble for evolution this will need to be weighed on 6 place balance |  | Normal | 14047/1 |  |  | 2012-08-24 | 711 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-29 | 210 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 334 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1333 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1968 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 126.8 | 89.1693 | 0.0003 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 |  |  |  | 51203 | 51202 | 51180 | 392 | 391 | 376 |  |  | jenny dahl | -998.29 | 0.08 ‰ | 2012-09-17 | 2012-09-17 | 6 | [5 empty] | 2.54‰ | 2.54‰ | 2.00‰ | 2.00‰ | 2.54 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0.04 | 0.8165 | -998.29 ‰ |  | 0.08 |  | -998.27 | 0.08 ‰ | -998.27 | 0.08 ‰ |  | Jocelyn Turnbull | 2012-09-27 | Jocelyn Turnbull | 17:27 | 0.0017053 | 0.00171 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 4.5 |  |  | 0.00171 | 8e-05 | 1335 |  |  |  | 191660 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.17 | 0.01% | 38 | 5 | 4 | ... | 0.00164 | 8e-05 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.00164 |  | 8e-05 |  | No | 1 | 1 | 0.00094 | 0.00094 | 0.00171 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00164 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-14 | jenny dahl | 14:48 |  | Completed |  |  |  |  | 60325 | 2000000 |  |  | 60293 60294 60295 60297 60303 60304 | 60293~60294~60295~60297~60303~60304 | 2791 |  | 2791 2791 2791 2791 2791 2791 |  |  |  |  | 0.43 mg |  | 0.43 | [5 empty] | 1.0021 | 0.00241 | 60293 |  | XCAMS | Both | 0.11% |  | This is the 1st time that the RRL team provided us an Alfa-Aesar graphite intended to be used for machine blank subtraction instead of doing this with a cathode prepped by XCAMS from the same stock material. The measurement shows that this different way of proceeding increased the 14C count rate by ca. a factor x4. This is a much higher count rate than we would expect from variability in the machine blank values we are accustomed to. Also, there were two other blanks (TP60325,60316) with a lower 14C count rate (0.41 14C cnt/s) than the new machine blank cathode (TP60326, 0.48 14C cnt/s). Because of this, we have not performed a machine blank subtraction in CalAMS. Uncertainties were adjusted to reflect overdispersion amongst runs and amongst Ox-I cathodes: global regression factor 1.14, system error 0.11%, resp. All runs with declining beam intensity due to sputtering through have been rejected. |  | RRL | [5 empty] | Background Inorganic |  | 2012-10-23 | Carbonate |  | JOB-191660 |  | JCT 5/4/13 Invoice GCT1. |  | Normal | 14047/1 |  |  | 2012-08-28 | 714 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99765 | 0.00225 | 60294 | [26 empty] | 2012-08-29 | 213 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99725 | 0.00225 | 60295 | [26 empty] | 2012-09-12 | 335 | Graphite | Completed
[168 empty] | 1.001 | 0.00233 | 60297 | [26 empty] | 2012-09-27 | 1335 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00235 | 60303 | [26 empty] | 2012-10-10 | 1970 | Close Job | Completed
[168 empty] | 0.99989 | 0.00224 | 60304
[7 empty] | 130.3 | 89.0339 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 52247 | 52247 | 52233 | 446 | 446 | 428 |  |  | jenny dahl | -998.5 | 0.08 ‰ | 2012-09-20 | 2012-09-20 | [6 empty] | 6.23‰ | 6.23‰ | 2.00‰ | 2.00‰ | 6.23 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0046 | 0.04 | 0.75593 | -998.50 ‰ |  | 0.08 |  | -998.48 | 0.08 ‰ | -998.48 | 0.08 ‰ |  | jenny dahl | 2012-09-24 | jenny dahl | 16:00 | 0.0014973 | 0.0015 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 0.8 |  |  | 0.0015 | 8e-05 | 1402 |  |  |  | 191618 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 28 | 5 | 6 | ... | 0.00144 | 8e-05 |  |  |  |  | 1.0067 | 1.0067 | Yes | [6 empty] | 0.00144 |  | 8e-05 |  | No | 1 | 1.0001 | 0.00117 | 0.00117 | 0.0015 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00144 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-19 | jenny dahl | 10:22 |  | Completed |  |  |  |  | 50711 | 2000000 |  |  | 50648 50675 50676 50677 50678 50733 50738 | 50648~50675~50676~50677~50678~50733~50738 | 2793 |  | 2793 2793 2793 2793 2793 2793 2793 |  |  |  |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99665 | 0.00306 | 50648 |  | XCAMS | Both | 0.26% |  | OxI in pos#0 (TP50643) came out 5% below the other OxIs and was therefore rejected as normalising cathode. We have a 0.26% system error and a global regression uncertainty multiplier of x1.05. In this wheel we measured 6 machine blanks from our own stock. The stdev in these 6 values is 24%, primarily due to cat#7. Cat#17 has been chosen to do the machine-blank 14C countrate subtraction with, because its countrate was the closest to the average count rate of all 6 MB's. Evaluation of TP 50743 and 50744 results: decided that they could be switched from background test to real blanks. Will average 3 blanks together. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-10-19 | Carbonate |  | JOB-191618 |  | This job is for a combusted marble standard load >15<25mg then combust gas Testing backgrounds - can be removed from wheel if necessary |  | Normal | 14047/1 |  |  | 2012-08-24 | 701 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0038 | 0.00308 | 50675 | [26 empty] | 2012-08-27 | 200 | Carbonate CO2 Evolution | Completed
[168 empty] | 1 | 0.00308 | 50676 | [26 empty] | 2012-09-04 | 319 | Sealed Tube Combustion | Completed
[168 empty] | 1.0001 | 0.00307 | 50677 | [26 empty] | 2012-09-05 | 711 | Graphite LG1 | Completed
[168 empty] | 1.0024 | 0.00308 | 50678 | [26 empty] | 2012-09-11 | 545 | C13 Measurement | Completed
[168 empty] | 1.0023 | 0.00336 | 50733 | [26 empty] | 2012-09-24 | 1402 | AMS Submission Results | Completed
[168 empty] | 0.99574 | 0.00306 | 50738 | [26 empty] | 2012-10-26 | 2282 | Close Job | Completed
[7 empty] | 138.9 | 88.9453 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 43367 | 43366 | 43372 | 222 | 221 | 213 |  |  | jenny dahl | -995.48 | 0.12 ‰ | 2012-09-20 | 2012-09-20 | [6 empty] | 7.23‰ | 7.23‰ | 2.00‰ | 2.00‰ | 7.23 ‰ | 2.00 ‰ | AMS | 2.07‰ | 0.2‰ | C13 | No | -19 | -19.0046 | 0.04 | 0.75593 | -995.48 ‰ |  | 0.12 |  | -995.42 | 0.13 ‰ | -995.42 | 0.13 ‰ |  | jenny dahl | 2012-09-24 | jenny dahl | 16:00 | 0.0045231 | 0.00452 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.3 |  |  | 0.00452 | 0.00012 | 1403 |  |  |  | 191621 | [5 empty] | 5 | 5/07/22 | 15:24:44 | Albert Zondervan | 0 |  | 0.45 | 0.01% | 18 | 5 | 5 | ... | 0.00435 | 0.00012 |  |  |  |  | 1.0067 | 1.0067 | Yes | [6 empty] | 0.00435 |  | 0.00012 |  | No | 1 | 1.0001 | 0.00117 | 0.00117 | 0.00452 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.00435 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-09-19 | jenny dahl | 10:22 |  | Completed |  |  |  |  | 50646 | 2000000 |  |  | 50648 50675 50676 50677 50678 50733 50738 | 50648~50675~50676~50677~50678~50733~50738 | 2793 |  | 2793 2793 2793 2793 2793 2793 2793 |  |  |  |  | 1.26 mg |  | 1.26 | [5 empty] | 0.99665 | 0.00306 | 50648 |  | XCAMS | Both | 0.26% |  | OxI in pos#0 (TP50643) came out 5% below the other OxIs and was therefore rejected as normalising cathode. We have a 0.26% system error and a global regression uncertainty multiplier of x1.05. In this wheel we measured 6 machine blanks from our own stock. The stdev in these 6 values is 24%, primarily due to cat#7. Cat#17 has been chosen to do the machine-blank 14C countrate subtraction with, because its countrate was the closest to the average count rate of all 6 MB's. Evaluation of TP 50743 and 50744 results: decided that they could be switched from background test to real blanks. Will average 3 blanks together. | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2012-10-19 | Carbonate |  | JOB-191621 |  | lrg std load >15 <25mg marble for evolution Testing backgrounds - can be removed from wheel if necessary |  | Normal | 14047/1 |  |  | 2012-08-24 | 704 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0038 | 0.00308 | 50675 | [26 empty] | 2012-08-27 | 203 | Carbonate CO2 Evolution | Completed
[168 empty] | 1 | 0.00308 | 50676 | [26 empty] | 2012-08-30 | 646 | Graphite LG1 | Completed
[168 empty] | 1.0001 | 0.00307 | 50677 | [26 empty] | 2012-09-07 | 514 | C13 Measurement | Completed
[168 empty] | 1.0024 | 0.00308 | 50678 | [26 empty] | 2012-09-24 | 1403 | AMS Submission Results | Completed
[168 empty] | 1.0023 | 0.00336 | 50733 | [26 empty] | 2012-10-26 | 2283 | Close Job | Completed
[168 empty] | 0.99574 | 0.00306 | 50738
[7 empty] | 150.1 | 89.0175 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 46062 | 46062 | 46071 | 284 | 284 | 274 |  |  | jenny dahl | -996.77 | 0.11 ‰ | 2012-10-11 | 2012-10-11 | 17 | [5 empty] | 3.83‰ | 3.83‰ | 2.00‰ | 2.00‰ | 3.83 ‰ | 2.00 ‰ | AMS | 1.98‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0.71 | 0.70711 | [6 empty] | -996.73 | 0.12 ‰ |  | Christine Prior | 2012-10-17 | Christine Prior | 14:02 | 0.0032338 | 0.00323 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.2 |  |  | 0.00323 | 0.00011 | 1745 |  |  |  | 191951 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 5 | ... | 0.00311 | 0.00011 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.00311 |  | 0.00011 |  | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00323 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00311 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-10-10 | jenny dahl | 11:08 |  | Completed |  |  |  |  | 60374 | 2000000 |  |  | 60364 60377 60378 60379 60380 60381 60382 60383 | 60364~60377~60378~60379~60380~60381~60382~60383 | 2801 |  | 2801 2801 2801 2801 2801 2801 2801 2801 |  |  |  |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99953 | 0.0022 | 60364 |  | XCAMS | Both | 0.12% |  | straightforward measurement system error = 0.12% The run-to-run variability for the whole of this wheel required a Poisson uncertainty multiplier of *1.08 (i.e. a correction for 8% over-dispersion). | 2 | RRL | [5 empty] | Background Inorganic |  | 2012-11-06 | Carbonate |  | JOB-191951 |  | RG20 start up known age materials |  | Normal | 14047/1 |  |  | 2012-09-14 | 788 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00222 | 60377 | [26 empty] | 2012-09-21 | 228 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0028 | 0.00228 | 60378 | [26 empty] | 2012-09-24 | 384 | Graphite | Completed
[168 empty] | 1.0017 | 0.00228 | 60379 | [26 empty] | 2012-10-19 | 599 | C13 Measurement | Completed
[168 empty] | 0.99751 | 0.00226 | 60380 | [26 empty] | 2012-10-17 | 1745 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00227 | 60381 | [26 empty] | 2012-10-26 | 2234 | Close Job | Completed
[168 empty] | 0.99802 | 0.00229 | 60382
[168 empty] | 0.99729 | 0.00226 | 60383
[7 empty] | 146.9 | 89.1138 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 54558 | 54558 | 54580 | 521 | 520 | 502 |  |  | jenny dahl | -998.88 | 0.07 ‰ | 2012-10-11 | 2012-10-11 | 6 | [5 empty] | 2.74‰ | 2.74‰ | 2.00‰ | 2.00‰ | 2.74 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0.71 | 0.70711 | [6 empty] | -998.86 | 0.07 ‰ |  | Christine Prior | 2012-10-17 | Christine Prior | 14:02 | 0.001123 | 0.00112 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.5 |  |  | 0.00112 | 7e-05 | 1757 |  |  |  | 192279 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 22 | 5 | 5 | ... | 0.00108 | 7e-05 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.00108 |  | 7e-05 |  | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00112 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00108 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-10-10 | jenny dahl | 11:09 |  | Completed |  |  |  |  | 60461 | 2000000 |  |  | 60364 60377 60378 60379 60380 60381 60382 60383 | 60364~60377~60378~60379~60380~60381~60382~60383 | 2801 |  | 2801 2801 2801 2801 2801 2801 2801 2801 |  |  |  |  | 0.80 mg |  | 0.8 | [5 empty] | 0.99953 | 0.0022 | 60364 |  | XCAMS | Both | 0.12% |  | straightforward measurement system error = 0.12% The run-to-run variability for the whole of this wheel required a Poisson uncertainty multiplier of *1.08 (i.e. a correction for 8% over-dispersion). | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-28 | Carbonate |  | JOB-192279 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard |  | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00222 | 60377 | [26 empty] | 2012-10-03 | 241 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0028 | 0.00228 | 60378 | [26 empty] | 2012-10-05 | 471 | Graphite | Completed
[168 empty] | 1.0017 | 0.00228 | 60379 | [26 empty] | 2012-10-11 | 641 | C13 Measurement | Completed
[168 empty] | 0.99751 | 0.00226 | 60380 | [26 empty] | 2012-10-17 | 1757 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00227 | 60381 | [26 empty] | 2012-10-26 | 2246 | Close Job | Completed
[168 empty] | 0.99802 | 0.00229 | 60382
[168 empty] | 0.99729 | 0.00226 | 60383
[7 empty] | 115 | 89.1202 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52648 | 52647 | 52673 | 410 | 410 | 396 |  |  | Johannes Kaiser | -998.58 | 0.07 ‰ | 2012-10-29 | 2012-10-29 | 24 | [5 empty] | 1.86‰ | 1.86‰ | 2.00‰ | 2.00‰ | 1.86 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0014 | 0.82 | 0.8165 | [6 empty] | -998.56 | 0.07 ‰ |  | Christine Prior | 2012-10-30 | Christine Prior | 14:30 | 0.0014245 | 0.00142 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.5 |  |  | 0.00142 | 7e-05 | 1983 |  |  |  | 192280 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 5 | ... | 0.00137 | 7e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.00137 |  | 7e-05 |  | No | 1 | 1.0001 | 0.00111 | 0.00111 | 0.00142 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00137 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-10-28 | Johannes Kaiser | 11:48 |  | Completed |  |  |  |  | 60462 | 2000000 |  |  | 60575 60576 60577 60587 60588 60589 | 60575~60576~60577~60587~60588~60589 | 2808 |  | 2808 2808 2808 2808 2808 2808 |  |  |  |  | 0.86 mg |  | 0.86 | [5 empty] | 1.0046 | 0.00278 | 60575 |  | XCAMS | Both | 0.21% |  | straightf/w meas't; sys-err = 0.21% During meas't there was an hour where the Pelletron was unstable. All those runs have been rejected. Debated averaging two marbles (TP60462 and TP60608) as blank, but the secondary standards made more sense if only TP60462 was used. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-28 | Carbonate |  | JOB-192280 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51397 | Normal | 14047/1 |  |  | 2012-09-14 | 788 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00273 | 60576 | [26 empty] | 2012-10-03 | 243 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99822 | 0.00271 | 60577 | [26 empty] | 2012-10-05 | 472 | Graphite | Completed
[168 empty] | 0.99686 | 0.00268 | 60587 | [26 empty] | 2012-10-11 | 642 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00271 | 60588 | [26 empty] | 2012-10-30 | 1983 | AMS Submission Results | Completed
[168 empty] | 0.99899 | 0.00273 | 60589 | [26 empty] | 2012-10-30 | 2402 | Close Job | Completed
[7 empty] | 97.5 | 89.2505 | 0.0005 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 46906 | 46905 | 46910 | 344 | 344 | 331 |  |  | Johannes Kaiser | -997.09 | 0.12 ‰ | 2012-10-29 | 2012-10-29 | 4 | [5 empty] | 0.40‰ | 0.40‰ | 2.00‰ | 2.00‰ | 0.40 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0014 | 0.82 | 0.8165 | [6 empty] | -997.05 | 0.13 ‰ |  | Christine Prior | 2012-10-30 | Christine Prior | 14:30 | 0.0029114 | 0.00291 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 10.2 |  |  | 0.00291 | 0.00012 | 1984 |  |  |  | 192913 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 5 | ... | 0.0028 | 0.00012 |  |  |  |  | 1.0033 | 1.0033 | Yes |  |  | Small marble blank intended to bracket Foram sample job 191560 (0.24mg) |  |  |  | 0.0028 |  | 0.00012 |  | No | 1 | 1.0001 | 0.00111 | 0.00111 | 0.00291 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.0028 | 0 | 0.00012 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-10-28 | Johannes Kaiser | 11:49 |  | Completed |  |  |  |  | 60609 | 2000000 |  |  | 60575 60576 60577 60587 60588 60589 | 60575~60576~60577~60587~60588~60589 | 2808 |  | 2808 2808 2808 2808 2808 2808 |  |  |  |  | 0.22 mg |  | 0.22 | [5 empty] | 1.0046 | 0.00278 | 60575 |  | XCAMS | Both | 0.21% |  | straightf/w meas't; sys-err = 0.21% During meas't there was an hour where the Pelletron was unstable. All those runs have been rejected. Debated averaging two marbles (TP60462 and TP60608) as blank, but the secondary standards made more sense if only TP60462 was used. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-21 | Carbonate |  | JOB-192913 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51398 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00273 | 60576 | [26 empty] | 2012-10-24 | 298 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99822 | 0.00271 | 60577 | [26 empty] | 2012-10-25 | 619 | Graphite | Completed
[168 empty] | 0.99686 | 0.00268 | 60587 | [26 empty] | 2012-10-26 | 698 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00271 | 60588 | [26 empty] | 2012-10-30 | 1984 | AMS Submission Results | Completed
[168 empty] | 0.99899 | 0.00273 | 60589 | [26 empty] | 2012-10-30 | 2411 | Close Job | Completed
[7 empty] | 97.5 | 89.3068 | 0.0004 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 48144 | 48143 | 48129 | 402 | 401 | 386 |  |  | Johannes Kaiser | -997.5 | 0.12 ‰ | 2012-10-29 | 2012-10-29 | 4 | [5 empty] | -0.23‰ | -0.23‰ | 2.00‰ | 2.00‰ | -0.23 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0014 | 0.82 | 0.8165 | [6 empty] | -997.47 | 0.13 ‰ |  | Christine Prior | 2012-10-30 | Christine Prior | 14:30 | 0.0024955 | 0.0025 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 7.3 |  |  | 0.0025 | 0.00012 | 1985 |  |  |  | 192902 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 24 | 5 | 5 | ... | 0.0024 | 0.00012 |  |  |  |  | 1.0033 | 1.0033 | Yes |  |  | Small marble blank intended to bracket Foram sample job 191560 (0.24mg) |  |  |  | 0.0024 |  | 0.00012 |  | No | 1 | 1.0001 | 0.00111 | 0.00111 | 0.0025 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.0024 | 0 | 0.00012 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-10-28 | Johannes Kaiser | 11:49 |  | Completed |  |  |  |  | 60608 | 2000000 |  |  | 60575 60576 60577 60587 60588 60589 | 60575~60576~60577~60587~60588~60589 | 2808 |  | 2808 2808 2808 2808 2808 2808 |  |  |  |  | 0.27 mg |  | 0.27 | [5 empty] | 1.0046 | 0.00278 | 60575 |  | XCAMS | Both | 0.21% |  | straightf/w meas't; sys-err = 0.21% During meas't there was an hour where the Pelletron was unstable. All those runs have been rejected. Debated averaging two marbles (TP60462 and TP60608) as blank, but the secondary standards made more sense if only TP60462 was used. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-20 | Carbonate |  | JOB-192902 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51399 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00273 | 60576 | [26 empty] | 2012-10-24 | 293 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99822 | 0.00271 | 60577 | [26 empty] | 2012-10-25 | 618 | Graphite | Completed
[168 empty] | 0.99686 | 0.00268 | 60587 | [26 empty] | 2012-10-26 | 697 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00271 | 60588 | [26 empty] | 2012-10-30 | 1985 | AMS Submission Results | Completed
[168 empty] | 0.99899 | 0.00273 | 60589 | [26 empty] | 2012-10-30 | 2410 | Close Job | Completed
[7 empty] | 126 | 88.9553 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52766 | 52765 | 52787 | 357 | 357 | 344 |  |  | jenny dahl | -998.6 | 0.06 ‰ | 2012-11-12 | 2012-11-12 | 38 | [5 empty] | 1.85‰ | 1.85‰ | 2.00‰ | 2.00‰ | 1.85 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.71 | 0.70711 | [6 empty] | -998.58 | 0.06 ‰ |  | Christine Prior | 2012-11-13 | Christine Prior | 11:32 | 0.0014037 | 0.0014 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.4 |  |  | 0.0014 | 6e-05 | 2156 |  |  |  | 191952 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 2 | 5 | 5 | ... | 0.00135 | 6e-05 |  |  |  |  | 1.0015 | 1.0015 | Yes | [6 empty] | 0.00135 |  | 6e-05 |  | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.0014 | 0 | 6e-05 | 0 | 0 |  |  |  |  | 0.00135 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-11-06 | jenny dahl | 15:07 |  | Completed |  |  |  |  | 60460 | 2000000 |  |  | 60717 60718 60719 60720 60737 60738 60739 60741 | 60717~60718~60719~60720~60737~60738~60739~60741 | 2813 |  | 2813 2813 2813 2813 2813 2813 2813 2813 |  |  |  |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99923 | 0.00171 | 60717 |  | XCAMS | ST | 0% |  | good measurement, no system error. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-06 | Carbonate |  | JOB-191952 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51681 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.0017 | 60718 | [26 empty] | 2012-10-03 | 242 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0013 | 0.00172 | 60719 | [26 empty] | 2012-10-05 | 470 | Graphite | Completed
[168 empty] | 1.0016 | 0.0017 | 60720 | [26 empty] | 2012-10-11 | 638 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00164 | 60737 | [26 empty] | 2012-11-13 | 2156 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.0017 | 60738 | [26 empty] | 2012-11-13 | 2647 | Close Job | Completed
[168 empty] | 0.9974 | 0.0017 | 60739
[168 empty] | 0.99899 | 0.00172 | 60741
[7 empty] | 125 | 88.9688 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 53449 | 53448 | 53444 | 389 | 388 | 374 |  |  | jenny dahl | -998.71 | 0.06 ‰ | 2012-11-12 | 2012-11-12 | 28 | [5 empty] | 1.70‰ | 1.70‰ | 2.00‰ | 2.00‰ | 1.70 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.71 | 0.70711 | [6 empty] | -998.69 | 0.06 ‰ |  | Christine Prior | 2012-11-13 | Christine Prior | 11:32 | 0.0012893 | 0.00129 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.8 |  |  | 0.00129 | 6e-05 | 2178 |  |  |  | 192613 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 22 | 5 | 5 | ... | 0.00124 | 6e-05 |  |  |  |  | 1.0015 | 1.0015 | Yes | [6 empty] | 0.00124 |  | 6e-05 |  | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00129 | 0 | 6e-05 | 0 | 0 |  |  |  |  | 0.00124 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-11-06 | jenny dahl | 15:09 |  | Completed |  |  |  |  | 60527 | 2000000 |  |  | 60717 60718 60719 60720 60737 60738 60739 60741 | 60717~60718~60719~60720~60737~60738~60739~60741 | 2813 |  | 2813 2813 2813 2813 2813 2813 2813 2813 |  |  |  |  | 0.76 mg |  | 0.76 | [5 empty] | 0.99923 | 0.00171 | 60717 |  | XCAMS | ST | 0% |  | good measurement, no system error. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2012-12-06 | Carbonate | Complete | JOB-192613 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51683 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.0017 | 60718 | [26 empty] | 2012-10-11 | 257 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0013 | 0.00172 | 60719 | [26 empty] | 2012-10-15 | 537 | Graphite | Completed
[168 empty] | 1.0016 | 0.0017 | 60720 | [26 empty] | 2012-10-19 | 660 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00164 | 60737 | [26 empty] | 2012-11-13 | 2178 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.0017 | 60738 | [26 empty] | 2012-11-13 | 2648 | Close Job | Completed
[168 empty] | 0.9974 | 0.0017 | 60739
[168 empty] | 0.99899 | 0.00172 | 60741
[7 empty] | 107.8 | 89.2739 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 51154 | 51153 | 51134 | 389 | 389 | 374 |  |  | jenny dahl | -998.28 | 0.08 ‰ | 2012-11-12 | 2012-11-12 | 31 | [5 empty] | -1.72‰ | -1.72‰ | 2.00‰ | 2.00‰ | -1.72 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.71 | 0.70711 | [6 empty] | -998.26 | 0.08 ‰ |  | Christine Prior | 2012-11-13 | Christine Prior | 11:32 | 0.0017157 | 0.00172 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 4.7 |  |  | 0.00172 | 8e-05 | 2179 |  |  |  | 192615 | [5 empty] | 5 | 5/07/22 | 15:24:45 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 6 | ... | 0.00165 | 8e-05 |  |  |  |  | 1.0015 | 1.0015 | Yes | [6 empty] | 0.00165 |  | 8e-05 |  | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00172 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00165 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-11-06 | jenny dahl | 15:09 |  | Completed |  |  |  |  | 60519 | 2000000 |  |  | 60717 60718 60719 60720 60737 60738 60739 60741 | 60717~60718~60719~60720~60737~60738~60739~60741 | 2813 |  | 2813 2813 2813 2813 2813 2813 2813 2813 |  |  |  |  | 0.51 mg |  | 0.51 | [5 empty] | 0.99923 | 0.00171 | 60717 |  | XCAMS | ST | 0% |  | good measurement, no system error. | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2012-12-06 | Carbonate |  | JOB-192615 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51682 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.0017 | 60718 | [26 empty] | 2012-10-11 | 259 | Carbonate CO2 Evolution | Cancelled
[168 empty] | 1.0013 | 0.00172 | 60719 | [26 empty] | 2012-10-11 | 264 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0017 | 60720 | [26 empty] | 2012-10-12 | 529 | Graphite | Completed
[168 empty] | 1.0022 | 0.00164 | 60737 | [26 empty] | 2012-10-19 | 655 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.0017 | 60738 | [26 empty] | 2012-11-13 | 2179 | AMS Submission Results | Completed
[168 empty] | 0.9974 | 0.0017 | 60739 | [26 empty] | 2012-11-13 | 2649 | Close Job | Completed
[168 empty] | 0.99899 | 0.00172 | 60741
[7 empty] | 126.3 | 88.9752 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49449 | 49449 | 49454 | 354 | 354 | 341 |  |  | john west | -997.88 | 0.09 ‰ | 2012-12-05 | 2012-12-05 | 51 | [5 empty] | 2.52‰ | 2.52‰ | 2.00‰ | 2.00‰ | 2.52 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0.76 | 0.75593 | [6 empty] | -997.85 | 0.09 ‰ |  | john west | 2012-12-06 | Christine Prior | 15:19 | 0.0021212 | 0.00212 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.5 |  |  | 0.00212 | 9e-05 | 2547 |  |  |  | 192614 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 7 | 5 | 5 | ... | 0.00204 | 9e-05 |  |  |  |  | 1.0024 | 1.0024 | Yes | [6 empty] | 0.00204 |  | 9e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00212 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00204 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-03 | john west | 11:09 |  | Completed |  |  |  |  | 60528 | 2000000 |  |  | 60936 60937 60942 60943 60944 60945 60975 | 60936~60937~60942~60943~60944~60945~60975 | 2821 |  | 2821 2821 2821 2821 2821 2821 2821 |  |  |  |  | 0.78 mg |  | 0.78 | [5 empty] | 1.0004 | 0.00187 | 60936 |  | XCAMS | ST | 0% |  | Straightforward measurement. No error adjustments needed. - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2012-12-06 | Carbonate |  | JOB-192614 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 51988 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00179 | 60937 | [26 empty] | 2012-10-11 | 258 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00179 | 60942 | [26 empty] | 2012-10-15 | 538 | Graphite | Completed
[168 empty] | 1.0009 | 0.00186 | 60943 | [26 empty] | 2012-10-19 | 661 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00175 | 60944 | [26 empty] | 2012-12-06 | 2547 | AMS Submission Results | Completed
[168 empty] | 0.99928 | 0.00178 | 60945 | [26 empty] | 2013-01-22 | 3562 | Close Job | Completed
[168 empty] | 0.99683 | 0.00182 | 60975
[7 empty] | 125.9 | 88.9409 | 0.0003 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 51203 | 51202 | 51180 | 343 | 342 | 329 |  |  | john west | -998.29 | 0.07 ‰ | 2012-12-05 | 2012-12-05 | 29 | [5 empty] | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0.76 | 0.75593 | [6 empty] | -998.27 | 0.07 ‰ |  | john west | 2012-12-06 | Christine Prior | 15:19 | 0.0017053 | 0.00171 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 4.6 |  |  | 0.00171 | 7e-05 | 2561 |  |  |  | 192904 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 5 | ... | 0.00164 | 7e-05 |  |  |  |  | 1.0024 | 1.0024 | Yes | [6 empty] | 0.00164 |  | 7e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00171 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00164 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-03 | john west | 11:10 |  | Completed |  |  |  |  | 60733 | 2000000 |  |  | 60936 60937 60942 60943 60944 60945 60975 | 60936~60937~60942~60943~60944~60945~60975 | 2821 |  | 2821 2821 2821 2821 2821 2821 2821 |  |  |  |  | 0.46 mg |  | 0.46 | [5 empty] | 1.0004 | 0.00187 | 60936 |  | XCAMS | ST | 0% |  | Straightforward measurement. No error adjustments needed. - 29/01/2015: filled System Error (%) box | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2012-11-21 | Carbonate |  | JOB-192904 |  | New subsample of standard material acid etched 25 October 2012. Create daughter jobs from this job using the physical pretreatment details for new jobs 0.8-1.0mg C | 51987 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00179 | 60937 | [26 empty] | 2012-11-05 | 316 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00179 | 60942 | [26 empty] | 2012-11-06 | 743 | Graphite | Completed
[168 empty] | 1.0009 | 0.00186 | 60943 | [26 empty] | 2012-11-07 | 746 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00175 | 60944 | [26 empty] | 2012-12-06 | 2561 | AMS Submission Results | Completed
[168 empty] | 0.99928 | 0.00178 | 60945 | [26 empty] | 2013-02-08 | 3717 | Close Job | Completed
[168 empty] | 0.99683 | 0.00182 | 60975
[7 empty] | 129.2 | 89.0003 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 53384 | 53384 | 53382 | 514 | 514 | 494 |  |  | john west | -998.7 | 0.08 ‰ | 2012-12-05 | 2012-12-05 | 27 | [5 empty] | 2.24‰ | 2.24‰ | 2.00‰ | 2.00‰ | 2.24 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0007 | 0.76 | 0.75593 | [6 empty] | -998.68 | 0.08 ‰ |  | john west | 2012-12-06 | Christine Prior | 15:19 | 0.0012997 | 0.0013 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.4 |  |  | 0.0013 | 8e-05 | 2563 |  |  |  | 193079 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 27 | 5 | 4 | ... | 0.00125 | 8e-05 |  |  |  |  | 1.0024 | 1.0024 | Yes | [6 empty] | 0.00125 |  | 8e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.0013 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00125 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-03 | john west | 11:10 |  | Completed |  |  |  |  | 60763 | 2000000 |  |  | 60936 60937 60942 60943 60944 60945 60975 | 60936~60937~60942~60943~60944~60945~60975 | 2821 |  | 2821 2821 2821 2821 2821 2821 2821 |  |  |  |  | 0.81 mg |  | 0.81 | [5 empty] | 1.0004 | 0.00187 | 60936 |  | XCAMS | ST | 0% |  | Straightforward measurement. No error adjustments needed. - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Inorganic |  | 2012-12-17 | Carbonate |  | JOB-193079 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard 0.8-1.0mg C | 51989 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00179 | 60937 | [26 empty] | 2012-11-05 | 320 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00179 | 60942 | [26 empty] | 2012-11-08 | 773 | Graphite | Completed
[168 empty] | 1.0009 | 0.00186 | 60943 | [26 empty] | 2012-12-06 | 2563 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00175 | 60944 | [26 empty] | 2013-02-08 | 3720 | Close Job | Completed
[168 empty] | 0.99928 | 0.00178 | 60945
[168 empty] | 0.99683 | 0.00182 | 60975
[7 empty] | 110.3 | 89.1177 | 0.0007 | RRL-Test-SM | Background Test |  |  |  |  | 44151 | 44150 | 44156 | 158 | 157 | 157 | 43994.8 | 44309.9 | jenny dahl | -995.9 | 0.08 ‰ | 2012-12-07 | 2012-12-07 | 49 | [5 empty] | 8.03‰ | 8.03‰ | 2.00‰ | 2.00‰ | 8.03 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -17.8 | -17.8149 | 0.53 | 0.53452 | [6 empty] | -995.84 | 0.08 ‰ |  | jenny dahl | 2012-12-17 | Christine Prior | 11:37 | 0.0041024 | 0.0041 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 5.1 |  |  | 0.0041 | 8e-05 | 2616 |  |  |  | 192616 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 2 | 5 | 4 | A.. | 0.00306 | 6e-05 |  |  |  |  | 1.0095 | 1.0095 | Yes |  |  | 49 days old |  |  |  | 0.00306 |  | 6e-05 |  | Yes | 1 | 1 | 0.00036 | 0.00036 | 0.0041 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00306 | 0 | 6e-05 | 0 | Regular |  | 1 | 0.7459 | 2 | 2012-12-05 | jenny dahl | 9:09 |  | Completed |  |  |  |  | 60573 | 2000000 |  |  | 2001017 2001018 2001019 2001020 2001021 2001022 2001023 2001024 2001025 2001026 2001027 2001028 2001047 2001048 | 2001017~2001018~2001019~2001020~2001021~2001022~2001023~2001024~2001025~2001026~2001027~2001028~2001047~2001048 | 2823 |  | 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 |  |  |  |  | 0.38 mg |  | 0.38 | [5 empty] | 0.99971 | 0.00148 | 2001017 |  | XCAMS | ST | 0% |  | Both low and high pressure pressings of OXII standards analysed as one set as no significant differences between the two pressings. The tune cathodes were included in the analysis. No error adjustments were required. - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Test |  | 2012-12-06 | Carbonate |  | JOB-192616 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 52066 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.00135 | 2001018 | [26 empty] | 2012-10-17 | 281 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99834 | 0.00146 | 2001019 | [26 empty] | 2012-10-19 | 583 | Graphite | Completed
[168 empty] | 1.0019 | 0.00132 | 2001020 | [26 empty] | 2012-12-17 | 2616 | AMS Submission Results | Completed
[168 empty] | 0.99963 | 0.00121 | 2001021 | [26 empty] | 2012-12-17 | 3181 | Close Job | Completed
[168 empty] | 0.99886 | 0.00141 | 2001022
[168 empty] | 1.0017 | 0.00126 | 2001023
[168 empty] | 1.0016 | 0.00138 | 2001024
[168 empty] | 1.0006 | 0.00128 | 2001025
[168 empty] | 0.99936 | 0.00147 | 2001026
[168 empty] | 1.0008 | 0.00132 | 2001027
[168 empty] | 1.0006 | 0.00149 | 2001028
[168 empty] | 0.99863 | 0.0013 | 2001047
[168 empty] | 0.99939 | 0.00136 | 2001048
[7 empty] | 94.6 | 89.2326 | 0.0004 | RRL-Test-SM | Background Test |  |  |  |  | 49016 | 49015 | 49012 | 241 | 240 | 251 | 48778.5 | 49259.7 | jenny dahl | -997.76 | 0.07 ‰ | 2012-12-07 | 2012-12-07 | 49 | [5 empty] | 6.74‰ | 6.74‰ | 2.00‰ | 2.00‰ | 6.74 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -17.8 | -17.8149 | 0.53 | 0.53452 | [6 empty] | -997.73 | 0.07 ‰ |  | jenny dahl | 2012-12-17 | Christine Prior | 11:37 | 0.0022389 | 0.00224 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 7.5 |  |  | 0.00224 | 7e-05 | 2617 |  |  |  | 192805 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 9 | 5 | 4 | ... | 0.00167 | 5e-05 |  |  |  |  | 1.0095 | 1.0095 | Yes | [6 empty] | 0.00167 |  | 5e-05 |  | Yes | 1 | 1 | 0.00036 | 0.00036 | 0.00224 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00167 | 0 | 5e-05 | 0 | Regular |  | 1 | 0.7459 | 2 | 2012-12-05 | jenny dahl | 9:09 |  | Completed |  |  |  |  | 60574 | 2000000 |  |  | 2001017 2001018 2001019 2001020 2001021 2001022 2001023 2001024 2001025 2001026 2001027 2001028 2001047 2001048 | 2001017~2001018~2001019~2001020~2001021~2001022~2001023~2001024~2001025~2001026~2001027~2001028~2001047~2001048 | 2823 |  | 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 |  |  |  |  | 0.30 mg |  | 0.3 | [5 empty] | 0.99971 | 0.00148 | 2001017 |  | XCAMS | ST | 0% |  | Both low and high pressure pressings of OXII standards analysed as one set as no significant differences between the two pressings. The tune cathodes were included in the analysis. No error adjustments were required. - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Test |  | 2012-12-05 | Carbonate |  | JOB-192805 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 52065 | Normal | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.00135 | 2001018 | [26 empty] | 2012-10-17 | 282 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99834 | 0.00146 | 2001019 | [26 empty] | 2012-10-19 | 584 | Graphite | Completed
[168 empty] | 1.0019 | 0.00132 | 2001020 | [26 empty] | 2012-12-17 | 2617 | AMS Submission Results | Completed
[168 empty] | 0.99963 | 0.00121 | 2001021 | [26 empty] | 2012-12-17 | 3182 | Close Job | Completed
[168 empty] | 0.99886 | 0.00141 | 2001022
[168 empty] | 1.0017 | 0.00126 | 2001023
[168 empty] | 1.0016 | 0.00138 | 2001024
[168 empty] | 1.0006 | 0.00128 | 2001025
[168 empty] | 0.99936 | 0.00147 | 2001026
[168 empty] | 1.0008 | 0.00132 | 2001027
[168 empty] | 1.0006 | 0.00149 | 2001028
[168 empty] | 0.99863 | 0.0013 | 2001047
[168 empty] | 0.99939 | 0.00136 | 2001048
[7 empty] | 107.1 | 89.2848 | 0.0006 | RRL-Test-SM | Background Test |  |  |  |  | 45337 | 45336 | 45335 | 183 | 182 | 182 | 45156.2 | 45521.4 | jenny dahl | -996.46 | 0.08 ‰ | 2012-12-07 | 2012-12-07 | 29 | [5 empty] | 6.15‰ | 6.15‰ | 2.00‰ | 2.00‰ | 6.15 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -17.8 | -17.8149 | 0.53 | 0.53452 | [6 empty] | -996.41 | 0.08 ‰ |  | jenny dahl | 2012-12-17 | Christine Prior | 11:37 | 0.0035393 | 0.00354 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 4.6 |  |  | 0.00354 | 8e-05 | 2618 |  |  |  | 193078 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 16 | 5 | 4 | ... | 0.00264 | 6e-05 |  |  |  |  | 1.0095 | 1.0095 | Yes | [6 empty] | 0.00264 |  | 6e-05 |  | Yes | 1 | 1 | 0.00036 | 0.00036 | 0.00354 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00264 | 0 | 6e-05 | 0 | Regular |  | 1 | 0.7459 | 2 | 2012-12-05 | jenny dahl | 9:09 |  | Completed |  |  |  |  | 60762 | 2000000 |  |  | 2001017 2001018 2001019 2001020 2001021 2001022 2001023 2001024 2001025 2001026 2001027 2001028 2001047 2001048 | 2001017~2001018~2001019~2001020~2001021~2001022~2001023~2001024~2001025~2001026~2001027~2001028~2001047~2001048 | 2823 |  | 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 2823 |  |  |  |  | 0.48 mg |  | 0.48 | [5 empty] | 0.99971 | 0.00148 | 2001017 |  | XCAMS | ST | 0% |  | Both low and high pressure pressings of OXII standards analysed as one set as no significant differences between the two pressings. The tune cathodes were included in the analysis. No error adjustments were required. - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Test |  | 2012-12-17 | Carbonate |  | JOB-193078 |  | parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard 0.8-1.0mg C | 52067 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9985 | 0.00135 | 2001018 | [26 empty] | 2012-11-05 | 319 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99834 | 0.00146 | 2001019 | [26 empty] | 2012-11-08 | 772 | Graphite | Completed
[168 empty] | 1.0019 | 0.00132 | 2001020 | [26 empty] | 2012-12-17 | 2618 | AMS Submission Results | Completed
[168 empty] | 0.99963 | 0.00121 | 2001021 | [26 empty] | 2012-12-17 | 3183 | Close Job | Completed
[168 empty] | 0.99886 | 0.00141 | 2001022
[168 empty] | 1.0017 | 0.00126 | 2001023
[168 empty] | 1.0016 | 0.00138 | 2001024
[168 empty] | 1.0006 | 0.00128 | 2001025
[168 empty] | 0.99936 | 0.00147 | 2001026
[168 empty] | 1.0008 | 0.00132 | 2001027
[168 empty] | 1.0006 | 0.00149 | 2001028
[168 empty] | 0.99863 | 0.0013 | 2001047
[168 empty] | 0.99939 | 0.00136 | 2001048
[7 empty] | 107.1 | 89.3627 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52474 | 52473 | 52450 | 402 | 401 | 385 |  |  | jenny dahl | -998.54 | 0.07 ‰ | 2012-12-12 | 2012-12-12 | 5 | [5 empty] | 0.20‰ | 0.20‰ | 2.00‰ | 2.00‰ | 0.20 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0.76 | 0.75593 | [6 empty] | -998.53 | 0.07 ‰ |  |  | 2012-12-17 | Christine Prior | 13:24 | 0.0014557 | 0.00146 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3 |  |  | 0.00146 | 7e-05 | 2797 |  |  |  | 193515 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 33 | 5 | 5 | ... | 0.0014 | 7e-05 |  |  |  |  | 1.0044 | 1.0044 | Yes | [6 empty] | 0.0014 |  | 7e-05 |  | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00146 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0014 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-10 | jenny dahl | 12:30 |  | Completed |  |  |  |  | 61063 | 2000000 |  |  | 61000 61001 61033 61037 61038 61039 61040 | 61000~61001~61033~61037~61038~61039~61040 | 2825 |  | 2825 2825 2825 2825 2825 2825 2825 |  |  |  |  | 0.70 mg |  | 0.7 | [5 empty] | 0.99872 | 0.00173 | 61000 |  | XCAMS | ST | 0% |  | The results were analysed using 2 regions. An error adjustment of 1.06 was necessary to account for global run-to-run variance. - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2012-12-18 | Carbonate |  | JOB-193515 |  |  | 52142 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00169 | 61001 | [26 empty] | 2012-12-04 | 350 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99739 | 0.00177 | 61033 | [26 empty] | 2012-12-07 | 1073 | Graphite | Completed
[168 empty] | 1.0017 | 0.00168 | 61037 | [26 empty] | 2012-12-10 | 966 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00172 | 61038 | [26 empty] | 2012-12-17 | 2797 | AMS Submission Results | Completed
[168 empty] | 0.99822 | 0.00169 | 61039 | [26 empty] | 2012-12-17 | 3202 | Close Job | Completed
[168 empty] | 1.0025 | 0.00169 | 61040
[7 empty] | 97.7 | 89.2898 | 0.0007 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 44754 | 44754 | 44745 | 329 | 329 | 337 |  |  | Kelly Lyons | -996.19 | 0.16 ‰ | 2012-12-17 | 2012-12-17 | 52 | [5 empty] | 0.55‰ | 0.55‰ | 2.00‰ | 2.00‰ | 0.55 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0.71 | 0.70711 | [6 empty] | -996.15 | 0.16 ‰ |  |  | 2012-12-20 | Christine Prior | 11:13 | 0.0038056 | 0.00381 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 8.3 |  |  | 0.00381 | 0.00016 | 2809 |  |  |  | 192903 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 33 | 5 | 5 | ... | 0.00366 | 0.00015 |  |  |  |  | 1.0039 | 1.0039 | Yes |  |  | An additional error factor of 1.4 was required. This is intended as blank to match unknown TP 61066 for DCC. |  |  |  | 0.00366 |  | 0.00015 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00381 | 0 | 0.00016 | 0 | 0 |  |  |  |  | 0.00366 | 0 | 0.00015 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2012-12-14 | Kelly Lyons | 11:27 |  | Completed |  |  |  |  | 60633 | 2000000 |  |  | 61050 61052 61053 61054 61055 61096 61098 61099 | 61050~61052~61053~61054~61055~61096~61098~61099 | 2828 |  | 2828 2828 2828 2828 2828 2828 2828 2828 |  |  |  |  | 0.26 mg |  | 0.26 | [5 empty] | 0.99634 | 0.00207 | 61050 |  | XCAMS | ST | 0.14% |  | A system error of 0.14% was needed to account for calibration pos-to-pos variance. One cathode (33) required an additional error factor. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2012-11-20 | Carbonate |  | JOB-192903 |  | Results comments and job notes appear different to each other. Older graphite, put in as comparison - not to use as std,parented from J 190973 (same acid etched material) all marble standards since that j have same physical details until we etch a new batch of marble standard | 52247 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00203 | 61052 | [26 empty] | 2012-10-24 | 294 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0006 | 0.00204 | 61053 | [26 empty] | 2012-10-26 | 643 | Graphite | Completed
[168 empty] | 0.99863 | 0.00207 | 61054 | [26 empty] | 2012-11-02 | 706 | C13 Measurement | Completed
[168 empty] | 1.0002 | 0.00205 | 61055 | [26 empty] | 2012-12-20 | 2809 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00205 | 61096 | [26 empty] | 2013-01-22 | 3564 | Close Job | Completed
[168 empty] | 0.99864 | 0.00204 | 61098
[168 empty] | 1.003 | 0.0021 | 61099
[7 empty] | 137.6 | 89.0076 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47784 | 47783 | 47784 | 224 | 224 | 215 |  |  | Kelly Lyons | -997.39 | 0.07 ‰ | 2012-12-17 | 2012-12-17 | 10 | [5 empty] | 3.73‰ | 3.73‰ | 2.00‰ | 2.00‰ | 3.73 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0.71 | 0.70711 | [6 empty] | -997.36 | 0.07 ‰ |  |  | 2012-12-20 | Christine Prior | 11:13 | 0.0026099 | 0.00261 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.6 |  |  | 0.00261 | 7e-05 | 2828 |  |  |  | 193516 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 5 | ... | 0.00251 | 7e-05 |  |  |  |  | 1.0039 | 1.0039 | Yes |  |  | Not nearly as good as usual carbonate blank, but probably a reasonable representation of what is going on with the unknowns on the carbonate line this week given the results for blanks TP61114 and TP60633 |  |  |  | 0.00251 |  | 7e-05 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00261 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00251 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-14 | Kelly Lyons | 11:32 |  | Completed |  |  |  |  | 61064 | 2000000 |  |  | 61050 61052 61053 61054 61055 61096 61098 61099 | 61050~61052~61053~61054~61055~61096~61098~61099 | 2828 |  | 2828 2828 2828 2828 2828 2828 2828 2828 |  |  |  |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99634 | 0.00207 | 61050 |  | XCAMS | ST | 0.14% |  | A system error of 0.14% was needed to account for calibration pos-to-pos variance. One cathode (33) required an additional error factor. | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2012-12-18 | Carbonate |  | JOB-193516 |  |  | 52246 |  | 14047/1 |  |  | 2012-09-14 | 789 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00203 | 61052 | [26 empty] | 2012-12-04 | 351 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0006 | 0.00204 | 61053 | [26 empty] | 2012-12-07 | 1074 | Graphite | Completed
[168 empty] | 0.99863 | 0.00207 | 61054 | [26 empty] | 2012-12-10 | 967 | C13 Measurement | Completed
[168 empty] | 1.0002 | 0.00205 | 61055 | [26 empty] | 2012-12-20 | 2828 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00205 | 61096 | [26 empty] | 2013-02-08 | 3784 | Close Job | Completed
[168 empty] | 0.99864 | 0.00204 | 61098
[168 empty] | 1.003 | 0.0021 | 61099
[7 empty] | 132.8 | 89.1012 | 0.0013 | RRL-Test-LG | Background Test |  |  |  |  | 38832 | 38831 | 38836 | 137 | 136 | 141 | 38696.4 | 38969.4 | Kelly Lyons | -992.05 | 0.14 ‰ | 2012-12-17 | 2012-12-17 | 5 | [5 empty] | 2.67‰ | 2.67‰ | 2.00‰ | 2.00‰ | 2.67 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0013 | 0.71 | 0.70711 | [6 empty] | -991.95 | 0.14 ‰ |  |  | 2012-12-20 | Christine Prior | 11:13 | 0.0079544 | 0.00795 | 0.00014 | 0.00014 | 0.00014 | 0.00013 | 1.7 |  |  | 0.00795 | 0.00014 | 2837 |  |  |  | 193637 | [5 empty] | 5 | 5/07/22 | 15:24:46 | Albert Zondervan | 0 |  |  |  | 22 | 5 | 3 | ... | 0.00765 | 0.00013 |  |  |  |  | 1.0039 | 1.0039 | Yes |  |  | Marble blanks are normally MUCH better than this; something must have gone wrong. Switch category from Inorganic Blank to background test and not use as blank for this wheel. |  |  |  | 0.00765 |  | 0.00013 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00795 | 0 | 0.00014 | 0 | 0 |  |  |  |  | 0.00765 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2012-12-14 | Kelly Lyons | 11:33 |  | Completed |  |  |  |  | 61114 | 2000000 |  |  | 61050 61052 61053 61054 61055 61096 61098 61099 | 61050~61052~61053~61054~61055~61096~61098~61099 | 2828 |  | 2828 2828 2828 2828 2828 2828 2828 2828 |  |  |  |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99634 | 0.00207 | 61050 |  | XCAMS | ST | 0.14% |  | A system error of 0.14% was needed to account for calibration pos-to-pos variance. One cathode (33) required an additional error factor. |  | RRL | [5 empty] | Background Test |  | 2012-12-25 | Carbonate |  | JOB-193637 |  |  | 52249 |  | 14047/1 |  |  | 2012-12-11 | 372 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00203 | 61052 | [26 empty] | 2012-12-12 | 1124 | Graphite | Completed
[168 empty] | 1.0006 | 0.00204 | 61053 | [26 empty] | 2012-12-20 | 2837 | AMS Submission Results | Completed
[168 empty] | 0.99863 | 0.00207 | 61054 | [26 empty] | 2013-02-08 | 3808 | Close Job | Completed
[168 empty] | 1.0002 | 0.00205 | 61055
[168 empty] | 1.0019 | 0.00205 | 61096
[168 empty] | 0.99864 | 0.00204 | 61098
[168 empty] | 1.003 | 0.0021 | 61099
[7 empty] | 148.9 | 88.4255 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48627 | 48626 | 48626 | 355 | 355 | 342 |  |  | Helen Zhang | -997.65 | 0.10 ‰ | 2013-01-23 | 2013-01-23 | 9 | [5 empty] | 4.56‰ | 4.56‰ | 2.00‰ | 2.00‰ | 4.56 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0.82 | 0.8165 | [6 empty] | -997.62 | 0.11 ‰ |  |  | 2013-01-24 | Christine Prior | 13:18 | 0.0023499 | 0.00235 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00235 | 0.0001 | 3026 |  |  |  | 193890 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 3 | 5 | 4 | ... | 0.00226 | 0.0001 |  |  |  |  | 0.99819 | 0.99819 | Yes |  |  | This cathode needed an additional factor of 1.3 to explain its individual run-to-run variance. |  |  |  | 0.00226 |  | 0.0001 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00235 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00226 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-01-18 | Jenny Dahl | 13:38 |  | Completed |  |  |  |  | 61141 | 2000000 |  |  | 61128 61129 61130 61132 61134 61135 | 61128~61129~61130~61132~61134~61135 | 2833 |  | 2833 2833 2833 2833 2833 2833 |  |  |  |  | 1.14 mg |  | 1.14 | [5 empty] | 1.0003 | 0.00163 | 61128 |  | XCAMS | ST | 0% |  | After initial Cathode potential problems (which resolved over time), we accumulated a good data set with a precision of 1.7 per mille. The calibration chisquare reduced = 0.38 and results in a probability of 88%. We needed to apply an additional factor of 1.04 to explain the global run-to-run variance. JCT - Blanks are all poor in this wheel. Everything was graphitized using the replacement RG5 pump - maybe this is the cause of poor blanks? - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-01-24 | Carbonate |  | JOB-193890 |  |  | 52437 |  | 14047/1 |  |  | 2013-01-10 | 380 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00167 | 61129 | [26 empty] | 2013-01-14 | 1151 | Graphite | Completed
[168 empty] | 0.99869 | 0.00165 | 61130 | [26 empty] | 2013-01-15 | 1029 | C13 Measurement | Completed
[168 empty] | 0.99902 | 0.00167 | 61132 | [26 empty] | 2013-01-24 | 3026 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00163 | 61134 | [26 empty] | 2013-01-24 | 2204 | Review Final Report | Completed
[168 empty] | 1.0013 | 0.00159 | 61135 | [26 empty] | 2013-03-19 | 4240 | Close Job | Completed
[7 empty] | 125 | 89.291 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49529 | 49528 | 49530 | 278 | 278 | 268 |  |  | jenny dahl | -997.9 | 0.07 ‰ | 2013-02-01 | 2013-02-01 | 14 | [5 empty] | 3.11‰ | 3.11‰ | 2.00‰ | 2.00‰ | 3.11 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0.63 | 0.63246 | [6 empty] | -997.87 | 0.07 ‰ |  |  | 2013-02-04 | christine prior | 11:16 | 0.0021004 | 0.0021 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.3 |  |  | 0.0021 | 7e-05 | 3200 |  |  |  | 194254 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 14 | 5 | 5 | ... | 0.00202 | 7e-05 |  |  |  |  | 1.0065 | 1.0065 | Yes | [6 empty] | 0.00202 |  | 7e-05 |  | Yes | 1 | 1 | 0.00053 | 0.00053 | 0.0021 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00202 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-01-24 | jenny dahl | 14:41 |  | Completed |  |  |  |  | 61206 | 2000000 |  |  | 61131 61133 61136 61137 61143 61144 61145 61146 61147 61158 | 61131~61133~61136~61137~61143~61144~61145~61146~61147~61158 | 2835 |  | 2835 2835 2835 2835 2835 2835 2835 2835 2835 2835 |  |  |  |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99826 | 0.00162 | 61131 |  | XCAMS | ST | 0% |  | * First attempt to measure this wheel resulted in 0.3% system error. After a complete retune, repeatability was consistent with counting statistics (400k 14C counts per Ox-I cathode). * Sputtering through the graphite seems to have affected run-to-run repeatabilities only marginally. * The results are a confidence booster for the upcoming WAI and HP wheels. JCT 30/1/15 Changed system error to 0%. System error was given as 0.3%, but clearly the second set of results had been used, as reported errors on OxIs were ~0.16% and therefore could not have included a 0.3% system error. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-02-28 | Carbonate |  | JOB-194254 |  | parented from J 192904 acid etched and ground material prepared 25/10/12 | 52477 |  | 14047/1 |  |  | 2013-01-17 | 1119 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99803 | 0.00161 | 61133 | [26 empty] | 2013-01-17 | 396 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00164 | 61136 | [26 empty] | 2013-01-18 | 1216 | Graphite | Completed
[168 empty] | 0.9998 | 0.00167 | 61137 | [26 empty] | 2013-01-23 | 1084 | C13 Measurement | Completed
[168 empty] | 0.99905 | 0.00159 | 61143 | [26 empty] | 2013-02-04 | 3200 | AMS Submission Results | Completed
[168 empty] | 1.0021 | 0.00162 | 61144 | [26 empty] | 2013-02-08 | 3826 | Close Job | Completed
[168 empty] | 1.0024 | 0.00162 | 61145
[168 empty] | 1.0003 | 0.00164 | 61146
[168 empty] | 0.99837 | 0.00211 | 61147
[168 empty] | 1.0002 | 0.00167 | 61158
[7 empty] | 138.1 | 88.8042 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48044 | 48044 | 48034 | 264 | 264 | 254 |  |  | jenny dahl | -997.47 | 0.08 ‰ | 2013-02-08 | 2013-02-08 | 4 | [5 empty] | 6.87‰ | 6.87‰ | 2.00‰ | 2.00‰ | 6.87 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0021 | 0.82 | 0.8165 | [6 empty] | -997.44 | 0.08 ‰ |  |  | 2013-02-11 | christine prior | 9:33 | 0.0025267 | 0.00253 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00253 | 8e-05 | 3319 |  |  |  | 194521 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 20 | 5 | 3 | ... | 0.00243 | 8e-05 |  |  |  |  | 1.0048 | 1.0048 | Yes | [6 empty] | 0.00243 |  | 8e-05 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00253 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00243 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-02-07 | jenny dahl | 16:02 |  | Completed |  |  |  |  | 61370 | 2000000 |  |  | 61210 61262 61289 61291 61293 61304 | 61210~61262~61289~61291~61293~61304 | 2839 |  | 2839 2839 2839 2839 2839 2839 |  |  |  |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99963 | 0.00155 | 61210 |  | XCAMS | Both | 0% |  | * zero system error * Kapuni cathode showed a sequence of runs with elevated 14C counts. These runs were rejected. This was not seen this with any of the other blanks or low-14C-counting cathodes, and thus think it must be some sort of contamination. However, we have seen such wildly variable count rates on blanks in earlier measurements. * furthermore, first 3 runs of all low-14C-counting cathodes rejected, to avoid any surface contamination * global run-to-run variance slightly above Poisson estimate |  | RRL | [5 empty] | Background Inorganic |  | 2013-02-11 | Carbonate |  | JOB-194521 |  | Taken from previously ground and acid etched material J192904. | 52659 |  | 14047/1 |  |  | 2013-01-29 | 426 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0002 | 0.00157 | 61262 | [26 empty] | 2013-02-05 | 1380 | Graphite | Completed
[168 empty] | 0.99815 | 0.00161 | 61289 | [26 empty] | 2013-02-11 | 3319 | AMS Submission Results | Completed
[168 empty] | 0.99901 | 0.00157 | 61291 | [26 empty] | 2013-02-12 | 3898 | Close Job | Completed
[168 empty] | 1.0004 | 0.00158 | 61293
[168 empty] | 1.0025 | 0.00156 | 61304
[7 empty] | 134.7 | 88.8789 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50021 | 50020 | 50003 | 296 | 295 | 284 |  |  | jenny dahl | -998.02 | 0.07 ‰ | 2013-02-11 | 2013-02-11 | 17 | [5 empty] | 1.16‰ | 1.16‰ | 2.00‰ | 2.00‰ | 1.16 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0007 | 0.82 | 0.8165 | [6 empty] | -998 | 0.07 ‰ |  |  | 2013-02-12 | christine prior | 10:55 | 0.0019756 | 0.00198 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3.2 |  |  | 0.00198 | 7e-05 | 3359 |  |  |  | 194474 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 4 | ... | 0.0019 | 7e-05 |  |  |  |  | 0.99991 | 0.99991 | Yes | [6 empty] | 0.0019 |  | 7e-05 |  | Yes | 1 | 0.99999 | 0.00067 | 0.00067 | 0.00198 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0019 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-02-08 | jenny dahl | 12:33 |  | Completed |  |  |  |  | 61266 | 2000000 |  |  | 61246 61263 61265 61279 61280 61287 | 61246~61263~61265~61279~61280~61287 | 2840 |  | 2840 2840 2840 2840 2840 2840 |  |  |  |  | 0.67 mg |  | 0.67 | [5 empty] | 0.99822 | 0.00156 | 61246 |  | XCAMS | Both | 0% |  | * zero system error * 3 cathodes had poor run-to-run repeatability, the worst of which was cat#13 (TP61252): it required an uncertainty multiplier of 2.4 (twice as large as that for cat#6 and #21); we also note that its beam intensity was more variable than that of the other cathodes |  | RRL | [5 empty] | Background Inorganic |  | 2013-02-07 | Carbonate | Complete | JOB-194474 |  |  | 52699 |  | 14047/1 |  |  | 2013-01-24 | 1155 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99997 | 0.0016 | 61263 | [26 empty] | 2013-01-24 | 412 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99779 | 0.00159 | 61265 | [26 empty] | 2013-01-28 | 1276 | Graphite | Completed
[168 empty] | 1.0016 | 0.00206 | 61279 | [26 empty] | 2013-02-12 | 3359 | AMS Submission Results | Completed
[168 empty] | 1.0017 | 0.00159 | 61280 | [26 empty] | 2013-02-12 | 3896 | Close Job | Completed
[168 empty] | 1.0013 | 0.00156 | 61287
[7 empty] | 129.8 | 88.4818 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49771 | 49770 | 49763 | 533 | 532 | 551 |  |  | christine prior | -997.96 | 0.14 ‰ | 2013-02-20 | 2013-02-20 | 12 | [5 empty] | 4.28‰ | 4.28‰ | 2.00‰ | 2.00‰ | 4.28 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -997.94 | 0.14 ‰ |  |  | 2013-02-21 | christine prior | 10:13 | 0.002038 | 0.00204 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00204 | 0.00014 | 3434 |  |  |  | 194545 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 30 | 5 | 4 | ... | 0.00196 | 0.00013 |  |  |  |  | 0.99855 | 0.99855 | Yes | [6 empty] | 0.00196 |  | 0.00013 |  | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00204 | 0 | 0.00014 | 0 | 0 |  |  |  |  | 0.00196 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-02-14 | christine prior | 15:28 |  | Completed |  |  |  |  | 61434 | 2000000 |  |  | 61447 61448 61449 61450 61451 61452 | 61447~61448~61449~61450~61451~61452 | 2844 |  | 2844 2844 2844 2844 2844 2844 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0025 | 0.00178 | 61447 |  | XCAMS | ST | 0% |  | - no system error - good measurement |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-04 | Carbonate |  | JOB-194545 |  |  | 52873 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99691 | 0.00177 | 61448 | [26 empty] | 2013-02-04 | 458 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99962 | 0.00176 | 61449 | [26 empty] | 2013-02-11 | 1444 | Graphite | Completed
[168 empty] | 0.99982 | 0.00177 | 61450 | [26 empty] | 2013-02-21 | 3434 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00178 | 61451 | [26 empty] | 2013-03-19 | 4261 | Close Job | Completed
[168 empty] | 1.0006 | 0.00178 | 61452
[7 empty] | 140.3 | 88.4747 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49103 | 49102 | 49120 | 302 | 301 | 291 |  |  | jenny dahl | -997.79 | 0.08 ‰ | 2013-02-22 | 2013-02-22 | 14 | [5 empty] | 3.19‰ | 3.19‰ | 2.00‰ | 2.00‰ | 3.19 ‰ | 2.00 ‰ | AMS | 2.76‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0.76 | 0.75593 | [6 empty] | -997.76 | 0.08 ‰ |  |  | 2013-02-25 | Christine Prior | 10:00 | 0.0022148 | 0.00221 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00221 | 8e-05 | 3473 |  |  |  | 194543 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 28 | 5 | 5 | ... | 0.00213 | 8e-05 |  |  |  |  | 0.99738 | 0.99738 | Yes | [6 empty] | 0.00213 |  | 8e-05 |  | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00221 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00213 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-02-21 | jenny dahl | 9:53 |  | Completed |  |  |  |  | 61432 | 2000000 |  |  | 61390 61440 61441 61442 61529 61530 61534 | 61390~61440~61441~61442~61529~61530~61534 | 2845 |  | 2845 2845 2845 2845 2845 2845 2845 |  |  |  |  | 0.95 mg |  | 0.95 | [5 empty] | 0.9989 | 0.00184 | 61390 |  | XCAMS | ST | 0% |  | * no system error * Kapuni cat#8 (TP61393) required uncertainty adjustment * correction for a small glitch during data acquisition might have improved overall consistency; data from the first measurement (first ~5 rotations) were therefore ignored. | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2013-03-04 | Carbonate |  | JOB-194543 |  |  | 52913 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99975 | 0.00183 | 61440 | [26 empty] | 2013-02-04 | 457 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99963 | 0.00184 | 61441 | [26 empty] | 2013-02-11 | 1442 | Graphite | Completed
[168 empty] | 1.0017 | 0.00184 | 61442 | [26 empty] | 2013-02-13 | 1207 | C13 Measurement | Completed
[168 empty] | 1.0029 | 0.00184 | 61529 | [26 empty] | 2013-02-25 | 3473 | AMS Submission Results | Completed
[168 empty] | 0.99866 | 0.00182 | 61530 | [26 empty] | 2013-04-08 | 4506 | Close Job | Completed
[168 empty] | 0.99846 | 0.00183 | 61534
[7 empty] | 98.2 | 89.3432 | 0.0006 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 45489 | 45488 | 45496 | 289 | 288 | 278 |  |  | christine prior | -996.53 | 0.12 ‰ | 2013-03-06 | 2013-03-06 | 21 | [5 empty] | 2.32‰ | 2.32‰ | 2.00‰ | 2.00‰ | 2.32 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0233 | 1.15 | 1.1547 | [6 empty] | -996.48 | 0.13 ‰ |  |  | 2013-03-08 | christine prior | 15:07 | 0.0034729 | 0.00347 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 9.2 |  |  | 0.00347 | 0.00012 | 3514 |  |  |  | 194568 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 5 | ... | 0.00334 | 0.00012 |  |  |  |  | 1.0063 | 1.0063 | Yes | [6 empty] | 0.00334 |  | 0.00012 |  | No | 1 | 0.99991 | 0.00293 | 0.00293 | 0.00347 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.00334 | 0 | 0.00012 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2013-03-01 | christine prior | 14:31 |  | Completed |  |  |  |  | 61482 | 2000000 |  |  | 61498 61499 61500 | 61498~61499~61500 | 2846.1 |  | 2846.1 2846.1 2846.1 |  |  |  |  | 0.23 mg |  | 0.23 | [5 empty] | 1.0056 | 0.00508 | 61498 |  | XCAMS | ST | 0.47% |  | * subset TW2846.1: 61455 <= TP <= 61519 * system error = 0.47% * one cathode had excessive run-to-run variability |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-07 | Carbonate | Complete | JOB-194568 |  | small blank to match speleothem samples. please use 1.7-1.9 mg of marble target size =0.19mg C low yield, need repeat. | 52990 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99736 | 0.00498 | 61499 | [26 empty] | 2013-02-11 | 478 | Carbonate CO2 Evolution | Cancelled
[168 empty] | 0.9967 | 0.00517 | 61500 | [26 empty] | 2013-02-12 | 487 | Carbonate CO2 Evolution | Completed
[197 empty] | 2013-02-13 | 1492 | Graphite | Completed
[197 empty] | 2013-03-08 | 3514 | AMS Submission Results | Completed
[197 empty] | 2013-04-12 | 4581 | Close Job | Completed
[7 empty] | 122.5 | 88.9753 | 0.0006 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 45182 | 45182 | 45178 | 255 | 254 | 245 |  |  | christine prior | -996.39 | 0.11 ‰ | 2013-03-06 | 2013-03-06 | 21 | [5 empty] | 6.46‰ | 6.46‰ | 2.00‰ | 2.00‰ | 6.46 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0233 | 1.15 | 1.1547 | [6 empty] | -996.35 | 0.12 ‰ |  |  | 2013-03-08 | christine prior | 15:07 | 0.0036081 | 0.00361 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 5.8 |  |  | 0.00361 | 0.00011 | 3515 |  |  |  | 194569 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 24 | 5 | 5 | ... | 0.00347 | 0.00011 |  |  |  |  | 1.0063 | 1.0063 | Yes | [6 empty] | 0.00347 |  | 0.00011 |  | No | 1 | 0.99991 | 0.00293 | 0.00293 | 0.00361 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00347 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-03-01 | christine prior | 14:31 |  | Completed |  |  |  |  | 61483 | 2000000 |  |  | 61498 61499 61500 | 61498~61499~61500 | 2846.1 |  | 2846.1 2846.1 2846.1 |  |  |  |  | 0.36 mg |  | 0.36 | [5 empty] | 1.0056 | 0.00508 | 61498 |  | XCAMS | ST | 0.47% |  | * subset TW2846.1: 61455 <= TP <= 61519 * system error = 0.47% * one cathode had excessive run-to-run variability |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-07 | Carbonate |  | JOB-194569 |  | small blank to match speleothem samples. please use 2.6-2.9 mg of marble target size =0.31mg C low yield, need repeat. | 52991 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99736 | 0.00498 | 61499 | [26 empty] | 2013-02-11 | 479 | Carbonate CO2 Evolution | Cancelled
[168 empty] | 0.9967 | 0.00517 | 61500 | [26 empty] | 2013-02-12 | 486 | Carbonate CO2 Evolution | Completed
[197 empty] | 2013-02-13 | 1493 | Graphite | Completed
[197 empty] | 2013-03-08 | 3515 | AMS Submission Results | Completed
[197 empty] | 2013-04-12 | 4582 | Close Job | Completed
[7 empty] | 114.2 | 89.4868 | 0.0007 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 44368 | 44368 | 44374 | 272 | 271 | 282 |  |  | christine prior | -996.01 | 0.14 ‰ | 2013-03-06 | 2013-03-06 | 20 | [5 empty] | 0.71‰ | 0.71‰ | 2.00‰ | 2.00‰ | 0.71 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0233 | 1.15 | 1.1547 | [6 empty] | -995.96 | 0.14 ‰ |  |  | 2013-03-08 | christine prior | 15:07 | 0.0039928 | 0.00399 | 0.00014 | 0.00014 | 0.00014 | 0.00013 | 4.6 |  |  | 0.00399 | 0.00014 | 3516 |  |  |  | 194603 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 4 | ... | 0.00384 | 0.00013 |  |  |  |  | 1.0063 | 1.0063 | Yes | [6 empty] | 0.00384 |  | 0.00013 |  | No | 1 | 0.99991 | 0.00293 | 0.00293 | 0.00399 | 0 | 0.00014 | 0 | 0 |  |  |  |  | 0.00384 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-03-01 | christine prior | 14:31 |  | Completed |  |  |  |  | 61501 | 2000000 |  |  | 61498 61499 61500 | 61498~61499~61500 | 2846.1 |  | 2846.1 2846.1 2846.1 |  |  |  |  | 0.50 mg |  | 0.5 | [5 empty] | 1.0056 | 0.00508 | 61498 |  | XCAMS | ST | 0.47% |  | * subset TW2846.1: 61455 <= TP <= 61519 * system error = 0.47% * one cathode had excessive run-to-run variability |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-12 | Carbonate | Complete | JOB-194603 |  |  | 52992 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99736 | 0.00498 | 61499 | [26 empty] | 2013-02-12 | 488 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9967 | 0.00517 | 61500 | [26 empty] | 2013-02-15 | 1511 | Graphite | Completed
[197 empty] | 2013-03-08 | 3516 | AMS Submission Results | Completed
[197 empty] | 2013-04-12 | 4583 | Close Job | Completed
[7 empty] | 142.6 | 88.8648 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50681 | 50681 | 50680 | 367 | 367 | 353 |  |  | jenny dahl | -998.18 | 0.08 ‰ | 2013-03-18 | 2013-03-18 | 12 | [5 empty] | 4.70‰ | 4.70‰ | 2.00‰ | 2.00‰ | 4.70 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -998.16 | 0.08 ‰ |  |  | 2013-03-19 | christine prior | 10:11 | 0.0018196 | 0.00182 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00182 | 8e-05 | 3758 |  |  |  | 194777 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 6 | ... | 0.00175 | 8e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.00175 |  | 8e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00182 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00175 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-03-13 | jenny dahl | 12:21 |  | Completed |  |  |  |  | 61684 | 2000000 |  |  | 61658 61659 61660 61661 61662 61695 | 61658~61659~61660~61661~61662~61695 | 2850 |  | 2850 2850 2850 2850 2850 2850 |  |  |  |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99941 | 0.00182 | 61658 |  | XCAMS | Both | 0% |  | Best measurement in a while!! no error adjustments needed :). - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-03-18 | Carbonate |  | JOB-194777 |  |  | 53113 |  | 14047/1 |  |  | 2013-03-04 | 529 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00182 | 61659 | [26 empty] | 2013-03-04 | 463 | Gas Splitting | Completed
[168 empty] | 1.0018 | 0.00177 | 61660 | [26 empty] | 2013-03-04 | 974 | Choose Gas Splitting | Completed
[168 empty] | 0.99981 | 0.00183 | 61661 | [26 empty] | 2013-03-06 | 1694 | Graphite | Completed
[168 empty] | 1.0004 | 0.00183 | 61662 | [26 empty] | 2013-03-11 | 1329 | C13 Measurement | Completed
[168 empty] | 0.99813 | 0.00176 | 61695 | [26 empty] | 2013-03-19 | 3758 | AMS Submission Results | Completed
[197 empty] | 2013-03-19 | 2647 | Review Final Report | Completed
[197 empty] | 2013-04-12 | 4592 | Close Job | Completed
[7 empty] | 143.1 | 88.8519 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52474 | 52473 | 52450 | 402 | 401 | 385 |  |  | jenny dahl | -998.54 | 0.07 ‰ | 2013-03-18 | 2013-03-18 | 12 | [5 empty] | 4.85‰ | 4.85‰ | 2.00‰ | 2.00‰ | 4.85 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -998.53 | 0.07 ‰ |  |  | 2013-03-19 | christine prior | 10:11 | 0.0014557 | 0.00146 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.2 |  |  | 0.00146 | 7e-05 | 3759 |  |  |  | 194778 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 24 | 5 | 3 | ... | 0.0014 | 7e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | [6 empty] | 0.0014 |  | 7e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00146 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0014 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-03-13 | jenny dahl | 12:21 |  | Completed |  |  |  |  | 61685 | 2000000 |  |  | 61658 61659 61660 61661 61662 61695 | 61658~61659~61660~61661~61662~61695 | 2850 |  | 2850 2850 2850 2850 2850 2850 |  |  |  |  | 1.03 mg |  | 1.03 | [5 empty] | 0.99941 | 0.00182 | 61658 |  | XCAMS | Both | 0% |  | Best measurement in a while!! no error adjustments needed :). - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-18 | Carbonate |  | JOB-194778 |  |  | 53112 |  | 14047/1 |  |  | 2013-03-04 | 530 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00182 | 61659 | [26 empty] | 2013-03-06 | 1695 | Graphite | Completed
[168 empty] | 1.0018 | 0.00177 | 61660 | [26 empty] | 2013-03-19 | 3759 | AMS Submission Results | Completed
[168 empty] | 0.99981 | 0.00183 | 61661 | [26 empty] | 2013-03-19 | 2648 | Review Final Report | Completed
[168 empty] | 1.0004 | 0.00183 | 61662 | [26 empty] | 2013-04-12 | 4593 | Close Job | Completed
[168 empty] | 0.99813 | 0.00176 | 61695
[7 empty] | 133.7 | 88.8755 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49895 | 49894 | 49882 | 291 | 291 | 280 |  |  | jenny dahl | -997.99 | 0.07 ‰ | 2013-03-26 | 2013-03-26 | 20 | [5 empty] | 1.37‰ | 1.37‰ | 2.00‰ | 2.00‰ | 1.37 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0002 | 0.82 | 0.8165 | [6 empty] | -997.97 | 0.07 ‰ |  |  | 2013-03-28 | christine prior | 11:59 | 0.0020068 | 0.00201 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.7 |  |  | 0.00201 | 7e-05 | 3821 |  |  |  | 194780 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 3 | 5 | 3 | ... | 0.00193 | 7e-05 |  |  |  |  | 1.0001 | 1.0001 | Yes | [6 empty] | 0.00193 |  | 7e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00201 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00193 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-03-22 | jenny dahl | 11:48 |  | Completed |  |  |  |  | 61687 | 2000000 |  |  | 61707 61708 61709 61710 61711 61723 | 61707~61708~61709~61710~61711~61723 | 2854 |  | 2854 2854 2854 2854 2854 2854 |  |  |  |  | 0.74 mg |  | 0.74 | [5 empty] | 0.99825 | 0.00183 | 61707 |  | XCAMS | ST | 0.09% |  | #NAME? |  | RRL | [5 empty] | Background Inorganic |  | 2013-03-18 | Carbonate | Complete | JOB-194780 |  |  | 53272 |  | 14047/1 |  |  | 2013-03-04 | 528 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00183 | 61708 | [26 empty] | 2013-03-18 | 1697 | Graphite | Completed
[168 empty] | 0.99762 | 0.00181 | 61709 | [26 empty] | 2013-03-28 | 3821 | AMS Submission Results | Completed
[168 empty] | 1.0025 | 0.0018 | 61710 | [26 empty] | 2013-05-27 | 4874 | Close Job | Completed
[168 empty] | 1.001 | 0.00181 | 61711
[168 empty] | 1.0005 | 0.00184 | 61723
[7 empty] | 135.5 | 88.8538 | 0.0006 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 45252 | 45251 | 45245 | 234 | 233 | 224 |  |  | jenny dahl | -996.42 | 0.10 ‰ | 2013-04-18 | 2013-04-18 | 27 | [5 empty] | 2.40‰ | 2.40‰ | 2.00‰ | 2.00‰ | 2.40 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -996.38 | 0.11 ‰ |  |  | 2013-04-19 | christine prior | 10:49 | 0.0035769 | 0.00358 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 3.4 |  |  | 0.00358 | 0.0001 | 3903 |  |  |  | 195013 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 4 | .D. | 0.00344 | 0.0001 |  |  |  |  | 1.0009 | 1.0009 | Yes | [6 empty] | 0.00344 |  | 0.0001 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00358 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00344 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-04-17 | jenny dahl | 9:38 |  | Completed |  |  |  |  | 61825 | 2000000 |  |  | 61863 61873 61874 61875 61879 61882 | 61863~61873~61874~61875~61879~61882 | 2856 |  | 2856 2856 2856 2856 2856 2856 |  |  |  |  | 0.69 mg |  | 0.69 | [5 empty] | 0.99953 | 0.00164 | 61863 |  | XCAMS | Both | 0% |  | * no system error or other uncertainty adjustments * measurement without problems CP: I'm inclined to go with the average of both microtubes as combusted blank for this wheel because neither the Kapuni nor the marble blanks were particularly good either. The blanks for this wheel may not be ideal, but the results of the C2 marble and the FIRI suggest that they are reasonable approximations of the real blank. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-04-03 | Carbonate | Complete | JOB-195013 |  | pre-bake yeild is high on this one too. yield is low. | 53357 |  | 14047/1 |  |  | 2013-03-20 | 542 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99749 | 0.00163 | 61873 | [26 empty] | 2013-03-22 | 1835 | Graphite | Completed
[168 empty] | 0.99979 | 0.00164 | 61874 | [26 empty] | 2013-04-03 | 1355 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00161 | 61875 | [26 empty] | 2013-04-19 | 3903 | AMS Submission Results | Completed
[168 empty] | 0.99991 | 0.00166 | 61879 | [26 empty] | 2013-05-27 | 4876 | Close Job | Completed
[168 empty] | 1.0019 | 0.00164 | 61882
[7 empty] | 94 | 88.9994 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52027 | 52027 | 52021 | 271 | 271 | 261 |  |  | jenny dahl | -998.46 | 0.05 ‰ | 2013-05-02 | 2013-05-02 | 13 | [5 empty] | 3.44‰ | 3.44‰ | 2.00‰ | 2.00‰ | 3.44 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0023 | 0.82 | 0.8165 | [6 empty] | -998.44 | 0.05 ‰ |  |  | 2013-05-06 | christine prior | 11:19 | 0.0015389 | 0.00154 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2 |  |  | 0.00154 | 5e-05 | 3977 |  |  |  | 195149 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 25 | 5 | 5 | ... | 0.00148 | 5e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00148 |  | 5e-05 |  | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00154 | 0 | 5e-05 | 0 | 0 |  |  |  |  | 0.00148 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-04-29 | jenny dahl | 11:27 |  | Completed |  |  |  |  | 61974 | 2000000 |  |  | 61936 61939 61942 61945 61955 61956 | 61936~61939~61942~61945~61955~61956 | 2858 |  | 2858 2858 2858 2858 2858 2858 |  |  |  |  | 1.00 mg |  | 1 | [5 empty] | 0.99987 | 0.00148 | 61936 |  | XCAMS | Both | 0% |  | - for this wheel measurement we had to retune twice - for the third tune we used the C-1 in pos#5 (TP61977) for the stable isotopes and the FIRI-H in pos#10 (TP61952), both samples are still reported - system error = 0 - some cathodes with little mass had | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2013-04-30 | Carbonate | Not Complete | JOB-195149 |  | sample was in B bank reactor furnaces stopped heating during prebake cancelled graphite | 53432 |  | 14047/1 |  |  | 2013-04-16 | 548 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00143 | 61939 | [26 empty] | 2013-04-18 | 1976 | Graphite | Cancelled
[168 empty] | 1.0007 | 0.0014 | 61942 | [26 empty] | 2013-04-19 | 1984 | Graphite | Completed
[168 empty] | 1.0019 | 0.00149 | 61945 | [26 empty] | 2013-04-24 | 1399 | C13 Measurement | Completed
[168 empty] | 0.99824 | 0.00134 | 61955 | [26 empty] | 2013-05-06 | 3977 | AMS Submission Results | Completed
[168 empty] | 0.99958 | 0.00137 | 61956 | [26 empty] | 2013-05-27 | 4885 | Close Job | Completed
[7 empty] | 94.7 | 88.8856 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48698 | 48698 | 48695 | 215 | 215 | 207 |  |  | jenny dahl | -997.67 | 0.06 ‰ | 2013-05-02 | 2013-05-02 | 13 | [5 empty] | 4.73‰ | 4.73‰ | 2.00‰ | 2.00‰ | 4.73 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0023 | 0.82 | 0.8165 | [6 empty] | -997.64 | 0.06 ‰ |  |  | 2013-05-06 | christine prior | 11:19 | 0.0023291 | 0.00233 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2 |  |  | 0.00233 | 6e-05 | 3978 |  |  |  | 195150 | [5 empty] | 5 | 5/07/22 | 15:24:47 | Albert Zondervan | 0 |  |  |  | 5 | 5 | 4 | ... | 0.00224 | 6e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00224 |  | 6e-05 |  | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00233 | 0 | 6e-05 | 0 | 0 |  |  |  |  | 0.00224 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-04-29 | jenny dahl | 11:27 |  | Completed |  |  |  |  | 61977 | 2000000 |  |  | 61936 61939 61942 61945 61955 61956 | 61936~61939~61942~61945~61955~61956 | 2858 |  | 2858 2858 2858 2858 2858 2858 |  |  |  |  | 1.03 mg |  | 1.03 | [5 empty] | 0.99987 | 0.00148 | 61936 |  | XCAMS | Both | 0% |  | - for this wheel measurement we had to retune twice - for the third tune we used the C-1 in pos#5 (TP61977) for the stable isotopes and the FIRI-H in pos#10 (TP61952), both samples are still reported - system error = 0 - some cathodes with little mass had |  | RRL | [5 empty] | Background Inorganic |  | 2013-04-30 | Carbonate |  | JOB-195150 |  | sample was in B bank reactor furnaces stopped heating during prebake cancelled graphite | 53433 |  | 14047/1 |  |  | 2013-04-16 | 549 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00143 | 61939 | [26 empty] | 2013-04-18 | 1977 | Graphite | Cancelled
[168 empty] | 1.0007 | 0.0014 | 61942 | [26 empty] | 2013-04-19 | 1987 | Graphite | Completed
[168 empty] | 1.0019 | 0.00149 | 61945 | [26 empty] | 2013-05-06 | 3978 | AMS Submission Results | Completed
[168 empty] | 0.99824 | 0.00134 | 61955 | [26 empty] | 2013-05-27 | 4886 | Close Job | Completed
[168 empty] | 0.99958 | 0.00137 | 61956
[7 empty] | 89.9 | 88.7838 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50322 | 50322 | 50334 | 307 | 307 | 296 |  |  | jenny dahl | -998.1 | 0.07 ‰ | 2013-05-17 | 2013-05-17 | 3 | [5 empty] | 1.65‰ | 1.65‰ | 2.00‰ | 2.00‰ | 1.65 ‰ | 2.00 ‰ | AMS | 2.38‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 1.15 | 1.1547 | [6 empty] | -998.07 | 0.07 ‰ |  |  | 2013-05-23 | christine prior | 13:05 | 0.0019028 | 0.0019 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3.4 |  |  | 0.0019 | 7e-05 | 4069 |  |  |  | 195304 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 15 | 5 | 7 | ... | 0.00183 | 7e-05 |  |  |  |  | 0.99933 | 0.99933 | Yes | [6 empty] | 0.00183 |  | 7e-05 |  | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.0019 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00183 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-05-16 | jenny dahl | 13:53 |  | Completed |  |  |  |  | 62080 | 2000000 |  |  | 62050 62067 62068 | 62050~62067~62068 | 2860 |  | 2860 2860 2860 |  |  |  |  | 0.60 mg |  | 0.6 | [5 empty] | 1.0002 | 0.00195 | 62050 |  | XCAMS | EA | 0% |  | * complex analysis due to issues with XCAMS tuning and possibly systematic differences amongst the Ox-I standards * this set contains results calibrated against EA-combusted OxIs: EA-comb, CO3-evol, and CO2 blanks+unknowns * the re-tune improved the cat-to-cat repeatability, but only on 2 out of 3 Ox-I sub-sets * the re-tune did not improve the run-to-run repeatability, which is explained by mode 5 taking care of fractionation in the source when sputtering proceeds past the sample depth * runs on (near-)blanks have been clumped and the first 4 have been rejected to avoid bias by surface contamination - 29/01/2015: filled System Error (%) box | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2013-05-31 | Carbonate | Not Complete | JOB-195304 |  | transferred from breakseal into bottle 12 | 53508 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99909 | 0.00192 | 62067 | [26 empty] | 2013-05-06 | 559 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0007 | 0.00194 | 62068 | [26 empty] | 2013-05-06 | 513 | Gas Splitting | Completed
[197 empty] | 2013-05-06 | 1092 | Choose Gas Splitting | Completed
[197 empty] | 2013-05-14 | 2090 | Graphite | Completed
[197 empty] | 2013-05-16 | 1425 | C13 Measurement | Completed
[197 empty] | 2013-05-23 | 4069 | AMS Submission Results | Completed
[197 empty] | 2013-05-27 | 4893 | Close Job | Completed
[7 empty] | 90.9 | 88.9426 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50727 | 50727 | 50724 | 369 | 369 | 355 |  |  | Margaret Norris | -998.19 | 0.08 ‰ | 2013-05-27 | 2013-05-27 | 12 | [5 empty] | 0.81‰ | 0.81‰ | 2.00‰ | 2.00‰ | 0.81 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | [6 empty] | -998.17 | 0.08 ‰ |  |  | 2013-05-29 | christine prior | 9:45 | 0.0018092 | 0.00181 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.3 |  |  | 0.00181 | 8e-05 | 4136 |  |  |  | 195441 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 9 | 5 | 5 | ... | 0.00174 | 8e-05 |  |  |  |  | 1.0003 | 1.0003 | Yes | [6 empty] | 0.00174 |  | 8e-05 |  | Yes | 1 | 1 | 0.00106 | 0.00106 | 0.00181 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00174 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-05-27 | Margaret Norris | 9:12 |  | Completed |  |  |  |  | 62097 | 2000000 |  |  | 62094 62098 62129 62130 62187 62190 | 62094~62098~62129~62130~62187~62190 | 2863 |  | 2863 2863 2863 2863 2863 2863 |  |  |  |  | 0.93 mg |  | 0.93 | [5 empty] | 0.99504 | 0.00256 | 62094 |  | XCAMS | ST | 0.18% |  | - system error = 0.18%, this is due to Ox-I in pos# 0 (TP# 62094) being low by ~5 per mille, Without this Ox there would be no system error but there was no hard reason to reject this std from the calibration - regression has a probability of 96% which is | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-06-10 | Carbonate |  | JOB-195441 |  |  | 53579 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99992 | 0.00262 | 62098 | [26 empty] | 2013-05-13 | 566 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00257 | 62129 | [26 empty] | 2013-05-15 | 2107 | Graphite | Completed
[168 empty] | 1.0009 | 0.00258 | 62130 | [26 empty] | 2013-05-16 | 1432 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00262 | 62187 | [26 empty] | 2013-05-29 | 4136 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00258 | 62190 | [26 empty] | 2013-07-16 | 5177 | Close Job | Completed
[7 empty] | 96.5 | 88.8244 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 53069 | 53068 | 53079 | 433 | 432 | 417 |  |  | Margaret Norris | -998.65 | 0.07 ‰ | 2013-05-27 | 2013-05-27 | 4 | [5 empty] | 2.15‰ | 2.15‰ | 2.00‰ | 2.00‰ | 2.15 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | [6 empty] | -998.63 | 0.07 ‰ |  |  | 2013-05-29 | christine prior | 9:45 | 0.0013517 | 0.00135 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.4 |  |  | 0.00135 | 7e-05 | 4144 |  |  |  | 195666 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 31 | 5 | 5 | ... | 0.0013 | 7e-05 |  |  |  |  | 1.0003 | 1.0003 | Yes | [6 empty] | 0.0013 |  | 7e-05 |  | Yes | 1 | 1 | 0.00106 | 0.00106 | 0.00135 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0013 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-05-27 | Margaret Norris | 9:12 |  | Completed |  |  |  |  | 62170 | 2000000 |  |  | 62094 62098 62129 62130 62187 62190 | 62094~62098~62129~62130~62187~62190 | 2863 |  | 2863 2863 2863 2863 2863 2863 |  |  |  |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99504 | 0.00256 | 62094 |  | XCAMS | ST | 0.18% |  | - system error = 0.18%, this is due to Ox-I in pos# 0 (TP# 62094) being low by ~5 per mille, Without this Ox there would be no system error but there was no hard reason to reject this std from the calibration - regression has a probability of 96% which is | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-06-18 | Carbonate |  | JOB-195666 |  |  | 53578 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99992 | 0.00262 | 62098 | [26 empty] | 2013-05-22 | 596 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00257 | 62129 | [26 empty] | 2013-05-23 | 2180 | Graphite | Completed
[168 empty] | 1.0009 | 0.00258 | 62130 | [26 empty] | 2013-05-27 | 1481 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00262 | 62187 | [26 empty] | 2013-05-29 | 4144 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00258 | 62190 | [26 empty] | 2013-08-05 | 5547 | Close Job | Completed
[7 empty] | 87.4 | 88.8932 | 0.0002 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52474 | 52473 | 52450 | 287 | 286 | 275 |  |  | jenny dahl | -998.54 | 0.05 ‰ | 2013-06-06 | 2013-06-06 | 14 | [5 empty] | 1.66‰ | 1.66‰ | 2.00‰ | 2.00‰ | 1.66 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -998.53 | 0.05 ‰ |  |  | 2013-06-10 | christine prior | 9:54 | 0.0014557 | 0.00146 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2.6 |  |  | 0.00146 | 5e-05 | 4214 |  |  |  | 195667 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 5 | 5 | 4 | ... | 0.0014 | 5e-05 |  |  |  |  | 1.0006 | 1.0006 | Yes | [6 empty] | 0.0014 |  | 5e-05 |  | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00146 | 0 | 5e-05 | 0 | 0 |  |  |  |  | 0.0014 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-05-31 | jenny dahl | 11:01 |  | Completed |  |  |  |  | 62171 | 2000000 |  |  | 62189 62191 62192 62220 62221 62223 | 62189~62191~62192~62220~62221~62223 | 2865 |  | 2865 2865 2865 2865 2865 2865 |  |  |  |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99828 | 0.00142 | 62189 |  | XCAMS | ST | 0% |  | * very good results for normal precision requirement * XCAMS was retuned due to poor pos-to-pos repeatability amongst OxI in measurement _1 * zero system error * only 1 cathode (TP62126) had high run-to-run variability |  | RRL | [5 empty] | Background Inorganic |  | 2013-06-18 | Carbonate |  | JOB-195667 |  |  | 53658 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99855 | 0.00143 | 62191 | [26 empty] | 2013-05-22 | 597 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0003 | 0.00142 | 62192 | [26 empty] | 2013-05-23 | 2181 | Graphite | Completed
[168 empty] | 1.0013 | 0.00143 | 62220 | [26 empty] | 2013-06-10 | 4214 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.0014 | 62221 | [26 empty] | 2013-07-16 | 5195 | Close Job | Completed
[168 empty] | 1.0002 | 0.0014 | 62223
[7 empty] | 92.9 | 88.6114 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49936 | 49936 | 49922 | 335 | 334 | 321 |  |  | jenny dahl | -998 | 0.08 ‰ | 2013-06-17 | 2013-06-17 | 17 | [5 empty] | 1.78‰ | 1.78‰ | 2.00‰ | 2.00‰ | 1.78 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | [6 empty] | -997.98 | 0.08 ‰ |  |  | 2013-06-18 | christine prior | 14:25 | 0.0019964 | 0.002 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.3 |  |  | 0.002 | 8e-05 | 4322 |  |  |  | 195769 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 5 | ... | 0.00192 | 8e-05 |  |  |  |  | 0.99752 | 0.99752 | Yes | [6 empty] | 0.00192 |  | 8e-05 |  | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.002 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00192 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-06-13 | jenny dahl | 13:33 |  | Completed |  |  |  |  | 62270 | 2000000 |  |  | 62287 62373 62388 62389 62390 62398 | 62287~62373~62388~62389~62390~62398 | 2868 |  | 2868 2868 2868 2868 2868 2868 |  |  |  |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0027 | 0.00248 | 62287 |  | XCAMS | ST | 0% |  | * straightforward measurement * no system error * ratio-to-standard values of 2 out of 3 OxI "std as unkn" cathodes deviate from 1: TP62323, TP62379 | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2013-06-24 | Carbonate |  | JOB-195769 |  |  | 53778 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99917 | 0.00182 | 62373 | [26 empty] | 2013-05-29 | 615 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99886 | 0.00187 | 62388 | [26 empty] | 2013-05-31 | 2280 | Graphite | Completed
[168 empty] | 0.99909 | 0.00184 | 62389 | [26 empty] | 2013-06-05 | 1503 | C13 Measurement | Completed
[168 empty] | 0.99975 | 0.00186 | 62390 | [26 empty] | 2013-06-18 | 4322 | AMS Submission Results | Completed
[168 empty] | 1.0017 | 0.00187 | 62398 | [26 empty] | 2013-07-16 | 5225 | Close Job | Completed
[7 empty] | 89.6 | 88.6103 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 51057 | 51057 | 51041 | 289 | 288 | 277 |  |  | jenny dahl | -998.26 | 0.06 ‰ | 2013-06-19 | 2013-06-19 | 12 | [5 empty] | 20.29‰ | 20.29‰ | 2.00‰ | 2.00‰ | 20.29 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.70711 | [6 empty] | -998.24 | 0.06 ‰ |  |  | 2013-07-01 | christine prior | 9:51 | 0.0017365 | 0.00174 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.7 |  |  | 0.00174 | 6e-05 | 4378 |  |  |  | 195934 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 4 | ... | 0.00167 | 6e-05 |  |  |  |  | 1.0159 | 1.0159 | Yes | Yes |  | only the runs from before SNICS clean were used since they gave lowest rts, understandably |  |  |  | 0.00167 |  | 6e-05 |  | Yes | 1 | 1 | 0.00052 | 0.00052 | 0.00174 | 0 | 6e-05 | 0 | 0 |  |  |  |  | 0.00167 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-06-18 | jenny dahl | 16:24 |  | Completed |  |  |  |  | 62345 | 2000000 |  |  | 62453 62454 62455 62456 62457 62458 62459 62460 | 62453~62454~62455~62456~62457~62458~62459~62460 | 2869 |  | 2869 2869 2869 2869 2869 2869 2869 2869 |  |  |  |  | 1.09 mg | 0.12 | 1.09 | [5 empty] | 1.002 | 0.00149 | 62453 |  | XCAMS | ST | 0.1% |  | * hi-prec goal achieved, despite problem with repeatability in beginning, prompting a full source clean * for each OxI: 900 k 14C counts, uncertainty = 0.15% * system error = 0.10% * wheel has been in SNICS 8.5 days !! | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-06-19 | Carbonate | Complete | JOB-195934 |  | Aiming it for 1.1mgC, to go with Hills paua samples | 53818 |  | 14047/1 |  |  | 2013-06-05 | 642 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99823 | 0.00148 | 62454 | [26 empty] | 2013-06-07 | 2355 | Graphite | Completed
[168 empty] | 1.0005 | 0.00149 | 62455 | [26 empty] | 2013-06-17 | 1542 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00148 | 62456 | [26 empty] | 2013-07-01 | 4378 | AMS Submission Results | Completed
[168 empty] | 0.99981 | 0.00148 | 62457 | [26 empty] | 2013-07-01 | 3122 | Review Final Report | Completed
[168 empty] | 0.99752 | 0.00147 | 62458 | [26 empty] | 2013-07-16 | 5239 | Close Job | Completed
[168 empty] | 1.0006 | 0.00149 | 62459
[168 empty] | 1.0011 | 0.00147 | 62460
[7 empty] | 107.4 | 89.0997 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49812 | 49811 | 49802 | 330 | 329 | 317 |  |  | jenny dahl | -997.97 | 0.08 ‰ | 2013-06-29 | 2013-06-29 | 23 | [5 empty] | 1.33‰ | 1.33‰ | 2.00‰ | 2.00‰ | 1.33 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | [6 empty] | -997.95 | 0.08 ‰ |  |  | 2013-07-01 | christine prior | 10:44 | 0.0020276 | 0.00203 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.00203 | 8e-05 | 4407 |  |  |  | 195768 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 5 | ... | 0.00195 | 8e-05 |  |  |  |  | 1.0026 | 1.0026 | Yes | [6 empty] | 0.00195 |  | 8e-05 |  | Yes | 1 | 0.99999 | 0.00092 | 0.00092 | 0.00203 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00195 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-06-20 | jenny dahl | 8:47 |  | Completed |  |  |  |  | 62324 | 2000000 |  |  | 62408 62409 62474 62475 62476 62477 | 62408~62409~62474~62475~62476~62477 | 2870 |  | 2870 2870 2870 2870 2870 2870 |  |  |  |  | 0.86 mg |  | 0.86 | [5 empty] | 1.0016 | 0.00226 | 62408 |  | XCAMS | EA | 0.15% |  | * system error = 0.15% * global uncertainty correction factor = 1.02 * cathode-specific uncertainty correction factors: 1.3 for cathode 4 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-06-24 | Carbonate |  | JOB-195768 |  | 0.9mgC, not big enough, make another one to go with Hills paua samples | 53858 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0011 | 0.00225 | 62409 | [26 empty] | 2013-06-04 | 637 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99647 | 0.00226 | 62474 | [26 empty] | 2013-06-06 | 2334 | Graphite | Completed
[168 empty] | 0.99897 | 0.00227 | 62475 | [26 empty] | 2013-06-12 | 1528 | C13 Measurement | Completed
[168 empty] | 1.0027 | 0.00225 | 62476 | [26 empty] | 2013-07-01 | 4407 | AMS Submission Results | Completed
[168 empty] | 0.9991 | 0.00225 | 62477 | [26 empty] | 2013-08-05 | 5558 | Close Job | Completed
[7 empty] | 116.6 | 88.963 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47816 | 47815 | 47814 | 321 | 321 | 309 |  |  | christine prior | -997.4 | 0.10 ‰ | 2013-07-02 | 2013-07-02 | 20 | [5 empty] | 2.85‰ | 2.85‰ | 2.00‰ | 2.00‰ | 2.85 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -997.37 | 0.11 ‰ |  |  | 2013-07-03 | christine prior | 15:16 | 0.0025995 | 0.0026 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.9 |  |  | 0.0026 | 0.0001 | 4474 |  |  |  | 195944 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 19 | 5 | 4 | ... | 0.0025 | 0.0001 |  |  |  |  | 1.0026 | 1.0025 | Yes | [6 empty] | 0.0025 |  | 0.0001 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.0026 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.0025 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-06-26 | christine prior | 14:31 |  | Completed |  |  |  |  | 62397 | 2000000 |  |  | 62415 62416 62417 62465 62482 62519 | 62415~62416~62417~62465~62482~62519 | 2871 |  | 2871 2871 2871 2871 2871 2871 |  |  |  |  | 0.73 mg |  | 0.73 | [5 empty] | 1.0022 | 0.00197 | 62415 |  | XCAMS | ST | 0% |  | * retune was necessary * no system error * 2 cathodes with uncertainty adjustment | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2013-06-25 | Carbonate |  | JOB-195944 |  |  | 53898 |  | 14047/1 |  |  | 2013-06-11 | 658 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9986 | 0.00195 | 62416 | [26 empty] | 2013-06-12 | 2407 | Graphite | Completed
[168 empty] | 0.99626 | 0.00258 | 62417 | [26 empty] | 2013-06-17 | 1543 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00198 | 62465 | [26 empty] | 2013-07-03 | 4474 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00195 | 62482 | [26 empty] | 2013-07-03 | 3176 | Review Final Report | Completed
[168 empty] | 0.99981 | 0.00201 | 62519 | [26 empty] | 2013-07-26 | 5311 | Close Job | Completed
[7 empty] | 119.5 | 88.7506 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48990 | 48990 | 48976 | 372 | 371 | 357 |  |  | christine prior | -997.75 | 0.10 ‰ | 2013-07-16 | 2013-07-16 | 26 | [5 empty] | 2.80‰ | 2.80‰ | 2.00‰ | 2.00‰ | 2.80 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -997.73 | 0.11 ‰ |  |  | 2013-07-17 | christine prior | 10:23 | 0.002246 | 0.00225 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.2 |  |  | 0.00225 | 0.0001 | 4544 |  |  |  | 196070 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 3 | 5 | 4 | ... | 0.00216 | 0.0001 |  |  |  |  | 1.0001 | 1.0001 | Yes | [6 empty] | 0.00216 |  | 0.0001 |  | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.00225 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00216 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-09 | christine prior | 13:08 |  | Completed |  |  |  |  | 62489 | 2000000 |  |  | 62569 62570 62598 62603 62606 62607 | 62569~62570~62598~62603~62606~62607 | 2873 |  | 2873 2873 2873 2873 2873 2873 |  |  |  |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99779 | 0.00197 | 62569 |  | XCAMS | EA | 0% |  | * no problems, despite lack of remote control magnet BM05 * zero system error * global uncertainty correction factor 1.06 * 2 cathodes with uncertainty correction factor 1.4 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-07-02 | Carbonate |  | JOB-196070 |  |  | 53978 |  | 14047/1 |  |  | 2013-06-18 | 671 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00198 | 62570 | [26 empty] | 2013-06-20 | 2499 | Graphite | Completed
[168 empty] | 0.99926 | 0.00198 | 62598 | [26 empty] | 2013-06-25 | 1587 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00197 | 62603 | [26 empty] | 2013-07-17 | 4544 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00194 | 62606 | [26 empty] | 2013-07-17 | 3227 | Review Final Report | Completed
[168 empty] | 1.0004 | 0.00202 | 62607 | [26 empty] | 2013-08-05 | 5584 | Close Job | Completed
[7 empty] | 119.6 | 88.9361 | 0.0007 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 43862 | 43861 | 43867 | 275 | 274 | 284 |  |  | christine prior | -995.75 | 0.15 ‰ | 2013-07-17 | 2013-07-17 | 19 | [5 empty] | 1.28‰ | 1.28‰ | 2.00‰ | 2.00‰ | 1.28 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | [6 empty] | -995.69 | 0.15 ‰ |  |  | 2013-07-18 | christine prior | 14:19 | 0.0042528 | 0.00425 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.4 |  |  | 0.00425 | 0.00015 | 4598 |  |  |  | 196127 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 4 | ... | 0.00409 | 0.00014 |  |  |  |  | 1.0007 | 1.0007 | Yes |  |  | Anomalously young marble blank in this wheel. Comparison of other blanks produced in June/July shows Fm average of 0.00224 ± .00009. Based on CO2 production dates and graphite dates of unknown carbonates in this wheel, decided to use assigned blank Fm average. |  |  |  | 0.00409 |  | 0.00014 |  | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00425 | 0 | 0.00015 | 0 | 0 |  |  |  |  | 0.00409 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-12 | christine prior | 16:04 |  | Completed |  |  |  |  | 62573 | 2000000 |  |  | 62520 62521 62522 62523 62543 62593 | 62520~62521~62522~62523~62543~62593 | 2874 |  | 2874 2874 2874 2874 2874 2874 |  |  |  |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0025 | 0.00174 | 62520 |  | XCAMS | ST | 0% |  | * good measurement * no error adjustments needed - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-07-10 | Carbonate |  | JOB-196127 |  |  | 54018 |  | 14047/1 |  |  | 2013-06-26 | 693 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00186 | 62521 | [26 empty] | 2013-06-28 | 2583 | Graphite | Completed
[168 empty] | 0.99958 | 0.00185 | 62522 | [26 empty] | 2013-07-02 | 1629 | C13 Measurement | Completed
[168 empty] | 0.9991 | 0.00187 | 62523 | [26 empty] | 2013-07-18 | 4598 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00183 | 62543 | [26 empty] | 2013-07-23 | 3266 | Review Final Report | Completed
[168 empty] | 0.99956 | 0.00183 | 62593 | [26 empty] | 2014-07-08 | 8950 | Close Job | Completed
[7 empty] | 127.2 | 88.7186 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48879 | 48879 | 48869 | 403 | 403 | 388 |  |  | jenny dahl | -997.72 | 0.11 ‰ | 2013-07-18 | 2013-07-18 | 13 | [5 empty] | 3.78‰ | 3.78‰ | 2.00‰ | 2.00‰ | 3.78 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0011 | 0 | 0.70711 | [6 empty] | -997.69 | 0.12 ‰ |  |  | 2013-07-19 | christine prior | 13:13 | 0.0022772 | 0.00228 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.8 |  |  | 0.00228 | 0.00011 | 4625 |  |  |  | 196180 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 37 | 5 | 3 | ... | 0.00219 | 0.00011 |  |  |  |  | 1.0007 | 1.0007 | Yes |  |  | Poisson uncertainty multiplied by factor 1.4 to account for excess run-to-run variability |  |  |  | 0.00219 |  | 0.00011 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00228 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00219 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-17 | jenny dahl | 10:16 |  | Completed |  |  |  |  | 62638 | 2000000 |  |  | 62614 62621 62622 62623 62629 62630 62636 62640 | 62614~62621~62622~62623~62629~62630~62636~62640 | 2875 |  | 2875 2875 2875 2875 2875 2875 2875 2875 |  |  |  |  | 1.15 mg |  | 1.15 | [5 empty] | 0.99798 | 0.00173 | 62614 |  | XCAMS | Both | 0% |  | * straightforward measurement * zero system error * 2 cathodes with error adjustment * HYO results imported separately JCT - 2019 - all are now assigned to TW2875 even though the HYO samples were analysed separately |  | RRL | [5 empty] | Background Inorganic |  | 2013-07-17 | Carbonate | Complete | JOB-196180 |  |  | 54073 |  | 14047/1 |  |  | 2013-07-03 | 714 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9993 | 0.00174 | 62621 | [26 empty] | 2013-07-05 | 2648 | Graphite | Completed
[168 empty] | 1.0026 | 0.00174 | 62622 | [26 empty] | 2013-07-19 | 4625 | AMS Submission Results | Completed
[168 empty] | 0.99958 | 0.00175 | 62623 | [26 empty] | 2013-07-23 | 3300 | Review Final Report | Completed
[168 empty] | 1.0013 | 0.00176 | 62629 | [26 empty] | 2013-07-26 | 5328 | Close Job | Completed
[168 empty] | 0.99838 | 0.00181 | 62630
[168 empty] | 0.99994 | 0.00175 | 62636
[168 empty] | 1.0018 | 0.00257 | 62640
[7 empty] | 57.7 | 90.0687 | 0.0008 | RRL-CBIn-TY | Background Inorganic |  |  |  |  | 43404 | 43403 | 43408 | 315 | 315 | 321 |  |  | jenny dahl | -995.5 | 0.18 ‰ | 2013-07-23 | 2013-07-23 | 21 | [5 empty] | -9.67‰ | -9.67‰ | 2.00‰ | 2.00‰ | -9.67 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0032 | 0 | 0.75593 | [6 empty] | -995.44 | 0.18 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.0045023 | 0.0045 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 14.3 |  |  | 0.0045 | 0.00018 | 4650 |  |  |  | 196128 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 3 | 5 | 3 | ... | 0.00433 | 0.00017 |  |  |  |  | 1.0023 | 1.0023 | Yes |  |  | Small blank to calculate DCC |  |  |  | 0.00433 |  | 0.00017 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.0045 | 0 | 0.00018 | 0 | 0 |  |  |  |  | 0.00433 | 0 | 0.00017 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62589 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.14 mg |  | 0.14 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 |  | RRL | [5 empty] | Background Inorganic |  | 2013-07-11 | Carbonate |  | JOB-196128 |  | small blank to match R40293/22 | 54098 |  | 14047/1 |  |  | 2013-07-01 | 709 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00181 | 62631 | [26 empty] | 2013-07-02 | 2599 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [26 empty] | 2013-07-24 | 4650 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [26 empty] | 2013-07-31 | 3321 | Review Final Report | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [26 empty] | 2014-07-08 | 8951 | Close Job | Completed
[168 empty] | 0.99895 | 0.00176 | 62645
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 41.9 | 91.2847 | 0.0025 | RRL-CBIn-TY | Background Inorganic |  |  |  |  | 33455 | 33454 | 33457 | 199 | 198 | 197 |  |  | jenny dahl | -984.47 | 0.38 ‰ | 2013-07-23 | 2013-07-23 | 19 | [5 empty] | -22.87‰ | -22.87‰ | 2.00‰ | 2.00‰ | -22.87 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0032 | 0 | 0.75593 | [6 empty] | -984.27 | 0.39 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.015535 | 0.01553 | 0.00038 | 0.00038 | 0.00038 | 0.00038 | 16.6 |  |  | 0.01553 | 0.00038 | 4651 |  |  |  | 196129 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 3 | ... | 0.01494 | 0.00037 |  |  |  |  | 1.0023 | 1.0023 | Yes |  |  | Small blank to calculate DCC |  |  |  | 0.01494 |  | 0.00037 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.01553 | 0 | 0.00038 | 0 | 0 |  |  |  |  | 0.01494 | 0 | 0.00037 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62618 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.14 mg |  | 0.14 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 |  | RRL | [5 empty] | Background Inorganic |  | 2013-07-11 | Carbonate |  | JOB-196129 |  | small blank to match R40293/22 | 54099 |  | 14047/1 |  |  | 2013-07-01 | 710 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00181 | 62631 | [26 empty] | 2013-07-04 | 2628 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [26 empty] | 2013-07-24 | 4651 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [26 empty] | 2013-07-31 | 3322 | Review Final Report | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [26 empty] | 2013-08-05 | 5592 | Close Job | Completed
[168 empty] | 0.99895 | 0.00176 | 62645
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 116 | 89.2584 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49689 | 49689 | 49685 | 325 | 324 | 312 |  |  | jenny dahl | -997.94 | 0.08 ‰ | 2013-07-23 | 2013-07-23 | 15 | [5 empty] | -0.68‰ | -0.68‰ | 2.00‰ | 2.00‰ | -0.68 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0032 | 0 | 0.75593 | [6 empty] | -997.92 | 0.08 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.0020588 | 0.00206 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00206 | 8e-05 | 4656 |  |  |  | 196181 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 22 | 5 | 4 | ... | 0.00198 | 8e-05 |  |  |  |  | 1.0023 | 1.0023 | Yes | [6 empty] | 0.00198 |  | 8e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00206 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00198 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62650 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.86 mg |  | 0.86 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-07-17 | Carbonate |  | JOB-196181 |  | normal size | 54100 |  | 14047/1 |  |  | 2013-07-03 | 715 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00181 | 62631 | [26 empty] | 2013-07-08 | 2660 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [26 empty] | 2013-07-11 | 1667 | C13 Measurement | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [26 empty] | 2013-07-24 | 4656 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [26 empty] | 2013-07-26 | 3324 | Review Final Report | Completed
[168 empty] | 0.99895 | 0.00176 | 62645 | [26 empty] | 2013-07-30 | 5357 | Close Job | Completed
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 122.9 | 88.9666 | 0.0003 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49608 | 49608 | 49607 | 321 | 321 | 309 | 49293.4 | 49936.4 | jenny dahl | -997.92 | 0.08 ‰ | 2013-07-23 | 2013-07-23 | 13 | [5 empty] | 2.59‰ | 2.59‰ | 2.00‰ | 2.00‰ | 2.59 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0032 | 0 | 0.75593 | -997.92 ‰ | -997.92 ‰ | 0.08 | 0.08 | -997.89 | 0.08 ‰ | -997.89 | 0.08 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.0020796 | 0.00208 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00208 | 8e-05 | 4662 |  |  |  | 196195 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  | 0.21 | 0.01% | 38 | 5 | 4 | .X. | 0.002 | 8e-05 |  |  |  |  | 1.0023 | 1.0023 | Yes |  |  | Good to see that blanks produced on larger water line are similar to blanks on smaller carbonate line. .X. to indicate that this is not really a water blank |  |  |  | 0.002 |  | 8e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00208 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.002 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62688 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.94 mg |  | 0.94 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2013-08-30 | Water | Complete | JOB-196195 |  | Water line evacuated like carbonate. | 54101 | Normal | 14047/11 |  |  | 2013-07-09 | 718 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00181 | 62631 | [26 empty] | 2013-07-10 | 2698 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [26 empty] | 2013-07-11 | 1669 | C13 Measurement | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [26 empty] | 2013-07-24 | 4662 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [26 empty] | 2013-07-26 | 3329 | Review Final Report | Completed
[168 empty] | 0.99895 | 0.00176 | 62645 | [26 empty] | 2013-07-30 | 5363 | Close Job | Completed
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 121.7 | 89.012 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47563 | 47562 | 47571 | 280 | 280 | 270 |  |  | jenny dahl | -997.32 | 0.09 ‰ | 2013-07-23 | 2013-07-23 | 13 | [5 empty] | 2.08‰ | 2.08‰ | 2.00‰ | 2.00‰ | 2.08 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0032 | 0 | 0.75593 | [6 empty] | -997.28 | 0.09 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.0026827 | 0.00268 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00268 | 9e-05 | 4663 |  |  |  | 196196 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 4 | ... | 0.00258 | 9e-05 |  |  |  |  | 1.0023 | 1.0023 | Yes | [6 empty] | 0.00258 |  | 9e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00268 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00258 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62686 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.89 mg |  | 0.89 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-07-19 | Carbonate |  | JOB-196196 |  | Normal size, process on carbonate line as normal to compare with water line evolved marble | 54102 |  | 14047/1 |  |  | 2013-07-09 | 719 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00181 | 62631 | [26 empty] | 2013-07-10 | 2696 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [26 empty] | 2013-07-11 | 1670 | C13 Measurement | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [26 empty] | 2013-07-24 | 4663 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [26 empty] | 2013-07-26 | 3330 | Review Final Report | Completed
[168 empty] | 0.99895 | 0.00176 | 62645 | [26 empty] | 2013-07-30 | 5364 | Close Job | Completed
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 110.9 | 89.1584 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49332 | 49332 | 49341 | 349 | 349 | 336 |  |  | jenny dahl | -997.85 | 0.09 ‰ | 2013-08-14 | 2013-08-14 | 13 | [5 empty] | 2.69‰ | 2.69‰ | 2.00‰ | 2.00‰ | 2.69 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.70711 | [6 empty] | -997.82 | 0.09 ‰ |  |  | 2013-08-15 | christine prior | 10:02 | 0.0021524 | 0.00215 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00215 | 9e-05 | 4821 |  |  |  | 196358 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 4 | ... | 0.00207 | 9e-05 |  |  |  |  | 1.0046 | 1.0046 | Yes | [6 empty] | 0.00207 |  | 9e-05 |  | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00215 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00207 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-08-07 | jenny dahl | 14:43 |  | Completed |  |  |  |  | 62782 | 2000000 |  |  | 62701 62702 62766 62788 62790 62792 62794 62796 | 62701~62702~62766~62788~62790~62792~62794~62796 | 2879 |  | 2879 2879 2879 2879 2879 2879 2879 2879 |  |  |  |  | 0.93 mg |  | 0.93 | [5 empty] | 0.99777 | 0.002 | 62701 |  | XCAMS | Both | 0% |  | - very good measurement for being the first after a SNICS opening - no system error - needed to add a minor global run-to-run variance factor of 1.04 - TP62773 needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-08-06 | Carbonate |  | JOB-196358 |  |  | 54218 |  | 14047/1 |  |  | 2013-07-30 | 737 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99954 | 0.00199 | 62702 | [26 empty] | 2013-08-02 | 2792 | Graphite | Completed
[168 empty] | 0.99937 | 0.00202 | 62766 | [26 empty] | 2013-08-08 | 1695 | C13 Measurement | Completed
[168 empty] | 1.0005 | 0.00201 | 62788 | [26 empty] | 2013-08-15 | 4821 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.002 | 62790 | [26 empty] | 2013-08-16 | 3393 | Review Final Report | Completed
[168 empty] | 0.99894 | 0.00202 | 62792 | [26 empty] | 2013-09-04 | 5822 | Close Job | Completed
[168 empty] | 1.0018 | 0.00204 | 62794
[168 empty] | 1.0019 | 0.00199 | 62796
[7 empty] | 112.8 | 89.1872 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 46679 | 46679 | 46692 | 335 | 334 | 322 |  |  | jenny dahl | -997.01 | 0.12 ‰ | 2013-08-23 | 2013-08-23 | 8 | [5 empty] | 2.55‰ | 2.55‰ | 2.00‰ | 2.00‰ | 2.55 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.70711 | [6 empty] | -996.97 | 0.13 ‰ |  |  | 2013-08-26 | jenny dahl | 13:50 | 0.0029946 | 0.00299 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00299 | 0.00012 | 4894 |  |  |  | 196451 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 6 | ... | 0.00288 | 0.00012 |  |  |  |  | 1.0048 | 1.0048 | Yes |  |  | 26/8/13-jd Younger marble, perhaps due to contamination. delta 13 and fm appear in keeping. 7 consecutive runs showed elevated rts consistent with contamination in graphite; these runs were rejected |  |  |  | 0.00288 |  | 0.00012 |  | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00299 | 0 | 0.00012 | 0 | 0 |  |  |  |  | 0.00288 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-08-21 | jenny dahl | 16:10 |  | Completed |  |  |  |  | 62844 | 2000000 |  |  | 62832 62845 62858 62859 62860 62861 62878 62879 | 62832~62845~62858~62859~62860~62861~62878~62879 | 2881 |  | 2881 2881 2881 2881 2881 2881 2881 2881 |  |  |  |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99895 | 0.00177 | 62832 |  | XCAMS | Both | 0% |  | good measurement with few uncertainty adjustments * zero system error as 8x OxIs repeat well * two cathodes have excess run-to-run variability * global error adjustment factor = 1.05 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-08-15 | Carbonate |  | JOB-196451 |  |  | 54325 |  | 14047/1 |  |  | 2013-08-01 | 749 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.0017 | 62845 | [26 empty] | 2013-08-14 | 628 | Gas Splitting | Completed
[168 empty] | 1.0015 | 0.00174 | 62858 | [26 empty] | 2013-08-14 | 1360 | Choose Gas Splitting | Completed
[168 empty] | 0.99952 | 0.00177 | 62859 | [26 empty] | 2013-08-15 | 2854 | Graphite | Completed
[168 empty] | 1.0001 | 0.00175 | 62860 | [26 empty] | 2013-08-23 | 1717 | C13 Measurement | Completed
[168 empty] | 0.99819 | 0.00179 | 62861 | [26 empty] | 2013-08-26 | 4894 | AMS Submission Results | Completed
[168 empty] | 0.99852 | 0.00176 | 62878 | [26 empty] | 2013-09-04 | 5844 | Close Job | Completed
[168 empty] | 1.0017 | 0.00176 | 62879
[7 empty] | 97.8 | 89.533 | 0.0003 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 50635 | 50635 | 50636 | 365 | 365 | 351 |  |  | Kelly Lyons | -998.17 | 0.08 ‰ | 2013-08-30 | 2013-08-30 | 10 | [5 empty] | 0.12‰ | 0.12‰ | 2.00‰ | 2.00‰ | 0.12 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19 | 0 | 0.8165 | [6 empty] | -998.15 | 0.08 ‰ |  |  | 2013-09-03 | christine prior | 15:59 | 0.00183 | 0.00183 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 4.5 |  |  | 0.00183 | 8e-05 | 4941 |  |  |  | 196511 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 3 | ... | 0.00176 | 8e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00176 |  | 8e-05 |  | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00183 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00176 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-08-28 | Kelly Lyons | 14:21 |  | Completed |  |  |  |  | 62870 | 2000000 |  |  | 62880 62881 62899 62900 62901 62902 | 62880~62881~62899~62900~62901~62902 | 2882 |  | 2882 2882 2882 2882 2882 2882 |  |  |  |  | 0.44 mg |  | 0.44 | [5 empty] | 1.0012 | 0.00172 | 62880 |  | XCAMS | EA | 0% |  | - good measurement after re-tune has been done - no system error or other error adjustments needed |  | RRL | [5 empty] | Background Inorganic |  | 2013-08-22 | Carbonate |  | JOB-196511 |  | too small for d13C | 54337 |  | 14047/1 |  |  | 2013-08-19 | 777 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99983 | 0.0017 | 62881 | [26 empty] | 2013-08-20 | 2880 | Graphite | Completed
[168 empty] | 0.99946 | 0.00171 | 62899 | [26 empty] | 2013-09-03 | 4941 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00169 | 62900 | [26 empty] | 2013-09-13 | 3457 | Review Final Report | Completed
[168 empty] | 1.0011 | 0.0017 | 62901 | [26 empty] | 2013-10-10 | 6093 | Close Job | Completed
[168 empty] | 0.99782 | 0.00169 | 62902
[7 empty] | 115.6 | 90.1483 | 0.0683 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7075 | 7075 | 7075 | 26 | 26 | 26 | 7049.3 | 7101.7 | jenny dahl | -585.55 | 1.35 ‰ | 2013-09-10 | 2013-09-10 | 20 | [5 empty] | -7.71‰ | -7.71‰ | 2.00‰ | 2.00‰ | -7.71 ‰ | 2.00 ‰ | AMS | -8.04‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | [6 empty] | -580.43 | 1.37 ‰ |  |  | 2013-09-12 | christine prior | 16:17 | 0.41445 | 0.41445 | 0.00135 | 0.00135 | 0.00135 | 0.00113 | 2 |  |  | 0.41445 | 0.00135 | 5020 |  |  |  | 196513 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 4 | A.. | 0.39859 | 0.0013 |  |  |  |  | 1.0052 | 1.0052 | Yes |  |  | This age and d13C are what would be expected from C2 travertine. Obvious that there was a switch with job 196512 |  |  |  | 0.39859 |  | 0.0013 |  | No | 1 | 0.99999 | 0.00102 | 0.00102 | 0.41445 | 0 | 0.00135 | 0 | 0 |  |  |  |  | 0.39859 | 0 | 0.0013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-09-05 | jenny dahl | 15:39 |  | Completed |  |  |  |  | 62889 | 2000000 |  |  | 63048 63052 63063 63064 63065 63082 | 63048~63052~63063~63064~63065~63082 | 2884 |  | 2884 2884 2884 2884 2884 2884 |  |  |  |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99955 | 0.00247 | 63048 |  | XCAMS | EA | 0.18% |  | * system error = 0.18% * no other error adjustments required * small block of runs rejected due to unstable Pelletron JCT - one OxI was a 2 sigma outlier, but this does not seem sufficient to exclude it. Excluding it would increase fm by 0.1% = 7 years. | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2013-08-22 | Carbonate |  | JOB-196513 |  | Confirmed swap with job 196512 This job was done directly after secondary sample that came out dead. This sample is possibly swapped with secondary material. | 54417 |  | 14047/1 |  |  | 2013-08-19 | 779 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99709 | 0.00251 | 63052 | [26 empty] | 2013-08-21 | 2899 | Graphite | Completed
[168 empty] | 0.99942 | 0.00251 | 63063 | [26 empty] | 2013-08-26 | 1730 | C13 Measurement | Completed
[168 empty] | 0.99828 | 0.0025 | 63064 | [26 empty] | 2013-09-12 | 5020 | AMS Submission Results | Completed
[168 empty] | 1.0042 | 0.0025 | 63065 | [26 empty] | 2013-09-13 | 3521 | Review Final Report | Completed
[168 empty] | 1.0014 | 0.00249 | 63082 | [26 empty] | 2013-10-10 | 6094 | Close Job | Completed
[7 empty] | 111.5 | 89.4367 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 52137 | 52136 | 52126 | 385 | 385 | 370 |  |  | jenny dahl | -998.48 | 0.07 ‰ | 2013-09-12 | 2013-09-12 | 8 | [5 empty] | 1.47‰ | 1.47‰ | 2.00‰ | 2.00‰ | 1.47 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -998.46 | 0.07 ‰ |  |  | 2013-09-13 | christine prior | 16:41 | 0.0015181 | 0.00152 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.3 |  |  | 0.00152 | 7e-05 | 5064 |  |  |  | 196658 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 4 | ... | 0.00146 | 7e-05 |  |  |  |  | 1.0065 | 1.0065 | Yes | [6 empty] | 0.00146 |  | 7e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00152 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00146 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-09-09 | jenny dahl | 16:07 |  | Completed |  |  |  |  | 63069 | 2000000 |  |  | 63083 63120 63122 63123 63124 63125 | 63083~63120~63122~63123~63124~63125 | 2885 |  | 2885 2885 2885 2885 2885 2885 |  |  |  |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99821 | 0.00181 | 63083 |  | XCAMS | EA | 0% |  | * error adjustments needed: NONE for cat-to-cat variability (i.e. system error = 0), factor 1.07 for global run-to-run variability, and factor 1.3 for run-to-run variability of cathode 36 (TP63061). * CalAMS results suggest that the ratio-to-standard of all of the cathodes fluctuates more than expected and that this is not due to poor XCAMS calibration during parts or the whole of the measurement period. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-09-17 | Carbonate |  | JOB-196658 |  |  | 54457 |  | 14047/1 |  |  | 2013-09-03 | 788 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99917 | 0.0018 | 63120 | [26 empty] | 2013-09-04 | 3079 | Graphite | Completed
[168 empty] | 1.0012 | 0.00184 | 63122 | [26 empty] | 2013-09-11 | 1787 | C13 Measurement | Completed
[168 empty] | 0.99912 | 0.00181 | 63123 | [26 empty] | 2013-09-13 | 5064 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 63124 | [26 empty] | 2013-09-17 | 3555 | Review Final Report | Completed
[168 empty] | 1.0013 | 0.00179 | 63125 | [26 empty] | 2014-07-10 | 9055 | Close Job | Completed
[7 empty] | 113.3 | 88.6329 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49936 | 49936 | 49922 | 335 | 334 | 321 |  |  | jenny dahl | -998 | 0.08 ‰ | 2013-09-23 | 2013-09-23 | 19 | [5 empty] | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0001 | 0 | 0.8165 | [6 empty] | -997.98 | 0.08 ‰ |  |  | 2013-09-25 | Jocelyn Turnbull | 8:27 | 0.0019964 | 0.002 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.002 | 8e-05 | 5224 | 2013-09-24 | 11:22 | Johannes Kaiser | 196659 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 5 | ... | 0.00192 | 8e-05 |  |  |  |  | 1.0005 | 1.0005 | Yes | [6 empty] | 0.00192 |  | 8e-05 |  | Yes | 1 | 1 | 0.00103 | 0.00103 | 0.002 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00192 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-09-17 | jenny dahl | 16:20 |  | Completed |  |  |  |  | 63070 | 2000000 |  |  | 63157 63158 63159 63160 63161 63164 | 63157~63158~63159~63160~63161~63164 | 2888 |  | 2888 2888 2888 2888 2888 2888 | 2013-09-24 | 11:54 | christine prior |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0027 | 0.00252 | 63157 |  | XCAMS | EA | 0.17% |  | - sys error = 0.17% due to Ox-I in pos# 28 (TP63161) being low by ~4 per mille, but there is no reason to reject it from the calibration - 13/12 are repeating very well (1 per mille spread) for the 6 calibration standards - Transmission > 43%, wide sputte |  | RRL | [5 empty] | Background Inorganic |  | 2013-09-17 | Carbonate |  | JOB-196659 |  |  | 54614 |  | 14047/1 |  |  | 2013-09-03 | 1772 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99838 | 0.00251 | 63158 | [26 empty] | 2013-09-03 | 789 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0008 | 0.00252 | 63159 | [26 empty] | 2013-09-04 | 3080 | Graphite | Completed
[168 empty] | 1.0016 | 0.00254 | 63160 | [26 empty] | 2013-09-04 | 1530 | Write Pretreatment Report | Completed
[168 empty] | 0.99575 | 0.00252 | 63161 | [26 empty] | 2013-09-25 | 5224 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.00251 | 63164 | [26 empty] | 2013-09-25 | 3654 | Review Final Report | Completed
[197 empty] | 2013-10-10 | 6174 | Close Job | Completed
[7 empty] | 118.8 | 88.7112 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49293 | 49293 | 49304 | 348 | 347 | 335 |  |  | jenny dahl | -997.84 | 0.09 ‰ | 2013-09-24 | 2013-09-24 | 6 | [5 empty] | 1.95‰ | 1.95‰ | 2.00‰ | 2.00‰ | 1.95 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | [6 empty] | -997.81 | 0.09 ‰ |  |  | 2013-09-25 | Jocelyn Turnbull | 16:30 | 0.0021628 | 0.00216 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00216 | 9e-05 | 5261 | 2013-09-25 | 10:02 | Johannes Kaiser | 196735 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 31 | 5 | 6 | ... | 0.00208 | 9e-05 |  |  |  |  | 0.99881 | 0.99881 | Yes | [6 empty] | 0.00208 |  | 9e-05 |  | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00216 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00208 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-09-19 | jenny dahl | 10:05 |  | Completed |  |  |  |  | 63247 | 2000000 |  |  | 63150 63206 63207 63210 63243 63248 | 63150~63206~63207~63210~63243~63248 | 2889 |  | 2889 2889 2889 2889 2889 2889 | 2013-09-25 | 16:17 | christine prior |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0026 | 0.00227 | 63150 |  | XCAMS | ST | 0.12% |  | - sys error = 0.12%, due to two Ox being high by 2.5 per mille and two being low by 2.5 per mille --> there is a hint of inhomogeneity for Ox-I in pos#21 (TP63210, reported rts = 0.99748)) as when taking out run 11 then it would be low by 4.3 per mille, | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-09-25 | Carbonate |  | JOB-196735 |  |  | 54616 |  | 14047/1 |  |  | 2013-09-11 | 819 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99926 | 0.00228 | 63206 | [26 empty] | 2013-09-16 | 704 | Gas Splitting | Completed
[168 empty] | 0.99769 | 0.00227 | 63207 | [26 empty] | 2013-09-16 | 1464 | Choose Gas Splitting | Completed
[168 empty] | 0.99748 | 0.00227 | 63210 | [26 empty] | 2013-09-18 | 3257 | Graphite | Completed
[168 empty] | 1.0026 | 0.00232 | 63243 | [26 empty] | 2013-09-20 | 1865 | C13 Measurement | Completed
[168 empty] | 1.0005 | 0.00228 | 63248 | [26 empty] | 2013-09-25 | 5261 | AMS Submission Results | Completed
[197 empty] | 2013-09-26 | 3687 | Review Final Report | Completed
[197 empty] | 2013-10-09 | 6073 | Close Job | Completed
[7 empty] | 117.6 | 88.8844 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 51604 | 51604 | 51615 | 360 | 360 | 347 |  |  | jenny dahl | -998.38 | 0.07 ‰ | 2013-09-30 | 2013-09-30 | 7 | [5 empty] | 1.47‰ | 1.47‰ | 2.00‰ | 2.00‰ | 1.47 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -998.36 | 0.07 ‰ |  |  | 2013-10-01 | Jocelyn Turnbull | 15:18 | 0.0016221 | 0.00162 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.4 |  |  | 0.00162 | 7e-05 | 5336 | 2013-10-01 | 11:37 | john west | 196736 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 6 | ... | 0.00156 | 7e-05 |  |  |  |  | 1.0003 | 1.0003 | Yes | [6 empty] | 0.00156 |  | 7e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00162 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00156 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-09-26 | jenny dahl | 15:35 |  | Completed |  |  |  |  | 63298 | 2000000 |  |  | 63190 63195 63196 63271 63299 63300 | 63190~63195~63196~63271~63299~63300 | 2891 |  | 2891 2891 2891 2891 2891 2891 | 2013-10-01 | 12:21 | christine prior |  | 0.73 mg |  | 0.73 | [5 empty] | 1 | 0.00169 | 63190 |  | XCAMS | EA | 0% |  | * no system error * two cathodes (TP63194,63325) Poisson errors multiplied by factor 1.3, because of excess run-to-run variability * of three blanks (TP63193,63286,63298) first 3 runs rejected, because of some surface contamination * of one unknown (TP63282) first 2 runs rejected, because of rising beam intensity | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2013-09-25 | Carbonate |  | JOB-196736 |  |  | 54696 |  | 14047/1 |  |  | 2013-09-11 | 820 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99987 | 0.00167 | 63195 | [26 empty] | 2013-09-20 | 716 | Gas Splitting | Completed
[168 empty] | 0.99751 | 0.00165 | 63196 | [26 empty] | 2013-09-20 | 1480 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00165 | 63271 | [26 empty] | 2013-09-23 | 3308 | Graphite | Completed
[168 empty] | 1.0029 | 0.00166 | 63299 | [26 empty] | 2013-09-26 | 1879 | C13 Measurement | Completed
[168 empty] | 0.99894 | 0.00166 | 63300 | [26 empty] | 2013-10-01 | 5336 | AMS Submission Results | Completed
[197 empty] | 2013-10-03 | 3727 | Review Final Report | Completed
[197 empty] | 2014-07-11 | 9076 | Close Job | Completed
[7 empty] | 124.1 | 88.7833 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49608 | 49608 | 49607 | 281 | 281 | 270 |  |  | jenny dahl | -997.92 | 0.07 ‰ | 2013-10-07 | 2013-10-07 | 13 | [5 empty] | 1.63‰ | 1.63‰ | 2.00‰ | 2.00‰ | 1.63 ‰ | 2.00 ‰ | AMS | 2.64‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | [6 empty] | -997.89 | 0.07 ‰ |  |  | 2013-10-09 | Jocelyn Turnbull | 9:25 | 0.0020796 | 0.00208 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.3 |  |  | 0.00208 | 7e-05 | 5418 | 2013-10-08 | 14:49 | albert zondervan | 196802 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 6 | ... | 0.002 | 7e-05 |  |  |  |  | 0.9993 | 0.9993 | Yes | [6 empty] | 0.002 |  | 7e-05 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00208 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.002 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-10-02 | jenny dahl | 16:46 |  | Completed |  |  |  |  | 63313 | 2000000 |  |  | 63380 63387 63388 63404 63405 63406 | 63380~63387~63388~63404~63405~63406 | 2893 |  | 2893 2893 2893 2893 2893 2893 | 2013-10-08 | 16:22 | christine prior |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99951 | 0.00159 | 63380 |  | XCAMS | ST | 0% |  | * 15 rotations giving OX-I average 14C counts greater than 400K * no global or system error * 2 cathodes TP63352, TP63355 required Poisson error multipliers, see comments | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-10-01 | Carbonate |  | JOB-196802 |  |  | 54735 |  | 14047/1 |  |  | 2013-09-17 | 839 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99781 | 0.00156 | 63387 | [26 empty] | 2013-09-23 | 718 | Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00158 | 63388 | [26 empty] | 2013-09-23 | 1482 | Choose Gas Splitting | Completed
[168 empty] | 0.99977 | 0.00158 | 63404 | [26 empty] | 2013-09-24 | 3323 | Graphite | Completed
[168 empty] | 1.0015 | 0.00154 | 63405 | [26 empty] | 2013-09-26 | 1880 | C13 Measurement | Completed
[168 empty] | 1.0002 | 0.00155 | 63406 | [26 empty] | 2013-10-09 | 5418 | AMS Submission Results | Completed
[197 empty] | 2013-10-09 | 3782 | Review Final Report | Completed
[197 empty] | 2014-07-11 | 9077 | Close Job | Completed
[7 empty] | 123.6 | 88.665 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48245 | 48244 | 48259 | 237 | 237 | 229 |  |  | jenny dahl | -997.54 | 0.07 ‰ | 2013-10-11 | 2013-10-11 | 10 | [5 empty] | 4.57‰ | 4.57‰ | 2.00‰ | 2.00‰ | 4.57 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -997.51 | 0.07 ‰ |  |  | 2013-10-14 | Jocelyn Turnbull | 12:50 | 0.0024643 | 0.00246 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00246 | 7e-05 | 5457 | 2013-10-14 | 10:07 | john west | 196870 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 4 | ... | 0.00237 | 7e-05 |  |  |  |  | 1.0009 | 1.0009 | Yes | [6 empty] | 0.00237 |  | 7e-05 |  | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00246 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00237 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-10-10 | jenny dahl | 15:23 |  | Completed |  |  |  |  | 63386 | 2000000 |  |  | 63414 63415 63416 63417 63461 63462 | 63414~63415~63416~63417~63461~63462 | 2894 |  | 2894 2894 2894 2894 2894 2894 | 2013-10-14 | 10:40 | christine prior |  | 1.02 mg |  | 1.02 | [5 empty] | 1.0015 | 0.00161 | 63414 |  | XCAMS | EA | 0% |  | * no error adjustments needed * 400K counts on OX-I - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-10-10 | Carbonate |  | JOB-196870 |  |  | 54816 |  | 14047/1 |  |  | 2013-09-26 | 865 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.0016 | 63415 | [26 empty] | 2013-10-01 | 3396 | Graphite | Completed
[168 empty] | 1.0002 | 0.0016 | 63416 | [26 empty] | 2013-10-03 | 1913 | C13 Measurement | Completed
[168 empty] | 0.99704 | 0.00159 | 63417 | [26 empty] | 2013-10-14 | 5457 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.0016 | 63461 | [26 empty] | 2013-10-14 | 3841 | Review Final Report | Completed
[168 empty] | 1.0006 | 0.00158 | 63462 | [26 empty] | 2014-07-11 | 9078 | Close Job | Completed
[7 empty] | 115.8 | 89.2874 | 0.0182 | RRL-UNSt-SM | Unknown Inorganic |  |  |  |  | 17731 | 17731 | 17731 | 44 | 44 | 44 | 17687.1 | 17775.2 | jenny dahl | -890 | 0.60 ‰ | 2013-11-04 | 2013-11-04 | 13 | [5 empty] | -2.34‰ | -2.34‰ | 2.00‰ | 2.00‰ | -2.34 ‰ | 2.00 ‰ | AMS | 0.72‰ | 0.2‰ | C13 | No | -19 | -19.0189 | 0 | 0.70711 | [6 empty] | -888.64 | 0.61 ‰ |  |  | 2013-11-12 | Jocelyn Turnbull | 14:32 | 0.11 | 0.11 | 0.0006 | 0.0006 | 0.0006 | 0.0006 | 4.8 |  |  | 0.11 | 0.0006 | 5623 | 2013-11-05 | 9:29 | Johannes Kaiser | 197031 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 7 | 5 | 4 | A.. | 0.10579 | 0.00058 |  |  |  |  | 1.001 | 1.001 | Yes |  |  | This is NOT a C1 marble. Looks like a mixup with the FIRI-C turbidite in this wheel. Have to manually assign blanks for this wheel from previous TWs. |  |  |  | 0.10579 |  | 0.00058 |  | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.11 | 0 | 0.0006 | 0 | 0 |  |  |  |  | 0.10579 | 0 | 0.00058 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-11-01 | jenny dahl | 12:09 |  | Completed |  |  |  |  | 63595 | 2000000 |  |  | 63427 63525 63526 63527 63528 63529 63532 63535 | 63427~63525~63526~63527~63528~63529~63532~63535 | 2898 |  | 2898 2898 2898 2898 2898 2898 2898 2898 | 2013-11-07 | 14:52 | christine prior |  | 0.42 mg |  | 0.42 | [5 empty] | 1.0005 | 0.00174 | 63427 |  | XCAMS | ST | 0% |  | * good measurement * no error adjustments needed - 29/01/2015: filled System Error (%) box | 0.7 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2013-11-04 | Carbonate |  | JOB-197031 |  | This sample has been switched with Job 197030. Tracking backwards through graphite to carbonate evolution, d13C indicates switch occurred before graphitization. % yield in CO2 evolution indicates switch probably occurred when rxn vessel was loaded or rxn vessels were switched when coming off line. | 54975 |  | 14047/1 |  |  | 2013-10-21 | 892 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0011 | 0.00175 | 63525 | [26 empty] | 2013-10-22 | 3605 | Graphite | Completed
[168 empty] | 0.99753 | 0.00173 | 63526 | [26 empty] | 2013-10-24 | 1966 | C13 Measurement | Completed
[168 empty] | 1.0024 | 0.00171 | 63527 | [26 empty] | 2013-11-12 | 5623 | AMS Submission Results | Completed
[168 empty] | 0.99978 | 0.00173 | 63528 | [26 empty] | 2013-12-24 | 6886 | Close Job | Completed
[168 empty] | 1.0007 | 0.00173 | 63529
[168 empty] | 1.0001 | 0.00189 | 63532
[168 empty] | 0.99683 | 0.00213 | 63535
[7 empty] | 128.8 | 88.7592 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50545 | 50544 | 50548 | 316 | 315 | 304 |  |  | jenny dahl | -998.15 | 0.07 ‰ | 2013-11-07 | 2013-11-07 | 6 | [5 empty] | 1.99‰ | 1.99‰ | 2.00‰ | 2.00‰ | 1.99 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -998.13 | 0.07 ‰ |  |  | 2013-11-12 | Jocelyn Turnbull | 14:37 | 0.0018508 | 0.00185 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.6 |  |  | 0.00185 | 7e-05 | 5662 | 2013-11-08 | 11:41 | albert zondervan | 197085 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 4 | ... | 0.00178 | 7e-05 |  |  |  |  | 0.99939 | 0.99939 | Yes | [6 empty] | 0.00178 |  | 7e-05 |  | No | 1 | 0.99999 | 0.00092 | 0.00092 | 0.00185 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00178 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-11-01 | jenny dahl | 16:42 |  | Completed |  |  |  |  | 63694 | 2000000 |  |  | 63507 63508 63609 63610 63620 63692 | 63507~63508~63609~63610~63620~63692 | 2899 |  | 2899 2899 2899 2899 2899 2899 | 2013-11-08 | 12:00 | christine prior |  | 0.75 mg |  | 0.75 | [5 empty] | 0.99892 | 0.00225 | 63507 |  | XCAMS | EA | 0.13% |  | * small system error * no other error adjustments | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2013-11-13 | Carbonate | Not Complete | JOB-197085 |  |  | 55013 |  | 14047/1 |  |  | 2013-10-31 | 906 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99643 | 0.00225 | 63508 | [26 empty] | 2013-11-01 | 3704 | Graphite | Completed
[168 empty] | 1.0007 | 0.00218 | 63609 | [26 empty] | 2013-11-04 | 2007 | C13 Measurement | Completed
[168 empty] | 1.0026 | 0.00224 | 63610 | [26 empty] | 2013-11-12 | 5662 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00225 | 63620 | [26 empty] | 2013-12-24 | 6887 | Close Job | Completed
[168 empty] | 0.9993 | 0.0023 | 63692
[7 empty] | 129.6 | 88.6598 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50021 | 50020 | 50003 | 296 | 295 | 284 |  |  | jenny dahl | -998.02 | 0.07 ‰ | 2013-11-07 | 2013-11-07 | 6 | [5 empty] | 3.11‰ | 3.11‰ | 2.00‰ | 2.00‰ | 3.11 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -998 | 0.07 ‰ |  |  | 2013-11-12 | Jocelyn Turnbull | 14:37 | 0.0019756 | 0.00198 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.2 |  |  | 0.00198 | 7e-05 | 5663 | 2013-11-08 | 11:41 | albert zondervan | 197086 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 4 | ... | 0.0019 | 7e-05 |  |  |  |  | 0.99939 | 0.99939 | Yes | [6 empty] | 0.0019 |  | 7e-05 |  | No | 1 | 0.99999 | 0.00092 | 0.00092 | 0.00198 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0019 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-11-01 | jenny dahl | 16:42 |  | Completed |  |  |  |  | 63695 | 2000000 |  |  | 63507 63508 63609 63610 63620 63692 | 63507~63508~63609~63610~63620~63692 | 2899 |  | 2899 2899 2899 2899 2899 2899 | 2013-11-08 | 12:00 | christine prior |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99892 | 0.00225 | 63507 |  | XCAMS | EA | 0.13% |  | * small system error * no other error adjustments | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-11-13 | Carbonate |  | JOB-197086 |  |  | 55014 |  | 14047/1 |  |  | 2013-10-31 | 907 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99643 | 0.00225 | 63508 | [26 empty] | 2013-11-01 | 3705 | Graphite | Completed
[168 empty] | 1.0007 | 0.00218 | 63609 | [26 empty] | 2013-11-04 | 2008 | C13 Measurement | Completed
[168 empty] | 1.0026 | 0.00224 | 63610 | [26 empty] | 2013-11-12 | 5663 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00225 | 63620 | [26 empty] | 2013-12-24 | 6888 | Close Job | Completed
[168 empty] | 0.9993 | 0.0023 | 63692
[7 empty] | 127.2 | 88.9327 | 0.0006 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 45959 | 45959 | 45948 | 204 | 204 | 196 |  |  | jenny dahl | -996.72 | 0.08 ‰ | 2013-11-22 | 2013-11-22 | 7 | [5 empty] | 2.18‰ | 2.18‰ | 2.00‰ | 2.00‰ | 2.18 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -996.68 | 0.08 ‰ |  |  | 2013-11-25 | Jocelyn Turnbull | 15:13 | 0.0032754 | 0.00328 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.6 |  |  | 0.00328 | 8e-05 | 5800 | 2013-11-25 | 10:17 | john west | 197174 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 4 | 5 | 4 | ... | 0.00315 | 8e-05 |  |  |  |  | 1.0015 | 1.0015 | Yes |  |  | Unusually young for a marble blank, but included because evolved same day as C2 travertine and both forams in this wheel. CO2 yield in evolution was lower than typical, though, so averaged in some recent marble blanks that might be more representative of carbonate background. |  |  |  | 0.00315 |  | 8e-05 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00328 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00315 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-11-20 | jenny dahl | 15:06 |  | Completed |  |  |  |  | 63802 | 2000000 |  |  | 63769 63851 63855 63856 63857 63858 | 63769~63851~63855~63856~63857~63858 | 2903 |  | 2903 2903 2903 2903 2903 2903 | 2013-11-25 | 13:14 | christine prior |  | 0.75 mg |  | 0.75 | [5 empty] | 1.0014 | 0.00154 | 63769 |  | XCAMS | ST | 0% |  | good measurement. no error adjustments needed. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2013-11-26 | Carbonate |  | JOB-197174 |  |  | 55164 |  | 14047/1 |  |  | 2013-11-13 | 912 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99844 | 0.00157 | 63851 | [26 empty] | 2013-11-15 | 3812 | Graphite | Completed
[168 empty] | 0.99933 | 0.00155 | 63855 | [26 empty] | 2013-11-21 | 2032 | C13 Measurement | Completed
[168 empty] | 0.99974 | 0.00156 | 63856 | [26 empty] | 2013-11-25 | 5800 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00157 | 63857 | [26 empty] | 2013-12-24 | 6920 | Close Job | Completed
[168 empty] | 1 | 0.00155 | 63858
[7 empty] | 138.7 | 88.6315 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48591 | 48591 | 48592 | 283 | 283 | 272 |  |  | Kelly Lyons | -997.64 | 0.08 ‰ | 2013-11-25 | 2013-11-25 | 5 | [5 empty] | 3.94‰ | 3.94‰ | 2.00‰ | 2.00‰ | 3.94 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | [6 empty] | -997.61 | 0.08 ‰ |  |  | 2013-11-26 | Jocelyn Turnbull | 11:56 | 0.0023603 | 0.00236 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.00236 | 8e-05 | 5842 | 2013-11-26 | 9:17 | john west | 197177 | [5 empty] | 5 | 5/07/22 | 15:24:49 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 6 | ... | 0.00227 | 8e-05 |  |  |  |  | 0.9999 | 0.9999 | Yes | [6 empty] | 0.00227 |  | 8e-05 |  | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00236 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00227 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-11-21 | Kelly Lyons | 14:56 |  | Completed |  |  |  |  | 63847 | 2000000 |  |  | 63865 63866 63867 63868 63876 63877 | 63865~63866~63867~63868~63876~63877 | 2904 |  | 2904 2904 2904 2904 2904 2904 | 2013-11-26 | 9:50 | christine prior |  | 0.94 mg |  | 0.94 | [5 empty] | 0.99901 | 0.00168 | 63865 |  | XCAMS | EA | 0% |  | - one OXI (TP63876) required an error factor of 1.3 to account for run - to - run variability. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-11-26 | Carbonate |  | JOB-197177 |  |  | 55202 |  | 14047/1 |  |  | 2013-11-13 | 915 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0023 | 0.00171 | 63866 | [26 empty] | 2013-11-19 | 771 | Gas Splitting | Completed
[168 empty] | 0.9997 | 0.00165 | 63867 | [26 empty] | 2013-11-19 | 1572 | Choose Gas Splitting | Completed
[168 empty] | 0.99921 | 0.0017 | 63868 | [26 empty] | 2013-11-20 | 3857 | Graphite | Completed
[168 empty] | 1.0008 | 0.00222 | 63876 | [26 empty] | 2013-11-25 | 2047 | C13 Measurement | Completed
[168 empty] | 0.99939 | 0.0017 | 63877 | [26 empty] | 2013-11-26 | 5842 | AMS Submission Results | Completed
[197 empty] | 2013-12-24 | 6861 | Close Job | Completed
[7 empty] | 105 | 89.2187 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50191 | 50191 | 50208 | 302 | 302 | 291 |  |  | Kelly Lyons | -998.07 | 0.07 ‰ | 2013-12-20 | 2013-12-20 | 16 | [5 empty] | 2.56‰ | 2.56‰ | 2.00‰ | 2.00‰ | 2.56 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | [6 empty] | -998.04 | 0.07 ‰ |  |  | 2013-12-23 | Christine Prior | 13:58 | 0.001934 | 0.00193 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00193 | 7e-05 | 6038 | 2013-12-21 | 12:47 | Albert Zondervan | 197589 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 37 | 5 | 4 | ... | 0.00186 | 7e-05 |  |  |  |  | 1.0051 | 1.0051 | Yes | [6 empty] | 0.00186 |  | 7e-05 |  | Yes | 1 | 0.99999 | 0.0007 | 0.0007 | 0.00193 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.00186 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-12-17 | Kelly Lyons | 14:29 |  | Completed |  |  |  |  | 63965 | 2000000 |  |  | 63959 63967 63968 63969 64008 64009 | 63959~63967~63968~63969~64008~64009 | 2909 |  | 2909 2909 2909 2909 2909 2909 | 2013-12-23 | 16:18 | Christine Prior |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0003 | 0.00168 | 63959 |  | XCAMS | Both | 0% |  | * good measurement * 14C results of 4x non-unknown cathodes do not match with expectation (this has been communicated with RRL team) * all other primary and secondary standards appear to be in the right position * it needs checking if unknowns have been misplaced, by comparing their AMS results (13/12, rts) with expected values * none of the identifiers have been changed by XCAMS team | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-12-16 | Carbonate | Complete | JOB-197589 |  |  | 55360 |  | 14047/1 |  |  | 2013-12-02 | 918 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0011 | 0.0017 | 63967 | [26 empty] | 2013-12-04 | 3975 | Graphite | Completed
[168 empty] | 1.0011 | 0.0017 | 63968 | [26 empty] | 2013-12-09 | 2081 | C13 Measurement | Completed
[168 empty] | 1.0012 | 0.00169 | 63969 | [26 empty] | 2013-12-23 | 6038 | AMS Submission Results | Completed
[168 empty] | 0.99761 | 0.00176 | 64008 | [26 empty] | 2014-05-06 | 8100 | Close Job | Completed
[168 empty] | 0.99841 | 0.00173 | 64009
[7 empty] | 120.2 | 88.8877 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47470 | 47470 | 47482 | 277 | 276 | 267 |  |  | Kelly Lyons | -997.29 | 0.09 ‰ | 2014-01-28 | 2014-01-28 | 14 | [5 empty] | 1.17‰ | 1.17‰ | 2.00‰ | 2.00‰ | 1.17 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | [6 empty] | -997.25 | 0.09 ‰ |  |  | 2014-02-03 | Jocelyn Turnbull | 12:08 | 0.0027139 | 0.00271 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.8 |  |  | 0.00271 | 9e-05 | 6200 | 2014-01-29 | 10:04 | Johannes Kaiser | 197870 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 4 | ... | 0.00261 | 9e-05 |  |  |  |  | 1 | 1 | Yes | [6 empty] | 0.00261 |  | 9e-05 |  | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00271 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00261 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-01-27 | Kelly Lyons | 10:59 |  | Completed |  |  |  |  | 64168 | 2000000 |  |  | 64190 64199 64200 64207 64208 64259 | 64190~64199~64200~64207~64208~64259 | 2913 |  | 2913 2913 2913 2913 2913 2913 | 2014-01-29 | 15:53 | christine prior |  | 0.69 mg |  | 0.69 | [5 empty] | 1.0003 | 0.00176 | 64190 |  | XCAMS | ST | 0% |  | - good measurement - pos#5 and 29 needed an additional factor to account for individual run-to-run variance - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-01-27 | Carbonate | Not Complete | JOB-197870 |  |  | 55525 |  | 14047/1 |  |  | 2014-01-13 | 935 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00177 | 64199 | [26 empty] | 2014-01-14 | 4178 | Graphite | Completed
[168 empty] | 0.99641 | 0.00178 | 64200 | [26 empty] | 2014-01-16 | 2164 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00176 | 64207 | [26 empty] | 2014-02-03 | 6200 | AMS Submission Results | Completed
[168 empty] | 1.0025 | 0.00178 | 64208 | [26 empty] | 2014-05-06 | 8046 | Close Job | Completed
[168 empty] | 1 | 0.00176 | 64259
[7 empty] | 111.5 | 89.3857 | 0.0007 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 44431 | 44431 | 44435 | 274 | 274 | 284 |  |  | Kelly Lyons | -996.04 | 0.14 ‰ | 2014-02-11 | 2014-02-11 | 19 | [5 empty] | 2.29‰ | 2.29‰ | 2.00‰ | 2.00‰ | 2.29 ‰ | 2.00 ‰ | AMS | 2.64‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | [6 empty] | -995.99 | 0.14 ‰ |  |  | 2014-02-19 | Jocelyn Turnbull | 16:32 | 0.0039616 | 0.00396 | 0.00014 | 0.00014 | 0.00014 | 0.00013 | 1.8 |  |  | 0.00396 | 0.00014 | 6212 | 2014-02-12 | 11:08 | Johannes Kaiser | 197590 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 5 | 5 | 4 | .X. | 0.00381 | 0.00013 |  |  |  |  | 1.0068 | 1.0068 | Yes |  |  | -first 4 runs are rejected - considerably younger than marble average, but this wheel was in and out of source several times & exposed to air. |  |  |  | 0.00381 |  | 0.00013 |  | No | 1 | 0.99999 | 0.00107 | 0.00107 | 0.00396 | 0 | 0.00014 | 0 | 0 |  |  |  |  | 0.00381 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-01-29 | Kelly Lyons | 14:43 |  | Completed |  |  |  |  | 64249 | 2000000 |  |  | 64260 64261 64262 64266 64270 64284 | 64260~64261~64262~64266~64270~64284 | 2914 |  | 2914 2914 2914 2914 2914 2914 | 2014-02-12 | 14:29 | Christine Prior |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99726 | 0.00266 | 64260 |  | XCAMS | ST | 0.19% |  | Mounted this wheel on Thursday the 31st of January. But had to abort tuning due to erratic cathode current (all samples have had contact with Caesium during the cooking round before tuning was started). We opened up the ion source 4 times to change pieces to get rid off the unwanted behaviour but were not successful. Those times the wheel was stored under vacuum and has not been in contact with ambient air. On Friday the 7th of February we had to take the wheel out of the ion source, has seen air!! - Kapuni in pos#36 has been used for warming up and searching for new ion source settings, still measured though - C-2 in pos#10 has been used to partially tune the system with, still measured though - had to apply a factor of 1.04 to account for the global run-to-run variance - system error = 0.19% had to reject many first runs of the blanks and low counting unknowns as there was a clear surface clean-up effect visible | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2013-12-16 | Carbonate |  | JOB-197590 |  | cracked 22/1/14 | 55600 |  | 14047/1 |  |  | 2013-12-02 | 919 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0036 | 0.00263 | 64261 | [26 empty] | 2014-01-23 | 4259 | Graphite | Completed
[168 empty] | 1.0027 | 0.00263 | 64262 | [26 empty] | 2014-01-24 | 2208 | C13 Measurement | Completed
[168 empty] | 0.9993 | 0.00264 | 64266 | [26 empty] | 2014-02-19 | 6212 | AMS Submission Results | Completed
[168 empty] | 0.99942 | 0.00258 | 64270 | [26 empty] | 2014-05-06 | 8028 | Close Job | Completed
[168 empty] | 0.9976 | 0.0026 | 64284
[7 empty] | 121.9 | 89.5354 | 0.0006 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 45833 | 45832 | 45827 | 176 | 175 | 169 |  |  | jenny dahl | -996.67 | 0.07 ‰ | 2014-02-20 | 2014-02-20 | 24 | [5 empty] | 0.97‰ | 0.97‰ | 2.00‰ | 2.00‰ | 0.97 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -996.63 | 0.07 ‰ |  | Christine Prior | 2014-02-25 | Jocelyn Turnbull | 14:23 | 0.0033273 | 0.00333 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00333 | 7e-05 | 6285 | 2014-02-24 | 10:54 | Johannes Kaiser | 198062 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 32 | 5 | 4 | .D. | 0.0032 | 7e-05 |  |  |  |  | 1.0071 | 1.0071 | Yes |  |  | not a great blank, but C2 result significantly younger as well. |  |  |  | 0.0032 |  | 7e-05 |  | Yes | 1 | 1 | 0.0005 | 0.0005 | 0.00333 | 0 | 7e-05 | 0 | 0 |  |  |  |  | 0.0032 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-02-12 | jenny dahl | 9:15 |  | Completed |  |  |  |  | 64288 | 2000000 |  |  | 64285 64286 64287 64337 64338 64339 | 64285~64286~64287~64337~64338~64339 | 2915 |  | 2915 2915 2915 2915 2915 2915 | 2014-02-24 | 15:09 | Christine Prior |  | 1.07 mg |  | 1.07 | [5 empty] | 1.0006 | 0.00123 | 64285 |  | XCAMS | ST | 0% |  | - first measurement after Pelletron maintenance - added a factor of 1.02 to explain the global run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-02-07 | Carbonate |  | JOB-198062 |  | C13 | 55637 |  | 14047/1 |  |  | 2014-01-24 | 977 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99927 | 0.00123 | 64286 | [26 empty] | 2014-01-27 | 4298 | Graphite | Completed
[168 empty] | 0.99766 | 0.00124 | 64287 | [26 empty] | 2014-01-28 | 2238 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00124 | 64337 | [26 empty] | 2014-02-25 | 6285 | AMS Submission Results | Completed
[168 empty] | 1.0009 | 0.00121 | 64338 | [26 empty] | 2014-05-06 | 8024 | Close Job | Completed
[168 empty] | 1.0002 | 0.00124 | 64339
[7 empty] | 123.1 | 88.5553 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47501 | 47500 | 47511 | 278 | 278 | 268 |  |  | jenny dahl | -997.3 | 0.09 ‰ | 2014-02-27 | 2014-02-27 | 31 | [5 empty] | 9.05‰ | 9.05‰ | 2.00‰ | 2.00‰ | 9.05 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.89443 | [6 empty] | -997.26 | 0.09 ‰ |  |  | 2014-03-04 | Jocelyn Turnbull | 16:06 | 0.0027035 | 0.0027 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.0027 | 9e-05 | 6320 | 2014-03-03 | 8:56 | Albert Zondervan | 198063 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 4 | 5 | 4 | ... | 0.0026 | 9e-05 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.0026 |  | 9e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.0027 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.0026 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-02-19 | jenny dahl | 10:13 |  | Completed |  |  |  |  | 64289 | 2000000 |  |  | 64336 64340 64439 64444 64445 | 64336~64340~64439~64444~64445 | 2917 |  | 2917 2917 2917 2917 2917 | 2014-03-03 | 10:33 | Christine Prior |  | 1.10 mg |  | 1.1 | [5 empty] | 1.002 | 0.00163 | 64336 |  | XCAMS | ST | 0% |  | * Measurement was delayed due to fluctuations in cooling water. Although not fully resolved, control has been improved such that measurements could be started. * The wheel was measured twice, with a retune in between, because of poor repeatability amongst Ox-I standards. While the retune has improved things, Ox-I in cathode position 0 is still too low by 2.5 sigma and was therefore rejected as standard. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-02-07 | Carbonate |  | JOB-198063 |  |  | 55719 |  | 14047/1 |  |  | 2014-01-24 | 978 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99862 | 0.00165 | 64340 | [26 empty] | 2014-01-27 | 4299 | Graphite | Completed
[168 empty] | 0.99995 | 0.00166 | 64439 | [26 empty] | 2014-01-28 | 2239 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00166 | 64444 | [26 empty] | 2014-03-04 | 6320 | AMS Submission Results | Completed
[168 empty] | 0.99943 | 0.00164 | 64445 | [26 empty] | 2014-05-06 | 7993 | Close Job | Completed
[7 empty] | 112.2 | 89.0852 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50366 | 50366 | 50376 | 397 | 397 | 383 |  |  | jenny dahl | -998.11 | 0.09 ‰ | 2014-02-28 | 2014-02-28 | 31 | [5 empty] | 3.05‰ | 3.05‰ | 2.00‰ | 2.00‰ | 3.05 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.89443 | -998.11 ‰ | -998.11 ‰ | 0.09 | 0.09 | -998.08 | 0.09 ‰ | -998.08 | 0.09 ‰ |  |  | 2014-03-04 | Jocelyn Turnbull | 16:06 | 0.0018924 | 0.00189 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00189 | 9e-05 | 6321 | 2014-03-03 | 8:56 | Albert Zondervan | 198067 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.19 | 0.01% | 10 | 5 | 4 | .X. | 0.00182 | 9e-05 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.00182 |  | 9e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00189 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00182 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-02-19 | jenny dahl | 10:13 |  | Completed |  |  |  |  | 64303 | 2000000 |  |  | 64336 64340 64439 64444 64445 | 64336~64340~64439~64444~64445 | 2917 |  | 2917 2917 2917 2917 2917 | 2014-03-03 | 10:33 | Christine Prior |  | 0.97 mg |  | 0.97 | [5 empty] | 1.002 | 0.00163 | 64336 |  | XCAMS | ST | 0% |  | * Measurement was delayed due to fluctuations in cooling water. Although not fully resolved, control has been improved such that measurements could be started. * The wheel was measured twice, with a retune in between, because of poor repeatability amongst Ox-I standards. While the retune has improved things, Ox-I in cathode position 0 is still too low by 2.5 sigma and was therefore rejected as standard. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-03-24 | Water | Complete | JOB-198067 | 54050101 $450 | Water line evacuated like carbonate. | 55729 | Normal | 14047/11 |  |  | 2014-01-27 | 982 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99862 | 0.00165 | 64340 | [26 empty] | 2014-01-28 | 4313 | Graphite | Completed
[168 empty] | 0.99995 | 0.00166 | 64439 | [26 empty] | 2014-01-31 | 2247 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00166 | 64444 | [26 empty] | 2014-03-04 | 6321 | AMS Submission Results | Completed
[168 empty] | 0.99943 | 0.00164 | 64445 | [26 empty] | 2014-04-30 | 7780 | Close Job | Completed
[197 empty] | 2014-04-30 | 387 | Invoice | Completed
[7 empty] | 121.3 | 89.1223 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49689 | 49689 | 49685 | 406 | 405 | 390 |  |  | jenny dahl | -997.94 | 0.10 ‰ | 2014-03-03 | 2014-03-03 | 26 | [5 empty] | 5.01‰ | 5.01‰ | 2.00‰ | 2.00‰ | 5.01 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | [6 empty] | -997.92 | 0.11 ‰ |  |  | 2014-03-04 | Jocelyn Turnbull | 15:12 | 0.0020588 | 0.00206 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00206 | 0.0001 | 6448 | 2014-03-04 | 13:23 | Albert Zondervan | 198136 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 30 | 5 | 4 | ... | 0.00198 | 0.0001 |  |  |  |  | 1.0065 | 1.0065 | Yes |  |  | * Poisson uncertainty multiplied by factor 1.4 to account for additional run-to-run variance |  |  |  | 0.00198 |  | 0.0001 |  | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00206 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.00198 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-02-24 | jenny dahl | 16:30 |  | Completed |  |  |  |  | 64390 | 2000000 |  |  | 64446 64493 64509 64510 64511 64512 | 64446~64493~64509~64510~64511~64512 | 2918 |  | 2918 2918 2918 2918 2918 2918 | 2014-03-04 | 14:17 | Christine Prior |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0011 | 0.00168 | 64446 |  | XCAMS | ST | 0% |  | * straightforward measurement * 2 cathodes with uncertainty multipliers: TP64450, TP64390. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-02-18 | Carbonate | Complete | JOB-198136 |  |  | 55715 |  | 14047/1 |  |  | 2014-02-04 | 1024 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00174 | 64493 | [26 empty] | 2014-02-05 | 4400 | Graphite | Completed
[168 empty] | 1.001 | 0.00179 | 64509 | [26 empty] | 2014-02-13 | 2303 | C13 Measurement | Completed
[168 empty] | 0.99946 | 0.00179 | 64510 | [26 empty] | 2014-03-04 | 6448 | AMS Submission Results | Completed
[168 empty] | 0.99832 | 0.00178 | 64511 | [26 empty] | 2014-05-06 | 7982 | Close Job | Completed
[168 empty] | 0.99922 | 0.00177 | 64512
[7 empty] | 146.2 | 89.1551 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 51105 | 51105 | 51087 | 290 | 290 | 279 |  |  | jenny dahl | -998.27 | 0.06 ‰ | 2014-03-14 | 2014-03-14 | 23 | [5 empty] | 3.62‰ | 3.62‰ | 2.00‰ | 2.00‰ | 3.62 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -998.25 | 0.06 ‰ |  |  | 2014-03-18 | Jocelyn Turnbull | 9:50 | 0.0017261 | 0.00173 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2 |  |  | 0.00173 | 6e-05 | 6477 | 2014-03-16 | 9:14 | Albert Zondervan | 198224 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 17 | 5 | 5 | ... | 0.00166 | 6e-05 |  |  |  |  | 1.0055 | 1.0055 | Yes |  |  | hurrah |  |  |  | 0.00166 |  | 6e-05 |  | No | 1 | 0.99998 | 0.00077 | 0.00077 | 0.00173 | 0 | 6e-05 | 0 | 0 |  |  |  |  | 0.00166 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-04 | jenny dahl | 15:46 |  | Completed |  |  |  |  | 64474 | 2000000 |  |  | 64562 64577 64578 64582 64609 64613 | 64562~64577~64578~64582~64609~64613 | 2919 |  | 2919 2919 2919 2919 2919 2919 | 2014-03-17 | 9:53 | Christine Prior |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99984 | 0.00187 | 64562 |  | XCAMS | ST | 0.11% |  | * good measurement * error corrections made for excess run-to-run variance and excess pos-to-pos variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-03-03 | Carbonate | Not Complete | JOB-198224 |  |  | 55838 |  | 14047/1 |  |  | 2014-02-17 | 2251 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99773 | 0.00192 | 64577 | [26 empty] | 2014-02-17 | 1076 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99951 | 0.00186 | 64578 | [26 empty] | 2014-02-19 | 4484 | Graphite | Completed
[168 empty] | 0.99834 | 0.00189 | 64582 | [26 empty] | 2014-02-24 | 2388 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00187 | 64609 | [26 empty] | 2014-03-18 | 6477 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00185 | 64613 | [26 empty] | 2014-05-06 | 7968 | Close Job | Completed
[7 empty] | 108.7 | 88.892 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 46271 | 46271 | 46273 | 239 | 238 | 230 |  |  | jenny dahl | -996.85 | 0.09 ‰ | 2014-03-17 | 2014-03-17 | 11 | [5 empty] | 4.92‰ | 4.92‰ | 2.00‰ | 2.00‰ | 4.92 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.89443 | [6 empty] | -996.81 | 0.09 ‰ |  |  | 2014-03-20 | Jocelyn Turnbull | 8:41 | 0.0031506 | 0.00315 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00315 | 9e-05 | 6537 | 2014-03-18 | 12:55 | John West | 198087 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 25 | 5 | 6 | .S. | 0.00303 | 9e-05 |  |  |  |  | 1.0038 | 1.0038 | Yes | [6 empty] | 0.00303 |  | 9e-05 |  | Yes | 1 | 1 | 0.00102 | 0.00102 | 0.00315 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00303 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-10 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 64650 | 2000000 |  |  | 64615 64625 64627 64634 64645 | 64615~64625~64627~64634~64645 | 2921 |  | 2921 2921 2921 2921 2921 | 2014-03-19 | 16:52 | Christine Prior |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99782 | 0.00227 | 64615 |  | XCAMS | ST | 0.14% |  | * one calibration Ox-I (TP64616) rejected as its rts is a 2.5 sigma outlier * further error adjustments made: - global factor 1.01 to account for excess run-to-run variance - factor 1.3 on cat#3 (TP64666) to account for excess run-to-run variance - global system error 0.14% to account for excess cat-to-cat variance on remaining 5 calibration Ox-Is * The two EA-combusted Ox-Is agree with these 5 calibration Ox-Is, in terms of rts. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-02-12 | Carbonate |  | JOB-198087 |  | seal for next carbonate water wheel | 55877 |  | 14047/1 |  |  | 2014-01-30 | 1003 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99965 | 0.00235 | 64625 | [26 empty] | 2014-03-04 | 904 | Gas Splitting | Completed
[168 empty] | 1.0031 | 0.0023 | 64627 | [26 empty] | 2014-03-04 | 1848 | Choose Gas Splitting | Completed
[168 empty] | 1.0014 | 0.00225 | 64634 | [26 empty] | 2014-03-06 | 4660 | Graphite | Completed
[168 empty] | 0.99797 | 0.00229 | 64645 | [26 empty] | 2014-03-10 | 2477 | C13 Measurement | Completed
[197 empty] | 2014-03-20 | 6537 | AMS Submission Results | Completed
[197 empty] | 2014-05-06 | 7940 | Close Job | Completed
[7 empty] | 108.1 | 88.9283 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48382 | 48381 | 48391 | 379 | 379 | 365 |  |  | jenny dahl | -997.58 | 0.11 ‰ | 2014-03-17 | 2014-03-17 | 10 | [5 empty] | 4.51‰ | 4.51‰ | 2.00‰ | 2.00‰ | 4.51 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.89443 | -997.58 ‰ | -997.58 ‰ | 0.11 | 0.11 | -997.55 | 0.12 ‰ | -997.55 | 0.12 ‰ |  |  | 2014-03-20 | Jocelyn Turnbull | 8:41 | 0.0024227 | 0.00242 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.3 |  |  | 0.00242 | 0.00011 | 6570 | 2014-03-18 | 12:55 | John West | 198394 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.24 | 0.01% | 3 | 5 | 4 | .X. | 0.00233 | 0.00011 |  |  |  |  | 1.0038 | 1.0038 | Yes | [6 empty] | 0.00233 |  | 0.00011 |  | Yes | 1 | 1 | 0.00102 | 0.00102 | 0.00242 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00233 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-10 | jenny dahl | 13:59 |  | Completed |  |  |  |  | 64666 | 2000000 |  |  | 64615 64625 64627 64634 64645 | 64615~64625~64627~64634~64645 | 2921 |  | 2921 2921 2921 2921 2921 | 2014-03-19 | 16:52 | Christine Prior |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99782 | 0.00227 | 64615 |  | XCAMS | ST | 0.14% |  | * one calibration Ox-I (TP64616) rejected as its rts is a 2.5 sigma outlier * further error adjustments made: - global factor 1.01 to account for excess run-to-run variance - factor 1.3 on cat#3 (TP64666) to account for excess run-to-run variance - global system error 0.14% to account for excess cat-to-cat variance on remaining 5 calibration Ox-Is * The two EA-combusted Ox-Is agree with these 5 calibration Ox-Is, in terms of rts. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-04-30 | Water | Complete | JOB-198394 |  | Water line evacuated like carbonate. | 55888 | Normal | 14047/11 |  |  | 2014-03-05 | 1152 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00235 | 64625 | [26 empty] | 2014-03-07 | 4676 | Graphite | Completed
[168 empty] | 1.0031 | 0.0023 | 64627 | [26 empty] | 2014-03-13 | 2490 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00225 | 64634 | [26 empty] | 2014-03-20 | 6570 | AMS Submission Results | Completed
[168 empty] | 0.99797 | 0.00229 | 64645 | [26 empty] | 2014-05-06 | 7949 | Close Job | Completed
[7 empty] | 61.5 | 90.4914 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48011 | 48011 | 48002 | 362 | 362 | 348 |  |  | jenny dahl | -997.46 | 0.11 ‰ | 2014-04-03 | 2014-04-03 | 21 | [5 empty] | -2.38‰ | -2.38‰ | 2.00‰ | 2.00‰ | -2.38 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | [6 empty] | -997.43 | 0.12 ‰ |  | Christine Prior | 2014-04-07 | Jocelyn Turnbull | 13:29 | 0.0025371 | 0.00254 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.7 |  |  | 0.00254 | 0.00011 | 6622 | 2014-04-07 | 11:05 | John West | 198169 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 6 | ... | 0.00244 | 0.00011 |  |  |  |  | 1.0145 | 1.0145 | Yes | [6 empty] | 0.00244 |  | 0.00011 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00254 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00244 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-17 | jenny dahl | 14:17 |  | Completed |  |  |  |  | 64716 | 2000000 |  |  | 64705 64738 64739 64740 64741 64742 | 64705~64738~64739~64740~64741~64742 | 2924 |  | 2924 2924 2924 2924 2924 2924 | 2014-04-07 | 12:03 | Christine Prior |  | 0.77 mg |  | 0.77 | [5 empty] | 1.0016 | 0.00206 | 64705 |  | XCAMS | ST | 0% |  | * good measurement, despite low currents * all runs of 1st rotation rejected as these were associated with voltage-clicks in SNICS | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-02-24 | Carbonate |  | JOB-198169 |  | Sealed for when needed - jessicas training | 55995 |  | 14047/1 |  |  | 2014-02-10 | 1034 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99893 | 0.00203 | 64738 | [26 empty] | 2014-03-12 | 910 | Gas Splitting | Completed
[168 empty] | 0.99977 | 0.00202 | 64739 | [26 empty] | 2014-03-12 | 1862 | Choose Gas Splitting | Completed
[168 empty] | 0.99929 | 0.00204 | 64740 | [26 empty] | 2014-03-13 | 4726 | Graphite | Completed
[168 empty] | 1.0004 | 0.00203 | 64741 | [26 empty] | 2014-03-18 | 2516 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00205 | 64742 | [26 empty] | 2014-04-07 | 6622 | AMS Submission Results | Completed
[197 empty] | 2014-04-07 | 4601 | Review Final Report | Completed
[197 empty] | 2014-07-11 | 9081 | Close Job | Completed
[7 empty] | 79.6 | 90.1402 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46651 | 46651 | 46638 | 306 | 305 | 294 |  |  | jenny dahl | -996.99 | 0.11 ‰ | 2014-04-03 | 2014-04-03 | 24 | [5 empty] | 1.51‰ | 1.51‰ | 2.00‰ | 2.00‰ | 1.51 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.001 | 0 | 0.8165 | -996.99 ‰ | -996.99 ‰ | 0.11 | 0.11 | -996.96 | 0.12 ‰ | -996.96 | 0.12 ‰ |  |  | 2014-04-07 | Jocelyn Turnbull | 13:29 | 0.003005 | 0.00301 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.2 |  |  | 0.00301 | 0.00011 | 6641 | 2014-04-07 | 11:05 | John West | 198407 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.3 | 0.01% | 25 | 5 | 3 | .X. | 0.00289 | 0.00011 |  |  |  |  | 1.0145 | 1.0145 | Yes |  |  | water line blanks have been much better, |  |  |  | 0.00289 |  | 0.00011 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00301 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00289 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-17 | jenny dahl | 14:18 |  | Completed |  |  |  |  | 64671 | 2000000 |  |  | 64705 64738 64739 64740 64741 64742 | 64705~64738~64739~64740~64741~64742 | 2924 |  | 2924 2924 2924 2924 2924 2924 | 2014-04-07 | 12:03 | Christine Prior |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0016 | 0.00206 | 64705 |  | XCAMS | ST | 0% |  | * good measurement, despite low currents * all runs of 1st rotation rejected as these were associated with voltage-clicks in SNICS |  | RRL | [5 empty] | Background Inorganic |  | 2014-05-01 | Water | Complete | JOB-198407 |  | Water line evacuated like carbonate. | 56011 | Normal | 14047/11 |  |  | 2014-03-06 | 1163 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99893 | 0.00203 | 64738 | [26 empty] | 2014-03-10 | 4681 | Graphite | Completed
[168 empty] | 0.99977 | 0.00202 | 64739 | [26 empty] | 2014-04-07 | 6641 | AMS Submission Results | Completed
[168 empty] | 0.99929 | 0.00204 | 64740 | [26 empty] | 2014-05-06 | 7920 | Close Job | Completed
[168 empty] | 1.0004 | 0.00203 | 64741
[168 empty] | 1.0001 | 0.00205 | 64742
[7 empty] | 103.6 | 89.1055 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48144 | 48143 | 48129 | 335 | 334 | 321 |  |  | Kelly lyons | -997.5 | 0.10 ‰ | 2014-04-01 | 2014-04-01 | 14 | [5 empty] | 2.76‰ | 2.76‰ | 2.00‰ | 2.00‰ | 2.76 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0015 | 0 | 0.8165 | -997.50 ‰ | -997.50 ‰ | 0.1 | 0.1 | -997.47 | 0.11 ‰ | -997.47 | 0.11 ‰ |  |  | 2014-04-02 | Jocelyn Turnbull | 14:24 | 0.0024955 | 0.0025 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.0025 | 0.0001 | 6686 | 2014-04-02 | 9:58 | Johannes Kaiser | 198501 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.25 | 0.01% | 24 | 5 | 4 | .X. | 0.0024 | 0.0001 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.0024 |  | 0.0001 |  | No | 1 | 1 | 0.00121 | 0.00121 | 0.0025 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.0024 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-21 | Kelly lyons | 13:24 |  | Completed |  |  |  |  | 64761 | 2000000 |  |  | 64706 64707 64723 64797 64803 64804 | 64706~64707~64723~64797~64803~64804 | 2925 |  | 2925 2925 2925 2925 2925 2925 | 2014-04-02 | 14:05 | Christine Prior |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0014 | 0.00291 | 64706 |  | XCAMS | ST | 0.22% |  | -first measurement after SNICS problems - needed to add a global run-to-run variance factor of 1.06 - Ox-I (TP64797) in pos#0 was high by 4 per mill -Ox-I (TP64707) in pos#14 was low by 4 per mill | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-05-12 | Carbonate |  | JOB-198501 |  | Water line evacuated like carbonate. | 55966 | Normal | 14047/11 |  |  | 2014-03-17 | 1188 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99591 | 0.00291 | 64707 | [26 empty] | 2014-03-18 | 4771 | Graphite | Completed
[168 empty] | 0.99915 | 0.00297 | 64723 | [26 empty] | 2014-03-19 | 2548 | C13 Measurement | Completed
[168 empty] | 1.0041 | 0.003 | 64797 | [26 empty] | 2014-04-02 | 6686 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00301 | 64803 | [26 empty] | 2014-05-06 | 7911 | Close Job | Completed
[168 empty] | 0.99802 | 0.00295 | 64804
[7 empty] | 78.5 | 89.5634 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50961 | 50961 | 50949 | 380 | 380 | 365 |  |  | Kelly lyons | -998.24 | 0.08 ‰ | 2014-04-07 | 2014-04-07 | 13 | [5 empty] | 2.14‰ | 2.14‰ | 2.00‰ | 2.00‰ | 2.14 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.24 ‰ | -998.24 ‰ | 0.08 | 0.08 | -998.22 | 0.08 ‰ | -998.22 | 0.08 ‰ |  |  | 2014-04-08 | Jocelyn Turnbull | 13:54 | 0.0017573 | 0.00176 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.7 |  |  | 0.00176 | 8e-05 | 6732 | 2014-04-08 | 12:17 | John West | 198577 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.18 | 0.01% | 4 | 5 | 4 | .X. | 0.00169 | 8e-05 |  |  |  |  | 1.0086 | 1.0086 | Yes | [6 empty] | 0.00169 |  | 8e-05 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00176 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00169 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-31 | Kelly lyons | 13:47 |  | Completed |  |  |  |  | 64857 | 2000000 |  |  | 64827 64832 64833 64834 64835 64889 | 64827~64832~64833~64834~64835~64889 | 2926 |  | 2926 2926 2926 2926 2926 2926 | 2014-04-08 | 13:38 | Christine Prior |  | 0.77 mg |  | 0.77 | [5 empty] | 0.99984 | 0.00203 | 64827 |  | XCAMS | ST | 0% |  | * good-quality measurement * not quite reached the 250k counts per Ox-I cathode | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-05-19 | Carbonate |  | JOB-198577 |  | Water line evacuated like carbonate. | 56045 | Normal | 14047/11 |  |  | 2014-03-24 | 1207 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99897 | 0.00201 | 64832 | [26 empty] | 2014-03-25 | 4867 | Graphite | Completed
[168 empty] | 1.001 | 0.00202 | 64833 | [26 empty] | 2014-03-26 | 2608 | C13 Measurement | Completed
[168 empty] | 0.99982 | 0.00204 | 64834 | [26 empty] | 2014-04-08 | 6732 | AMS Submission Results | Completed
[168 empty] | 0.99985 | 0.00203 | 64835 | [26 empty] | 2014-05-06 | 7899 | Close Job | Completed
[168 empty] | 1.0005 | 0.00201 | 64889
[7 empty] | 82.5 | 89.1397 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48698 | 48698 | 48695 | 323 | 322 | 310 |  |  | Kelly lyons | -997.67 | 0.09 ‰ | 2014-04-07 | 2014-04-07 | 39 | [5 empty] | 6.91‰ | 6.91‰ | 2.00‰ | 2.00‰ | 6.91 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -997.64 | 0.09 ‰ |  |  | 2014-04-08 | Jocelyn Turnbull | 13:54 | 0.0023291 | 0.00233 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.9 |  |  | 0.00233 | 9e-05 | 6733 | 2014-04-08 | 12:17 | John West | 198274 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 4 | ... | 0.00224 | 9e-05 |  |  |  |  | 1.0086 | 1.0086 | Yes | [6 empty] | 0.00224 |  | 9e-05 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00233 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00224 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-31 | Kelly lyons | 13:47 |  | Completed |  |  |  |  | 64563 | 2000000 |  |  | 64827 64832 64833 64834 64835 64889 | 64827~64832~64833~64834~64835~64889 | 2926 |  | 2926 2926 2926 2926 2926 2926 | 2014-04-08 | 13:38 | Christine Prior |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99984 | 0.00203 | 64827 |  | XCAMS | ST | 0% |  | * good-quality measurement * not quite reached the 250k counts per Ox-I cathode | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-03-10 | Carbonate |  | JOB-198274 |  |  | 56036 |  | 14047/1 |  |  | 2014-02-24 | 1110 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99897 | 0.00201 | 64832 | [26 empty] | 2014-02-27 | 4573 | Graphite | Completed
[168 empty] | 1.001 | 0.00202 | 64833 | [26 empty] | 2014-02-28 | 2446 | C13 Measurement | Completed
[168 empty] | 0.99982 | 0.00204 | 64834 | [26 empty] | 2014-04-08 | 6733 | AMS Submission Results | Completed
[168 empty] | 0.99985 | 0.00203 | 64835 | [26 empty] | 2014-05-06 | 7890 | Close Job | Completed
[168 empty] | 1.0005 | 0.00201 | 64889
[7 empty] | 81.4 | 90.1071 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51401 | 51400 | 51419 | 251 | 251 | 242 |  |  | Kelly lyons | -998.34 | 0.05 ‰ | 2014-04-14 | 2014-04-14 | 3 | [5 empty] | 2.04‰ | 2.04‰ | 2.00‰ | 2.00‰ | 2.04 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.70711 | -998.34 ‰ | -998.34 ‰ | 0.05 | 0.05 | -998.32 | 0.05 ‰ | -998.32 | 0.05 ‰ |  | Christine Prior | 2014-04-17 | Jocelyn Turnbull | 13:02 | 0.0016637 | 0.00166 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 1.9 |  |  | 0.00166 | 5e-05 | 6929 | 2014-04-17 | 10:21 | John West | 198718 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.17 | 0.01% | 22 | 5 | 4 | .X. | 0.0016 | 5e-05 |  |  |  |  | 1.0146 | 1.0146 | Yes | [6 empty] | 0.0016 |  | 5e-05 |  | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00166 | 0 | 5e-05 | 0 | 0 |  |  |  |  | 0.0016 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-04-11 | Kelly lyons | 13:39 |  | Completed |  |  |  |  | 65070 | 2000000 |  |  | 64901 64903 64918 64996 64999 65034 65037 65038 | 64901~64903~64918~64996~64999~65034~65037~65038 | 2928 |  | 2928 2928 2928 2928 2928 2928 2928 2928 | 2014-04-17 | 11:27 | Christine Prior |  | 0.98 mg |  | 0.98 | [5 empty] | 1.001 | 0.00215 | 64901 |  | XCAMS | ST | 0.18% |  | * high-prec not achieved due to one Ox-I cathode (TP64903, cat#5) coming out low by 3*sigma_Poisson; found no reason to reject it; system error = 0.18% * global run-to-run variability correction factor = 1.03 * one cathode (TP65036, cat#2) corrected for excess run-to-run variability | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-06-04 | Carbonate | Complete | JOB-198718 |  | Water line evacuated like carbonate. | 56127 | Normal | 14047/11 |  |  | 2014-04-10 | 1242 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99584 | 0.00214 | 64903 | [26 empty] | 2014-04-11 | 5080 | Graphite | Completed
[168 empty] | 1.0024 | 0.00214 | 64918 | [26 empty] | 2014-04-15 | 2725 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00217 | 64996 | [26 empty] | 2014-04-17 | 6929 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.00216 | 64999 | [26 empty] | 2014-05-06 | 7864 | Close Job | Completed
[168 empty] | 1.0001 | 0.00215 | 65034
[168 empty] | 0.99805 | 0.00214 | 65037
[168 empty] | 0.99986 | 0.00214 | 65038
[7 empty] | 48 | 90.167 | 0.0009 | RRL-CBIn-TY | Background Inorganic |  |  |  |  | 41410 | 41409 | 41411 | 289 | 289 | 292 |  |  | Kelly lyons | -994.23 | 0.21 ‰ | 2014-04-28 | 2014-04-28 | 25 | [5 empty] | -9.15‰ | -9.15‰ | 2.00‰ | 2.00‰ | -9.15 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0007 | 0 | 0.89443 | [6 empty] | -994.16 | 0.21 ‰ |  |  | 2014-05-07 | Jocelyn Turnbull | 8:32 | 0.0057709 | 0.00577 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 13 |  |  | 0.00577 | 0.00021 | 6999 | 2014-04-29 | 11:53 | John West | 198622 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 8 | 5 | 3 | ... | 0.00555 | 0.0002 |  |  |  |  | 1.004 | 1.004 | Yes |  |  | Some runs on this blank show 14C contamination, not only at the surface (1st run rejected; not unexpected) but also deeper in (runs not rejected; unexpected but effect seen before). The reported result underestimates the actual uncertainty by ca. factor 2. |  |  |  | 0.00555 |  | 0.0002 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00577 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.00555 | 0 | 0.0002 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2014-04-17 | Kelly lyons | 14:19 |  | Completed |  |  |  |  | 64967 | 2000000 |  |  | 64887 64904 64905 64906 65066 | 64887~64904~64905~64906~65066 | 2931 |  | 2931 2931 2931 2931 2931 |  |  |  |  | 0.14 mg |  | 0.14 | [5 empty] | 0.99824 | 0.00181 | 64887 |  | XCAMS | ST | 0% |  | * staightforward measurement * XCAMS calibration done with 5 large Ox-Is only; the smalls set to "standard as unknown" * error adjustments: - no system error - global run-to-run variance consistent with Poisson stats - 3 cathodes with adjustment for excess run-to-run variance: TP64899 (cat#2), TP64969 (cat#12), TP64793 (cat#21) - IAEA-C1 blank (TP64967, cat#8) is reported with an underestimated uncertainty d=0.001 based on performance of small oxalics in this wheel inorganic m=0.0004 organic m=0.003 - kauris were much worse, but we think there is a problem with the kauri pretreatment, so used something based on previous smalls and recent larger kauris |  | RRL | [5 empty] | Background Inorganic |  | 2014-04-10 | Carbonate |  | JOB-198622 |  | notes in results comment regarding poorer quality measurement | 56276 |  | 14047/1 |  |  | 2014-03-31 | 1225 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00186 | 64904 | [26 empty] | 2014-04-03 | 4977 | Graphite | Completed
[168 empty] | 1.0003 | 0.00184 | 64905 | [26 empty] | 2014-05-07 | 6999 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00182 | 64906 | [26 empty] | 2014-05-13 | 8265 | Close Job | Completed
[168 empty] | 1.0009 | 0.00201 | 65066
[7 empty] | 69 | 89.624 | 0.0004 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 47978 | 47978 | 47970 | 361 | 360 | 347 |  |  | Kelly lyons | -997.45 | 0.11 ‰ | 2014-04-28 | 2014-04-28 | 25 | [5 empty] | -3.14‰ | -3.14‰ | 2.00‰ | 2.00‰ | -3.14 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0007 | 0 | 0.89443 | [6 empty] | -997.42 | 0.12 ‰ |  |  | 2014-05-07 | Jocelyn Turnbull | 8:32 | 0.0025475 | 0.00255 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 8.9 |  |  | 0.00255 | 0.00011 | 7002 | 2014-04-29 | 11:53 | John West | 198638 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 25 | 5 | 3 | ... | 0.00245 | 0.00011 |  |  |  |  | 1.004 | 1.004 | Yes | [6 empty] | 0.00245 |  | 0.00011 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00255 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00245 | 0 | 0.00011 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2014-04-17 | Kelly lyons | 14:19 |  | Completed |  |  |  |  | 64973 | 2000000 |  |  | 64887 64904 64905 64906 65066 | 64887~64904~64905~64906~65066 | 2931 |  | 2931 2931 2931 2931 2931 |  |  |  |  | 0.23 mg |  | 0.23 | [5 empty] | 0.99824 | 0.00181 | 64887 |  | XCAMS | ST | 0% |  | * staightforward measurement * XCAMS calibration done with 5 large Ox-Is only; the smalls set to "standard as unknown" * error adjustments: - no system error - global run-to-run variance consistent with Poisson stats - 3 cathodes with adjustment for excess run-to-run variance: TP64899 (cat#2), TP64969 (cat#12), TP64793 (cat#21) - IAEA-C1 blank (TP64967, cat#8) is reported with an underestimated uncertainty d=0.001 based on performance of small oxalics in this wheel inorganic m=0.0004 organic m=0.003 - kauris were much worse, but we think there is a problem with the kauri pretreatment, so used something based on previous smalls and recent larger kauris |  | RRL | [5 empty] | Background Inorganic |  | 2014-04-15 | Carbonate | Complete | JOB-198638 |  |  | 56277 |  | 14047/1 |  |  | 2014-04-01 | 1232 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00186 | 64904 | [26 empty] | 2014-04-03 | 4983 | Graphite | Completed
[168 empty] | 1.0003 | 0.00184 | 64905 | [26 empty] | 2014-05-07 | 7002 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00182 | 64906 | [26 empty] | 2014-05-13 | 8268 | Close Job | Completed
[168 empty] | 1.0009 | 0.00201 | 65066
[7 empty] | 76.4 | 88.949 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50021 | 50020 | 50003 | 338 | 338 | 325 |  |  | Christine Prior | -998.02 | 0.08 ‰ | 2014-05-05 | 2014-05-05 | 47 | [5 empty] | 2.78‰ | 2.78‰ | 2.00‰ | 2.00‰ | 2.78 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -998 | 0.08 ‰ |  |  | 2014-05-07 | Jocelyn Turnbull | 10:43 | 0.0019756 | 0.00198 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.7 |  |  | 0.00198 | 8e-05 | 7012 | 2014-05-06 | 12:20 | Albert Zondervan | 198517 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  |  |  | 12 | 5 | 4 | ... | 0.0019 | 8e-05 |  |  |  |  | 1.0023 | 1.0023 | Yes | [6 empty] | 0.0019 |  | 8e-05 |  | Yes | 1 | 0.99999 | 0.0012 | 0.0012 | 0.00198 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0019 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-02 | Christine Prior | 11:59 |  | Completed |  |  |  |  | 64787 | 2000000 |  |  | 65095 65101 65102 65103 65104 65159 | 65095~65101~65102~65103~65104~65159 | 2934 |  | 2934 2934 2934 2934 2934 2934 | 2014-05-07 | 10:36 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 0.99777 | 0.00295 | 65095 |  | XCAMS | ST | 0.22% |  | * somewhat low C-beam intensities * error adjustments: cathode pos#2 (TP65027) and system error * some early runs rejected because of likely surface contamination * despite the good tune, it remains to be sorted why we haven't achieved a lower system error | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-04-01 | Carbonate |  | JOB-198517 |  |  | 56353 |  | 14047/1 |  |  | 2014-03-18 | 1198 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0033 | 0.00294 | 65101 | [26 empty] | 2014-03-19 | 4797 | Graphite | Completed
[168 empty] | 0.99954 | 0.00293 | 65102 | [26 empty] | 2014-03-21 | 2568 | C13 Measurement | Completed
[168 empty] | 0.99988 | 0.00292 | 65103 | [26 empty] | 2014-05-07 | 7012 | AMS Submission Results | Completed
[168 empty] | 0.99609 | 0.00292 | 65104 | [26 empty] | 2014-06-12 | 8420 | Close Job | Completed
[168 empty] | 1.0034 | 0.00295 | 65159
[7 empty] | 76.1 | 88.9932 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49178 | 49178 | 49193 | 381 | 380 | 367 |  |  | Christine Prior | -997.81 | 0.10 ‰ | 2014-05-05 | 2014-05-05 | 21 | [5 empty] | 2.28‰ | 2.28‰ | 2.00‰ | 2.00‰ | 2.28 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -997.81 ‰ | -997.81 ‰ | 0.1 | 0.1 | -997.78 | 0.11 ‰ | -997.78 | 0.11 ‰ |  |  | 2014-05-07 | Jocelyn Turnbull | 10:43 | 0.002194 | 0.00219 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.9 |  |  | 0.00219 | 0.0001 | 7037 | 2014-05-06 | 12:20 | Albert Zondervan | 198737 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.22 | 0.01% | 22 | 5 | 4 | .X. | 0.00211 | 0.0001 |  |  |  |  | 1.0023 | 1.0023 | Yes | [6 empty] | 0.00211 |  | 0.0001 |  | Yes | 1 | 0.99999 | 0.0012 | 0.0012 | 0.00219 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00211 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-02 | Christine Prior | 11:59 |  | Completed |  |  |  |  | 65085 | 2000000 |  |  | 65095 65101 65102 65103 65104 65159 | 65095~65101~65102~65103~65104~65159 | 2934 |  | 2934 2934 2934 2934 2934 2934 | 2014-05-07 | 10:36 | Christine Prior |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99777 | 0.00295 | 65095 |  | XCAMS | ST | 0.22% |  | * somewhat low C-beam intensities * error adjustments: cathode pos#2 (TP65027) and system error * some early runs rejected because of likely surface contamination * despite the good tune, it remains to be sorted why we haven't achieved a lower system error | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-06-06 | Carbonate | Complete | JOB-198737 |  | Water line evacuated like carbonate. | 56362 | Normal | 14047/11 |  |  | 2014-04-11 | 1249 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0033 | 0.00294 | 65101 | [26 empty] | 2014-04-14 | 5095 | Graphite | Completed
[168 empty] | 0.99954 | 0.00293 | 65102 | [26 empty] | 2014-04-15 | 2736 | C13 Measurement | Completed
[168 empty] | 0.99988 | 0.00292 | 65103 | [26 empty] | 2014-05-07 | 7037 | AMS Submission Results | Completed
[168 empty] | 0.99609 | 0.00292 | 65104 | [26 empty] | 2014-06-12 | 8543 | Close Job | Completed
[168 empty] | 1.0034 | 0.00295 | 65159
[7 empty] | 90.8 | 89.1722 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52082 | 52081 | 52074 | 437 | 437 | 420 |  |  | jenny dahl | -998.47 | 0.08 ‰ | 2014-05-13 | 2014-05-13 | 13 | [5 empty] | -0.68‰ | -0.68‰ | 2.00‰ | 2.00‰ | -0.68 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -998.47 ‰ | -998.47 ‰ | 0.08 | 0.08 | -998.45 | 0.08 ‰ | -998.45 | 0.08 ‰ |  |  | 2014-05-14 | Jocelyn Turnbull | 16:25 | 0.0015285 | 0.00153 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00153 | 8e-05 | 7113 | 2014-05-14 | 9:58 | Albert Zondervan | 198814 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 32 | 5 | 4 | .X. | 0.00147 | 8e-05 |  |  |  |  | 1.0014 | 1.0014 | Yes | [6 empty] | 0.00147 |  | 8e-05 |  | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00153 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00147 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-08 | jenny dahl | 16:16 |  | Completed |  |  |  |  | 65169 | 2000000 |  |  | 65188 65189 65190 65216 65217 65218 | 65188~65189~65190~65216~65217~65218 | 2936 |  | 2936 2936 2936 2936 2936 2936 | 2014-05-14 | 12:37 | Christine Prior |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0007 | 0.00185 | 65188 |  | XCAMS | ST | 0% |  | * excellent 1stC-14 measurement after Be-10 * no error adjustments * some first runs on blanks rejected: usual degree of surface contamination | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-06-23 | Carbonate |  | JOB-198814 |  | Water line evacuated like carbonate. | 56441 | Normal | 14047/11 |  |  | 2014-04-28 | 1264 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99958 | 0.00188 | 65189 | [26 empty] | 2014-04-30 | 5179 | Graphite | Completed
[168 empty] | 1.0005 | 0.00185 | 65190 | [26 empty] | 2014-05-02 | 2771 | C13 Measurement | Completed
[168 empty] | 0.99707 | 0.00185 | 65216 | [26 empty] | 2014-05-14 | 7113 | AMS Submission Results | Completed
[168 empty] | 0.99982 | 0.00189 | 65217 | [26 empty] | 2014-06-12 | 8558 | Close Job | Completed
[168 empty] | 1.0024 | 0.00187 | 65218
[7 empty] | 96.3 | 88.9614 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48662 | 48662 | 48661 | 321 | 321 | 309 |  |  | jenny dahl | -997.66 | 0.09 ‰ | 2014-05-13 | 2014-05-13 | 6 | [5 empty] | 1.69‰ | 1.69‰ | 2.00‰ | 2.00‰ | 1.69 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | [6 empty] | -997.63 | 0.09 ‰ |  |  | 2014-05-14 | Jocelyn Turnbull | 16:25 | 0.0023395 | 0.00234 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.5 |  |  | 0.00234 | 9e-05 | 7122 | 2014-05-14 | 9:58 | Albert Zondervan | 198837 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  |  |  | 25 | 5 | 5 | ... | 0.00225 | 9e-05 |  |  |  |  | 1.0014 | 1.0014 | Yes | [6 empty] | 0.00225 |  | 9e-05 |  | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00234 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00225 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-08 | jenny dahl | 16:17 |  | Completed |  |  |  |  | 65246 | 2000000 |  |  | 65188 65189 65190 65216 65217 65218 | 65188~65189~65190~65216~65217~65218 | 2936 |  | 2936 2936 2936 2936 2936 2936 | 2014-05-14 | 12:37 | Christine Prior |  | 0.89 mg |  | 0.89 | [5 empty] | 1.0007 | 0.00185 | 65188 |  | XCAMS | ST | 0% |  | * excellent 1stC-14 measurement after Be-10 * no error adjustments * some first runs on blanks rejected: usual degree of surface contamination | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-05-14 | Carbonate |  | JOB-198837 |  | carbonate line, normal size | 56432 |  | 14047/1 |  |  | 2014-05-06 | 2461 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99958 | 0.00188 | 65189 | [26 empty] | 2014-05-06 | 1271 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00185 | 65190 | [26 empty] | 2014-05-07 | 5256 | Graphite | Completed
[168 empty] | 0.99707 | 0.00185 | 65216 | [26 empty] | 2014-05-08 | 2799 | C13 Measurement | Completed
[168 empty] | 0.99982 | 0.00189 | 65217 | [26 empty] | 2014-05-14 | 7122 | AMS Submission Results | Completed
[168 empty] | 1.0024 | 0.00187 | 65218 | [26 empty] | 2014-06-12 | 8481 | Close Job | Completed
[7 empty] | 103.3 | 88.7989 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47257 | 47257 | 47248 | 330 | 329 | 317 |  |  | jenny dahl | -997.21 | 0.11 ‰ | 2014-05-19 | 2014-05-19 | 11 | [5 empty] | 3.58‰ | 3.58‰ | 2.00‰ | 2.00‰ | 3.58 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.82 | 0.8165 | [6 empty] | -997.18 | 0.12 ‰ |  |  | 2014-05-21 | Jocelyn Turnbull | 16:19 | 0.0027866 | 0.00279 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.3 |  |  | 0.00279 | 0.00011 | 7188 | 2014-05-20 | 14:19 | Johannes Kaiser | 198838 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 5 | ... | 0.00268 | 0.00011 |  |  |  |  | 1.0014 | 1.0014 | Yes | [6 empty] | 0.00268 |  | 0.00011 |  | No | 1 | 0.99998 | 0.00096 | 0.00096 | 0.00279 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00268 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-16 | jenny dahl | 16:19 |  | Completed |  |  |  |  | 65261 | 2000000 |  |  | 65272 65273 65274 65306 65307 65308 | 65272~65273~65274~65306~65307~65308 | 2938 |  | 2938 2938 2938 2938 2938 2938 | 2014-05-22 | 9:33 | jenny dahl |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99978 | 0.00233 | 65272 |  | XCAMS | EA | 0.14% |  | - normal measurement water blank (TP 65331) was a strong outlier and was not included in the blank correction | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-05-14 | Carbonate | Not Complete | JOB-198838 |  | carbonate line, normal size | 56473 |  | 14047/1 |  |  | 2014-05-06 | 2461 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0021 | 0.00231 | 65273 | [26 empty] | 2014-05-07 | 1279 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99836 | 0.00231 | 65274 | [26 empty] | 2014-05-08 | 5271 | Graphite | Completed
[168 empty] | 0.99986 | 0.00233 | 65306 | [26 empty] | 2014-05-12 | 2809 | C13 Measurement | Completed
[168 empty] | 0.99639 | 0.00249 | 65307 | [26 empty] | 2014-05-21 | 7188 | AMS Submission Results | Completed
[168 empty] | 1.0031 | 0.00241 | 65308 | [26 empty] | 2014-06-12 | 8482 | Close Job | Completed
|  |  | 0 | 0 |  |  | 100.2 | 88.7845 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41266 | 41266 | 41273 | 199 | 199 | 205 | 41069.5 | 41467.7 | jenny dahl | -994.13 | 0.15 ‰ | 2014-05-19 | 2014-05-19 | 3 | [5 empty] | 3.74‰ | 3.74‰ | 2.00‰ | 2.00‰ | 3.74 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.82 | 0.8165 | -994.13 ‰ | -994.13 ‰ | 0.15 | 0.15 | -994.05 | 0.15 ‰ | -994.05 | 0.15 ‰ |  |  | 2014-05-21 | Jocelyn Turnbull | 16:21 | 0.0058748 | 0.00587 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.3 |  |  | 0.00587 | 0.00015 | 7203 | 2014-05-20 | 14:19 | Johannes Kaiser | 198966 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.59 | 0.01% | 24 | 5 | 5 | .X. | 0.00565 | 0.00014 |  |  |  |  | 1.0014 | 1.0014 | Yes |  |  | really terrible water blank |  |  |  | 0.00565 |  | 0.00014 |  | No | 1 | 0.99998 | 0.00096 | 0.00096 | 0.00587 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00565 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-16 | jenny dahl | 16:20 |  | Completed |  |  |  |  | 65331 | 2000000 |  |  | 65272 65273 65274 65306 65307 65308 | 65272~65273~65274~65306~65307~65308 | 2938 |  | 2938 2938 2938 2938 2938 2938 | 2014-05-22 | 9:33 | jenny dahl |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99978 | 0.00233 | 65272 |  | XCAMS | EA | 0.14% |  | - normal measurement water blank (TP 65331) was a strong outlier and was not included in the blank correction | 2.6 | RRL | [5 empty] | Background Test |  | 2014-07-07 | Carbonate |  | JOB-198966 |  | Water line evacuated like carbonate. | 56484 | Normal | 14047/11 |  |  | 2014-05-13 | 1287 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0021 | 0.00231 | 65273 | [26 empty] | 2014-05-13 | 1288 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99836 | 0.00231 | 65274 | [26 empty] | 2014-05-16 | 5341 | Graphite | Completed
[168 empty] | 0.99986 | 0.00233 | 65306 | [26 empty] | 2014-05-19 | 2827 | C13 Measurement | Completed
[168 empty] | 0.99639 | 0.00249 | 65307 | [26 empty] | 2014-05-21 | 7203 | AMS Submission Results | Completed
[168 empty] | 1.0031 | 0.00241 | 65308 | [26 empty] | 2014-06-12 | 8568 | Close Job | Completed
[7 empty] | 66.3 | 89.3092 | 0.0005 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 46192 | 46192 | 46197 | 341 | 341 | 354 |  |  | jenny dahl | -996.82 | 0.14 ‰ | 2014-06-06 | 2014-06-06 | 15 | [5 empty] | -2.68‰ | -2.68‰ | 2.00‰ | 2.00‰ | -2.68 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -996.78 | 0.14 ‰ |  |  | 2014-06-09 | Jocelyn Turnbull | 13:46 | 0.0031818 | 0.00318 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 6.6 |  |  | 0.00318 | 0.00014 | 7315 | 2014-06-09 | 10:29 | Albert Zondervan | 198934 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  |  |  | 13 | 5 | 6 | ... | 0.00306 | 0.00013 |  |  |  |  | 1.0009 | 1.0009 | Yes | [6 empty] | 0.00306 |  | 0.00013 |  | Yes | 1 | 1 | 0.00107 | 0.00107 | 0.00318 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00306 | 0 | 0.00013 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2014-05-30 | jenny dahl | 16:07 |  | Completed |  |  |  |  | 65396 | 2000000 |  |  | 65398 65399 65400 65401 65414 65415 | 65398~65399~65400~65401~65414~65415 | 2942 |  | 2942 2942 2942 2942 2942 2942 | 2014-06-09 | 13:38 | Jocelyn Turnbull |  | 0.28 mg |  | 0.28 | [5 empty] | 0.99703 | 0.0026 | 65398 |  | XCAMS | ST | 0.2% |  | * 1st wheel after SNICS source clean * measurement interrupted by trivial indexer error, not affecting carbon-isotopes ratio measurement data * on all blanks, the first few runs show elevated 14C count rate due to surface contamination and were thus rejected * cat#16 (TP65460) has runs showing a rise and fall of 14C count rate during measurement and these were thus rejected * all Ox-Is <= 0.32 mg C were set to "standard as unknown" * all 3 types of uncertainty adjustment were needed: - system error (from 0.8-1.0 mg C Ox-Is) = 0.20% - global run-to-run repeatability adjustment factor = 1.09 - cat#22 (TP65404) run-to-run repeatability adjustment factor = 1.3 |  | RRL | [5 empty] | Background Inorganic |  | 2014-05-21 | Carbonate |  | JOB-198934 |  | small std to go with Chang foram R 40504/5 | 56649 |  | 14047/1 |  |  | 2014-05-06 | 2461 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00261 | 65399 | [26 empty] | 2014-05-07 | 1284 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99638 | 0.00261 | 65400 | [26 empty] | 2014-05-21 | 1003 | Gas Splitting | Completed
[168 empty] | 1.0018 | 0.0026 | 65401 | [26 empty] | 2014-05-21 | 2062 | Choose Gas Splitting | Completed
[168 empty] | 1.002 | 0.00263 | 65414 | [26 empty] | 2014-05-22 | 5406 | Graphite | Completed
[168 empty] | 1.002 | 0.00263 | 65415 | [26 empty] | 2014-06-09 | 7315 | AMS Submission Results | Completed
[197 empty] | 2014-06-12 | 8498 | Close Job | Completed
[7 empty] | 87.3 | 88.9463 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50914 | 50913 | 50903 | 378 | 378 | 363 |  |  | jenny dahl | -998.23 | 0.08 ‰ | 2014-06-06 | 2014-06-06 | 15 | [5 empty] | 1.39‰ | 1.39‰ | 2.00‰ | 2.00‰ | 1.39 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -998.21 | 0.08 ‰ |  |  | 2014-06-09 | Jocelyn Turnbull | 13:46 | 0.0017677 | 0.00177 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.3 |  |  | 0.00177 | 8e-05 | 7335 | 2014-06-09 | 10:29 | Albert Zondervan | 199117 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  |  |  | 23 | 5 | 4 | ... | 0.0017 | 8e-05 |  |  |  |  | 1.0009 | 1.0009 | Yes | [6 empty] | 0.0017 |  | 8e-05 |  | Yes | 1 | 1 | 0.00107 | 0.00107 | 0.00177 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0017 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-30 | jenny dahl | 16:10 |  | Completed |  |  |  |  | 65402 | 2000000 |  |  | 65398 65399 65400 65401 65414 65415 | 65398~65399~65400~65401~65414~65415 | 2942 |  | 2942 2942 2942 2942 2942 2942 | 2014-06-09 | 13:38 | Jocelyn Turnbull |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99703 | 0.0026 | 65398 |  | XCAMS | ST | 0.2% |  | * 1st wheel after SNICS source clean * measurement interrupted by trivial indexer error, not affecting carbon-isotopes ratio measurement data * on all blanks, the first few runs show elevated 14C count rate due to surface contamination and were thus rejected * cat#16 (TP65460) has runs showing a rise and fall of 14C count rate during measurement and these were thus rejected * all Ox-Is <= 0.32 mg C were set to "standard as unknown" * all 3 types of uncertainty adjustment were needed: - system error (from 0.8-1.0 mg C Ox-Is) = 0.20% - global run-to-run repeatability adjustment factor = 1.09 - cat#22 (TP65404) run-to-run repeatability adjustment factor = 1.3 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-06-04 | Carbonate |  | JOB-199117 |  | Reg InO Blank | 56648 |  | 14047/1 |  |  | 2014-05-21 | 1297 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00261 | 65399 | [26 empty] | 2014-05-22 | 5412 | Graphite | Completed
[168 empty] | 0.99638 | 0.00261 | 65400 | [26 empty] | 2014-05-26 | 2867 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.0026 | 65401 | [26 empty] | 2014-06-09 | 7335 | AMS Submission Results | Completed
[168 empty] | 1.002 | 0.00263 | 65414 | [26 empty] | 2014-06-12 | 8538 | Close Job | Completed
[168 empty] | 1.002 | 0.00263 | 65415
[7 empty] | 77.7 | 89.2551 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47913 | 47912 | 47908 | 260 | 260 | 250 |  |  | Kelly lyons | -997.43 | 0.08 ‰ | 2014-06-23 | 2014-06-23 | 24 | [5 empty] | 3.24‰ | 3.24‰ | 2.00‰ | 2.00‰ | 3.24 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.82 | 0.8165 | -997.43 ‰ | -997.43 ‰ | 0.08 | 0.08 | -997.4 | 0.08 ‰ | -997.4 | 0.08 ‰ |  |  | 2014-06-25 | Jocelyn Turnbull | 15:10 | 0.0025683 | 0.00257 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.00257 | 8e-05 | 7491 | 2014-06-25 | 10:13 | Johannes Kaiser | 199226 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.26 | 0.01% | 17 | 5 | 4 | .X. | 0.00247 | 8e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00247 |  | 8e-05 |  | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00257 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00247 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-19 | Kelly lyons | 9:36 |  | Completed |  |  |  |  | 65481 | 2000000 |  |  | 65525 65526 65527 65569 65570 65571 | 65525~65526~65527~65569~65570~65571 | 2947 |  | 2947 2947 2947 2947 2947 2947 | 2014-06-25 | 11:00 | Christine Prior |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0013 | 0.00162 | 65525 |  | XCAMS | ST | 0% |  | - overall a good measurement - needed to add a small factor of 1.03 to account for the global run-to-run variance - had to add an _2 measurement as 15 rotations of _1 did not give 250k counts on Ox-I std due to low beam intensity - 29/01/2015: filled Sys | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-07-23 | Carbonate |  | JOB-199226 |  | Water line evacuated like carbonate. | 56854 | Normal | 14047/11 |  |  | 2014-05-28 | 1313 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99924 | 0.00156 | 65526 | [26 empty] | 2014-05-30 | 5491 | Graphite | Completed
[168 empty] | 0.99881 | 0.00155 | 65527 | [26 empty] | 2014-06-09 | 2883 | C13 Measurement | Completed
[168 empty] | 1.0025 | 0.00167 | 65569 | [26 empty] | 2014-06-25 | 7491 | AMS Submission Results | Completed
[168 empty] | 0.99871 | 0.00165 | 65570 | [26 empty] | 2014-06-25 | 5143 | Review Final Report | Completed
[168 empty] | 0.99973 | 0.00162 | 65571 | [26 empty] | 2014-09-16 | 9688 | Close Job | Completed
[7 empty] | 77.9 | 89.2195 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 51553 | 51553 | 51565 | 307 | 306 | 296 |  |  | Kelly lyons | -998.37 | 0.06 ‰ | 2014-06-23 | 2014-06-23 | 17 | [5 empty] | 3.64‰ | 3.64‰ | 2.00‰ | 2.00‰ | 3.64 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.82 | 0.8165 | [6 empty] | -998.35 | 0.06 ‰ |  |  | 2014-06-25 | Jocelyn Turnbull | 15:09 | 0.0016325 | 0.00163 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.4 |  |  | 0.00163 | 6e-05 | 7496 | 2014-06-25 | 10:13 | Johannes Kaiser | 199491 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  |  |  | 4 | 5 | 4 | ... | 0.00157 | 6e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00157 |  | 6e-05 |  | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00163 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00157 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-19 | Kelly lyons | 9:36 |  | Completed |  |  |  |  | 65566 | 2000000 |  |  | 65525 65526 65527 65569 65570 65571 | 65525~65526~65527~65569~65570~65571 | 2947 |  | 2947 2947 2947 2947 2947 2947 | 2014-06-25 | 11:00 | Christine Prior |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0013 | 0.00162 | 65525 |  | XCAMS | ST | 0% |  | - overall a good measurement - needed to add a small factor of 1.03 to account for the global run-to-run variance - had to add an _2 measurement as 15 rotations of _1 did not give 250k counts on Ox-I std due to low beam intensity - 29/01/2015: filled Sys | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-06-18 | Carbonate |  | JOB-199491 |  | material ground up under job 198224 | 56844 |  | 14047/1 |  |  | 2014-06-04 | 1317 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99924 | 0.00156 | 65526 | [26 empty] | 2014-06-06 | 5576 | Graphite | Completed
[168 empty] | 0.99881 | 0.00155 | 65527 | [26 empty] | 2014-06-10 | 2924 | C13 Measurement | Completed
[168 empty] | 1.0025 | 0.00167 | 65569 | [26 empty] | 2014-06-25 | 7496 | AMS Submission Results | Completed
[168 empty] | 0.99871 | 0.00165 | 65570 | [26 empty] | 2014-06-30 | 8858 | Close Job | Completed
[168 empty] | 0.99973 | 0.00162 | 65571
[7 empty] | 72 | 89.2508 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49489 | 49488 | 49492 | 277 | 276 | 266 |  |  | Kelly lyons | -997.89 | 0.07 ‰ | 2014-06-23 | 2014-06-23 | 17 | [5 empty] | 3.29‰ | 3.29‰ | 2.00‰ | 2.00‰ | 3.29 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.82 | 0.8165 | -997.89 ‰ | -997.89 ‰ | 0.07 | 0.07 | -997.86 | 0.07 ‰ | -997.86 | 0.07 ‰ |  |  | 2014-06-25 | Jocelyn Turnbull | 15:11 | 0.0021108 | 0.00211 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3 |  |  | 0.00211 | 7e-05 | 7497 | 2014-06-25 | 10:13 | Johannes Kaiser | 199492 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.21 | 0.01% | 10 | 5 | 4 | .X. | 0.00203 | 7e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00203 |  | 7e-05 |  | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00211 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00203 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-19 | Kelly lyons | 9:36 |  | Completed |  |  |  |  | 65567 | 2000000 |  |  | 65525 65526 65527 65569 65570 65571 | 65525~65526~65527~65569~65570~65571 | 2947 |  | 2947 2947 2947 2947 2947 2947 | 2014-06-25 | 11:00 | Christine Prior |  | 0.65 mg |  | 0.65 | [5 empty] | 1.0013 | 0.00162 | 65525 |  | XCAMS | ST | 0% |  | - overall a good measurement - needed to add a small factor of 1.03 to account for the global run-to-run variance - had to add an _2 measurement as 15 rotations of _1 did not give 250k counts on Ox-I std due to low beam intensity - 29/01/2015: filled Sys | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-07-30 | Carbonate | Complete | JOB-199492 |  | Water nitrogen flush but carbonate vessel | 56855 | Normal | 14047/11 |  |  | 2014-06-04 | 1319 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99924 | 0.00156 | 65526 | [26 empty] | 2014-06-06 | 5577 | Graphite | Completed
[168 empty] | 0.99881 | 0.00155 | 65527 | [26 empty] | 2014-06-10 | 2925 | C13 Measurement | Completed
[168 empty] | 1.0025 | 0.00167 | 65569 | [26 empty] | 2014-06-25 | 7497 | AMS Submission Results | Completed
[168 empty] | 0.99871 | 0.00165 | 65570 | [26 empty] | 2014-06-25 | 5145 | Review Final Report | Completed
[168 empty] | 0.99973 | 0.00162 | 65571 | [26 empty] | 2014-09-16 | 9689 | Close Job | Completed
[7 empty] | 97.3 | 88.7754 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52474 | 52473 | Background | 344 | 344 |  |  |  | Kelly lyons | -998.54 | 0.06 ‰ | 2014-07-04 | 2014-07-04 | 11 | [5 empty] | 7.89‰ | 7.89‰ | 2.00‰ | 2.00‰ | 7.89 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0015 | 0 | 0.8165 | -998.54 ‰ | -1000.00 ‰ | 0.06 |  | -998.53 | 0.06 ‰ | -998.53 | 0.06 ‰ |  |  | 2014-07-07 | Jocelyn Turnbull | 9:50 | 0.0014557 |  | 6e-05 | 0 | 6e-05 | 6e-05 | 1.6 |  |  | 0.00146 | 6e-05 | 7658 | 2014-07-07 | 8:50 | Johannes Kaiser | 199596 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 30 | 5 | 4 | .X. | 0.0014 | 6e-05 |  |  |  |  | 1.0055 | 1.0055 | Yes | [6 empty] | 0.0014 |  | 6e-05 |  | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00146 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0014 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-30 | Kelly lyons | 15:34 |  | Completed |  |  |  |  | 65737 | 2000000 |  |  | 65572 65573 65656 65659 65660 65662 | 65572~65573~65656~65659~65660~65662 | 2950 |  | 2950 2950 2950 2950 2950 2950 | 2014-07-07 | 9:27 | Christine Prior |  | 1.21 mg |  | 1.21 | [5 empty] | 0.9998 | 0.00173 | 65572 |  | XCAMS | ST | 0% |  | - good measurement - no system error |  | RRL | [5 empty] | Background Inorganic |  | 2014-08-13 | Carbonate | Complete | JOB-199596 |  | Water line evacuated like carbonate. | 56973 | Normal | 14047/11 |  |  | 2014-06-18 | 1354 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00169 | 65573 | [26 empty] | 2014-06-19 | 1358 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99868 | 0.00171 | 65656 | [26 empty] | 2014-06-23 | 5747 | Graphite | Completed
[168 empty] | 0.99966 | 0.0017 | 65659 | [26 empty] | 2014-07-07 | 7658 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00188 | 65660 | [26 empty] | 2014-07-10 | 9001 | Close Job | Completed
[168 empty] | 0.99914 | 0.00186 | 65662
[7 empty] | 82.7 | 89.2642 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 49255 | 49254 |  | 346 | 345 |  |  |  | Kelly lyons | -997.83 | 0.09 ‰ | 2014-07-07 | 2014-07-07 | 5 | [5 empty] | 4.94‰ | 4.94‰ | 2.00‰ | 2.00‰ | 4.94 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -997.83 ‰ |  | 0.09 |  | -997.8 | 0.09 ‰ | -997.8 | 0.09 ‰ |  |  | 2014-07-09 | Jocelyn Turnbull | 8:32 | 0.0021732 |  | 9e-05 |  | 9e-05 | 9e-05 | 1.7 |  |  | 0.00217 | 9e-05 | 7693 | 2014-07-08 | 16:43 | Johannes Kaiser | 199750 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.22 | 0.01% | 30 | 5 | 4 | .X. | 0.00209 | 9e-05 |  |  |  |  | 1.008 | 1.008 | Yes | [6 empty] | 0.00209 |  | 9e-05 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00217 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00209 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-03 | Kelly lyons | 15:25 |  | Completed |  |  |  |  | 65834 | 2000000 |  |  | 65826 65835 65836 65837 65838 65840 | 65826~65835~65836~65837~65838~65840 | 2951 |  | 2951 2951 2951 2951 2951 2951 | 2014-07-09 | 6:22 | Christine Prior |  | 1.12 mg |  | 1.12 | [5 empty] | 1.0012 | 0.00204 | 65826 |  | XCAMS | ST | 0% |  | - had to let it go for 15 rotations in order to get at least 250k counts on Ox-I standards - needed a small error adjustment of 1.04 to explain the global run-to-run variance Both secondaries in this wheel are old - Oxs are biased young, perhaps? | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-08-26 | Carbonate |  | JOB-199750 |  | Water line evacuated like carbonate. | 57014 | Normal | 14047/11 |  |  | 2014-07-01 | 1373 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0026 | 0.00208 | 65835 | [26 empty] | 2014-07-02 | 5844 | Graphite | Completed
[168 empty] | 0.99672 | 0.00201 | 65836 | [26 empty] | 2014-07-07 | 3020 | C13 Measurement | Completed
[168 empty] | 0.99962 | 0.00197 | 65837 | [26 empty] | 2014-07-09 | 7693 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.002 | 65838 | [26 empty] | 2014-07-10 | 9009 | Close Job | Completed
[168 empty] | 0.99983 | 0.00206 | 65840
[7 empty] | 75.6 | 89.4977 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48556 | 48555 |  | 352 | 352 |  |  |  | jenny dahl | -997.63 | 0.10 ‰ | 2014-07-14 | 2014-07-14 | 7 | [5 empty] | 1.74‰ | 1.74‰ | 2.00‰ | 2.00‰ | 1.74 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0 | 0.8165 | [6 empty] | -997.6 | 0.11 ‰ |  | Christine Prior | 2014-07-16 | Jocelyn Turnbull | 9:17 | 0.0023707 |  | 0.0001 |  | 0.0001 | 0.0001 | 2.4 |  |  | 0.00237 | 0.0001 | 7774 | 2014-07-15 | 15:44 | Johannes Kaiser | 199771 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  |  |  | 4 | 5 | 4 | ... | 0.00228 | 0.0001 |  |  |  |  | 1.0074 | 1.0075 | Yes | [6 empty] | 0.00228 |  | 0.0001 |  | Yes | 1 | 1 | 0.00082 | 0.00082 | 0.00237 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00228 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-10 | jenny dahl | 16:07 |  | Completed |  |  |  |  | 65882 | 2000000 |  |  | 65858 65909 65910 65911 65925 65926 | 65858~65909~65910~65911~65925~65926 | 2953 |  | 2953 2953 2953 2953 2953 2953 | 2014-07-15 | 16:23 | Christine Prior |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99899 | 0.00211 | 65858 |  | XCAMS | EA | 0% |  | - good measurement - no error adjustment needed = ~7000 counts short to reach 1.5 million counts on Ox-I (6 * 250k counts = 1.5 million counts) | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-07-16 | Carbonate |  | JOB-199771 |  |  | 57085 |  | 14047/1 |  |  | 2014-07-03 | 1389 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.002 | 0.00208 | 65909 | [26 empty] | 2014-07-07 | 5892 | Graphite | Completed
[168 empty] | 0.99756 | 0.00201 | 65910 | [26 empty] | 2014-07-11 | 3041 | C13 Measurement | Completed
[168 empty] | 0.99996 | 0.00194 | 65911 | [26 empty] | 2014-07-16 | 7774 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00201 | 65925 | [26 empty] | 2014-08-01 | 9225 | Close Job | Completed
[168 empty] | 1.0003 | 0.00198 | 65926
[7 empty] | 80.4 | 89.7009 | 0.0006 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 45883 | 45883 | 45875 | 303 | 303 | 291 |  |  | Kelly lyons | -996.69 | 0.12 ‰ | 2014-07-30 | 2014-07-30 | 16 | [5 empty] | 2.06‰ | 2.06‰ | 2.00‰ | 2.00‰ | 2.06 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | [6 empty] | -996.65 | 0.13 ‰ |  |  | 2014-08-01 | Jocelyn Turnbull | 13:27 | 0.0033065 | 0.00331 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00331 | 0.00012 | 7819 | 2014-07-31 | 16:35 | Johannes Kaiser | 198170 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  |  |  | 10 | 5 | 6 | .S. | 0.00318 | 0.00012 |  |  |  |  | 1.0101 | 1.0101 | Yes |  |  | poorer than average of blanks for carbonate line in last 3 months |  |  |  | 0.00318 |  | 0.00012 |  | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00331 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00318 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-24 | Kelly lyons | 15:34 |  | Completed |  |  |  |  | 65945 | 2000000 |  |  | 65265 65661 65852 65853 65855 65856 | 65265~65661~65852~65853~65855~65856 | 2954 |  | 2954 2954 2954 2954 2954 2954 | 2014-08-01 | 13:07 | Christine Prior |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0006 | 0.00193 | 65265 |  | XCAMS | ST | 0% |  | - good measurement - no error adjustments needed | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-02-24 | Carbonate |  | JOB-198170 |  | Sealed for when needed, good blank to go with sealtube water | 57124 |  | 14047/1 |  |  | 2014-02-11 | 1042 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99917 | 0.00199 | 65661 | [26 empty] | 2014-07-11 | 1060 | Gas Splitting | Completed
[168 empty] | 1.001 | 0.00201 | 65852 | [26 empty] | 2014-07-11 | 2159 | Choose Gas Splitting | Completed
[168 empty] | 0.99869 | 0.00189 | 65853 | [26 empty] | 2014-07-14 | 5955 | Graphite | Completed
[168 empty] | 1.0013 | 0.00191 | 65855 | [26 empty] | 2014-07-18 | 3065 | C13 Measurement | Completed
[168 empty] | 0.99926 | 0.00209 | 65856 | [26 empty] | 2014-08-01 | 7819 | AMS Submission Results | Completed
[197 empty] | 2014-08-01 | 5311 | Review Final Report | Completed
[197 empty] | 2014-08-01 | 9193 | Close Job | Completed
[7 empty] | 88 | 89.485 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47287 | 47287 | 47277 | 301 | 300 | 289 |  |  | Kelly lyons | -997.22 | 0.10 ‰ | 2014-07-30 | 2014-07-30 | 6 | [5 empty] | 4.48‰ | 4.48‰ | 2.00‰ | 2.00‰ | 4.48 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -997.22 ‰ | -997.22 ‰ | 0.1 | 0.1 | -997.19 | 0.11 ‰ | -997.19 | 0.11 ‰ |  |  | 2014-08-01 | Jocelyn Turnbull | 13:27 | 0.0027763 | 0.00278 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.1 |  |  | 0.00278 | 0.0001 | 7856 | 2014-07-31 | 16:35 | Johannes Kaiser | 199812 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.28 | 0.01% | 38 | 5 | 4 | .X. | 0.00267 | 0.0001 |  |  |  |  | 1.0101 | 1.0101 | Yes |  |  | poorer than average of blanks for water line in last 3 months |  |  |  | 0.00267 |  | 0.0001 |  | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00278 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00267 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-24 | Kelly lyons | 15:34 |  | Completed |  |  |  |  | 66072 | 2000000 |  |  | 65265 65661 65852 65853 65855 65856 | 65265~65661~65852~65853~65855~65856 | 2954 |  | 2954 2954 2954 2954 2954 2954 | 2014-08-01 | 13:07 | Christine Prior |  | 0.86 mg |  | 0.86 | [5 empty] | 1.0006 | 0.00193 | 65265 |  | XCAMS | ST | 0% |  | - good measurement - no error adjustments needed | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2014-09-02 | Carbonate | Not Complete | JOB-199812 |  | Water nitrogen flush but carbonate vessel | 57134 | Normal | 14047/11 |  |  | 2014-07-23 | 1449 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99917 | 0.00199 | 65661 | [26 empty] | 2014-07-24 | 6082 | Graphite | Completed
[168 empty] | 1.001 | 0.00201 | 65852 | [26 empty] | 2014-07-28 | 3107 | C13 Measurement | Completed
[168 empty] | 0.99869 | 0.00189 | 65853 | [26 empty] | 2014-08-01 | 7856 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00191 | 65855 | [26 empty] | 2014-08-01 | 9239 | Close Job | Completed
[168 empty] | 0.99926 | 0.00209 | 65856
[7 empty] | 97.2 | 89.2327 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51553 | 51553 | 51565 | 358 | 358 | 345 |  |  | jenny dahl | -998.37 | 0.07 ‰ | 2014-08-11 | 2014-08-11 | 17 | [5 empty] | 2.46‰ | 2.46‰ | 2.00‰ | 2.00‰ | 2.46 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.37 ‰ | -998.37 ‰ | 0.07 | 0.07 | -998.35 | 0.07 ‰ | -998.35 | 0.07 ‰ |  |  | 2014-08-14 | Christine Prior | 14:55 | 0.0016325 | 0.00163 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.00163 | 7e-05 | 7957 | 2014-08-12 | 11:06 | Johannes Kaiser | 199811 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.16 | 0.01% | 17 | 5 | 4 | .X. | 0.00157 | 7e-05 |  |  |  |  | 1.0052 | 1.0052 | Yes | [6 empty] | 0.00157 |  | 7e-05 |  | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00163 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00157 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-08-06 | jenny dahl | 13:56 |  | Completed |  |  |  |  | 66087 | 2000000 |  |  | 66010 66022 66023 66027 66028 66142 | 66010~66022~66023~66027~66028~66142 | 2958 |  | 2958 2958 2958 2958 2958 2958 | 2014-08-12 | 11:37 | Christine Prior |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99714 | 0.0019 | 66010 |  | XCAMS | ST | 0% |  | - pos# 22 and 33 needed an additional factor of 1.3 to explain its individual run-to-run variance - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-09-02 | Carbonate |  | JOB-199811 |  | Water line evacuated like carbonate. | 57254 | Normal | 14047/11 |  |  | 2014-07-23 | 1448 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00191 | 66022 | [26 empty] | 2014-07-25 | 6097 | Graphite | Completed
[168 empty] | 1.0002 | 0.00187 | 66023 | [26 empty] | 2014-07-28 | 3112 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00187 | 66027 | [26 empty] | 2014-08-14 | 7957 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00189 | 66028 | [26 empty] | 2014-08-14 | 9360 | Close Job | Completed
[168 empty] | 0.99756 | 0.00192 | 66142
[7 empty] | 101.3 | 89.1338 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50455 | 50454 | 50462 | 357 | 357 | 344 |  |  | jenny dahl | -998.13 | 0.08 ‰ | 2014-08-11 | 2014-08-11 | 18 | [5 empty] | 3.57‰ | 3.57‰ | 2.00‰ | 2.00‰ | 3.57 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0003 | 0 | 0.8165 | [6 empty] | -998.11 | 0.08 ‰ |  |  | 2014-08-14 | Christine Prior | 14:55 | 0.0018716 | 0.00187 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00187 | 8e-05 | 7964 | 2014-08-12 | 11:06 | Johannes Kaiser | 199841 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  |  |  | 31 | 5 | 6 | ... | 0.0018 | 8e-05 |  |  |  |  | 1.0052 | 1.0052 | Yes | [6 empty] | 0.0018 |  | 8e-05 |  | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00187 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0018 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-08-06 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 66075 | 2000000 |  |  | 66010 66022 66023 66027 66028 66142 | 66010~66022~66023~66027~66028~66142 | 2958 |  | 2958 2958 2958 2958 2958 2958 | 2014-08-12 | 11:37 | Christine Prior |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99714 | 0.0019 | 66010 |  | XCAMS | ST | 0% |  | - pos# 22 and 33 needed an additional factor of 1.3 to explain its individual run-to-run variance - 29/01/2015: filled System Error (%) box |  | RRL | [5 empty] | Background Inorganic |  | 2014-07-28 | Carbonate |  | JOB-199841 |  | cracked on 22/7/14 for graphite 24/7/14; size verified (CG) 0.97mg Sample sealed in breakseal 15/7/14 until needed, KL. Sample size differs from value on bottle. CF base pressure before measuring sample was -4.7 and after was -233.17 Check sample size before sealing. | 57245 |  | 14047/1 |  |  | 2014-07-14 | 1425 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00191 | 66022 | [26 empty] | 2014-07-15 | 2161 | Choose Gas Splitting | Cancelled
[168 empty] | 1.0002 | 0.00187 | 66023 | [26 empty] | 2014-07-15 | 1062 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00187 | 66027 | [26 empty] | 2014-07-15 | 2162 | Choose Gas Splitting | Completed
[168 empty] | 1.0026 | 0.00189 | 66028 | [26 empty] | 2014-07-24 | 6085 | Graphite | Completed
[168 empty] | 0.99756 | 0.00192 | 66142 | [26 empty] | 2014-08-14 | 7964 | AMS Submission Results | Completed
[197 empty] | 2014-08-14 | 9361 | Close Job | Completed
[7 empty] | 98.8 | 89.2914 | 0.0006 | RRL-Test-LG | Background Test |  |  |  |  | 45513 | 45513 | 45519 | 241 | 241 | 232 | 45275.4 | 45758 | Kelly lyons | -996.54 | 0.10 ‰ | 2014-08-20 | 2014-08-20 | 2 | [5 empty] | 1.66‰ | 1.66‰ | 2.00‰ | 2.00‰ | 1.66 ‰ | 2.00 ‰ | AMS | 2.82‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.8165 | [6 empty] | -996.49 | 0.11 ‰ |  |  | 2014-08-27 | Jocelyn Turnbull | 10:33 | 0.0034625 | 0.00346 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.6 |  |  | 0.00346 | 0.0001 | 8097 | 2014-08-21 | 15:24 | Johannes Kaiser | 200025 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 3 | 5 | 4 | ... | 0.00333 | 0.0001 |  |  |  |  | 1.005 | 1.0051 | Yes |  |  | lousy blank, especially compared to SIRI doublespar. Not representative - something went wrong in carbonate evolution Chose blanks for this wheel from blanks in recent wheels |  |  |  | 0.00333 |  | 0.0001 |  | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00346 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00333 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-08-19 | Kelly lyons | 15:55 |  | Completed |  |  |  |  | 66273 | 2000000 |  |  | 66143 66144 66145 66146 66187 66195 | 66143~66144~66145~66146~66187~66195 | 2961 |  | 2961 2961 2961 2961 2961 2961 | 2014-08-26 | 15:39 | Christine Prior |  | 0.79 mg |  | 0.79 | [5 empty] | 0.99845 | 0.00185 | 66143 |  | XCAMS | ST | 0% |  | - pos# 24 and 26 needed an additional factor of 1.4 to explain its individual run-to-run variance | 2.8 | RRL | [5 empty] | Background Test |  | 2014-08-25 | Carbonate | Not Complete | JOB-200025 |  | low yield | 57398 |  | 14047/1 |  |  | 2014-08-14 | 1475 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.001 | 0.00191 | 66144 | [26 empty] | 2014-08-18 | 6283 | Graphite | Completed
[168 empty] | 1.0006 | 0.00195 | 66145 | [26 empty] | 2014-08-26 | 3167 | C13 Measurement | Completed
[168 empty] | 0.99775 | 0.00185 | 66146 | [26 empty] | 2014-08-27 | 8097 | AMS Submission Results | Completed
[168 empty] | 1.0032 | 0.00182 | 66187 | [26 empty] | 2014-08-27 | 9549 | Close Job | Completed
[168 empty] | 0.99891 | 0.00191 | 66195
[7 empty] | 117.3 | 89.0334 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51252 | 51251 | 51275 | 345 | 344 | 333 |  |  | jenny dahl | -998.31 | 0.07 ‰ | 2014-09-02 | 2014-09-02 | 7 | [5 empty] | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.31 ‰ | -998.31 ‰ | 0.07 | 0.07 | -998.28 | 0.07 ‰ | -998.28 | 0.07 ‰ |  |  | 2014-09-04 | Jocelyn Turnbull | 13:14 | 0.0016949 | 0.00169 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00169 | 7e-05 | 8184 | 2014-09-03 | 16:36 | Johannes Kaiser | 200154 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 23 | 5 | 6 | .X. | 0.00163 | 7e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00163 |  | 7e-05 |  | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.00169 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00163 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-01 | jenny dahl | 16:09 |  | Completed |  |  |  |  | 66332 | 2000000 |  |  | 66289 66295 66296 66297 66298 66299 | 66289~66295~66296~66297~66298~66299 | 2964 |  | 2964 2964 2964 2964 2964 2964 | 2014-09-04 | 10:05 | Christine Prior |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0018 | 0.00274 | 66289 |  | XCAMS | ST | 0.21% |  | - needed a system error of 0.21% - OX-I with TP 66297 and 66298 were ~ 0.3% low and TP 66299 was high by ~ 0.3% | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-10-20 | Carbonate |  | JOB-200154 | charge to 540GCT21 - test of waterline extraction method | Water nitrogen flush but carbonate vessel | 57494 | Normal | 14047/11 |  |  | 2014-08-25 | 1483 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00274 | 66295 | [26 empty] | 2014-08-25 | 2777 | Physical Pretreatment | Completed
[168 empty] | 0.99916 | 0.00276 | 66296 | [26 empty] | 2014-08-25 | 1485 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99732 | 0.00272 | 66297 | [26 empty] | 2014-08-26 | 6342 | Graphite | Completed
[168 empty] | 0.9966 | 0.00272 | 66298 | [26 empty] | 2014-08-27 | 3193 | C13 Measurement | Completed
[168 empty] | 1.0033 | 0.00275 | 66299 | [26 empty] | 2014-09-04 | 8184 | AMS Submission Results | Completed
[197 empty] | 2014-09-10 | 9619 | Close Job | Completed
[197 empty] | 2014-11-10 | 2088 | Invoice | Completed
[7 empty] | 115.7 | 88.8515 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49529 | 49528 | 49530 | 318 | 318 | 306 |  |  | jenny dahl | -997.9 | 0.08 ‰ | 2014-09-02 | 2014-09-02 | 7 | [5 empty] | 5.05‰ | 5.05‰ | 2.00‰ | 2.00‰ | 5.05 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -997.90 ‰ | -997.90 ‰ | 0.08 | 0.08 | -997.87 | 0.08 ‰ | -997.87 | 0.08 ‰ |  |  | 2014-09-04 | Jocelyn Turnbull | 13:14 | 0.0021004 | 0.0021 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.8 |  |  | 0.0021 | 8e-05 | 8185 | 2014-09-03 | 16:36 | Johannes Kaiser | 200155 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.21 | 0.01% | 3 | 5 | 4 | .X. | 0.00202 | 8e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00202 |  | 8e-05 |  | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.0021 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00202 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-01 | jenny dahl | 16:09 |  | Completed |  |  |  |  | 66333 | 2000000 |  |  | 66289 66295 66296 66297 66298 66299 | 66289~66295~66296~66297~66298~66299 | 2964 |  | 2964 2964 2964 2964 2964 2964 | 2014-09-04 | 10:05 | Christine Prior |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0018 | 0.00274 | 66289 |  | XCAMS | ST | 0.21% |  | - needed a system error of 0.21% - OX-I with TP 66297 and 66298 were ~ 0.3% low and TP 66299 was high by ~ 0.3% | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-10-20 | Carbonate | Complete | JOB-200155 | charge to 540GCT21 - test of waterline extraction method | Water nitrogen flush but carbonate vessel | 57495 | Normal | 14047/11 |  |  | 2014-08-25 | 1484 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00274 | 66295 | [26 empty] | 2014-08-26 | 6343 | Graphite | Completed
[168 empty] | 0.99916 | 0.00276 | 66296 | [26 empty] | 2014-08-27 | 3194 | C13 Measurement | Completed
[168 empty] | 0.99732 | 0.00272 | 66297 | [26 empty] | 2014-09-04 | 8185 | AMS Submission Results | Completed
[168 empty] | 0.9966 | 0.00272 | 66298 | [26 empty] | 2014-09-10 | 9620 | Close Job | Completed
[168 empty] | 1.0033 | 0.00275 | 66299 | [26 empty] | 2014-11-10 | 2089 | Invoice | Completed
[7 empty] | 115.7 | 88.956 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38727 | 38727 | 38726 | 124 | 124 | 120 | 38604 | 38852.8 | Kelly Lyons | -991.94 | 0.12 ‰ | 2014-09-18 | 2014-09-18 | 9 | [5 empty] | 2.56‰ | 2.56‰ | 2.00‰ | 2.00‰ | 2.56 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -991.94 ‰ | -991.94 ‰ | 0.12 | 0.12 | -991.84 | 0.13 ‰ | -991.84 | 0.13 ‰ |  | Christine Prior | 2014-09-23 | Jocelyn Turnbull | 10:27 | 0.0080584 | 0.00806 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00806 | 0.00012 | 8345 | 2014-09-22 | 9:48 | John West | 200243 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.81 | 0.01% | 18 | 5 | 4 | .X. | 0.00775 | 0.00012 |  |  |  |  | 1.0022 | 1.0022 | Yes |  |  | Horrible blank. Unacceptable - need to check conditions. Going to add marble blanks from recent wheels at BG correction for this wheel. |  |  |  | 0.00775 |  | 0.00012 |  | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00806 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00775 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-17 | Kelly Lyons | 8:09 |  | Completed |  |  |  |  | 66448 | 2000000 |  |  | 66301 66370 66371 66372 66373 66454 | 66301~66370~66371~66372~66373~66454 | 2968 |  | 2968 2968 2968 2968 2968 2968 | 2014-09-22 | 14:40 | Christine Prior |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99874 | 0.0014 | 66301 |  | XCAMS | ST | 0% |  | - 2 measurements done on this wheel - 1st measurement Judge=ON - all rotations with Judge=ON and cathodes skipped have been rejected due to unexplained run-to-run variance - no error corrections were needed - later rotations for cathode 36 (TP66338) rejec | 2.5 | RRL | [5 empty] | Background Test |  | 2014-11-03 | Carbonate | Complete | JOB-200243 |  | Water nitrogen flush but carbonate vessel forgot to write down total extraction time | 57662 | Normal | 14047/11 |  |  | 2014-09-08 | 1491 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00139 | 66370 | [26 empty] | 2014-09-09 | 6458 | Graphite | Completed
[168 empty] | 0.99934 | 0.00142 | 66371 | [26 empty] | 2014-09-11 | 3245 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00139 | 66372 | [26 empty] | 2014-09-23 | 8345 | AMS Submission Results | Completed
[168 empty] | 0.99961 | 0.00138 | 66373 | [26 empty] | 2014-09-26 | 9970 | Close Job | Completed
[168 empty] | 1.0007 | 0.00139 | 66454
[5 empty] | 5.4777e-05 | 0.0018707 | 121 | 89.7419 | 0.0681 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7153 | 7153 | 7152 | 29 | 29 | 20 | 7124.3 | 7182.3 | Kelly Lyons | -589.54 | 1.48 ‰ | 2014-09-18 | 2014-09-18 | 7 | [5 empty] | -6.22‰ | -6.22‰ | 2.00‰ | 2.00‰ | -6.22 ‰ | 2.00 ‰ | AMS | -8.09‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | [6 empty] | -584.47 | 1.50 ‰ |  |  | 2014-09-23 | Jocelyn Turnbull | 10:27 | 0.41046 | 0.41051 | 0.00148 | 0.00102 | 0.00103 | 0.00103 | 2 |  |  | 0.41161 | 0.00141 | 8346 | 2014-09-22 | 9:48 | John West | 200248 | 0.00183 |  | 0.00073 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 4 | ... | 0.39586 | 0.00088 |  |  |  |  | 1.0022 | 1.0022 | Yes |  |  | Also done on water line with N2 flush - so vac line can't be explanation for bad blank. | 0.00183 | 0.0018 | 0.00028 | 0.39475 |  | 0.00099 |  | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.41046 | 0 | 0.00148 | 0 | 0 | 0.00183 | 0 | 0.00073 | 0 | 0.39475 | 0 | 0.00099 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-17 | Kelly Lyons | 8:09 |  | Completed | 0.41047 | 0.41048 | 0.39476 | 0.39477 | 66463 | 2000000 | 66332 66333 | 66332~66333 | 66301 66370 66371 66372 66373 66454 | 66301~66370~66371~66372~66373~66454 | 2968 | 2964 2964 | 2968 2968 2968 2968 2968 2968 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.00 mg | 0.26 | 1 | 2014-09-02 | 0.00163 | 66332 | 2964 | 0.99 | 0.99874 | 0.0014 | 66301 | ST | XCAMS | ST | 0% |  | - 2 measurements done on this wheel - 1st measurement Judge=ON - all rotations with Judge=ON and cathodes skipped have been rejected due to unexplained run-to-run variance - no error corrections were needed - later rotations for cathode 36 (TP66338) rejec | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2014-11-05 | Water |  | JOB-200248 |  | waterline, nitrogen flush | 57665 | Normal | 14047/12 |  |  | 2014-09-10 | 1502 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-09-02 | 0.00202 | 66333 | 2964 | 1.1 | 1.0002 | 0.00139 | 66370 | [26 empty] | 2014-09-11 | 6473 | Graphite | Completed
[168 empty] | 0.99934 | 0.00142 | 66371 | [26 empty] | 2014-09-15 | 3258 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00139 | 66372 | [26 empty] | 2014-09-23 | 8346 | AMS Submission Results | Completed
[168 empty] | 0.99961 | 0.00138 | 66373 | [26 empty] | 2014-09-26 | 9971 | Close Job | Completed
[168 empty] | 1.0007 | 0.00139 | 66454
[7 empty] | 110.3 | 89.1032 | 0.0004 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 48990 | 48990 | 48976 | 298 | 297 | 286 |  |  | Kelly lyons | -997.75 | 0.08 ‰ | 2014-09-26 | 2014-09-26 | 9 | [5 empty] | 1.25‰ |  | 2.00‰ |  | 1.25 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0006 | 1.15 | 1.1547 | [6 empty] | -997.73 | 0.08 ‰ |  |  | 2014-10-02 | Jocelyn Turnbull | 13:37 | 0.002246 | 0.00225 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 4.5 |  |  | 0.00225 | 8e-05 | 8416 | 2014-10-02 | 14:22 | Johannes Kaiser | 200302 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 31 | 5 | 3 | ... | 0.00216 | 8e-05 |  |  |  |  | 1.0025 | 1.0025 | Yes | [6 empty] | 0.00216 |  | 8e-05 |  | Yes | 1 | 1 | 0.00096 | 0.00096 | 0.00225 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00216 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-24 | Kelly lyons | 13:17 |  | Completed |  |  |  |  | 66516 | 2000000 |  |  | 66506 66511 66512 | 66506~66511~66512 | 2970 |  | 2970 2970 2970 | 2014-10-02 | 14:41 | jenny dahl |  | 0.49 mg |  | 0.49 | [5 empty] | 1.0001 | 0.00166 | 66506 |  | XCAMS | EA | 0% |  | - this wheel is calibrated with Ox-I standards from EA310 only as the other Ox-I standards from EA312 are regarded as suspect to C-14 contamination --> 13/12 ratios of EA312 Ox-I were OK but TP66559 had its rts high by 0.5%, TP66560 had its rts high by 0. |  | RRL | [5 empty] | Background Inorganic |  | 2014-09-29 | Carbonate |  | JOB-200302 |  |  | 57732 |  | 14047/1 |  |  | 2014-09-16 | 1504 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0002 | 0.00165 | 66511 | [26 empty] | 2014-09-17 | 6526 | Graphite | Completed
[168 empty] | 0.99976 | 0.0017 | 66512 | [26 empty] | 2014-10-02 | 8416 | AMS Submission Results | Completed
[197 empty] | 2014-10-03 | 10043 | Close Job | Completed
[7 empty] | 104.4 | 88.9946 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41556 | 41555 | 41551 | 192 | 191 | 198 |  |  | Kelly Lyons | -994.33 | 0.14 ‰ | 2014-10-08 | 2014-10-08 | 20 | [5 empty] | 2.68‰ | 2.68‰ | 2.00‰ | 2.00‰ | 2.68 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | -994.33 ‰ | -994.33 ‰ | 0.14 | 0.14 | -994.26 | 0.14 ‰ | -994.26 | 0.14 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.0056669 | 0.00567 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 3.2 |  |  | 0.00567 | 0.00014 | 8476 | 2014-10-09 | 13:55 | Johannes Kaiser | 200296 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 32 | 5 | 4 | .X. | 0.00545 | 0.00013 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | not a great result since waters are fairly old in wheel and job notes indicate this sample had a low yield, I have changed category to Background test and will use TW 2964 TP 66333 for correction with other InO blank in TW 2972 |  |  |  | 0.00545 |  | 0.00013 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00567 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00545 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed |  |  |  |  | 66536 | 2000000 |  |  | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 |  | 2972 2972 2972 2972 2972 2972 | 2014-10-10 | 10:47 | jenny dahl |  | 0.59 mg |  | 0.59 | [5 empty] | 1.0011 | 0.00186 | 66592 |  | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-11-10 | Carbonate |  | JOB-200296 |  | Water nitrogen flush but carbonate vessel | 57814 | Normal | 14047/11 |  |  | 2014-09-17 | 1512 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00179 | 66639 | [26 empty] | 2014-09-18 | 6546 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [26 empty] | 2014-09-19 | 3271 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [26 empty] | 2014-10-10 | 8476 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [26 empty] | 2014-10-10 | 10134 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[5 empty] | 9.9054e-05 | 0.0051517 | 114.1 | 90.0078 | 0.0685 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7159 | 7159 | 7158 | 39 | 39 | 33 | 7119.8 | 7198.6 | Kelly Lyons | -589.84 | 2.01 ‰ | 2014-10-08 | 2014-10-08 | 20 | [5 empty] | -8.60‰ | -8.60‰ | 2.00‰ | 2.00‰ | -8.60 ‰ | 2.00 ‰ | AMS | -8.03‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -584.78 | 2.04 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.41016 | 0.41022 | 0.00201 | 0.00169 | 0.00171 | 0.00171 | 2 |  |  | 0.41326 | 0.00158 | 8477 | 2014-10-09 | 13:55 | Johannes Kaiser | 200297 | 0.00492 |  | 0.00197 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 4 | ... | 0.39744 | 0.00112 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | bit young, looks okay with correction second one done on waterline in a bottle with N flush | 0.00492 | 0.00495 | 0.00076 | 0.39446 |  | 0.00164 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.41016 | 0 | 0.00201 | 0 | 0 | 0.00492 | 0 | 0.00197 | 0 | 0.39446 | 0 | 0.00164 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed | 0.41016 | 0.41014 | 0.39446 | 0.39444 | 66537 | 2000000 | 66536 66541 | 66536~66541 | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 | 2972 2972 | 2972 2972 2972 2972 2972 2972 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.93 mg | 0.26 | 0.93 | 2014-10-08 | 0.00545 | 66536 | 2972 | 0.59 | 1.0011 | 0.00186 | 66592 | EA | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2014-11-10 | Water | Complete | JOB-200297 |  | Waterline std - low yeild, backup sample added | 57815 | Normal | 14047/12 |  |  | 2014-09-17 | 1513 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-10-08 | 0.00438 | 66541 | 2972 | 1.09 | 1.0015 | 0.00179 | 66639 | [26 empty] | 2014-09-18 | 6547 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [26 empty] | 2014-09-19 | 3272 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [26 empty] | 2014-10-10 | 8477 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [26 empty] | 2014-10-10 | 10135 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[7 empty] | 116.6 | 88.9177 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43311 | 43311 | 43319 | 257 | 256 | 265 |  |  | Kelly Lyons | -995.45 | 0.15 ‰ | 2014-10-08 | 2014-10-08 | 16 | [5 empty] | 3.55‰ | 3.55‰ | 2.00‰ | 2.00‰ | 3.55 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | -995.45 ‰ | -995.45 ‰ | 0.15 | 0.15 | -995.39 | 0.15 ‰ | -995.39 | 0.15 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.0045543 | 0.00455 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00455 | 0.00015 | 8482 | 2014-10-09 | 13:55 | Johannes Kaiser | 200316 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 18 | 5 | 4 | .X. | 0.00438 | 0.00014 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | not super great |  |  |  | 0.00438 |  | 0.00014 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00455 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00438 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed |  |  |  |  | 66541 | 2000000 |  |  | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 |  | 2972 2972 2972 2972 2972 2972 | 2014-10-10 | 10:47 | jenny dahl |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0011 | 0.00186 | 66592 |  | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-11-12 | Carbonate |  | JOB-200316 |  | Water line not sure how std prepared marked as test MN | 57820 | Normal | 14047/11 |  |  | 2014-09-18 | 1515 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00179 | 66639 | [26 empty] | 2014-09-22 | 6551 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [26 empty] | 2014-09-24 | 3277 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [26 empty] | 2014-10-10 | 8482 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [26 empty] | 2014-10-10 | 10136 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[5 empty] | 9.9054e-05 | 0.0051517 | 115.7 | 89.8682 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7070 | 7070 | 7069 | 39 | 39 | 33 | 7031.4 | 7109.6 | Kelly Lyons | -585.29 | 2.02 ‰ | 2014-10-08 | 2014-10-08 | 16 | [5 empty] | -7.06‰ | -7.06‰ | 2.00‰ | 2.00‰ | -7.06 ‰ | 2.00 ‰ | AMS | -8.10‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | [6 empty] | -580.17 | 2.04 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.41471 | 0.41478 | 0.00202 | 0.00168 | 0.00171 | 0.00171 | 2 |  |  | 0.41779 | 0.00159 | 8483 | 2014-10-09 | 13:55 | Johannes Kaiser | 200317 | 0.00492 |  | 0.00197 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 24 | 5 | 4 | ... | 0.4018 | 0.00112 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | bit young, with or without correction third one done on water line in bottle with N flush | 0.00492 | 0.00495 | 0.00076 | 0.39884 |  | 0.00164 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.41471 | 0 | 0.00202 | 0 | 0 | 0.00492 | 0 | 0.00197 | 0 | 0.39884 | 0 | 0.00164 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed | 0.41472 | 0.41469 | 0.39885 | 0.39882 | 66542 | 2000000 | 66536 66541 | 66536~66541 | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 | 2972 2972 | 2972 2972 2972 2972 2972 2972 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.08 mg | 0.26 | 1.08 | 2014-10-08 | 0.00545 | 66536 | 2972 | 0.59 | 1.0011 | 0.00186 | 66592 | EA | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2014-11-12 | Water |  | JOB-200317 |  | waterline | 57821 | Normal | 14047/12 |  |  | 2014-09-18 | 1516 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-10-08 | 0.00438 | 66541 | 2972 | 1.09 | 1.0015 | 0.00179 | 66639 | [26 empty] | 2014-09-22 | 6552 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [26 empty] | 2014-09-24 | 3278 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [26 empty] | 2014-10-10 | 8483 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [26 empty] | 2014-10-10 | 10137 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[5 empty] | 9.3582e-05 | 0.0029426 | 124.8 | 89.6998 | 0.0684 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7153 | 7153 | 7152 | 33 | 32 | 25 | 7120.3 | 7185.8 | Kelly Lyons | -589.53 | 1.67 ‰ | 2014-10-15 | 2014-10-15 | 12 | [5 empty] | -6.35‰ | -6.35‰ | 2.00‰ | 2.00‰ | -6.35 ‰ | 2.00 ‰ | AMS | -8.14‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0.82 | 0.8165 | [6 empty] | -584.46 | 1.69 ‰ |  |  | 2014-10-20 | Jocelyn Turnbull | 14:38 | 0.41047 | 0.41051 | 0.00167 | 0.00128 | 0.00129 | 0.00129 | 1.9 |  |  | 0.41225 | 0.00152 | 8567 | 2014-10-16 | 13:49 | Johannes Kaiser | 200397 | 0.00283 |  | 0.00113 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 4 | 5 | 4 | ... | 0.39647 | 0.00103 |  |  |  |  | 1.0016 | 1.0016 | Yes |  |  | No diff between waterline and carbonate line... until BG is chosen. | 0.00283 | 0.00283 | ? | 0.39476 |  | 0.00124 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.41047 | 0 | 0.00167 | 0 | 0 | 0.00283 | 0 | 0.00113 | 0 | 0.39476 | 0 | 0.00124 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-14 | Kelly Lyons | 10:28 |  | Completed | 0.41047 | 0.41047 | 0.39476 | 0.39476 | 66666 | 2000000 | 66667 | 66667 | 66694 66697 66699 66700 66703 66704 | 66694~66697~66699~66700~66703~66704 | 2974 | 2974 | 2974 2974 2974 2974 2974 2974 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.09 mg | 0.26 | 1.09 | 2014-10-15 | 0.00283 | 66667 | 2974 | 1.11 | 1.0033 | 0.00166 | 66694 | EA | XCAMS | EA | 0% |  | - good measurement - needed to apply a factor of 1.04 to explain the global run-to-run variance - cathode 14 (TP66712) needed a factor of 1.4 to explain its individual run-to-run variance | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2014-11-25 | Water |  | JOB-200397 |  | Waterline secondary | 57902 | Normal | 14047/12 |  |  | 2014-10-01 | 1544 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00168 | 66697 | [26 empty] | 2014-10-03 | 6676 | Graphite | Completed
[168 empty] | 0.99853 | 0.00165 | 66699 | [26 empty] | 2014-10-07 | 3305 | C13 Measurement | Completed
[168 empty] | 0.99903 | 0.00164 | 66700 | [26 empty] | 2014-10-20 | 8567 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00169 | 66703 | [26 empty] | 2014-10-22 | 10190 | Close Job | Completed
[168 empty] | 1 | 0.00162 | 66704
[7 empty] | 126.4 | 88.7293 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46820 | 46819 | 46827 | 255 | 255 | 246 |  |  | Kelly Lyons | -997.06 | 0.09 ‰ | 2014-10-15 | 2014-10-15 | 12 | [5 empty] | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0.82 | 0.8165 | -997.06 ‰ | -997.06 ‰ | 0.09 | 0.09 | -997.02 | 0.09 ‰ | -997.02 | 0.09 ‰ |  |  | 2014-10-20 | Jocelyn Turnbull | 14:38 | 0.0029426 | 0.00294 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00294 | 9e-05 | 8568 | 2014-10-16 | 13:49 | Johannes Kaiser | 200398 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.29 | 0.01% | 20 | 5 | 4 | .X. | 0.00283 | 9e-05 |  |  |  |  | 1.0016 | 1.0016 | Yes | [6 empty] | 0.00283 |  | 9e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00294 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00283 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-14 | Kelly Lyons | 10:28 |  | Completed |  |  |  |  | 66667 | 2000000 |  |  | 66694 66697 66699 66700 66703 66704 | 66694~66697~66699~66700~66703~66704 | 2974 |  | 2974 2974 2974 2974 2974 2974 | 2014-10-17 | 12:15 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 1.0033 | 0.00166 | 66694 |  | XCAMS | EA | 0% |  | - good measurement - needed to apply a factor of 1.04 to explain the global run-to-run variance - cathode 14 (TP66712) needed a factor of 1.4 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-11-25 | Carbonate | Complete | JOB-200398 |  | Water line not sure how std prepared | 57899 | Normal | 14047/11 |  |  | 2014-10-01 | 1545 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00168 | 66697 | [26 empty] | 2014-10-03 | 6677 | Graphite | Completed
[168 empty] | 0.99853 | 0.00165 | 66699 | [26 empty] | 2014-10-07 | 3306 | C13 Measurement | Completed
[168 empty] | 0.99903 | 0.00164 | 66700 | [26 empty] | 2014-10-20 | 8568 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00169 | 66703 | [26 empty] | 2014-10-22 | 10191 | Close Job | Completed
[168 empty] | 1 | 0.00162 | 66704
[7 empty] | 119 | 88.8634 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 48916 | 48916 | 48905 | 295 | 294 | 283 |  |  | Kelly Lyons | -997.73 | 0.08 ‰ | 2014-10-15 | 2014-10-15 | 14 | [5 empty] | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0.82 | 0.8165 | [6 empty] | -997.71 | 0.08 ‰ |  |  | 2014-10-20 | Jocelyn Turnbull | 14:38 | 0.0022668 | 0.00227 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00227 | 8e-05 | 8581 | 2014-10-16 | 13:49 | Johannes Kaiser | 200411 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 28 | 5 | 5 | ... | 0.00218 | 8e-05 |  |  |  |  | 1.0016 | 1.0016 | Yes | [6 empty] | 0.00218 |  | 8e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00227 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00218 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-14 | Kelly Lyons | 10:28 |  | Completed |  |  |  |  | 66633 | 2000000 |  |  | 66694 66697 66699 66700 66703 66704 | 66694~66697~66699~66700~66703~66704 | 2974 |  | 2974 2974 2974 2974 2974 2974 | 2014-10-17 | 12:15 | Christine Prior |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0033 | 0.00166 | 66694 |  | XCAMS | EA | 0% |  | - good measurement - needed to apply a factor of 1.04 to explain the global run-to-run variance - cathode 14 (TP66712) needed a factor of 1.4 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-10-29 | Carbonate |  | JOB-200411 |  | recreated job 200395 as R14047/1 so that R number and line evolution matched correctly | 57888 |  | 14047/1 |  |  | 2012-10-25 | 961 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00168 | 66697 | [26 empty] | 2014-10-01 | 1543 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99853 | 0.00165 | 66699 | [26 empty] | 2014-10-01 | 6643 | Graphite | Completed
[168 empty] | 0.99903 | 0.00164 | 66700 | [26 empty] | 2014-10-07 | 3301 | C13 Measurement | Completed
[168 empty] | 0.99904 | 0.00169 | 66703 | [26 empty] | 2014-10-20 | 8581 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00162 | 66704 | [26 empty] | 2014-10-22 | 10189 | Close Job | Completed
[7 empty] | 106.2 | 89.1196 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 50914 | 50913 | 50903 | 378 | 378 | 363 |  |  | Kelly lyons | -998.23 | 0.08 ‰ | 2014-11-09 | 2014-11-09 | 24 | [5 empty] | 0.08‰ | 0.08‰ | 2.00‰ | 2.00‰ | 0.08 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | [6 empty] | -998.21 | 0.08 ‰ |  | Christine Prior | 2014-11-11 | Jocelyn Turnbull | 12:43 | 0.0017677 | 0.00177 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00177 | 8e-05 | 8628 | 2014-11-11 | 9:26 | Johannes Kaiser | 200396 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 4 | 5 | 5 | ... | 0.0017 | 8e-05 |  |  |  |  | 1.0015 | 1.0015 | Yes | [6 empty] | 0.0017 |  | 8e-05 |  | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00177 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0017 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-04 | Kelly lyons | 15:52 |  | Completed |  |  |  |  | 66748 | 2000000 |  |  | 66455 66456 66457 66458 66791 66792 | 66455~66456~66457~66458~66791~66792 | 2976 |  | 2976 2976 2976 2976 2976 2976 | 2014-11-11 | 12:08 | Christine Prior |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99828 | 0.00232 | 66455 |  | XCAMS | ST | 0.14% |  | - first wheel after SNICS clean - needed an additonal factor of 1.05 to explain the global run-to-run variance - system error is 0.14% | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-10-14 | Carbonate |  | JOB-200396 |  | New subsample of standard material acid etched 30 September 2014. Create daughter jobs from this job using the physical pretreatment details for new jobs | 57963 |  | 14047/1 |  |  | 2014-10-01 | 2868 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99884 | 0.00228 | 66456 | [26 empty] | 2014-10-15 | 1549 | Carbonate CO2 Evolution | Completed
[168 empty] | 1 | 0.00229 | 66457 | [26 empty] | 2014-10-16 | 6758 | Graphite | Completed
[168 empty] | 1.0016 | 0.00228 | 66458 | [26 empty] | 2014-10-20 | 3313 | C13 Measurement | Completed
[168 empty] | 1.0036 | 0.00228 | 66791 | [26 empty] | 2014-11-11 | 8628 | AMS Submission Results | Completed
[168 empty] | 0.99759 | 0.00226 | 66792 | [26 empty] | 2014-11-12 | 10310 | Close Job | Completed
[7 empty] | 119.1 | 88.992 | 0.0003 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 49568 | 49568 | 49568 | 200 | 199 | 192 |  |  | Kelly lyons | -997.91 | 0.05 ‰ | 2014-11-10 | 2014-11-10 | 20 | [5 empty] | 0.66‰ | 0.66‰ | 2.00‰ | 2.00‰ | 0.66 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.70711 | [6 empty] | -997.88 | 0.05 ‰ |  |  | 2014-11-13 | Jocelyn Turnbull | 15:51 | 0.00209 | 0.00209 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 1.7 |  |  | 0.00209 | 5e-05 | 8698 | 2014-11-13 | 13:39 | Johannes Kaiser | 200511 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 22 | 5 | 5 | ... | 0.00201 | 5e-05 |  |  |  |  | 1.0007 | 1.0007 | Yes | [6 empty] | 0.00201 |  | 5e-05 |  | Yes | 1 | 1 | 0.00042 | 0.00042 | 0.00209 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.00201 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-06 | Kelly lyons | 12:08 |  | Completed |  |  |  |  | 66777 | 2000000 |  |  | 66787 66789 66790 66819 66820 66844 66847 66850 | 66787~66789~66790~66819~66820~66844~66847~66850 | 2977 |  | 2977 2977 2977 2977 2977 2977 2977 2977 | 2014-11-13 | 15:14 | Christine Prior |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99961 | 0.00122 | 66787 |  | XCAMS | EA | 0% |  | - good measurement - reached ~ 750k counts on average on each Ox-I - OX-I in pos#25 (TP66819) needed an additional factor of 1.2 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-11-03 | Carbonate |  | JOB-200511 |  |  | 58001 |  | 14047/1 |  |  | 2014-10-01 | 2868 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00116 | 66789 | [26 empty] | 2014-10-20 | 1555 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99951 | 0.00117 | 66790 | [26 empty] | 2014-10-21 | 6787 | Graphite | Completed
[168 empty] | 1.0018 | 0.00141 | 66819 | [26 empty] | 2014-10-24 | 3325 | C13 Measurement | Completed
[168 empty] | 1.001 | 0.00117 | 66820 | [26 empty] | 2014-11-13 | 8698 | AMS Submission Results | Completed
[168 empty] | 0.9995 | 0.00116 | 66844 | [26 empty] | 2014-11-13 | 10351 | Close Job | Completed
[168 empty] | 0.99793 | 0.00115 | 66847
[168 empty] | 0.99975 | 0.00117 | 66850
[5 empty] | 6.7274e-05 | 0.0025228 | 124.6 | 89.7122 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7043 | 7043 | 7040 | 38 | 38 | 32 | 7004.9 | 7081.8 | jenny dahl | -583.88 | 1.99 ‰ | 2014-11-24 | 2014-11-24 | 10 | [5 empty] | -6.97‰ | -6.97‰ | 2.00‰ | 2.00‰ | -6.97 ‰ | 2.00 ‰ | AMS | -8.04‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | [6 empty] | -578.75 | 2.02 ‰ |  |  | 2014-12-01 | Jocelyn Turnbull | 15:49 | 0.41612 | 0.41631 | 0.00199 | 0.00164 | 0.00167 | 0.0014 | 2.2 |  |  | 0.41778 | 0.00188 | 8747 | 2014-11-27 | 9:56 | Johannes Kaiser | 200698 | 0.00267 |  | 0.00107 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 24 | 5 | 4 | ... | 0.40179 | 0.00147 |  |  |  |  | 1.0011 | 1.0011 | Yes |  |  |  | 0.00267 | 0.00243 | 0.0011 | 0.40019 |  | 0.00161 |  | Yes | 1 | 1 | 0.00118 | 0.00118 | 0.41612 | 0 | 0.00199 | 0 | 0 | 0.00267 | 0 | 0.00107 | 0 | 0.40019 | 0 | 0.00161 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-21 | jenny dahl | 14:28 |  | Completed | 0.41612 | 0.41626 | 0.40019 | 0.40033 | 66932 | 2000000 | 66887 66921 | 66887~66921 | 66817 66854 66908 66911 66914 66917 | 66817~66854~66908~66911~66914~66917 | 2979 | 2979 2979 | 2979 2979 2979 2979 2979 2979 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.97 mg | 0.26 | 0.97 | 2014-11-24 | 0.00189 | 66887 | 2979 | 0.9 | 0.99761 | 0.00285 | 66817 | EA | XCAMS | EA | 0.22% |  | - stopped _1 measurement after ~4.5 rotations due to poor repeatability amongst the calibration standards - only _2 measurement data were used for reporting - system error = 0.22% - cat#5 (TP66822) needed an additional factor of 1.4 to explain its individ | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-01-07 | Water | Complete | JOB-200698 | standard, no charge | in water bottle with N flush material dried to attempt to improve yield | 58054 | Normal | 14047/12 |  |  | 2014-11-13 | 1601 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-11-24 | 0.00344 | 66921 | 2979 | 0.62 | 1.0024 | 0.00287 | 66854 | [26 empty] | 2014-11-14 | 6942 | Graphite | Completed
[168 empty] | 1.0024 | 0.00291 | 66908 | [26 empty] | 2014-12-01 | 3377 | C13 Measurement | Completed
[168 empty] | 0.99609 | 0.00289 | 66911 | [26 empty] | 2014-12-01 | 8747 | AMS Submission Results | Completed
[168 empty] | 0.99882 | 0.00287 | 66914 | [26 empty] | 2014-12-01 | 10412 | Close Job | Completed
[168 empty] | 1.0028 | 0.00289 | 66917
[7 empty] | 113.8 | 89.0032 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50063 | 50062 | 50043 | 340 | 340 | 326 |  |  | jenny dahl | -998.03 | 0.08 ‰ | 2014-11-24 | 2014-11-24 | 13 | [5 empty] | 0.94‰ | 0.94‰ | 2.00‰ | 2.00‰ | 0.94 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -998.03 ‰ | -998.03 ‰ | 0.08 | 0.08 | -998.01 | 0.08 ‰ | -998.01 | 0.08 ‰ |  |  | 2014-12-01 | Jocelyn Turnbull | 15:49 | 0.0019652 | 0.00197 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00197 | 8e-05 | 8765 | 2014-11-27 | 9:56 | Johannes Kaiser | 200651 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 31 | 5 | 4 | .X. | 0.00189 | 8e-05 |  |  |  |  | 1.0011 | 1.0011 | Yes | [6 empty] | 0.00189 |  | 8e-05 |  | Yes | 1 | 1 | 0.00118 | 0.00118 | 0.00197 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00189 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-21 | jenny dahl | 14:29 |  | Completed |  |  |  |  | 66887 | 2000000 |  |  | 66817 66854 66908 66911 66914 66917 | 66817~66854~66908~66911~66914~66917 | 2979 |  | 2979 2979 2979 2979 2979 2979 | 2014-11-27 | 11:10 | Christine Prior |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99761 | 0.00285 | 66817 |  | XCAMS | EA | 0.22% |  | - stopped _1 measurement after ~4.5 rotations due to poor repeatability amongst the calibration standards - only _2 measurement data were used for reporting - system error = 0.22% - cat#5 (TP66822) needed an additional factor of 1.4 to explain its individ | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-12-31 | Carbonate |  | JOB-200651 | standard, no charge | Water nitrogen flush but carbonate vessel | 58050 | Normal | 14047/11 |  |  | 2014-11-10 | 1587 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00287 | 66854 | [26 empty] | 2014-11-11 | 6897 | Graphite | Completed
[168 empty] | 1.0024 | 0.00291 | 66908 | [26 empty] | 2014-12-01 | 3371 | C13 Measurement | Completed
[168 empty] | 0.99609 | 0.00289 | 66911 | [26 empty] | 2014-12-01 | 8765 | AMS Submission Results | Completed
[168 empty] | 0.99882 | 0.00287 | 66914 | [26 empty] | 2014-12-01 | 10404 | Close Job | Completed
[168 empty] | 1.0028 | 0.00289 | 66917
[7 empty] | 120.5 | 88.8482 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45252 | 45251 | 45245 | 257 | 256 | 247 |  |  | jenny dahl | -996.42 | 0.11 ‰ | 2014-11-24 | 2014-11-24 | 11 | [5 empty] | 2.69‰ | 2.69‰ | 2.00‰ | 2.00‰ | 2.69 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.42 ‰ | -996.42 ‰ | 0.11 | 0.11 | -996.38 | 0.12 ‰ | -996.38 | 0.12 ‰ |  |  | 2014-12-01 | Jocelyn Turnbull | 15:49 | 0.0035769 | 0.00358 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3 |  |  | 0.00358 | 0.00011 | 8766 | 2014-11-27 | 9:56 | Johannes Kaiser | 200652 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 19 | 5 | 6 | ... | 0.00344 | 0.00011 |  |  |  |  | 1.0011 | 1.0011 | Yes |  |  | Interesting difference in blank run through water bottle versus carbonate vessel, even though on the same vacuum line. |  |  |  | 0.00344 |  | 0.00011 |  | Yes | 1 | 1 | 0.00118 | 0.00118 | 0.00358 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00344 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-21 | jenny dahl | 14:29 |  | Completed |  |  |  |  | 66921 | 2000000 |  |  | 66817 66854 66908 66911 66914 66917 | 66817~66854~66908~66911~66914~66917 | 2979 |  | 2979 2979 2979 2979 2979 2979 | 2014-11-27 | 11:10 | Christine Prior |  | 0.62 mg |  | 0.62 | [5 empty] | 0.99761 | 0.00285 | 66817 |  | XCAMS | EA | 0.22% |  | - stopped _1 measurement after ~4.5 rotations due to poor repeatability amongst the calibration standards - only _2 measurement data were used for reporting - system error = 0.22% - cat#5 (TP66822) needed an additional factor of 1.4 to explain its individ | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2014-12-31 | Carbonate |  | JOB-200652 | standard, no charge | water line nitrogen flush water vessel | 58051 | Normal | 14047/11 |  |  | 2014-11-11 | 1595 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00287 | 66854 | [26 empty] | 2014-11-11 | 1139 | Gas Splitting | Completed
[168 empty] | 1.0024 | 0.00291 | 66908 | [26 empty] | 2014-11-11 | 2299 | Choose Gas Splitting | Completed
[168 empty] | 0.99609 | 0.00289 | 66911 | [26 empty] | 2014-11-13 | 6931 | Graphite | Completed
[168 empty] | 0.99882 | 0.00287 | 66914 | [26 empty] | 2014-12-01 | 3376 | C13 Measurement | Completed
[168 empty] | 1.0028 | 0.00289 | 66917 | [26 empty] | 2014-12-01 | 8766 | AMS Submission Results | Completed
[197 empty] | 2014-12-01 | 10405 | Close Job | Completed
[7 empty] | 122.8 | 88.8922 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43516 | 43515 | 43516 | 207 | 206 | 199 |  |  | jenny dahl | -995.56 | 0.11 ‰ | 2014-12-11 | 2014-12-11 | 20 | [5 empty] | 2.83‰ | 2.83‰ | 2.00‰ | 2.00‰ | 2.83 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | -995.56 ‰ | -995.56 ‰ | 0.11 | 0.11 | -995.51 | 0.12 ‰ | -995.51 | 0.12 ‰ |  | Christine Prior | 2014-12-12 | Jocelyn Turnbull | 11:55 | 0.0044399 | 0.00444 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00444 | 0.00011 | 8912 | 2014-12-12 | 10:49 | Johannes Kaiser | 200767 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.44 | 0.01% | 3 | 5 | 4 | ... | 0.00427 | 0.00011 |  |  |  |  | 1.0017 | 1.0017 | Yes |  |  | hmmm, not so good. Blank from water bottle with N2 flush. Cf job 200652 TP66921; may be realistic when use of water vessels compared with carbonate vessels. |  |  |  | 0.00427 |  | 0.00011 |  | Yes | 1 | 1 | 0.00103 | 0.00103 | 0.00444 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00427 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-12-10 | jenny dahl | 14:19 |  | Completed |  |  |  |  | 66980 | 2000000 |  |  | 66903 66922 67039 67040 67041 67042 | 66903~66922~67039~67040~67041~67042 | 2983 |  | 2983 2983 2983 2983 2983 2983 | 2014-12-12 | 11:43 | Christine Prior |  | 1.08 mg |  | 1.08 | [5 empty] | 0.99785 | 0.00249 | 66903 |  | XCAMS | EA | 0.18% |  | #NAME? | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-01-14 | Carbonate | Complete | JOB-200767 | standard, no charge | water line nitrogen flush water vessel | 58233 | Normal | 14047/11 |  |  | 2014-11-19 | 1620 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99715 | 0.0025 | 66922 | [26 empty] | 2014-11-21 | 6990 | Graphite | Completed
[168 empty] | 1.0004 | 0.00257 | 67039 | [26 empty] | 2014-12-01 | 3385 | C13 Measurement | Completed
[168 empty] | 1.0042 | 0.00253 | 67040 | [26 empty] | 2014-12-12 | 8912 | AMS Submission Results | Completed
[168 empty] | 1.0007 | 0.00252 | 67041 | [26 empty] | 2014-12-12 | 10618 | Close Job | Completed
[168 empty] | 0.99991 | 0.00256 | 67042
[5 empty] | 8.0877e-05 | 0.0040084 | 121.2 | 89.8726 | 0.0683 | RRL-UNSt-LG | Unknown Organic |  |  |  |  | 7152 | 7151 | 7150 | 39 | 38 | 32 | 7113 | 7190.7 | jenny dahl | -589.47 | 1.99 ‰ | 2014-12-11 | 2014-12-11 | 20 | [5 empty] | -8.11‰ | -8.11‰ | 2.00‰ | 2.00‰ | -8.11 ‰ | 2.00 ‰ | AMS | -8.09‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | [6 empty] | -584.4 | 2.01 ‰ |  |  | 2014-12-12 | Jocelyn Turnbull | 11:55 | 0.41053 | 0.4106 | 0.00199 | 0.00166 | 0.00167 | 0.0015 | 1.8 |  |  | 0.41296 | 0.00173 | 8913 | 2014-12-12 | 10:49 | Johannes Kaiser | 200768 | 0.00386 |  | 0.00154 |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 4 | ... | 0.39716 | 0.00131 |  |  |  |  | 1.0017 | 1.0017 | Yes |  |  |  | 0.00386 | 0.00386 | 0.00059 | 0.39482 |  | 0.00161 |  | Yes | 1 | 1 | 0.00103 | 0.00103 | 0.41053 | 0 | 0.00199 | 0 | 0 | 0.00386 | 0 | 0.00154 | 0 | 0.39482 | 0 | 0.00161 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-12-10 | jenny dahl | 14:19 |  | Completed | 0.41054 | 0.41053 | 0.39483 | 0.39482 | 66981 | 2000000 | 66980 66921 | 66980~66921 | 66903 66922 67039 67040 67041 67042 | 66903~66922~67039~67040~67041~67042 | 2983 | 2983 2979 | 2983 2983 2983 2983 2983 2983 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.12 mg | 0.26 | 1.12 | 2014-12-11 | 0.00427 | 66980 | 2983 | 1.08 | 0.99785 | 0.00249 | 66903 | EA | XCAMS | EA | 0.18% |  | #NAME? | -8.1 | RRL | [5 empty] | Standard as Unknown |  | 2015-01-14 | Water |  | JOB-200768 |  | In water bottle with nitrogen flush | 58248 | Normal | 14047/12 |  |  | 2014-11-20 | 1621 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-11-24 | 0.00344 | 66921 | 2979 | 0.62 | 0.99715 | 0.0025 | 66922 | EA | [25 empty] | 2014-11-21 | 6991 | Graphite | Completed
[168 empty] | 1.0004 | 0.00257 | 67039 | [26 empty] | 2014-12-01 | 3386 | C13 Measurement | Completed
[168 empty] | 1.0042 | 0.00253 | 67040 | [26 empty] | 2014-12-12 | 8913 | AMS Submission Results | Completed
[168 empty] | 1.0007 | 0.00252 | 67041 | [26 empty] | 2014-12-12 | 10619 | Close Job | Completed
[168 empty] | 0.99991 | 0.00256 | 67042
[7 empty] | 116.9 | 89.1259 | 0.0004 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 47816 | 47815 | 47814 | 257 | 257 | 247 |  |  | Cathy Ginnane | -997.4 | 0.08 ‰ | 2015-01-26 | 2015-01-26 | 13 | [5 empty] | 1.91‰ | 1.91‰ | 2.00‰ | 2.00‰ | 1.91 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | [6 empty] | -997.37 | 0.08 ‰ |  | Christine Prior | 2015-01-28 | Jocelyn Turnbull | 15:21 | 0.0025995 | 0.0026 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.5 |  |  | 0.0026 | 8e-05 | 9001 | 2015-01-27 | 14:07 | Johannes Kaiser | 200765 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  |  |  | 4 | 5 | 5 | ... | 0.0025 | 8e-05 |  |  |  |  | 1.0034 | 1.0034 | Yes |  |  | a tad younger than pre-christmas carbonate blanks, but probably representative |  |  |  | 0.0025 |  | 8e-05 |  | No | 1 | 1 | 0.00096 | 0.00096 | 0.0026 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0025 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-01-21 | Cathy Ginnane | 9:20 |  | Completed |  |  |  |  | 67143 | 2000000 |  |  | 67102 67103 67106 67107 67108 67109 | 67102~67103~67106~67107~67108~67109 | 2986 |  | 2986 2986 2986 2986 2986 2986 | 2015-01-28 | 15:11 | Christine Prior |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99566 | 0.00235 | 67102 |  | XCAMS | ST | 0.17% |  | - rejected first rotation from analysis - needed a system error of 0.17% due to Ox-I in pos#0 (TP67102) being low by almost 3 sigma | 2.5 | RRL | [5 empty] | Background Inorganic | 2.5 | 2014-12-03 | Carbonate | Complete | JOB-200765 |  | carbonate line blank carbonate line not calibrated so measured on carbonate line and then measured on processing line to check mgC. 0.05mgC off. | 58339 |  | 14047/1 |  |  | 2015-01-12 | 1640 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0018 | 0.00236 | 67103 | [26 empty] | 2015-01-12 | 251 | Combine Water CO2 Evolution | Completed
[168 empty] | 1.0002 | 0.00235 | 67106 | [26 empty] | 2015-01-13 | 7153 | Graphite | Completed
[168 empty] | 0.99974 | 0.00235 | 67107 | [26 empty] | 2015-01-20 | 3417 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00235 | 67108 | [26 empty] | 2015-01-28 | 9001 | AMS Submission Results | Completed
[168 empty] | 1.0024 | 0.00241 | 67109 | [26 empty] | 2015-01-28 | 10781 | Close Job | Completed
[7 empty] | 123.9 | 88.9051 | 0.0018 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 36576 | 36576 | 36578 | 127 | 126 | 130 | 36450.2 | 36704 | Cathy Ginnane | -989.47 | 0.17 ‰ | 2015-01-26 | 2015-01-26 | 19 | [5 empty] | 4.40‰ | 4.40‰ | 2.00‰ | 2.00‰ | 4.40 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -989.47 ‰ | -989.47 ‰ | 0.17 | 0.17 | -989.34 | 0.17 ‰ | -989.34 | 0.17 ‰ |  |  | 2015-01-28 | Jocelyn Turnbull | 15:21 | 0.010533 | 0.01053 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.01053 | 0.00017 | 9030 | 2015-01-27 | 14:07 | Johannes Kaiser | 201062 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 24 | 5 | 4 | ... | 0.01013 | 0.00016 |  |  |  |  | 1.0034 | 1.0034 | Yes |  |  | so much younger than previous water line blanks, even considering N2 flush. Chose to use three most recent water line blanks with N2 flush instead First use of water line after holiday - maybe N2 flush not complete? |  |  |  | 0.01013 |  | 0.00016 |  | No | 1 | 1 | 0.00096 | 0.00096 | 0.01053 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01013 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-01-21 | Cathy Ginnane | 9:20 |  | Completed |  |  |  |  | 67105 | 2000000 |  |  | 67102 67103 67106 67107 67108 67109 | 67102~67103~67106~67107~67108~67109 | 2986 |  | 2986 2986 2986 2986 2986 2986 | 2015-01-28 | 15:11 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 0.99566 | 0.00235 | 67102 |  | XCAMS | ST | 0.17% |  | - rejected first rotation from analysis - needed a system error of 0.17% due to Ox-I in pos#0 (TP67102) being low by almost 3 sigma | 2.4 | RRL | [5 empty] | Background Test |  | 2015-02-16 | Carbonate |  | JOB-201062 |  | water line nitrogen flush water vessel | 58352 | Normal | 14047/11 |  |  | 2015-01-06 | 1641 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0018 | 0.00236 | 67103 | [26 empty] | 2015-01-07 | 7115 | Graphite | Completed
[168 empty] | 1.0002 | 0.00235 | 67106 | [26 empty] | 2015-01-13 | 3413 | C13 Measurement | Completed
[168 empty] | 0.99974 | 0.00235 | 67107 | [26 empty] | 2015-01-28 | 9030 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00235 | 67108 | [26 empty] | 2015-01-28 | 10788 | Close Job | Completed
[168 empty] | 1.0024 | 0.00241 | 67109
[5 empty] | 5.7987e-05 | 0.0030155 | 127.5 | 89.7089 | 0.0685 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7066 | 7066 | 7062 | 51 | 50 | 42 | 7015.7 | 7117 | Cathy Ginnane | -585.07 | 2.61 ‰ | 2015-01-26 | 2015-01-26 | 13 | [5 empty] | -4.60‰ | -4.60‰ | 2.00‰ | 2.00‰ | -4.60 ‰ | 2.00 ‰ | AMS | -8.19‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | [6 empty] | -579.95 | 2.65 ‰ |  |  | 2015-01-28 | Jocelyn Turnbull | 15:21 | 0.41493 | 0.41516 | 0.00261 | 0.00219 | 0.00238 | 0.00227 | 1.7 |  |  | 0.41693 | 0.00168 | 9031 | 2015-01-27 | 14:07 | Johannes Kaiser | 201063 | 0.0032 |  | 0.0032 |  | 100% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 4 | ... | 0.40097 | 0.00123 |  |  |  |  | 1.0034 | 1.0034 | Yes |  |  | still too young. JCT - if I use the poor water line blank, I get a much better number for this one. Increased error in blank to 100% because we are not very certain about the water line blank for this wheel | 0.0032 | 0.0029 | 0.00121 | 0.39905 |  | 0.00229 |  | No | 1 | 1 | 0.00096 | 0.00096 | 0.41493 | 0 | 0.00261 | 0 | 0 | 0.0032 | 0 | 0.0032 | 0 | 0.39905 | 0 | 0.00229 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-01-21 | Cathy Ginnane | 9:20 |  | Completed | 0.41493 | 0.41512 | 0.39905 | 0.39923 | 67149 | 2000000 | 66887 66921 66980 | 66887~66921~66980 | 67102 67103 67106 67107 67108 67109 | 67102~67103~67106~67107~67108~67109 | 2986 | 2979 2979 2983 | 2986 2986 2986 2986 2986 2986 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.17 mg | 0.26 | 1.17 | 2014-12-11 | 0.00427 | 66980 | 2983 | 1.08 | 0.99566 | 0.00235 | 67102 | EA | XCAMS | ST | 0.17% |  | - rejected first rotation from analysis - needed a system error of 0.17% due to Ox-I in pos#0 (TP67102) being low by almost 3 sigma | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-02-16 | Water |  | JOB-201063 |  | water vessel, nitrogen flush | 58355 | Normal | 14047/12 |  |  | 2015-01-12 | 1652 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2014-11-24 | 0.00189 | 66887 | 2979 | 0.9 | 1.0018 | 0.00236 | 67103 | EA | [25 empty] | 2015-01-13 | 7159 | Graphite | Completed
[163 empty] | 2014-11-24 | 0.00344 | 66921 | 2979 | 0.62 | 1.0002 | 0.00235 | 67106 | [26 empty] | 2015-01-20 | 3426 | C13 Measurement | Completed
[168 empty] | 0.99974 | 0.00235 | 67107 | [26 empty] | 2015-01-28 | 9031 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00235 | 67108 | [26 empty] | 2015-01-28 | 10789 | Close Job | Completed
[168 empty] | 1.0024 | 0.00241 | 67109
[7 empty] | 128.2 | 88.9732 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43348 | 43348 | 43354 | 203 | 202 | 195 |  |  | jenny dahl | -995.47 | 0.11 ‰ | 2015-02-11 | 2015-02-11 | 15 | [5 empty] | 3.22‰ | 3.22‰ | 2.00‰ | 2.00‰ | 3.22 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -995.47 ‰ | -995.47 ‰ | 0.11 | 0.11 | -995.41 | 0.12 ‰ | -995.41 | 0.12 ‰ |  |  | 2015-02-13 | Jocelyn Turnbull | 11:38 | 0.0045335 | 0.00453 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.2 |  |  | 0.00453 | 0.00011 | 9108 | 2015-02-12 | 16:27 | Johannes Kaiser | 201195 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.45 | 0.01% | 10 | 5 | 6 | ... | 0.00436 | 0.00011 |  |  |  |  | 1.003 | 1.003 | Yes |  |  | Blank looks like crap, but... compare this to the recent blanks-in-water-bottle-with-N2-flush. I think this is realistic blank |  |  |  | 0.00436 |  | 0.00011 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00453 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00436 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-05 | jenny dahl | 16:03 |  | Completed |  |  |  |  | 67228 | 2000000 |  |  | 67240 67241 67248 67263 67300 67303 | 67240~67241~67248~67263~67300~67303 | 2988 |  | 2988 2988 2988 2988 2988 2988 | 2015-02-13 | 10:51 | Christine Prior |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00154 | 67240 |  | XCAMS | EA | 0% |  | -very good measurement - pos#26 needed an additional factor of 1.3 due to individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-03-12 | Carbonate | Not Complete | JOB-201195 |  | water line nitrogen flush water vessel | 58417 | Normal | 14047/11 |  |  | 2015-01-26 | 1689 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00154 | 67241 | [27 empty] | 1164 | Gas Splitting | Completed
[168 empty] | 0.99923 | 0.00156 | 67248 | [26 empty] | 2015-01-27 | 2333 | Choose Gas Splitting | Completed
[168 empty] | 0.99904 | 0.00168 | 67263 | [26 empty] | 2015-01-27 | 7238 | Graphite | Completed
[168 empty] | 1.0006 | 0.00155 | 67300 | [26 empty] | 2015-02-02 | 3450 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00156 | 67303 | [26 empty] | 2015-02-13 | 9108 | AMS Submission Results | Completed
[197 empty] | 2015-02-13 | 10865 | Close Job | Completed
[5 empty] | 0.00011438 | 0.0045335 | 124.8 | 89.9895 | 0.0679 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7137 | 7137 | 7136 | 38 | 37 | 31 | 7099.3 | 7175.2 | jenny dahl | -588.72 | 1.94 ‰ | 2015-02-11 | 2015-02-11 | 15 | [5 empty] | -8.11‰ | -8.11‰ | 2.00‰ | 2.00‰ | -8.11 ‰ | 2.00 ‰ | AMS | -8.06‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | [6 empty] | -583.64 | 1.97 ‰ |  |  | 2015-02-13 | Jocelyn Turnbull | 11:38 | 0.41128 | 0.41136 | 0.00194 | 0.0016 | 0.00162 | 0.00162 | 1.8 |  |  | 0.41403 | 0.0015 | 9109 | 2015-02-12 | 16:27 | Johannes Kaiser | 201196 | 0.00436 |  | 0.00196 |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 30 | 5 | 4 | ... | 0.39818 | 0.001 |  |  |  |  | 1.003 | 1.003 | Yes |  |  | supports view that young blank for water bottles is realistic | 0.00436 | 0.00436 | ? | 0.39554 |  | 0.00156 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41128 | 0 | 0.00194 | 0 | 0 | 0.00436 | 0 | 0.00196 | 0 | 0.39554 | 0 | 0.00156 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-05 | jenny dahl | 16:04 |  | Completed | 0.41128 | 0.41128 | 0.39554 | 0.39554 | 67229 | 2000000 | 67228 | 67228 | 67240 67241 67248 67263 67300 67303 | 67240~67241~67248~67263~67300~67303 | 2988 | 2988 | 2988 2988 2988 2988 2988 2988 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.14 mg | 0.26 | 1.14 | 2015-02-11 | 0.00436 | 67228 | 2988 | 0.99 | 1.0003 | 0.00154 | 67240 | EA | XCAMS | EA | 0% |  | -very good measurement - pos#26 needed an additional factor of 1.3 due to individual run-to-run variance | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-03-12 | Water | Complete | JOB-201196 |  | water bottle w nitrogen flush | 58418 | Normal | 14047/12 |  |  | 2015-01-26 | 1690 | Carbonate CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00154 | 67241 | [26 empty] | 2015-01-27 | 7239 | Graphite | Completed
[168 empty] | 0.99923 | 0.00156 | 67248 | [26 empty] | 2015-02-02 | 3456 | C13 Measurement | Completed
[168 empty] | 0.99904 | 0.00168 | 67263 | [26 empty] | 2015-02-13 | 9109 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00155 | 67300 | [26 empty] | 2015-02-13 | 10866 | Close Job | Completed
[168 empty] | 1.0004 | 0.00156 | 67303
[7 empty] | 127.4 | 88.8635 | 0.0007 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 43822 | 43822 | 43829 | 235 | 234 | 226 |  |  | Margaret Norris | -995.73 | 0.12 ‰ | 2015-02-17 | 2015-02-17 | 14 | [5 empty] | 4.14‰ | 4.14‰ | 2.00‰ | 2.00‰ | 4.14 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0026 | 0 | 0.8165 | [6 empty] | -995.67 | 0.13 ‰ |  |  | 2015-03-04 | jenny dahl | 15:31 | 0.0042736 | 0.00427 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00427 | 0.00012 | 9152 | 2015-02-18 | 9:37 | Johannes Kaiser | 201317 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 18 | 5 | 6 | .X. | 0.00411 | 0.00012 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | What? 0.3mg blank in this wheel is better than 1mg blank? |  |  |  | 0.00411 |  | 0.00012 |  | No | 1 | 0.99987 | 0.00142 | 0.00142 | 0.00427 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00411 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-13 | Margaret Norris | 16:23 |  | Completed |  |  |  |  | 67287 | 2000000 |  |  | 67169 67185 67188 67204 67346 67351 | 67169~67185~67188~67204~67346~67351 | 2989 |  | 2989 2989 2989 2989 2989 2989 | 2015-03-02 | 13:31 | jenny dahl |  | 1.07 mg |  | 1.07 | [5 empty] | 1.0026 | 0.00346 | 67169 |  | XCAMS | EA | 0.29% |  | - system error = 0.29% - needed an additional error adjustment factor of 1.02 to explain the global run-to-run variance - needed an additional error adjustment factor of 1.3 for pos# 31 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-02-13 | Carbonate | Complete | JOB-201317 |  | regular sized C13 tuning split- job 201328 | 58530 |  | 14047/1 |  |  | 2015-02-02 | 1718 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99715 | 0.0035 | 67185 | [26 empty] | 2015-02-02 | 1166 | Gas Splitting | Completed
[168 empty] | 0.99569 | 0.0035 | 67188 | [26 empty] | 2015-02-02 | 2335 | Choose Gas Splitting | Completed
[168 empty] | 0.99727 | 0.00352 | 67204 | [26 empty] | 2015-02-03 | 7297 | Graphite | Completed
[168 empty] | 1.0032 | 0.00347 | 67346 | [26 empty] | 2015-02-09 | 3468 | C13 Measurement | Completed
[168 empty] | 1.0031 | 0.00347 | 67351 | [26 empty] | 2015-03-04 | 9152 | AMS Submission Results | Completed
[197 empty] | 2015-03-12 | 10953 | Close Job | Completed
[7 empty] | 106.8 | 89.182 | 0.0007 | RRL-CBIn-SM | Background Inorganic |  |  |  |  | 44061 | 44060 | 44058 | 242 | 241 | 232 |  |  | Margaret Norris | -995.85 | 0.12 ‰ | 2015-02-17 | 2015-02-17 | 14 | [5 empty] | 0.55‰ | 0.55‰ | 2.00‰ | 2.00‰ | 0.55 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0026 | 0 | 0.8165 | [6 empty] | -995.8 | 0.13 ‰ |  |  | 2015-03-04 | jenny dahl | 15:31 | 0.0041488 | 0.00415 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 7.2 |  |  | 0.00415 | 0.00012 | 9153 | 2015-02-18 | 9:37 | Johannes Kaiser | 201318 |  |  |  |  | 45% | 5 | 9/10/19 | 13:11:09 | jenny dahl | 0 |  |  |  | 30 | 5 | 3 | ... | 0.00399 | 0.00012 |  |  |  |  | 1.0027 | 1.0027 | Yes | [6 empty] | 0.00399 |  | 0.00012 |  | No | 1 | 0.99987 | 0.00142 | 0.00142 | 0.00415 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00399 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-13 | Margaret Norris | 16:23 |  | Completed |  |  |  |  | 67278 | 2000000 |  |  | 67169 67185 67188 67204 67346 67351 | 67169~67185~67188~67204~67346~67351 | 2989 |  | 2989 2989 2989 2989 2989 2989 | 2015-03-02 | 13:31 | jenny dahl |  | 0.31 mg |  | 0.31 | [5 empty] | 1.0026 | 0.00346 | 67169 |  | XCAMS | EA | 0.29% |  | - system error = 0.29% - needed an additional error adjustment factor of 1.02 to explain the global run-to-run variance - needed an additional error adjustment factor of 1.3 for pos# 31 to explain its individual run-to-run variance |  | RRL | [5 empty] | Background Inorganic |  | 2015-02-13 | Carbonate |  | JOB-201318 |  | smaller~0.2 | 58531 |  | 14047/1 |  |  | 2015-02-02 | 1719 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99715 | 0.0035 | 67185 | [26 empty] | 2015-02-03 | 7288 | Graphite | Completed
[168 empty] | 0.99569 | 0.0035 | 67188 | [26 empty] | 2015-03-04 | 9153 | AMS Submission Results | Completed
[168 empty] | 0.99727 | 0.00352 | 67204 | [26 empty] | 2015-03-12 | 10954 | Close Job | Completed
[168 empty] | 1.0032 | 0.00347 | 67346
[168 empty] | 1.0031 | 0.00347 | 67351
[7 empty] | 44.9 | 89.8176 | 0.0015 | RRL-CBIn-TY | Background Inorganic |  |  |  |  | 37882 | 37882 | 37884 | 261 | 261 | 260 |  |  | Margaret Norris | -991.05 | 0.29 ‰ | 2015-02-17 | 2015-02-17 | 12 | [5 empty] | -6.53‰ | -6.53‰ | 2.00‰ | 2.00‰ | -6.53 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0026 | 0 | 0.8165 | [6 empty] | -990.94 | 0.29 ‰ |  |  | 2015-03-04 | jenny dahl | 15:31 | 0.0089526 | 0.00895 | 0.00029 | 0.00029 | 0.00029 | 0.00029 | 18.4 |  |  | 0.00895 | 0.00029 | 9155 | 2015-02-18 | 9:37 | Johannes Kaiser | 201333 |  |  |  |  | 45% | 5 | 9/10/19 | 13:11:35 | jenny dahl | 0 |  |  |  | 10 | 5 | 3 | ... | 0.00861 | 0.00028 |  |  |  |  | 1.0027 | 1.0027 | Yes | [6 empty] | 0.00861 |  | 0.00028 |  | No | 1 | 0.99987 | 0.00142 | 0.00142 | 0.00895 | 0 | 0.00029 | 0 | 0 |  |  | 0 | 0 | 0.00861 | 0 | 0.00028 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-02-13 | Margaret Norris | 16:23 |  | Completed |  |  |  |  | 67313 | 2000000 |  |  | 67169 67185 67188 67204 67346 67351 | 67169~67185~67188~67204~67346~67351 | 2989 |  | 2989 2989 2989 2989 2989 2989 | 2015-03-02 | 13:31 | jenny dahl |  | 0.12 mg |  | 0.12 | [5 empty] | 1.0026 | 0.00346 | 67169 |  | XCAMS | EA | 0.29% |  | - system error = 0.29% - needed an additional error adjustment factor of 1.02 to explain the global run-to-run variance - needed an additional error adjustment factor of 1.3 for pos# 31 to explain its individual run-to-run variance |  | RRL | [5 empty] | Background Inorganic |  | 2015-02-18 | Carbonate |  | JOB-201333 |  | SMALL had lower than normal graphite yield. Chris says okay to run | 58533 |  | 14047/1 |  |  | 2015-02-04 | 1737 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99715 | 0.0035 | 67185 | [26 empty] | 2015-02-05 | 7323 | Graphite | Completed
[168 empty] | 0.99569 | 0.0035 | 67188 | [26 empty] | 2015-03-04 | 9155 | AMS Submission Results | Completed
[168 empty] | 0.99727 | 0.00352 | 67204 | [26 empty] | 2015-03-12 | 10955 | Close Job | Completed
[168 empty] | 1.0032 | 0.00347 | 67346
[168 empty] | 1.0031 | 0.00347 | 67351
[7 empty] | 127.9 | 88.8574 | 0.0005 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 46298 | 46297 | 46298 | 213 | 212 | 205 |  |  | Christine Prior | -996.86 | 0.08 ‰ | 2015-02-20 | 2015-02-20 | 7 | [5 empty] | 2.35‰ | 2.35‰ | 2.00‰ | 2.00‰ | 2.35 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | [6 empty] | -996.82 | 0.08 ‰ |  |  | 2015-02-27 | jenny dahl | 9:13 | 0.0031402 | 0.00314 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.4 |  |  | 0.00314 | 8e-05 | 9184 | 2015-02-25 | 11:18 | Johannes Kaiser | 201334 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 4 | 5 | 6 | ... | 0.00302 | 8e-05 |  |  |  |  | 1.0009 | 1.0009 | Yes | [6 empty] | 0.00302 |  | 8e-05 |  | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00314 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00302 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-19 | Christine Prior | 15:54 |  | Completed |  |  |  |  | 67360 | 2000000 |  |  | 67354 67387 67411 67412 67413 67414 | 67354~67387~67411~67412~67413~67414 | 2990 |  | 2990 2990 2990 2990 2990 2990 | 2015-02-25 | 14:57 | jenny dahl |  | 0.81 mg |  | 0.81 | [5 empty] | 0.9963 | 0.00206 | 67354 |  | XCAMS | EA | 0.13% |  | - system error = 0.13% and this is mainly due to Ox-I in pos# 7 (TP67354) being low by ~2.5 sigma - cat# 24 (TP67355) and cat# 27 (TP67335) needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP6730 | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-02-18 | Carbonate | Complete | JOB-201334 |  | No prebake yield? | 58482 |  | 14047/1 |  |  | 2015-02-11 | 1758 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0019 | 0.00209 | 67387 | [26 empty] | 2015-02-11 | 1173 | Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00206 | 67411 | [26 empty] | 2015-02-11 | 2345 | Choose Gas Splitting | Completed
[168 empty] | 0.99951 | 0.00204 | 67412 | [26 empty] | 2015-02-13 | 7370 | Graphite | Completed
[168 empty] | 0.99978 | 0.00209 | 67413 | [26 empty] | 2015-02-16 | 3509 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00205 | 67414 | [26 empty] | 2015-02-27 | 9184 | AMS Submission Results | Completed
[197 empty] | 2015-02-27 | 10912 | Close Job | Completed
[7 empty] | 131.5 | 88.8213 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45067 | 45067 | 45067 | 251 | 251 | 241 |  |  | Christine Prior | -996.34 | 0.11 ‰ | 2015-02-20 | 2015-02-20 | 8 | [5 empty] | 2.76‰ | 2.76‰ | 2.00‰ | 2.00‰ | 2.76 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.34 ‰ | -996.34 ‰ | 0.11 | 0.11 | -996.29 | 0.12 ‰ | -996.29 | 0.12 ‰ |  |  | 2015-02-27 | jenny dahl | 9:13 | 0.0036601 | 0.00366 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.1 |  |  | 0.00366 | 0.00011 | 9190 | 2015-02-25 | 11:18 | Johannes Kaiser | 201353 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 24 | 5 | 6 | ... | 0.00352 | 0.00011 |  |  |  |  | 1.0009 | 1.0009 | Yes |  |  | - needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance |  |  |  | 0.00352 |  | 0.00011 |  | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00366 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00352 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-19 | Christine Prior | 15:54 |  | Completed |  |  |  |  | 67355 | 2000000 |  |  | 67354 67387 67411 67412 67413 67414 | 67354~67387~67411~67412~67413~67414 | 2990 |  | 2990 2990 2990 2990 2990 2990 | 2015-02-25 | 14:57 | jenny dahl |  | 1.08 mg |  | 1.08 | [5 empty] | 0.9963 | 0.00206 | 67354 |  | XCAMS | EA | 0.13% |  | - system error = 0.13% and this is mainly due to Ox-I in pos# 7 (TP67354) being low by ~2.5 sigma - cat# 24 (TP67355) and cat# 27 (TP67335) needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP6730 | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-02 | Water | Complete | JOB-201353 |  | water line nitrogen flush water vessel | 58488 | Normal | 14047/11 |  |  | 2015-02-11 | 1763 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00209 | 67387 | [26 empty] | 2015-02-11 | 1174 | Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00206 | 67411 | [26 empty] | 2015-02-11 | 2347 | Choose Gas Splitting | Completed
[168 empty] | 0.99951 | 0.00204 | 67412 | [26 empty] | 2015-02-12 | 7365 | Graphite | Completed
[168 empty] | 0.99978 | 0.00209 | 67413 | [26 empty] | 2015-02-13 | 3492 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00205 | 67414 | [26 empty] | 2015-02-27 | 9190 | AMS Submission Results | Completed
[197 empty] | 2015-02-27 | 10904 | Close Job | Completed
[5 empty] | 6.7274e-05 | 0.00332 | 128.2 | 89.7902 | 0.0678 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7128 | 7128 | 7126 | 36 | 35 | 29 | 7092.6 | 7164.4 | Christine Prior | -588.27 | 1.84 ‰ | 2015-02-20 | 2015-02-20 | 8 | [5 empty] | -8.06‰ | -8.06‰ | 2.00‰ | 2.00‰ | -8.06 ‰ | 2.00 ‰ | AMS | -8.13‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | [6 empty] | -583.19 | 1.86 ‰ |  |  | 2015-02-27 | jenny dahl | 9:13 | 0.41173 | 0.41183 | 0.00184 | 0.00147 | 0.0015 | 0.0014 | 2 |  |  | 0.41379 | 0.00159 | 9191 | 2015-02-25 | 11:18 | Johannes Kaiser | 201354 | 0.00327 |  | 0.00147 |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 30 | 5 | 6 | ... | 0.39795 | 0.00113 |  |  |  |  | 1.0009 | 1.0009 | Yes |  |  |  | 0.00327 | 0.00319 | 0.00035 | 0.39597 |  | 0.00144 |  | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.41173 | 0 | 0.00184 | 0 | 0 | 0.00327 | 0 | 0.00147 | 0 | 0.39597 | 0 | 0.00144 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-19 | Christine Prior | 15:54 |  | Completed | 0.41173 | 0.41178 | 0.39597 | 0.39602 | 67352 | 2000000 | 67355 67360 | 67355~67360 | 67354 67387 67411 67412 67413 67414 | 67354~67387~67411~67412~67413~67414 | 2990 | 2990 2990 | 2990 2990 2990 2990 2990 2990 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.06 mg | 0.26 | 1.06 | 2015-02-20 | 0.00352 | 67355 | 2990 | 1.08 | 0.9963 | 0.00206 | 67354 | EA | XCAMS | EA | 0.13% |  | - system error = 0.13% and this is mainly due to Ox-I in pos# 7 (TP67354) being low by ~2.5 sigma - cat# 24 (TP67355) and cat# 27 (TP67335) needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP6730 | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-04-02 | Water |  | JOB-201354 |  |  | 58489 | Normal | 14047/12 |  |  | 2015-02-11 | 1765 | Water CO2 Evolution | Completed | 14047/1 | IAEA-C1: Carrara Marble | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-02-20 | 0.00302 | 67360 | 2990 | 0.81 | 1.0019 | 0.00209 | 67387 | [26 empty] | 2015-02-11 | 1175 | Gas Splitting | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE
[168 empty] | 1.0017 | 0.00206 | 67411 | [26 empty] | 2015-02-11 | 2349 | Choose Gas Splitting | Completed
[168 empty] | 0.99951 | 0.00204 | 67412 | [26 empty] | 2015-02-12 | 7362 | Graphite | Completed
[168 empty] | 0.99978 | 0.00209 | 67413 | [26 empty] | 2015-02-13 | 3496 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00205 | 67414 | [26 empty] | 2015-02-27 | 9191 | AMS Submission Results | Completed
[197 empty] | 2015-02-27 | 10905 | Close Job | Completed
[7 empty] | 82.9 | 89.0554 | 0.0009 | RRL-Test-TY | Background Test |  |  |  |  | 41810 | 41810 | 41810 | 289 | 289 | 293 | 41526.1 | 42104.8 | Christine Prior | -994.51 | 0.20 ‰ | 2015-02-25 | 2015-02-25 | 14 | [5 empty] | 2.34‰ | 2.34‰ | 2.00‰ | 2.00‰ | 2.34 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0002 | 0 | 0.75593 | [6 empty] | -994.44 | 0.20 ‰ |  |  | 2015-03-11 | Jocelyn Turnbull | 15:21 | 0.0054901 | 0.00549 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 18.3 |  |  | 0.00549 | 0.0002 | 9238 | 2015-02-27 | 10:58 | Johannes Kaiser | 201352 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 23 | 5 | 3 | ... | 0.00528 | 0.00019 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes | [5 empty] | 0.00528 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00047 | 0.00047 | 0.00549 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.00528 | 0 | 0.00019 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-02-20 | Christine Prior | 10:53 |  | Completed |  |  |  |  | 67336 | 2000000 |  |  | 67342 67353 67381 67382 67383 67405 67415 | 67342~67353~67381~67382~67383~67405~67415 | 2991 |  | 2991 2991 2991 2991 2991 2991 2991 | 2015-02-27 | 15:46 | jenny dahl |  | 0.11 mg | 0.12 | 0.11 | [5 empty] | 0.99901 | 0.00125 | 67342 |  | XCAMS | EA | 0% |  | - good measurement - calibration Ox-I in pos#15 (TP67380) was rejected from calibration brecause being high by 3.6 sigma/4.9 per mille - rejected all runs with le12C current being lower than 30 uA - rejected the last 3 rotations to avoid accuracy loss in |  | RRL | [5 empty] | Background Test |  | 2015-02-19 | Carbonate | Complete | JOB-201352 | Small test - charge to 540GCT21 | confirmed test sample make small ~0.1ish | 58537 |  | 14047/1 |  |  | 2015-02-10 | 1748 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0002 | 0.00127 | 67353 | [26 empty] | 2015-02-11 | 7346 | Graphite | Completed
[168 empty] | 1.0015 | 0.00126 | 67381 | [26 empty] | 2015-03-11 | 9238 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00125 | 67382 | [26 empty] | 2015-03-12 | 10973 | Close Job | Completed
[168 empty] | 1.0004 | 0.00126 | 67383 | [26 empty] | 2015-06-10 | 2990 | Invoice | Completed
[168 empty] | 0.99772 | 0.00125 | 67405
[168 empty] | 1.0008 | 0.00123 | 67415
[7 empty] | 94.4 | 89.4691 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46624 | 46623 | 46611 | 249 | 249 | 239 |  |  | jenny dahl | -996.98 | 0.09 ‰ | 2015-03-06 | 2015-03-06 | 29 | [5 empty] | -5.21‰ | -5.21‰ | 2.00‰ | 2.00‰ | -5.21 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | -996.98 ‰ | -996.98 ‰ | 0.09 | 0.09 | -996.95 | 0.09 ‰ | -996.95 | 0.09 ‰ |  |  | 2015-03-11 | Jocelyn Turnbull | 15:24 | 0.0030154 | 0.00302 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.5 |  |  | 0.00302 | 9e-05 | 9262 | 2015-03-09 | 11:26 | Johannes Kaiser | 201236 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.3 | 0.01% | 4 | 5 | 6 | ... | 0.0029 | 9e-05 |  |  |  |  | 1.0001 | 1.0001 | Yes | [6 empty] | 0.0029 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00302 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0029 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-05 | jenny dahl | 15:41 |  | Completed |  |  |  |  | 67304 | 2000000 |  |  | 67457 67458 67513 67514 67515 | 67457~67458~67513~67514~67515 | 2992 |  | 2992 2992 2992 2992 2992 | 2015-03-11 | 11:50 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 1.001 | 0.00154 | 67457 |  | XCAMS | EA | 0% |  | - no system error - Ox-I (TP67494) was rejected from calibration as it was low by 4.3 sigma - cat# 18 (TP67465) needed an additional factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP67377) needed an additional factor of 1.4 to expl | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-03-19 | Carbonate |  | JOB-201236 |  | water line nitrogen flush water vessel | 58588 | Normal | 14047/11 |  |  | 2015-02-03 | 1725 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99995 | 0.00153 | 67458 | [26 empty] | 2015-02-03 | 1168 | Gas Splitting | Completed
[168 empty] | 0.99914 | 0.00152 | 67513 | [26 empty] | 2015-02-03 | 2338 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 67514 | [26 empty] | 2015-02-05 | 7314 | Graphite | Completed
[168 empty] | 0.99974 | 0.00155 | 67515 | [26 empty] | 2015-02-10 | 3476 | C13 Measurement | Completed
[197 empty] | 2015-03-11 | 9262 | AMS Submission Results | Completed
[197 empty] | 2015-03-12 | 10974 | Close Job | Completed
[5 empty] | 6.6172e-05 | 0.0033949 | 133 | 89.6061 | 0.0673 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7181 | 7181 | 7180 | 29 | 29 | 27 | 7152.1 | 7210.3 | jenny dahl | -590.97 | 1.48 ‰ | 2015-03-06 | 2015-03-06 | 29 | [5 empty] | -6.74‰ | -6.74‰ | 2.00‰ | 2.00‰ | -6.74 ‰ | 2.00 ‰ | AMS | -8.04‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | [6 empty] | -585.92 | 1.50 ‰ |  |  | 2015-03-11 | Jocelyn Turnbull | 15:24 | 0.40903 | 0.40909 | 0.00148 | 0.00137 | 0.00138 | 0.00138 | 1.7 |  |  | 0.41109 | 0.00115 | 9263 | 2015-03-09 | 11:26 | Johannes Kaiser | 201237 | 0.00327 | 0.0034 | 0.00147 | 0.0015 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 6 | ... | 0.39536 | 0.00098 |  |  |  |  | 1.0001 | 1.0001 | Yes |  |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 34.2613386391903102. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00327 | 0.00327 | 0.00052 | 0.39338 | 0.39‰ | 0.00133 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.40903 | 0 | 0.00148 | 0 | 0 | 0.00327 | 0 | 0.00147 | 0 | 0.39338 | 0 | 0.00133 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-05 | jenny dahl | 15:41 |  | Completed | 0.40903 | 0.40903 | 0.39338 | 0.39338 | 67305 | 2000000 | 67304 67428 | 67304~67428 | 67457 67458 67513 67514 67515 | 67457~67458~67513~67514~67515 | 2992 | 2992 2992 | 2992 2992 2992 2992 2992 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.07 mg | 0.13 | 1.07 | 2015-03-06 | 0.0029 | 67304 | 2992 | 1.11 | 1.001 | 0.00154 | 67457 | EA | XCAMS | EA | 0% |  | - no system error - Ox-I (TP67494) was rejected from calibration as it was low by 4.3 sigma - cat# 18 (TP67465) needed an additional factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP67377) needed an additional factor of 1.4 to expl | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-03-19 | Water | Complete | JOB-201237 |  |  | 58591 | Normal | 14047/12 |  |  | 2015-02-03 | 1727 | Water CO2 Evolution | Completed | 14047/1 | IAEA-C1: Carrara Marble | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-03-06 | 0.00363 | 67428 | 2992 | 1.19 | 0.99995 | 0.00153 | 67458 | [26 empty] | 2015-02-03 | 1169 | Gas Splitting | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE
[168 empty] | 0.99914 | 0.00152 | 67513 | [26 empty] | 2015-02-03 | 2340 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 67514 | [26 empty] | 2015-02-05 | 7315 | Graphite | Completed
[168 empty] | 0.99974 | 0.00155 | 67515 | [26 empty] | 2015-02-10 | 3477 | C13 Measurement | Completed
[197 empty] | 2015-03-11 | 9263 | AMS Submission Results | Completed
[197 empty] | 2015-03-12 | 10975 | Close Job | Completed
[7 empty] | 133.4 | 88.6478 | 0.0006 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 44820 | 44820 | 44830 | 199 | 199 | 192 |  |  | jenny dahl | -996.23 | 0.09 ‰ | 2015-03-06 | 2015-03-06 | 11 | [5 empty] | 4.00‰ | 4.00‰ | 2.00‰ | 2.00‰ | 4.00 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0016 | 0 | 0.89443 | [6 empty] | -996.18 | 0.09 ‰ |  | Christine Prior | 2015-03-11 | Jocelyn Turnbull | 15:23 | 0.0037745 | 0.00377 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.9 |  |  | 0.00377 | 9e-05 | 9276 | 2015-03-09 | 11:26 | Johannes Kaiser | 201385 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 17 | 5 | 5 | ... | 0.00363 | 9e-05 |  |  |  |  | 1.0001 | 1.0001 | Yes | [6 empty] | 0.00363 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00377 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00363 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-05 | jenny dahl | 15:42 |  | Completed |  |  |  |  | 67428 | 2000000 |  |  | 67457 67458 67513 67514 67515 | 67457~67458~67513~67514~67515 | 2992 |  | 2992 2992 2992 2992 2992 | 2015-03-11 | 11:50 | Christine Prior |  | 1.19 mg |  | 1.19 | [5 empty] | 1.001 | 0.00154 | 67457 |  | XCAMS | EA | 0% |  | - no system error - Ox-I (TP67494) was rejected from calibration as it was low by 4.3 sigma - cat# 18 (TP67465) needed an additional factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP67377) needed an additional factor of 1.4 to expl |  | RRL | [5 empty] | Background Inorganic |  | 2015-03-11 | Carbonate |  | JOB-201385 |  |  | 58578 |  | 14047/1 |  |  | 2015-02-19 | 1793 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99995 | 0.00153 | 67458 | [26 empty] | 2015-02-19 | 1181 | Gas Splitting | Completed
[168 empty] | 0.99914 | 0.00152 | 67513 | [26 empty] | 2015-02-19 | 2356 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 67514 | [26 empty] | 2015-02-23 | 7438 | Graphite | Completed
[168 empty] | 0.99974 | 0.00155 | 67515 | [26 empty] | 2015-03-11 | 9276 | AMS Submission Results | Completed
[197 empty] | 2015-03-12 | 10986 | Close Job | Completed
[7 empty] | 76.1 | 88.8027 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44495 | 44494 | 44495 | 254 | 254 | 254 |  |  | Christine Prior | -996.07 | 0.12 ‰ | 2015-03-18 | 2015-03-06 | 19 | 12 |  |  |  |  | 1.77‰ | 1.77‰ | 2.00‰ | 2.00‰ | 1.77 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -996.07 ‰ | -996.07 ‰ | 0.12 | 0.12 | -996.02 | 0.13 ‰ | -996.02 | 0.13 ‰ |  |  | 2015-03-20 | Jocelyn Turnbull | 12:15 | 0.0039304 | 0.0039304 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00393 | 0.00012 | 9343 | 2015-03-19 | 15:52 | Johannes Kaiser | 201402 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.39 | 0.01% | 23 | 5 | 6 | ... | 0.00378 | 0.00012 |  |  |  |  | 0.99966 | 0.99966 | Yes | Yes | [5 empty] | 0.00378 | 0.00‰ | 0.00012 | 0.00‰ | No | 1 | 0.99996 | 0.00125 | 0.00125 | 0.00393 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00378 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-17 | Christine Prior | 9:12 |  | Completed |  |  |  |  | 67468 | 2000000 |  |  | 67519 67540 67541 67543 67544 67558 | 67519~67540~67541~67543~67544~67558 | 2994 |  | 2994 2994 2994 2994 2994 2994 | 2015-03-20 | 10:11 | Christine Prior |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99664 | 0.00306 | 67519 |  | XCAMS | EA | 0.24% |  | - first measurement after major source clean - system error = 0.24% - cat in pos #25 (TP67469) needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-16 | Water |  | JOB-201402 |  | water line nitrogen flush water vessel | 58668 | Normal | 14047/11 |  |  | 2015-02-26 | 1807 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9989 | 0.00305 | 67540 | [26 empty] | 2015-02-26 | 1187 | Gas Splitting | Completed
[168 empty] | 0.99716 | 0.00305 | 67541 | [26 empty] | 2015-02-26 | 2369 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00305 | 67543 | [26 empty] | 2015-02-27 | 7478 | Graphite | Completed
[168 empty] | 1.0046 | 0.00305 | 67544 | [26 empty] | 2015-03-03 | 3545 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00306 | 67558 | [26 empty] | 2015-03-20 | 9343 | AMS Submission Results | Completed
[197 empty] | 2015-03-20 | 11113 | Close Job | Completed
[5 empty] | 3.4282e-05 | 0.0020601 | 75.9 | 89.7949 | 0.0684 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7140 | 7140 | 7136 | 35 | 35 | 35 | 7105.2 | 7175.3 | Christine Prior | -588.88 | 1.79 ‰ | 2015-03-18 | 2015-03-06 | 19 | 12 |  |  |  |  | -9.30‰ | -9.30‰ | 2.00‰ | 2.00‰ | -9.30 ‰ | 2.00 ‰ | AMS | -8.16‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | [6 empty] | -583.8 | 1.82 ‰ |  |  | 2015-03-20 | Jocelyn Turnbull | 12:15 | 0.41112 | 0.41134 | 0.00179 | 0.00182 | 0.00171 | 0.00171 | 2.1 |  |  | 0.41259 | 0.00166 | 9344 | 2015-03-19 | 15:52 | Johannes Kaiser | 201403 | 0.00234 | 0.002 | 0.00105 | 0.0009 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  |  |  | 25 | 5 | 6 | ... | 0.3968 | 0.00179 |  |  |  |  | 0.99966 | 0.99966 | Yes | Yes |  | - needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 39.428609198611949. | 0.00234 | 0.00198 | 0.00107 | 0.39539 | 0.40‰ | 0.0019 | 0.00‰ | No | 1 | 0.99996 | 0.00125 | 0.00125 | 0.41112 | 0 | 0.00179 | 0 | 0 | 0.00234 | 0 | 0.00105 | 0 | 0.39539 | 0 | 0.0019 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-17 | Christine Prior | 9:12 |  | Completed | 0.41111 | 0.41134 | 0.39538 | 0.3956 | 67469 | 2000000 | 67530 67559 67447 67435 67512 67517 | 67530~67559~67447~67435~67512~67517 | 67519 67540 67541 67543 67544 67558 | 67519~67540~67541~67543~67544~67558 | 2994 | 2994 2994 2993 2993 2992 2992 | 2994 2994 2994 2994 2994 2994 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.94 mg | 0.13 | 0.94 | 2015-03-18 | 0.002 | 67530 | 2994 | 1.04 | 0.99664 | 0.00306 | 67519 | EA | XCAMS | EA | 0.24% |  | - first measurement after major source clean - system error = 0.24% - cat in pos #25 (TP67469) needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-04-16 | Water |  | JOB-201403 |  | water line standard | 58683 | Normal | 14047/12 |  |  | 2015-02-26 | 1808 | Water CO2 Evolution | Completed | 40142/1 | Kauri Renton Road | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-03-18 | 0.00189 | 67559 | 2994 | 1.06 | 0.9989 | 0.00305 | 67540 | EA | [25 empty] | 2015-02-26 | 1188 | Gas Splitting | Completed | 40142/2 | Kauri Renton Road
[163 empty] | 2015-03-09 | 0.00202 | 67435 | 2993 | 0.85 | 0.99716 | 0.00305 | 67541 | EA | [25 empty] | 2015-02-26 | 2371 | Choose Gas Splitting | Completed | 40142/3 | Kauri Renton Road
[163 empty] | 2015-03-09 | 0.00451 | 67447 | 2993 | 0.84 | 1.0004 | 0.00305 | 67543 | [26 empty] | 2015-02-27 | 7479 | Graphite | Completed
[163 empty] | 2015-03-06 | 0.00186 | 67512 | 2992 | 1.16 | 1.0046 | 0.00305 | 67544 | [26 empty] | 2015-03-03 | 3546 | C13 Measurement | Completed
[163 empty] | 2015-03-06 | 0.00178 | 67517 | 2992 | 1.07 | 1.0021 | 0.00306 | 67558 | [26 empty] | 2015-03-20 | 9344 | AMS Submission Results | Completed
[197 empty] | 2015-03-20 | 11114 | Close Job | Completed
[7 empty] | 78.1 | 88.7635 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42711 | 42710 | 42711 | 203 | 203 | 203 |  |  | Christine Prior | -995.09 | 0.12 ‰ | 2015-03-18 | 2015-03-06 | 16 | 12 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -995.09 ‰ | -995.09 ‰ | 0.12 | 0.12 | -995.03 | 0.13 ‰ | -995.03 | 0.13 ‰ |  |  | 2015-03-20 | Jocelyn Turnbull | 12:15 | 0.0049078 | 0.0049078 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00491 | 0.00012 | 9350 | 2015-03-19 | 15:52 | Johannes Kaiser | 201418 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 2 | 5 | 6 | ... | 0.00472 | 0.00012 |  |  |  |  | 0.99966 | 0.99966 | Yes | Yes | [5 empty] | 0.00472 | 0.00‰ | 0.00012 | 0.00‰ | No | 1 | 0.99996 | 0.00125 | 0.00125 | 0.00491 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00472 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-17 | Christine Prior | 9:12 |  | Completed |  |  |  |  | 67472 | 2000000 |  |  | 67519 67540 67541 67543 67544 67558 | 67519~67540~67541~67543~67544~67558 | 2994 |  | 2994 2994 2994 2994 2994 2994 | 2015-03-20 | 10:11 | Christine Prior |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99664 | 0.00306 | 67519 |  | XCAMS | EA | 0.24% |  | - first measurement after major source clean - system error = 0.24% - cat in pos #25 (TP67469) needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2015-04-20 | Water |  | JOB-201418 |  | water line nitrogen flush water vessel | 58669 | Normal | 14047/11 |  |  | 2015-02-27 | 1812 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9989 | 0.00305 | 67540 | [26 empty] | 2015-02-27 | 1189 | Gas Splitting | Completed
[168 empty] | 0.99716 | 0.00305 | 67541 | [26 empty] | 2015-02-27 | 2373 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00305 | 67543 | [26 empty] | 2015-03-02 | 7482 | Graphite | Completed
[168 empty] | 1.0046 | 0.00305 | 67544 | [26 empty] | 2015-03-10 | 3550 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00306 | 67558 | [26 empty] | 2015-03-20 | 9350 | AMS Submission Results | Completed
[197 empty] | 2015-03-20 | 11120 | Close Job | Completed
[7 empty] | 75.7 | 88.8391 | 0.0008 | RRL-CBIn-LG | Background Inorganic |  |  |  |  | 43238 | 43238 | 43238 | 200 | 199 | 200 |  |  | jenny dahl | -995.4 | 0.11 ‰ | 2015-03-20 | 2015-02-28 | 16 | 20 |  |  |  |  | 5.90‰ | 5.90‰ | 2.00‰ | 2.00‰ | 5.90 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | [6 empty] | -995.35 | 0.12 ‰ |  | Christine Prior | 2015-03-23 | Jocelyn Turnbull | 13:00 | 0.0045959 | 0.0045959 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.9 |  |  | 0.0046 | 0.00011 | 9379 | 2015-03-23 | 10:46 | Johannes Kaiser | 201422 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 10 | 5 | 5 | ... | 0.00442 | 0.00011 |  |  |  |  | 1.0042 | 1.0042 | Yes | Yes | [5 empty] | 0.00442 | 0.00‰ | 0.00011 | 0.00‰ | No | 1 | 1.0001 | 0.00094 | 0.00094 | 0.0046 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00442 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-18 | jenny dahl | 16:41 |  | Completed |  |  |  |  | 67503 | 2000000 |  |  | 67470 67471 67476 67477 67478 67479 | 67470~67471~67476~67477~67478~67479 | 2995 |  | 2995 2995 2995 2995 2995 2995 | 2015-03-23 | 12:09 | Christine Prior |  | 1.15 mg |  | 1.15 | [5 empty] | 0.99958 | 0.00231 | 67470 |  | XCAMS | ST | 0.16% |  | - 0.16% system error - needed an additional uncertainty adjustment factor of 1.03 to explain its global run-to-run variance - cathode in pos# 23 (TP67460) needed an uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance |  | RRL | [5 empty] | Background Inorganic |  | 2015-03-17 | Carbonate |  | JOB-201422 |  |  | 58694 |  | 14047/1 |  |  | 2015-03-03 | 1818 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99755 | 0.00228 | 67471 | [26 empty] | 2015-03-03 | 1190 | Gas Splitting | Completed
[168 empty] | 1.0042 | 0.00239 | 67476 | [26 empty] | 2015-03-03 | 2375 | Choose Gas Splitting | Completed
[168 empty] | 1.0016 | 0.00236 | 67477 | [26 empty] | 2015-03-04 | 7513 | Graphite | Completed
[168 empty] | 0.99894 | 0.00227 | 67478 | [26 empty] | 2015-03-23 | 9379 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00228 | 67479 | [26 empty] | 2015-03-23 | 11159 | Close Job | Completed
[7 empty] | 119.4 | 88.8274 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 49489 | 49488 | 49489 | 317 | 316 | 317 |  |  | jenny dahl | -997.89 | 0.08 ‰ | 2015-04-29 | 2015-02-28 | 20 | 60 |  |  |  |  | 3.58‰ | 3.58‰ | 2.00‰ | 2.00‰ | 3.58 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0047 | 0 | 0.8165 | -997.89 ‰ |  | 0.08 |  | -997.86 | 0.08 ‰ | -997.86 | 0.08 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.0021108 | 0.0021108 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00211 | 8e-05 | 9588 | 2015-04-30 | 15:12 | Johannes Kaiser | 201947 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.21 | 0.01% | 9 | 5 | 7 | ... | 0.00203 | 8e-05 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes | [5 empty] | 0.00203 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00211 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00203 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67734 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-16 | Carbonate |  | JOB-201947 |  | reg sized blank- fresh ground from fresh acid etch | 58905 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-08 | 1860 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-08 | 1222 | Gas Splitting | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-04-08 | 2428 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-04-09 | 7744 | Graphite | Completed
[168 empty] | 0.99964 | 0.00156 | 67802 | [26 empty] | 2015-04-13 | 3605 | C13 Measurement | Completed
[197 empty] | 2015-05-04 | 9588 | AMS Submission Results | Completed
[197 empty] | 2015-05-12 | 11418 | Close Job | Completed
[7 empty] | 8.8 | 91.2784 | 0.0039 | RRL-Test-TY | Background Test | 1950-01-01 |  |  |  | 30317 | 30317 | 30317 | 313 | 312 | 313 | 30010.1 | 30636.2 | jenny dahl | -977.04 | 0.89 ‰ | 2015-04-29 | 2015-02-28 | 8 | 60 |  |  |  |  | -23.37‰ | -23.37‰ | 2.00‰ | 2.00‰ | -23.37 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | -977.04 ‰ |  | 0.89 |  | -976.76 | 0.91 ‰ | -976.76 | 0.91 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.022959 | 0.022959 | 0.00089 | 0.00089 | 0.00089 | 0.00089 | 141.3 |  |  | 0.02296 | 0.00089 | 9589 | 2015-04-30 | 15:12 | Johannes Kaiser | 201948 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 2.3 | 0.09% | 5 | 5 | 4 | ... | 0.02208 | 0.00086 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes | [5 empty] | 0.02208 | 0.02‰ | 0.00086 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.02296 | 0 | 0.00089 | 0 | 0 |  |  | 0 | 0 | 0.02208 | 0 | 0.00086 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67803 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.02 mg |  | 0.02 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) |  | RRL | [5 empty] | Background Test |  | 2015-04-16 | Carbonate |  | JOB-201948 |  | small test sample 0.01 mg | 58906 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-20 | 1883 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-21 | 7813 | Graphite | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-05-04 | 9589 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-05-12 | 11419 | Close Job | Completed
[168 empty] | 0.99964 | 0.00156 | 67802
[7 empty] | 23.3 | 90.4185 | 0.0044 | RRL-Test-TY | Background Test |  |  |  |  | 29163 | 29163 | 29163 | 180 | 179 | 180 | 28985.7 | 29345 | jenny dahl | -973.5 | 0.59 ‰ | 2015-04-29 | 2015-02-28 | 8 | 60 |  |  |  |  | -14.08‰ | -14.08‰ | 2.00‰ | 2.00‰ | -14.08 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | [6 empty] | -973.17 | 0.60 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.026504 | 0.026504 | 0.00059 | 0.00059 | 0.00059 | 0.00059 | 40.6 |  |  | 0.0265 | 0.00059 | 9590 | 2015-04-30 | 15:12 | Johannes Kaiser | 201949 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 16 | 5 | 4 | ... | 0.02549 | 0.00057 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes |  | high outlier - but the only carbonate that looks this way |  |  |  | 0.02549 | 0.03‰ | 0.00057 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.0265 | 0 | 0.00059 | 0 | 0 |  |  | 0 | 0 | 0.02549 | 0 | 0.00057 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67804 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.06 mg |  | 0.06 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) |  | RRL | [5 empty] | Background Test |  | 2015-04-16 | Carbonate |  | JOB-201949 |  | small test sample 0.05 mg | 58907 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-20 | 1884 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-21 | 7814 | Graphite | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-05-04 | 9590 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-05-12 | 11420 | Close Job | Completed
[168 empty] | 0.99964 | 0.00156 | 67802
[7 empty] | 57.4 | 89.5485 | 0.002 | RRL-Test-TY | Background Test |  |  |  |  | 35627 | 35627 | 35627 | 176 | 176 | 176 | 35453 | 35805.4 | jenny dahl | -988.15 | 0.26 ‰ | 2015-04-29 | 2015-02-28 | 8 | 60 |  |  |  |  | -4.50‰ | -4.50‰ | 2.00‰ | 2.00‰ | -4.50 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | [6 empty] | -988 | 0.26 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.011854 | 0.011854 | 0.00026 | 0.00026 | 0.00026 | 0.00026 | 18.2 |  |  | 0.01185 | 0.00026 | 9591 | 2015-04-30 | 15:12 | Johannes Kaiser | 201950 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 18 | 5 | 4 | ... | 0.0114 | 0.00025 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes | [5 empty] | 0.0114 | 0.01‰ | 0.00025 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.01185 | 0 | 0.00026 | 0 | 0 |  |  | 0 | 0 | 0.0114 | 0 | 0.00025 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67805 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.13 mg |  | 0.13 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) |  | RRL | [5 empty] | Background Test |  | 2015-04-16 | Carbonate |  | JOB-201950 |  | small test sample 0.1 mg | 58908 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-20 | 1885 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-21 | 7815 | Graphite | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-05-04 | 9591 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-05-12 | 11421 | Close Job | Completed
[168 empty] | 0.99964 | 0.00156 | 67802
[7 empty] | 95.4 | 89.3288 | 0.0003 | RRL-Test-SM | Background Test |  |  |  |  | 51351 | 51350 | 51351 | 399 | 399 | 399 | 50961.4 | 51760.3 | jenny dahl | -998.33 | 0.08 ‰ | 2015-04-29 | 2015-02-28 | 8 | 60 |  |  |  |  | -2.05‰ | -2.05‰ | 2.00‰ | 2.00‰ | -2.05 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | [6 empty] | -998.31 | 0.08 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.0016741 | 0.0016741 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 9.7 |  |  | 0.00167 | 8e-05 | 9592 | 2015-04-30 | 15:12 | Johannes Kaiser | 201951 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 25 | 5 | 4 | ... | 0.00161 | 8e-05 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes |  | wow, 0.25 mg is better than 0.37 mg |  |  |  | 0.00161 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00167 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00161 | 0 | 8e-05 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67806 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.25 mg |  | 0.25 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) |  | RRL | [5 empty] | Background Test |  | 2015-04-16 | Carbonate |  | JOB-201951 |  | small test sample 0.2 mg | 58909 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-20 | 1886 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-21 | 7816 | Graphite | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-05-04 | 9592 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-05-12 | 11422 | Close Job | Completed
[168 empty] | 0.99964 | 0.00156 | 67802
[7 empty] | 108.6 | 89.3075 | 0.0008 | RRL-Test-SM | Background Test |  |  |  |  | 42918 | 42917 | 42918 | 192 | 192 | 192 | 42727.9 | 43112.1 | jenny dahl | -995.22 | 0.11 ‰ | 2015-04-29 | 2015-02-28 | 8 | 60 |  |  |  |  | -1.81‰ | -1.81‰ | 2.00‰ | 2.00‰ | -1.81 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0047 | 0 | 0.8165 | [6 empty] | -995.16 | 0.12 ‰ |  |  | 2015-05-04 | Jocelyn Turnbull | 13:54 | 0.0047831 | 0.0047831 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 5.6 |  |  | 0.00478 | 0.00011 | 9593 | 2015-04-30 | 15:12 | Johannes Kaiser | 201952 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  |  |  | 32 | 5 | 4 | ... | 0.0046 | 0.00011 |  |  |  |  | 1.0017 | 1.0017 | Yes | Yes | [5 empty] | 0.0046 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00478 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.0046 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-04-24 | jenny dahl | 15:25 |  | Completed |  |  |  |  | 67807 | 2000000 |  |  | 67746 67747 67793 67794 67801 67802 | 67746~67747~67793~67794~67801~67802 | 3001 |  | 3001 3001 3001 3001 3001 3001 |  |  |  |  | 0.37 mg |  | 0.37 | [5 empty] | 0.99868 | 0.00162 | 67746 |  | XCAMS | EA | 0% |  | - good measurement - needed to add a factor of 1.3 to explain the individual run-to-run variance for cat#4 (TP67769), 11 (TP67820) and 21 (TP67747) |  | RRL | [5 empty] | Background Test |  | 2015-04-16 | Carbonate | Not Complete | JOB-201952 |  | small test sample 0.3 mg | 58910 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00201 | 67747 | [26 empty] | 2015-04-20 | 1887 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00155 | 67793 | [26 empty] | 2015-04-21 | 7817 | Graphite | Completed
[168 empty] | 0.99991 | 0.00162 | 67794 | [26 empty] | 2015-05-04 | 9593 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00156 | 67801 | [26 empty] | 2015-05-12 | 11423 | Close Job | Completed
[168 empty] | 0.99964 | 0.00156 | 67802
[7 empty] | 111.2 | 89.2667 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49853 | 49853 | 49853 | 331 | 330 | 331 |  |  | Margaret Norris | -997.98 | 0.08 ‰ | 2015-05-08 | 2015-03-18 | 21 | 51 |  |  |  |  | -1.96‰ | -1.96‰ | 2.00‰ | 2.00‰ | -1.96 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -997.98 ‰ | -997.98 ‰ | 0.08 | 0.08 | -997.96 | 0.08 ‰ | -997.96 | 0.08 ‰ |  | Christine Prior | 2015-05-12 | Jocelyn Turnbull | 10:55 | 0.0020172 | 0.0020172 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.6 |  |  | 0.00202 | 8e-05 | 9608 | 2015-05-11 | 10:15 | Johannes Kaiser | 201994 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 12 | 5 | 6 | ... | 0.00194 | 8e-05 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.00194 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 0.99998 | 0.00092 | 0.00092 | 0.00202 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00194 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-07 | Margaret Norris | 15:30 |  | Completed |  |  |  |  | 67781 | 2000000 |  |  | 67726 67727 67894 67895 67898 67899 | 67726~67727~67894~67895~67898~67899 | 3002 |  | 3002 3002 3002 3002 3002 3002 | 2015-05-11 | 15:09 | Christine Prior |  | 0.76 mg |  | 0.76 | [5 empty] | 0.99675 | 0.00227 | 67726 |  | XCAMS | EA | 0.17% |  | - interesting fact to note that Kauri (TP67855) has roughly 10x times more beam intensity than Ox-I (TP67860), Ox-I (TP657844) and FIRI-I (TP67847) and all are ~ 0.05mgC -It is noted that there is a correlation in rts differences based on the two TP grou | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-04 | Water |  | JOB-201994 |  | Took some of the acid etched and ground material from stock bottle and labelled as R14047/11 | 58945 | Normal | 14047/11 |  |  | 2015-04-14 | 1866 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99785 | 0.00227 | 67727 | [26 empty] | 2015-04-14 | 1224 | Gas Splitting | Completed
[168 empty] | 1.0019 | 0.00225 | 67894 | [26 empty] | 2015-04-14 | 2431 | Choose Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00225 | 67895 | [26 empty] | 2015-04-17 | 7791 | Graphite | Completed
[168 empty] | 1.0025 | 0.00227 | 67898 | [26 empty] | 2015-04-20 | 3622 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00224 | 67899 | [26 empty] | 2015-05-12 | 9608 | AMS Submission Results | Completed
[197 empty] | 2015-05-12 | 11479 | Close Job | Completed
[5 empty] | 6.0393e-05 | 0.0031513 | 123.3 | 90.0112 | 0.0683 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7089 | 7088 | 7083 | 30 | 29 | 34 | 7058.8 | 7118.4 | Margaret Norris | -586.22 | 1.54 ‰ | 2015-05-08 | 2015-03-18 | 21 | 51 |  |  |  |  | -10.21‰ | -10.21‰ | 2.00‰ | 2.00‰ | -10.21 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | [6 empty] | -581.12 | 1.55 ‰ |  |  | 2015-05-12 | Jocelyn Turnbull | 10:55 | 0.41378 | 0.41407 | 0.00154 | 0.00173 | 0.00144 | 0.00144 | 1.9 |  |  | 0.41596 | 0.00117 | 9631 | 2015-05-11 | 10:15 | Johannes Kaiser | 201995 | 0.00348 | 0.003 | 0.00157 | 0.0014 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  |  |  | 20 | 5 | 7 | ... | 0.40004 | 0.00121 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes |  | Had to go back 50 days to get blanks run on water line wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 34.8731758800840638. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00348 | 0.00303 | 0.00141 | 0.39794 | 0.40‰ | 0.00154 | 0.00‰ | No | 1 | 0.99998 | 0.00092 | 0.00092 | 0.41378 | 0 | 0.00154 | 0 | 0 | 0.00348 | 0 | 0.00157 | 0 | 0.39794 | 0 | 0.00154 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-07 | Margaret Norris | 15:30 |  | Completed | 0.41378 | 0.41407 | 0.39794 | 0.39822 | 67792 | 2000000 | 67781 67468 67472 | 67781~67468~67472 | 67726 67727 67894 67895 67898 67899 | 67726~67727~67894~67895~67898~67899 | 3002 | 3002 2994 2994 | 3002 3002 3002 3002 3002 3002 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.99 mg | 0.13 | 0.99 | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 0.99675 | 0.00227 | 67726 | EA | XCAMS | EA | 0.17% |  | - interesting fact to note that Kauri (TP67855) has roughly 10x times more beam intensity than Ox-I (TP67860), Ox-I (TP657844) and FIRI-I (TP67847) and all are ~ 0.05mgC -It is noted that there is a correlation in rts differences based on the two TP grou | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-06-04 | Water | Not Complete | JOB-201995 |  |  | 58946 | Normal | 14047/12 |  |  | 2015-04-14 | 3306 | Physical Pretreatment | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-03-18 | 0.00378 | 67468 | 2994 | 0.98 | 0.99785 | 0.00227 | 67727 | EA | [25 empty] | 2015-04-14 | 1865 | Water CO2 Evolution | Completed
[163 empty] | 2015-03-18 | 0.00472 | 67472 | 2994 | 1.01 | 1.0019 | 0.00225 | 67894 | [26 empty] | 2015-04-14 | 1225 | Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00225 | 67895 | [26 empty] | 2015-04-14 | 2433 | Choose Gas Splitting | Completed
[168 empty] | 1.0025 | 0.00227 | 67898 | [26 empty] | 2015-04-17 | 7802 | Graphite | Completed
[168 empty] | 1.0009 | 0.00224 | 67899 | [26 empty] | 2015-04-20 | 3623 | C13 Measurement | Completed
[197 empty] | 2015-05-12 | 9631 | AMS Submission Results | Completed
[197 empty] | 2015-05-12 | 11480 | Close Job | Completed
[5 empty] | 5.5783e-05 | 0.0036214 | 133.8 | 89.825 | 0.0688 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7070 | 7069 | 7062 | 33 | 33 | 36 | 7036.6 | 7103.2 | jenny dahl | -585.26 | 1.72 ‰ | 2015-05-12 | 2015-03-18 | 11 | 55 |  |  |  |  | -8.89‰ | -8.89‰ | 2.00‰ | 2.00‰ | -8.89 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | [6 empty] | -580.14 | 1.74 ‰ |  |  | 2015-05-14 | Jocelyn Turnbull | 10:57 | 0.41474 | 0.41516 | 0.00172 | 0.00185 | 0.00163 | 0.00163 | 2 |  |  | 0.41733 | 0.00126 | 9680 | 2015-05-13 | 10:47 | Johannes Kaiser | 202155 | 0.00414 | 0.0035 | 0.00186 | 0.0016 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  |  |  | 3 | 5 | 7 | ... | 0.40136 | 0.00109 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 37.8941711357073823. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00414 | 0.00348 | 0.00175 | 0.39887 | 0.40‰ | 0.00157 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.41474 | 0 | 0.00172 | 0 | 0 | 0.00414 | 0 | 0.00186 | 0 | 0.39887 | 0 | 0.00157 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-08 | jenny dahl | 16:07 |  | Completed | 0.41474 | 0.41516 | 0.39887 | 0.39927 | 67850 | 2000000 | 67849 67781 67468 67472 | 67849~67781~67468~67472 | 67885 67886 67887 67888 67914 67918 | 67885~67886~67887~67888~67914~67918 | 3003 | 3003 3002 2994 2994 | 3003 3003 3003 3003 3003 3003 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.92 mg | 0.13 | 0.92 | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 0.99931 | 0.00167 | 67885 | EA | XCAMS | ST | 0% |  | - good measurement - rejected first rotation from reporting - cat# 16 (TP67877) needed an additional factor of 1.3 to explain its individual run-to-run variance | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-06-25 | Water |  | JOB-202155 |  |  | 58986 | Normal | 14047/12 |  |  | 2015-04-14 | 3306 | Physical Pretreatment | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 1.0024 | 0.00168 | 67886 | EA | [25 empty] | 2015-04-30 | 1914 | Water CO2 Evolution | Completed
[163 empty] | 2015-03-18 | 0.00378 | 67468 | 2994 | 0.98 | 0.99865 | 0.00168 | 67887 | ST | [25 empty] | 2015-04-30 | 1237 | Gas Splitting | Completed
[163 empty] | 2015-03-18 | 0.00472 | 67472 | 2994 | 1.01 | 0.99929 | 0.0017 | 67888 | [26 empty] | 2015-04-30 | 2447 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.0017 | 67914 | [26 empty] | 2015-05-01 | 7860 | Graphite | Completed
[168 empty] | 0.99935 | 0.0017 | 67918 | [26 empty] | 2015-05-05 | 3642 | C13 Measurement | Completed
[197 empty] | 2015-05-14 | 9680 | AMS Submission Results | Completed
[197 empty] | 2015-05-14 | 11526 | Close Job | Completed
[7 empty] | 111.6 | 89.2316 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40637 | 40637 | 40637 | 184 | 184 | 184 |  |  | jenny dahl | -993.65 | 0.15 ‰ | 2015-05-12 | 2015-03-18 | 11 | 55 |  |  |  |  | -2.30‰ | -2.30‰ | 2.00‰ | 2.00‰ | -2.30 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -993.65 ‰ | -993.65 ‰ | 0.15 | 0.15 | -993.57 | 0.15 ‰ | -993.57 | 0.15 ‰ |  |  | 2015-05-14 | Jocelyn Turnbull | 10:57 | 0.0063531 | 0.0063531 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.2 |  |  | 0.00635 | 0.00015 | 9681 | 2015-05-13 | 10:47 | Johannes Kaiser | 202156 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.64 | 0.01% | 17 | 5 | 6 | ... | 0.00611 | 0.00014 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes | [5 empty] | 0.00611 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00635 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00611 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-08 | jenny dahl | 16:07 |  | Completed |  |  |  |  | 67849 | 2000000 |  |  | 67885 67886 67887 67888 67914 67918 | 67885~67886~67887~67888~67914~67918 | 3003 |  | 3003 3003 3003 3003 3003 3003 | 2015-05-13 | 12:01 | Christine Prior |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99931 | 0.00167 | 67885 |  | XCAMS | ST | 0% |  | - good measurement - rejected first rotation from reporting - cat# 16 (TP67877) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-25 | Water |  | JOB-202156 |  | water line nitrogen flush water vessel | 58985 | Normal | 14047/11 |  |  | 2015-04-30 | 1912 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00168 | 67886 | [26 empty] | 2015-04-30 | 1236 | Gas Splitting | Completed
[168 empty] | 0.99865 | 0.00168 | 67887 | [26 empty] | 2015-04-30 | 2445 | Choose Gas Splitting | Completed
[168 empty] | 0.99929 | 0.0017 | 67888 | [26 empty] | 2015-05-01 | 7859 | Graphite | Completed
[168 empty] | 1.001 | 0.0017 | 67914 | [26 empty] | 2015-05-05 | 3643 | C13 Measurement | Completed
[168 empty] | 0.99935 | 0.0017 | 67918 | [26 empty] | 2015-05-14 | 9681 | AMS Submission Results | Completed
[197 empty] | 2015-05-14 | 11527 | Close Job | Completed
[7 empty] | 123.3 | 89.1526 | 0.004 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 30014 | 30013 | 30014 | 81 | 80 | 81 | 29933.4 | 30094.6 | Margaret Norris | -976.16 | 0.24 ‰ | 2015-05-13 | 2015-05-05 | 26 | 8 |  |  |  |  | 0.84‰ | 0.84‰ | 2.00‰ | 2.00‰ | 0.84 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.0084 | 0 | 0.8165 | -976.16 ‰ | -976.16 ‰ | 0.24 | 0.24 | -975.86 | 0.24 ‰ | -975.86 | 0.24 ‰ |  |  | 2015-05-19 | Jocelyn Turnbull | 11:32 | 0.023842 | 0.023842 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.4 |  |  | 0.02384 | 0.00024 | 9716 | 2015-05-15 | 9:54 | Johannes Kaiser | 202041 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 2.38 | 0.02% | 19 | 5 | 7 | ... | 0.02293 | 0.00023 |  |  |  |  | 1.0027 | 1.0027 | Yes | Yes |  | Dreadful blank!! no obvious explanation: CO2 production and graphitization look okay. In fact, water blank graphitised on same day in next reactor as job 201994 came out great. Decided this was anomalous enough to reject as blank and change category |  |  |  | 0.02293 | 0.02‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.02384 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.02293 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-12 | Margaret Norris | 13:54 |  | Completed |  |  |  |  | 67782 | 2000000 |  |  | 67896 67897 67907 67908 67930 67931 | 67896~67897~67907~67908~67930~67931 | 3004 |  | 3004 3004 3004 3004 3004 3004 | 2015-05-18 | 12:33 | Christine Prior |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99798 | 0.0015 | 67896 |  | XCAMS | EA | 0% |  | #NAME? | 2.3 | RRL | [5 empty] | Background Test |  | 2015-06-10 | Water |  | JOB-202041 |  | test status due to CRA of ~30014/81 | 59025 | Normal | 14047/11 |  |  | 2015-04-15 | 1882 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00148 | 67897 | [26 empty] | 2015-04-15 | 1226 | Gas Splitting | Cancelled
[168 empty] | 1.0018 | 0.00149 | 67907 | [26 empty] | 2015-04-15 | 1227 | Gas Splitting | Completed
[168 empty] | 0.9999 | 0.00152 | 67908 | [26 empty] | 2015-04-15 | 2435 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00161 | 67930 | [26 empty] | 2015-04-17 | 7792 | Graphite | Completed
[168 empty] | 1.0002 | 0.00171 | 67931 | [26 empty] | 2015-04-20 | 3624 | C13 Measurement | Completed
[197 empty] | 2015-05-19 | 9716 | AMS Submission Results | Completed
[197 empty] | 2015-06-02 | 11623 | Close Job | Completed
|  |  | 0.01 | 0.01 | 7.2177e-05 | 0.0030831 | 107.4 | 90.3092 | 0.0678 | RRL-UNSt-LG | Unknown Inorganic |  |  |  |  | 7127 | 7127 | 7114 | 33 | 32 | 34 | 7094.1 | 7160 | Margaret Norris | -588.2 | 1.69 ‰ | 2015-05-13 | 2015-05-05 | 5 | 8 |  |  |  |  | -11.97‰ | -11.97‰ | 2.00‰ | 2.00‰ | -11.97 ‰ | 2.00 ‰ | AMS | -8.22‰ | 0.2‰ | C13 | No | -19 | -19.0084 | 0 | 0.8165 | [6 empty] | -583.12 | 1.71 ‰ |  |  | 2015-05-19 | Jocelyn Turnbull | 11:32 | 0.4118 | 0.41247 | 0.00169 | 0.00177 | 0.0016 | 0.0016 | 1.8 |  |  | 0.41433 | 0.00125 | 9724 | 2015-05-15 | 9:54 | Johannes Kaiser | 202158 | 0.00403 | 0.003 | 0.00181 | 0.0013 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  |  |  | 12 | 5 | 7 | ... | 0.39847 | 0.00108 |  |  |  |  | 1.0027 | 1.0027 | Yes | Yes |  | If marble blank j202041 in this wheel used in average, wildly skews age much older than plausible. wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 37.5755559986303995. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00403 | 0.00297 | 0.00295 | 0.39604 | 0.40‰ | 0.00154 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.4118 | 0 | 0.00169 | 0 | 0 | 0.00403 | 0 | 0.00181 | 0 | 0.39604 | 0 | 0.00154 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-12 | Margaret Norris | 13:54 |  | Completed | 0.4118 | 0.41247 | 0.39604 | 0.39668 | 67928 | 2000000 | 67849 67781 | 67849~67781 | 67896 67897 67907 67908 67930 67931 | 67896~67897~67907~67908~67930~67931 | 3004 | 3003 3002 | 3004 3004 3004 3004 3004 3004 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.05 mg | 0.13 | 1.05 | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 0.99798 | 0.0015 | 67896 | EA | XCAMS | EA | 0% |  | #NAME? | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-06-25 | Water |  | JOB-202158 |  |  | 59026 | Normal | 14047/12 |  |  | 2015-04-14 | 3306 | Physical Pretreatment | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 1 | 0.00148 | 67897 | ST | [25 empty] | 2015-05-07 | 1923 | Water CO2 Evolution | Completed
[168 empty] | 1.0018 | 0.00149 | 67907 | [26 empty] | 2015-05-07 | 1246 | Gas Splitting | Completed
[168 empty] | 0.9999 | 0.00152 | 67908 | [26 empty] | 2015-05-07 | 2463 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00161 | 67930 | [26 empty] | 2015-05-08 | 7938 | Graphite | Completed
[168 empty] | 1.0002 | 0.00171 | 67931 | [26 empty] | 2015-05-13 | 3683 | C13 Measurement | Completed
[197 empty] | 2015-05-19 | 9724 | AMS Submission Results | Completed
[197 empty] | 2015-06-02 | 11625 | Close Job | Completed
[7 empty] | 142.9 | 88.8692 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48734 | 48734 | 48734 | 288 | 288 | 288 |  |  | Christine Prior | -997.68 | 0.08 ‰ | 2015-05-28 | 2015-05-12 | 17 | 16 |  |  |  |  | 4.50‰ | 4.50‰ | 2.00‰ | 2.00‰ | 4.50 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0059 | 0 | 0.8165 | -997.68 ‰ | -997.68 ‰ | 0.08 | 0.08 | -997.65 | 0.08 ‰ | -997.65 | 0.08 ‰ |  |  | 2015-05-29 | Jocelyn Turnbull | 13:48 | 0.0023187 | 0.0023187 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.00232 | 8e-05 | 9767 | 2015-05-29 | 9:59 | Johannes Kaiser | 201946 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.23 | 0.01% | 3 | 5 | 5 | ... | 0.00223 | 8e-05 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes | [5 empty] | 0.00223 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00232 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00223 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-22 | Christine Prior | 17:58 |  | Completed |  |  |  |  | 67935 | 2000000 |  |  | 67916 67917 68032 68033 68034 68035 | 67916~67917~68032~68033~68034~68035 | 3006 |  | 3006 3006 3006 3006 3006 3006 | 2015-05-29 | 12:58 | Christine Prior |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99988 | 0.00162 | 67916 |  | XCAMS | ST | 0% |  | - good measurement - needed to add a minor factor of 1.02 due to excess global run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-16 | Carbonate | Not Complete | JOB-201946 |  |  | 59105 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0021 | 0.00164 | 67917 | [26 empty] | 2015-05-08 | 1934 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99899 | 0.00184 | 68032 | [26 empty] | 2015-05-11 | 7945 | Graphite | Completed
[168 empty] | 0.99916 | 0.0018 | 68033 | [26 empty] | 2015-05-13 | 3674 | C13 Measurement | Completed
[168 empty] | 1.0008 | 0.00206 | 68034 | [26 empty] | 2015-05-29 | 9767 | AMS Submission Results | Completed
[168 empty] | 0.99899 | 0.00166 | 68035 | [26 empty] | 2015-06-02 | 11636 | Close Job | Completed
[7 empty] | 133.5 | 88.9585 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51708 | 51708 | 51708 | 313 | 312 | 313 |  |  | jenny dahl | -998.4 | 0.06 ‰ | 2015-05-29 | 2015-04-29 | 3 | 30 |  |  |  |  | 2.43‰ | 2.43‰ | 2.00‰ | 2.00‰ | 2.43 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0065 | 0 | 0.8165 | -998.40 ‰ | -998.40 ‰ | 0.06 | 0.06 | -998.38 | 0.06 ‰ | -998.38 | 0.06 ‰ |  | Christine Prior | 2015-06-03 | Jocelyn Turnbull | 11:23 | 0.0016013 | 0.0016013 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.7 |  |  | 0.0016 | 6e-05 | 9841 | 2015-06-02 | 15:06 | Albert Zondervan | 202297 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.16 | 0.01% | 34 | 5 | 7 | ... | 0.00154 | 6e-05 |  |  |  |  | 1.0021 | 1.0021 | Yes | Yes |  | Looking at trends in recent *carbonate* line C1 blanks, I think this is real. |  |  |  | 0.00154 | 0.00‰ | 6e-05 | 0.00‰ | No | 1 | 1.0001 | 0.00096 | 0.00096 | 0.0016 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00154 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-28 | jenny dahl | 14:20 |  | Completed |  |  |  |  | 68043 | 2000000 |  |  | 67968 67969 67981 67982 68005 68006 | 67968~67969~67981~67982~68005~68006 | 3007 |  | 3007 3007 3007 3007 3007 3007 | 2015-06-03 | 11:09 | Christine Prior |  | 0.73 mg |  | 0.73 | [5 empty] | 1.0021 | 0.00233 | 67968 |  | XCAMS | EA | 0.17% |  | * The ion source was somewhat unstable during first few rotations, but this has not affected quality of the measurement data. * The first several runs on all cathodes with low RTS (blanks, others) are clear outliers and have been rejected: Surface contamination is the usual explanation. * System error due to excess variance amongst calibration Ox-Is above Poisson expectation. * No other error adjustments necessary. |  | RRL | [5 empty] | Background Inorganic |  | 2015-06-03 | Carbonate |  | JOB-202297 |  | botched first attempt at making a tuning split- repeated | 59145 |  | 14047/1 |  |  | 2015-05-22 | 1946 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0024 | 0.00235 | 67969 | [26 empty] | 2015-05-22 | 1253 | Gas Splitting | Cancelled
[168 empty] | 1.0018 | 0.0025 | 67981 | [26 empty] | 2015-05-22 | 2474 | Choose Gas Splitting | Cancelled
[168 empty] | 0.998 | 0.00233 | 67982 | [26 empty] | 2015-05-22 | 1254 | Gas Splitting | Completed
[168 empty] | 0.99932 | 0.00231 | 68005 | [26 empty] | 2015-05-22 | 2475 | Choose Gas Splitting | Completed
[168 empty] | 0.99698 | 0.00231 | 68006 | [26 empty] | 2015-05-26 | 8053 | Graphite | Completed
[197 empty] | 2015-06-03 | 9841 | AMS Submission Results | Completed
[197 empty] | 2015-06-04 | 11679 | Close Job | Completed
[7 empty] | 135.2 | 88.8502 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52082 | 52081 | 52082 | 328 | 327 | 328 |  |  | Margaret Norris | -998.47 | 0.06 ‰ | 2015-06-15 | 2015-06-02 | 5 | 13 |  |  |  |  | 2.12‰ | 2.12‰ | 2.00‰ | 2.00‰ | 2.12 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -998.47 ‰ | -998.47 ‰ | 0.06 | 0.06 | -998.45 | 0.06 ‰ | -998.45 | 0.06 ‰ |  |  | 2015-06-18 | Jocelyn Turnbull | 12:22 | 0.0015285 | 0.0015285 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.2 |  |  | 0.00153 | 6e-05 | 9893 | 2015-06-16 | 11:23 | Johannes Kaiser | 202157 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 3 | 5 | 6 | ... | 0.00147 | 6e-05 |  |  |  |  | 1.0005 | 1.0005 | Yes | Yes |  | Great blank. Probably unrealistically good. Will average other recent water blanks. |  |  |  | 0.00147 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00153 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00147 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-06-11 | Margaret Norris | 16:29 |  | Completed |  |  |  |  | 68127 | 2000000 |  |  | 68120 68121 68144 68145 68146 68147 | 68120~68121~68144~68145~68146~68147 | 3009 |  | 3009 3009 3009 3009 3009 3009 |  |  |  |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99907 | 0.00162 | 68120 |  | XCAMS | EA | 0% |  | - good measurement - need to add a global factor of 1.02 to account for global run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-25 | Water | Complete | JOB-202157 |  | water line nitrogen flush water vessel | 59231 | Normal | 14047/11 |  |  | 2015-06-09 | 1966 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99998 | 0.00161 | 68121 | [26 empty] | 2015-06-09 | 1257 | Gas Splitting | Completed
[168 empty] | 0.99981 | 0.00158 | 68144 | [26 empty] | 2015-06-09 | 2480 | Choose Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00158 | 68145 | [26 empty] | 2015-06-10 | 8137 | Graphite | Completed
[168 empty] | 0.99851 | 0.00158 | 68146 | [26 empty] | 2015-06-15 | 3716 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00157 | 68147 | [26 empty] | 2015-06-18 | 9893 | AMS Submission Results | Completed
[197 empty] | 2015-06-18 | 11783 | Close Job | Completed
[5 empty] | 4.7199e-05 | 0.0021933 | 140.1 | 89.7342 | 0.0683 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7146 | 7145 | 7133 | 29 | 28 | 31 | 7117 | 7174.9 | Margaret Norris | -589.17 | 1.48 ‰ | 2015-06-15 | 2015-06-02 | 5 | 13 |  |  |  |  | -7.75‰ | -7.75‰ | 2.00‰ | 2.00‰ | -7.75 ‰ | 2.00 ‰ | AMS | -8.07‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -589.17 ‰ | -588.49 ‰ | 1.48 | 1.61 | -584.09 | 1.50 ‰ | -584.09 | 1.50 ‰ |  |  | 2015-06-18 | Jocelyn Turnbull | 12:22 | 0.41083 | 0.41151 | 0.00148 | 0.00161 | 0.00138 | 0.00138 | 2.2 |  |  | 0.41283 | 0.00117 | 9894 | 2015-06-16 | 11:23 | Johannes Kaiser | 202159 | 0.00317 | 0.0021 | 0.00143 | 0.001 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.08 | 0.15% | 10 | 5 | 7 | ... | 0.39703 | 0.001 |  |  |  |  | 1.0005 | 1.0005 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 34.1671489047049599. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00317 | 0.00211 | 0.00255 | 0.39511 | 0.40‰ | 0.00133 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41083 | 0 | 0.00148 | 0 | 0 | 0.00317 | 0 | 0.00143 | 0 | 0.39511 | 0 | 0.00133 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-06-11 | Margaret Norris | 16:29 |  | Completed | 0.41083 | 0.41151 | 0.39511 | 0.39576 | 68128 | 2000000 | 68127 67781 67849 | 68127~67781~67849 | 68120 68121 68144 68145 68146 68147 | 68120~68121~68144~68145~68146~68147 | 3009 | 3009 3002 3003 | 3009 3009 3009 3009 3009 3009 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.01 mg | 0.13 | 1.01 | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 0.99907 | 0.00162 | 68120 | EA | XCAMS | EA | 0% |  | - good measurement - need to add a global factor of 1.02 to account for global run-to-run variance | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-06-25 | Water |  | JOB-202159 |  | To go with next water samples | 59232 | Normal | 14047/12 |  |  | 2015-04-14 | 3306 | Physical Pretreatment | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 0.99998 | 0.00161 | 68121 | ST | [25 empty] | 2015-06-09 | 1967 | Water CO2 Evolution | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 0.99981 | 0.00158 | 68144 | EA | [25 empty] | 2015-06-09 | 1260 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00158 | 68145 | [26 empty] | 2015-06-09 | 2484 | Choose Gas Splitting | Completed
[168 empty] | 0.99851 | 0.00158 | 68146 | [26 empty] | 2015-06-10 | 8138 | Graphite | Completed
[168 empty] | 1.0014 | 0.00157 | 68147 | [26 empty] | 2015-06-15 | 3717 | C13 Measurement | Completed
[197 empty] | 2015-06-18 | 9894 | AMS Submission Results | Completed
[197 empty] | 2015-06-18 | 11784 | Close Job | Completed
[7 empty] | 143.3 | 88.8921 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51301 | 51301 | 51301 | 297 | 297 | 297 |  |  | jenny dahl | -998.32 | 0.06 ‰ | 2015-07-06 | 2015-05-27 | 12 | 40 |  |  |  |  | 0.74‰ | 0.74‰ | 2.00‰ | 2.00‰ | 0.74 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -998.32 ‰ | -998.32 ‰ | 0.06 | 0.06 | -998.29 | 0.06 ‰ | -998.29 | 0.06 ‰ |  |  | 2015-07-10 | Jocelyn Turnbull | 13:35 | 0.0016845 | 0.0016845 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.3 |  |  | 0.00168 | 6e-05 | 10001 | 2015-07-08 | 16:49 | albert zondervan | 201828 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 24 | 5 | 7 | ... | 0.00162 | 6e-05 |  |  |  |  | 0.99963 | 0.99963 | Yes | Yes | [5 empty] | 0.00162 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00097 | 0.00097 | 0.00168 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00162 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-06-26 | jenny dahl | 15:48 |  | Completed |  |  |  |  | 68230 | 2000000 |  |  | 68205 68206 68207 68258 68259 68260 | 68205~68206~68207~68258~68259~68260 | 3012 |  | 3012 3012 3012 3012 3012 3012 | 2015-07-09 | 14:29 | Christine Prior |  | 0.84 mg |  | 0.84 | [5 empty] | 0.99666 | 0.00237 | 68205 |  | XCAMS | EA | 0.18% |  | - as in TW3011, the Cs delivery in the ion source was different than normal, but this is now well understood and just requires a bit more monitoring - apart from 0.18% system error, all runs demonstrate good run-to-run repeatability; but - the clear excep | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-04-03 | Carbonate | Not Complete | JOB-201828 |  | Fresh acid etch, make daughter jobs from this job | 59339 |  | 14047/1 |  |  | 2015-03-20 | 3271 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0016 | 0.00238 | 68206 | [26 empty] | 2015-06-22 | 1981 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0009 | 0.00239 | 68207 | [26 empty] | 2015-06-22 | 1263 | Gas Splitting | Completed
[168 empty] | 0.99764 | 0.0024 | 68258 | [26 empty] | 2015-06-22 | 2488 | Choose Gas Splitting | Completed
[168 empty] | 1.0029 | 0.00242 | 68259 | [26 empty] | 2015-06-24 | 8240 | Graphite | Completed
[168 empty] | 1.0005 | 0.00237 | 68260 | [26 empty] | 2015-06-26 | 3737 | C13 Measurement | Completed
[197 empty] | 2015-07-10 | 10001 | AMS Submission Results | Completed
[197 empty] | 2015-07-10 | 11915 | Close Job | Completed
[5 empty] | 4.4898e-05 | 0.002406 | 78.1 | 90.1716 | 0.0679 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7202 | 7202 | 7188 | 34 | 33 | 36 | 7168.1 | 7236.1 | Margaret Norris | -592.03 | 1.73 ‰ | 2015-08-11 | 2015-05-03 | 13 | 100 |  |  |  |  | -8.39‰ | -8.39‰ | 2.00‰ | 2.00‰ | -8.39 ‰ | 2.00 ‰ | AMS | -8.07‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -592.03 ‰ | -591.30 ‰ | 1.73 | 1.81 | -586.99 | 1.75 ‰ | -586.99 | 1.75 ‰ |  |  | 2015-08-13 | Jocelyn Turnbull | 15:20 | 0.40797 | 0.4087 | 0.00173 | 0.00181 | 0.00164 | 0.00164 | 1.9 |  |  | 0.41016 | 0.00141 | 10290 | 2015-08-13 | 11:29 | Albert Zondervan | 202250 | 0.00345 | 0.0023 | 0.00155 | 0.001 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.8 | 0.17% | 10 | 5 | 6 | ... | 0.39446 | 0.00126 |  |  |  |  | 1.0048 | 1.0048 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 38.5048402901871561. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00345 | 0.00231 | 0.00216 | 0.39236 | 0.39‰ | 0.00158 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.40797 | 0 | 0.00173 | 0 | 0 | 0.00345 | 0 | 0.00155 | 0 | 0.39236 | 0 | 0.00158 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-10 | Margaret Norris | 10:53 |  | Completed | 0.40798 | 0.4087 | 0.39237 | 0.39306 | 68508 | 2000000 | 68516 68127 67849 67781 | 68516~68127~67849~67781 | 68544 68545 68547 68548 68549 68550 | 68544~68545~68547~68548~68549~68550 | 3020 | 3020 3009 3003 3002 | 3020 3020 3020 3020 3020 3020 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.06 mg | 0.13 | 1.06 | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 0.99999 | 0.0019 | 68544 | EA | XCAMS | ST | 0% |  | - required quality achieved through close-up scrutiny and adjustment - very last TW measurement before SNICS source clean - error adjustments only needed for run-to-run repeatability: global factor 1.04 and cathode-specific factor 1.3 for pos#13 and #36 - | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-07-13 | Water | Complete | JOB-202250 |  |  | 59630 | Normal | 14047/12 |  |  | 2015-07-28 | 2035 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 1.0013 | 0.00203 | 68545 | ST | [25 empty] | 2015-07-28 | 1281 | Gas Splitting | Completed
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 1.0012 | 0.00198 | 68547 | EA | [25 empty] | 2015-07-28 | 2523 | Choose Gas Splitting | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 0.99837 | 0.00197 | 68548 | ST | [25 empty] | 2015-07-29 | 8518 | Graphite | Completed
[168 empty] | 1.0015 | 0.00196 | 68549 | [26 empty] | 2015-08-03 | 3840 | C13 Measurement | Completed
[168 empty] | 0.99767 | 0.00199 | 68550 | [26 empty] | 2015-08-13 | 10290 | AMS Submission Results | Completed
[197 empty] | 2015-08-13 | 12239 | Close Job | Completed
[7 empty] | 69.2 | 89.3663 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50235 | 50234 | 50235 | 434 | 434 | 434 |  |  | Margaret Norris | -998.08 | 0.10 ‰ | 2015-08-11 | 2015-05-03 | 18 | 100 |  |  |  |  | 0.55‰ | 0.55‰ | 2.00‰ | 2.00‰ | 0.55 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -998.08 ‰ | -998.08 ‰ | 0.1 | 0.1 | -998.05 | 0.11 ‰ | -998.05 | 0.11 ‰ |  |  | 2015-08-13 | Jocelyn Turnbull | 15:20 | 0.0019236 | 0.0019236 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.00192 | 0.0001 | 10314 | 2015-08-13 | 11:29 | Albert Zondervan | 202609 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 12 | 5 | 7 | ... | 0.00185 | 0.0001 |  |  |  |  | 1.0048 | 1.0048 | Yes | Yes | [5 empty] | 0.00185 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00192 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00185 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-10 | Margaret Norris | 10:53 |  | Completed |  |  |  |  | 68473 | 2000000 |  |  | 68544 68545 68547 68548 68549 68550 | 68544~68545~68547~68548~68549~68550 | 3020 |  | 3020 3020 3020 3020 3020 3020 | 2015-08-13 | 14:29 | Albert Zondervan |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99999 | 0.0019 | 68544 |  | XCAMS | ST | 0% |  | - required quality achieved through close-up scrutiny and adjustment - very last TW measurement before SNICS source clean - error adjustments only needed for run-to-run repeatability: global factor 1.04 and cathode-specific factor 1.3 for pos#13 and #36 - | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-07-24 | Carbonate | Not Complete | JOB-202609 |  | make a c13 std from this job | 59628 |  | 14047/1 |  |  | 2015-07-22 | 3496 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00203 | 68545 | [26 empty] | 2015-07-23 | 2024 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00198 | 68547 | [26 empty] | 2015-07-23 | 1277 | Gas Splitting | Completed
[168 empty] | 0.99837 | 0.00197 | 68548 | [26 empty] | 2015-07-23 | 2513 | Choose Gas Splitting | Completed
[168 empty] | 1.0015 | 0.00196 | 68549 | [26 empty] | 2015-07-24 | 8483 | Graphite | Completed
[168 empty] | 0.99767 | 0.00199 | 68550 | [26 empty] | 2015-08-03 | 3816 | C13 Measurement | Completed
[197 empty] | 2015-08-13 | 10314 | AMS Submission Results | Completed
[197 empty] | 2015-08-13 | 12240 | Close Job | Completed
[7 empty] | 80.2 | 89.1527 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43535 | 43534 | 43535 | 264 | 263 | 264 |  |  | Margaret Norris | -995.57 | 0.15 ‰ | 2015-08-11 | 2015-05-03 | 13 | 100 |  |  |  |  | 2.95‰ | 2.95‰ | 2.00‰ | 2.00‰ | 2.95 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -995.57 ‰ | -995.57 ‰ | 0.15 | 0.15 | -995.52 | 0.15 ‰ | -995.52 | 0.15 ‰ |  |  | 2015-08-13 | Jocelyn Turnbull | 15:20 | 0.0044295 | 0.0044295 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00443 | 0.00015 | 10318 | 2015-08-13 | 11:29 | Albert Zondervan | 202675 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.44 | 0.01% | 32 | 5 | 6 | ... | 0.00426 | 0.00014 |  |  |  |  | 1.0048 | 1.0048 | Yes | Yes | [5 empty] | 0.00426 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00443 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00426 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-10 | Margaret Norris | 10:53 |  | Completed |  |  |  |  | 68516 | 2000000 |  |  | 68544 68545 68547 68548 68549 68550 | 68544~68545~68547~68548~68549~68550 | 3020 |  | 3020 3020 3020 3020 3020 3020 | 2015-08-13 | 14:29 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99999 | 0.0019 | 68544 |  | XCAMS | ST | 0% |  | - required quality achieved through close-up scrutiny and adjustment - very last TW measurement before SNICS source clean - error adjustments only needed for run-to-run repeatability: global factor 1.04 and cathode-specific factor 1.3 for pos#13 and #36 - | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-09-09 | Water |  | JOB-202675 |  | water line nitrogen flush water vessel | 59629 | Normal | 14047/11 |  |  | 2015-07-27 | 2027 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00203 | 68545 | [26 empty] | 2015-07-27 | 1279 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00198 | 68547 | [26 empty] | 2015-07-27 | 2516 | Choose Gas Splitting | Completed
[168 empty] | 0.99837 | 0.00197 | 68548 | [26 empty] | 2015-07-29 | 8526 | Graphite | Completed
[168 empty] | 1.0015 | 0.00196 | 68549 | [26 empty] | 2015-08-03 | 3848 | C13 Measurement | Completed
[168 empty] | 0.99767 | 0.00199 | 68550 | [26 empty] | 2015-08-13 | 10318 | AMS Submission Results | Completed
[197 empty] | 2015-08-13 | 12244 | Close Job | Completed
[7 empty] | 118.5 | 88.8922 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49895 | 49894 | 49895 | 416 | 416 | 416 |  |  | jenny dahl | -997.99 | 0.10 ‰ | 2015-08-27 | 2015-08-27 | 6 | [5 empty] | 1.19‰ | 1.19‰ | 2.00‰ | 2.00‰ | 1.19 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -997.99 ‰ | -997.99 ‰ | 0.1 | 0.1 | -997.97 | 0.11 ‰ | -997.97 | 0.11 ‰ |  |  | 2015-09-01 | jenny dahl | 17:11 | 0.0020068 | 0.0020068 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.4 |  |  | 0.00201 | 0.0001 | 10427 | 2015-08-31 | 9:18 | Albert Zondervan | 202674 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 4 | 5 | 6 | ... | 0.00193 | 0.0001 |  |  |  |  | 1.0001 | 1.0001 | Yes | Yes |  | shows unexpected steep rise in 14C/C towards later runs, which can't be explained by sputtering through; these runs were rejected |  |  |  | 0.00193 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00201 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00193 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-26 | jenny dahl | 16:47 |  | Completed |  |  |  |  | 68687 | 2000000 |  |  | 68551 68552 68652 68655 68656 68657 | 68551~68552~68652~68655~68656~68657 | 3023 |  | 3023 3023 3023 3023 3023 3023 | 2015-08-31 | 11:53 | Albert Zondervan |  | 0.80 mg |  | 0.8 | [5 empty] | 1.0009 | 0.00165 | 68551 |  | XCAMS | ST | 0% |  | - good measurement overall - global uncertainty multiplier is 1.05 to account for excess run-to-run variability - no system error - some questions arose regarding time-patterns of cathodes with (near-) blank 14C/C ratio values: * usual surface contaminati | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-07-29 | Carbonate |  | JOB-202674 |  | spare | 59765 |  | 14047/1 |  |  | 2015-08-18 | 2078 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00166 | 68552 | [26 empty] | 2015-08-18 | 1295 | Gas Splitting | Completed
[168 empty] | 1 | 0.00164 | 68652 | [26 empty] | 2015-08-18 | 2548 | Choose Gas Splitting | Completed
[168 empty] | 0.99921 | 0.00165 | 68655 | [26 empty] | 2015-08-21 | 8697 | Graphite | Completed
[168 empty] | 1.0014 | 0.00163 | 68656 | [26 empty] | 2015-08-24 | 3905 | C13 Measurement | Completed
[168 empty] | 0.99773 | 0.00177 | 68657 | [26 empty] | 2015-09-01 | 10427 | AMS Submission Results | Completed
[197 empty] | 2015-09-03 | 12423 | Close Job | Completed
[7 empty] | 124.3 | 88.7803 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47109 | 47108 | 47109 | 294 | 294 | 294 |  |  | jenny dahl | -997.16 | 0.10 ‰ | 2015-08-27 | 2015-08-27 | 20 | [5 empty] | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -997.16 ‰ | -997.16 ‰ | 0.1 | 0.1 | -997.13 | 0.11 ‰ | -997.13 | 0.11 ‰ |  |  | 2015-09-01 | jenny dahl | 17:11 | 0.0028386 | 0.0028386 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00284 | 0.0001 | 10437 | 2015-08-31 | 9:18 | Albert Zondervan | 202777 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 24 | 5 | 4 | ... | 0.00273 | 0.0001 |  |  |  |  | 1.0001 | 1.0001 | Yes | Yes | [5 empty] | 0.00273 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00284 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00273 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-26 | jenny dahl | 16:47 |  | Completed |  |  |  |  | 68609 | 2000000 |  |  | 68551 68552 68652 68655 68656 68657 | 68551~68552~68652~68655~68656~68657 | 3023 |  | 3023 3023 3023 3023 3023 3023 | 2015-08-31 | 11:53 | Albert Zondervan |  | 1.21 mg |  | 1.21 | [5 empty] | 1.0009 | 0.00165 | 68551 |  | XCAMS | ST | 0% |  | - good measurement overall - global uncertainty multiplier is 1.05 to account for excess run-to-run variability - no system error - some questions arose regarding time-patterns of cathodes with (near-) blank 14C/C ratio values: * usual surface contaminati | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-09-25 | Water |  | JOB-202777 |  | water line nitrogen flush water vessel | 59775 | Normal | 14047/11 |  |  | 2015-08-05 | 2058 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00166 | 68552 | [26 empty] | 2015-08-07 | 8619 | Graphite | Completed
[168 empty] | 1 | 0.00164 | 68652 | [26 empty] | 2015-08-11 | 3869 | C13 Measurement | Completed
[168 empty] | 0.99921 | 0.00165 | 68655 | [26 empty] | 2015-09-01 | 10437 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00163 | 68656 | [26 empty] | 2015-09-03 | 12432 | Close Job | Completed
[168 empty] | 0.99773 | 0.00177 | 68657
[5 empty] | 3.4666e-05 | 0.0028509 | 118.9 | 89.9558 | 0.0679 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7168 | 7168 | 7159 | 31 | 31 | 34 | 7137.2 | 7199.5 | jenny dahl | -590.31 | 1.59 ‰ | 2015-08-27 | 2015-02-28 | 20 | 180 |  |  |  |  | -10.65‰ | -10.65‰ | 2.00‰ | 2.00‰ | -10.65 ‰ | 2.00 ‰ | AMS | -8.03‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -590.31 ‰ | -589.84 ‰ | 1.59 | 1.73 | -585.25 | 1.61 ‰ | -585.25 | 1.61 ‰ |  |  | 2015-09-01 | jenny dahl | 17:11 | 0.40969 | 0.41016 | 0.00159 | 0.00173 | 0.0015 | 0.0015 | 2.2 |  |  | 0.41188 | 0.00123 | 10438 | 2015-08-31 | 9:18 | Albert Zondervan | 202778 | 0.00349 | 0.0027 | 0.00157 | 0.0012 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.97 | 0.16% | 30 | 5 | 6 | ... | 0.39612 | 0.00107 |  |  |  |  | 1.0001 | 1.0001 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 36.0296377747981716. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00349 | 0.00274 | 0.00153 | 0.39401 | 0.39‰ | 0.00144 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.40969 | 0 | 0.00159 | 0 | 0 | 0.00349 | 0 | 0.00157 | 0 | 0.39401 | 0 | 0.00144 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-26 | jenny dahl | 16:47 |  | Completed | 0.40969 | 0.41016 | 0.39401 | 0.39446 | 68610 | 2000000 | 68609 68516 68127 67849 67781 67468 67472 67304 | 68609~68516~68127~67849~67781~67468~67472~67304 | 68551 68552 68652 68655 68656 68657 | 68551~68552~68652~68655~68656~68657 | 3023 | 3023 3020 3009 3003 3002 2994 2994 2992 | 3023 3023 3023 3023 3023 3023 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.99 mg | 0.13 | 0.99 | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 1.0009 | 0.00165 | 68551 | EA | XCAMS | ST | 0% |  | - good measurement overall - global uncertainty multiplier is 1.05 to account for excess run-to-run variability - no system error - some questions arose regarding time-patterns of cathodes with (near-) blank 14C/C ratio values: * usual surface contaminati | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-09-25 | Water |  | JOB-202778 |  |  | 59776 | Normal | 14047/12 |  |  | 2015-08-05 | 2059 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 1.0004 | 0.00166 | 68552 | EA | [25 empty] | 2015-08-05 | 1286 | Gas Splitting | Completed
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 1 | 0.00164 | 68652 | EA | [25 empty] | 2015-08-05 | 2535 | Choose Gas Splitting | Completed
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 0.99921 | 0.00165 | 68655 | ST | [25 empty] | 2015-08-07 | 8620 | Graphite | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 1.0014 | 0.00163 | 68656 | EA | [25 empty] | 2015-08-11 | 3870 | C13 Measurement | Completed
[163 empty] | 2015-03-18 | 0.00378 | 67468 | 2994 | 0.98 | 0.99773 | 0.00177 | 68657 | ST | [25 empty] | 2015-09-01 | 10438 | AMS Submission Results | Completed
[163 empty] | 2015-03-18 | 0.00472 | 67472 | 2994 | 1.01 |  |  |  | ST | [25 empty] | 2015-09-03 | 12433 | Close Job | Completed
[163 empty] | 2015-03-06 | 0.0029 | 67304 | 2992 | 1.11
[7 empty] | 124.2 | 89.008 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44327 | 44326 | 44327 | 229 | 228 | 229 |  |  | jenny dahl | -995.99 | 0.11 ‰ | 2015-09-17 | 2015-03-21 | 9 | 180 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -995.99 ‰ | -995.99 ‰ | 0.11 | 0.11 | -995.94 | 0.12 ‰ | -995.94 | 0.12 ‰ |  |  | 2015-09-21 | jenny dahl | 13:06 | 0.0040136 | 0.0040136 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.9 |  |  | 0.00401 | 0.00011 | 10565 | 2015-09-18 | 13:24 | john west | 202831 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 4 | 5 | 6 | ... | 0.00386 | 0.00011 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes | [5 empty] | 0.00386 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00401 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00386 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-11 | jenny dahl | 16:10 |  | Completed |  |  |  |  | 68799 | 2000000 |  |  | 68752 68755 68820 68821 68845 68846 | 68752~68755~68820~68821~68845~68846 | 3026 |  | 3026 3026 3026 3026 3026 3026 | 2015-09-18 | 14:41 | john west |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0015 | 0.00185 | 68752 |  | XCAMS | EA | 0% |  | - 1st measurement after tank opening - first 3 rotations rejected due to gas-stripper stabilisation - only error adjustment is on cat#7, which is an Ox-I | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-10-07 | Water |  | JOB-202831 |  | water line nitrogen flush water vessel | 59868 | Normal | 14047/11 |  |  | 2015-09-04 | 2106 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00236 | 68755 | [26 empty] | 2015-09-04 | 1304 | Gas Splitting | Completed
[168 empty] | 0.99824 | 0.00177 | 68820 | [26 empty] | 2015-09-04 | 2561 | Choose Gas Splitting | Completed
[168 empty] | 0.99697 | 0.00177 | 68821 | [26 empty] | 2015-09-08 | 8809 | Graphite | Completed
[168 empty] | 1.0009 | 0.0018 | 68845 | [26 empty] | 2015-09-14 | 3932 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.0018 | 68846 | [26 empty] | 2015-09-21 | 10565 | AMS Submission Results | Completed
[197 empty] | 2015-09-21 | 12580 | Close Job | Completed
[5 empty] | 3.8778e-05 | 0.002651 | 124.2 | 89.9745 | 0.0684 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7086 | 7085 | 7075 | 31 | 31 | 34 | 7054.3 | 7116.9 | jenny dahl | -586.07 | 1.61 ‰ | 2015-09-17 | 2015-03-21 | 9 | 180 |  |  |  |  | -7.88‰ | -7.88‰ | 2.00‰ | 2.00‰ | -7.88 ‰ | 2.00 ‰ | AMS | -8.12‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -586.07 ‰ | -585.53 ‰ | 1.61 | 1.75 | -580.96 | 1.63 ‰ | -580.96 | 1.63 ‰ |  |  | 2015-09-21 | jenny dahl | 13:06 | 0.41393 | 0.41447 | 0.00161 | 0.00175 | 0.00152 | 0.00152 | 2.2 |  |  | 0.41606 | 0.00129 | 10566 | 2015-09-18 | 13:24 | john west | 202832 | 0.0034 | 0.0026 | 0.00153 | 0.0012 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.39 | 0.16% | 30 | 5 | 6 | ... | 0.40014 | 0.00113 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 36.1133581757484852. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0034 | 0.00255 | 0.00171 | 0.39809 | 0.40‰ | 0.00146 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.41393 | 0 | 0.00161 | 0 | 0 | 0.0034 | 0 | 0.00153 | 0 | 0.39809 | 0 | 0.00146 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-11 | jenny dahl | 16:10 |  | Completed | 0.41394 | 0.41447 | 0.3981 | 0.39861 | 68800 | 2000000 | 68799 68609 68516 68127 67849 67781 | 68799~68609~68516~68127~67849~67781 | 68752 68755 68820 68821 68845 68846 | 68752~68755~68820~68821~68845~68846 | 3026 | 3026 3023 3020 3009 3003 3002 | 3026 3026 3026 3026 3026 3026 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.82 mg | 0.13 | 0.82 | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 1.0015 | 0.00185 | 68752 | EA | XCAMS | EA | 0% |  | - 1st measurement after tank opening - first 3 rotations rejected due to gas-stripper stabilisation - only error adjustment is on cat#7, which is an Ox-I | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-10-07 | Water |  | JOB-202832 |  |  | 59869 | Normal | 14047/12 |  |  | 2015-09-07 | 2107 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 1.0019 | 0.00236 | 68755 | ST | [25 empty] | 2015-09-07 | 1305 | Gas Splitting | Completed
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 0.99824 | 0.00177 | 68820 | EA | [25 empty] | 2015-09-07 | 2563 | Choose Gas Splitting | Completed
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 0.99697 | 0.00177 | 68821 | ST | [25 empty] | 2015-09-08 | 8810 | Graphite | Completed
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 1.0009 | 0.0018 | 68845 | ST | [25 empty] | 2015-09-14 | 3933 | C13 Measurement | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 1.0016 | 0.0018 | 68846 | EA | [25 empty] | 2015-09-21 | 10566 | AMS Submission Results | Completed
[197 empty] | 2015-09-21 | 12581 | Close Job | Completed
[7 empty] | 100.5 | 89.3455 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43822 | 43822 | 43822 | 235 | 234 | 235 |  |  | jenny dahl | -995.73 | 0.12 ‰ | 2015-09-17 | 2015-03-21 | 14 | 180 |  |  |  |  | -0.90‰ | -0.90‰ | 2.00‰ | 2.00‰ | -0.90 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -995.73 ‰ | -995.73 ‰ | 0.12 | 0.12 | -995.67 | 0.13 ‰ | -995.67 | 0.13 ‰ |  |  | 2015-09-21 | jenny dahl | 13:06 | 0.0042736 | 0.0042736 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.6 |  |  | 0.00427 | 0.00012 | 10567 | 2015-09-18 | 13:24 | john west | 202833 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.43 | 0.01% | 24 | 5 | 6 | ... | 0.00411 | 0.00012 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes | [5 empty] | 0.00411 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00427 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00411 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-11 | jenny dahl | 16:10 |  | Completed |  |  |  |  | 68776 | 2000000 |  |  | 68752 68755 68820 68821 68845 68846 | 68752~68755~68820~68821~68845~68846 | 3026 |  | 3026 3026 3026 3026 3026 3026 | 2015-09-18 | 14:41 | john west |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0015 | 0.00185 | 68752 |  | XCAMS | EA | 0% |  | - 1st measurement after tank opening - first 3 rotations rejected due to gas-stripper stabilisation - only error adjustment is on cat#7, which is an Ox-I | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-08-26 | Carbonate | Complete | JOB-202833 |  |  | 59867 |  | 14047/1 |  |  | 2015-09-01 | 2099 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0019 | 0.00236 | 68755 | [26 empty] | 2015-09-01 | 1303 | Gas Splitting | Completed
[168 empty] | 0.99824 | 0.00177 | 68820 | [26 empty] | 2015-09-01 | 2559 | Choose Gas Splitting | Completed
[168 empty] | 0.99697 | 0.00177 | 68821 | [26 empty] | 2015-09-03 | 8786 | Graphite | Completed
[168 empty] | 1.0009 | 0.0018 | 68845 | [26 empty] | 2015-09-07 | 3918 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.0018 | 68846 | [26 empty] | 2015-09-21 | 10567 | AMS Submission Results | Completed
[197 empty] | 2015-09-21 | 12571 | Close Job | Completed
[5 empty] | 3.8778e-05 | 0.002651 | 94 | 90.5944 | 0.068 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7072 | 7071 | 7061 | 32 | 31 | 34 | 7039.9 | 7103.5 | jenny dahl | -585.35 | 1.64 ‰ | 2015-09-30 | 2015-04-03 | 2 | 180 |  |  |  |  | -8.35‰ | -8.35‰ | 2.00‰ | 2.00‰ | -8.35 ‰ | 2.00 ‰ | AMS | -8.07‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -585.35 ‰ | -584.82 ‰ | 1.64 | 1.78 | -580.23 | 1.66 ‰ | -580.23 | 1.66 ‰ |  |  | 2015-10-08 | jenny dahl | 11:44 | 0.41465 | 0.41518 | 0.00164 | 0.00178 | 0.00155 | 0.00155 | 2.3 |  |  | 0.41677 | 0.00133 | 10676 | 2015-10-01 | 12:54 | john west | 203020 | 0.0034 | 0.0026 | 0.00153 | 0.0012 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.46 | 0.16% | 37 | 5 | 6 | ... | 0.40082 | 0.00117 |  |  |  |  | 1.0095 | 1.0095 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 36.5661728766724057. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0034 | 0.00255 | 0.00171 | 0.39878 | 0.40‰ | 0.00149 | 0.00‰ | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.41465 | 0 | 0.00164 | 0 | 0 | 0.0034 | 0 | 0.00153 | 0 | 0.39878 | 0 | 0.00149 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-28 | jenny dahl | 14:47 |  | Completed | 0.41465 | 0.41518 | 0.39878 | 0.39929 | 68964 | 2000000 | 68799 68609 68516 68127 67849 67781 | 68799~68609~68516~68127~67849~67781 | 68916 68917 68918 68919 68940 68946 | 68916~68917~68918~68919~68940~68946 | 3029 | 3026 3023 3020 3009 3003 3002 | 3029 3029 3029 3029 3029 3029 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.89 mg | 0.13 | 0.89 | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 1.0019 | 0.00183 | 68916 | EA | XCAMS | ST | 0% |  | - 1st measurement aborted after 6 rotations - 2nd measurement finished after 15 rotations - good measurement, no error adjustments required - IAEA-C1 cathode (TP68963) gave high result (>4x other blanks) consistently throughout measurement | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-11-06 | Water | Complete | JOB-203020 |  |  | 59948 | Normal | 14047/12 |  |  | 2015-09-25 | 2150 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 0.99909 | 0.00179 | 68917 | ST | [25 empty] | 2015-09-25 | 1319 | Gas Splitting | Completed
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 1.0013 | 0.00181 | 68918 | EA | [25 empty] | 2015-09-25 | 2585 | Choose Gas Splitting | Completed
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 0.99896 | 0.00181 | 68919 | ST | [25 empty] | 2015-09-28 | 8974 | Graphite | Completed
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 | 0.99698 | 0.00189 | 68940 | ST | [25 empty] | 2015-10-06 | 3977 | C13 Measurement | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 | 1.0016 | 0.00181 | 68946 | EA | [25 empty] | 2015-10-08 | 10676 | AMS Submission Results | Completed
[197 empty] | 2015-10-08 | 12634 | Close Job | Completed
[7 empty] | 90.8 | 89.7276 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37168 | 37168 | 37168 | 145 | 145 | 145 |  |  | jenny dahl | -990.22 | 0.18 ‰ | 2015-09-30 | 2015-04-03 | 2 | 180 |  |  |  |  | 1.23‰ | 1.23‰ | 2.00‰ | 2.00‰ | 1.23 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -990.22 ‰ | -990.22 ‰ | 0.18 | 0.18 | -990.09 | 0.18 ‰ | -990.09 | 0.18 ‰ |  |  | 2015-10-08 | jenny dahl | 11:44 | 0.0097845 | 0.0097845 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.3 |  |  | 0.00978 | 0.00018 | 10678 | 2015-10-01 | 12:54 | john west | 203037 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.98 | 0.02% | 30 | 5 | 6 | ... | 0.00941 | 0.00017 |  |  |  |  | 1.0095 | 1.0095 | Yes | Yes |  | between the horrid CRA and the notes from John about measurement we have decided to discard this sample as a blank. the secondary done same day is a bit off in CRA but still in keeping with past samples and didn't have the measurement error so believe it was not an effect in the waterline- 8/10/15-jd Measured RTS > 4x other blanks consistent results throughout measurement |  |  |  | 0.00941 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00978 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00941 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-28 | jenny dahl | 14:47 |  | Completed |  |  |  |  | 68963 | 2000000 |  |  | 68916 68917 68918 68919 68940 68946 | 68916~68917~68918~68919~68940~68946 | 3029 |  | 3029 3029 3029 3029 3029 3029 |  |  |  |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0019 | 0.00183 | 68916 |  | XCAMS | ST | 0% |  | - 1st measurement aborted after 6 rotations - 2nd measurement finished after 15 rotations - good measurement, no error adjustments required - IAEA-C1 cathode (TP68963) gave high result (>4x other blanks) consistently throughout measurement | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-11-09 | Water | Complete | JOB-203037 |  | dreadful blank, nothing unusual about its processing, have changed category so it doesn't come up in blank selection. John has alerted us that this sample is behaving badly in measurement. In looking through the evolution, splitting and graphite I can't see any issues. material was freshly etched and ground-jd From job 203101 | 59947 | Normal | 14047/11 |  |  | 2015-09-25 | 2151 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99909 | 0.00179 | 68917 | [26 empty] | 2015-09-25 | 1318 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00181 | 68918 | [26 empty] | 2015-09-25 | 2583 | Choose Gas Splitting | Completed
[168 empty] | 0.99896 | 0.00181 | 68919 | [26 empty] | 2015-09-28 | 8973 | Graphite | Completed
[168 empty] | 0.99698 | 0.00189 | 68940 | [26 empty] | 2015-10-06 | 3976 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.00181 | 68946 | [26 empty] | 2015-10-08 | 10678 | AMS Submission Results | Completed
[197 empty] | 2015-10-08 | 12635 | Close Job | Completed
[5 empty] | 3.6808e-05 | 0.003218 | 91.7 | 90.2468 | 0.0683 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7185 | 7185 | 7167 | 37 | 36 | 33 | 7148.7 | 7222 | jenny dahl | -591.17 | 1.87 ‰ | 2015-10-16 | 2015-04-19 | 8 | 180 |  |  |  |  | -6.05‰ | -6.05‰ | 2.00‰ | 2.00‰ | -6.05 ‰ | 2.00 ‰ | AMS | -8.10‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.8165 | -591.17 ‰ | -590.23 ‰ | 1.87 | 1.66 | -586.13 | 1.89 ‰ | -586.13 | 1.89 ‰ |  |  | 2015-10-19 | Jocelyn Turnbull | 15:16 | 0.40883 | 0.40977 | 0.00187 | 0.00166 | 0.00179 | 0.00179 | 2.2 |  |  | 0.41173 | 0.00133 | 10766 | 2015-10-18 | 14:25 | Albert Zondervan | 203021 | 0.0046 | 0.0031 | 0.00207 | 0.0014 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.88 | 0.19% | 37 | 5 | 6 | ... | 0.39597 | 0.00117 |  |  |  |  | 1.008 | 1.008 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.8792170580251316. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0046 | 0.0031 | 0.00273 | 0.39318 | 0.39‰ | 0.00172 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.40883 | 0 | 0.00187 | 0 | 0 | 0.0046 | 0 | 0.00207 | 0 | 0.39318 | 0 | 0.00172 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-12 | jenny dahl | 10:11 |  | Completed | 0.40883 | 0.40977 | 0.39318 | 0.39409 | 69033 | 2000000 | 69032 68963 68799 68609 68516 68127 67849 67781 | 69032~68963~68799~68609~68516~68127~67849~67781 | 68974 68977 68988 68989 69028 69029 | 68974~68977~68988~68989~69028~69029 | 3031 | 3031 3029 3026 3023 3020 3009 3003 3002 | 3031 3031 3031 3031 3031 3031 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.84 mg | 0.13 | 0.84 | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 0.9991 | 0.00179 | 68974 | EA | XCAMS | EA | 0% |  | * good measurement TW3031_2 * all cathodes ran well | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-11-06 | Water | Complete | JOB-203021 |  | started C13 to allow AMS to be started Spare | 60053 | Normal | 14047/12 |  |  | 2015-10-07 | 2170 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-09-30 | 0.00941 | 68963 | 3029 | 0.84 | 0.99801 | 0.00187 | 68977 | ST | [25 empty] | 2015-10-07 | 1324 | Gas Splitting | Completed
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 1.0039 | 0.00194 | 68988 | EA | [25 empty] | 2015-10-07 | 2593 | Choose Gas Splitting | Completed
[163 empty] | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 0.99862 | 0.00188 | 68989 | ST | [25 empty] | 2015-10-08 | 9043 | Graphite | Completed
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 0.99989 | 0.00189 | 69028 | ST | [25 empty] | 2015-10-16 | 3998 | C13 Measurement | Completed
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 1.0016 | 0.0026 | 69029 | EA | [25 empty] | 2015-10-19 | 10766 | AMS Submission Results | Completed
[163 empty] | 2015-05-12 | 0.00611 | 67849 | 3003 | 0.83 |  |  |  | ST | [25 empty] | 2015-10-19 | 12759 | Close Job | Completed
[163 empty] | 2015-05-08 | 0.00194 | 67781 | 3002 | 0.76 |  |  |  | EA
[7 empty] | 92.8 | 89.2758 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39545 | 39545 | 39545 | 172 | 172 | 172 |  |  | jenny dahl | -992.72 | 0.16 ‰ | 2015-10-16 | 2015-10-16 | 8 | 0 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.8165 | -992.72 ‰ | -992.72 ‰ | 0.16 | 0.16 | -992.63 | 0.16 ‰ | -992.63 | 0.16 ‰ |  |  | 2015-10-19 | Jocelyn Turnbull | 15:16 | 0.0072786 | 0.0072786 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00728 | 0.00016 | 10767 | 2015-10-18 | 14:25 | Albert Zondervan | 203038 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.73 | 0.02% | 34 | 5 | 6 | ... | 0.007 | 0.00015 |  |  |  |  | 1.008 | 1.008 | Yes | Yes |  | Ugh! But fits with other recent waters. |  |  |  | 0.007 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00728 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.007 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-12 | jenny dahl | 10:11 |  | Completed |  |  |  |  | 69032 | 2000000 |  |  | 68974 68977 68988 68989 69028 69029 | 68974~68977~68988~68989~69028~69029 | 3031 |  | 3031 3031 3031 3031 3031 3031 | 2015-10-19 | 14:30 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 0.9991 | 0.00179 | 68974 |  | XCAMS | EA | 0% |  | * good measurement TW3031_2 * all cathodes ran well | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2015-11-09 | Water |  | JOB-203038 |  | water line nitrogen flush water vessel | 60052 | Normal | 14047/11 |  |  | 2015-10-07 | 2171 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99801 | 0.00187 | 68977 | [26 empty] | 2015-10-07 | 1323 | Gas Splitting | Completed
[168 empty] | 1.0039 | 0.00194 | 68988 | [26 empty] | 2015-10-07 | 2591 | Choose Gas Splitting | Completed
[168 empty] | 0.99862 | 0.00188 | 68989 | [26 empty] | 2015-10-08 | 9042 | Graphite | Completed
[168 empty] | 0.99989 | 0.00189 | 69028 | [26 empty] | 2015-10-16 | 3999 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.0026 | 69029 | [26 empty] | 2015-10-19 | 10767 | AMS Submission Results | Completed
[197 empty] | 2015-10-19 | 12761 | Close Job | Completed
[7 empty] | 84.5 | 89.4909 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39957 | 39957 | 39957 | 217 | 217 | 217 |  |  | jenny dahl | -993.09 | 0.19 ‰ | 2015-10-22 | 2015-07-24 | 6 | 90 |  |  |  |  | -0.25‰ | -0.25‰ | 2.00‰ | 2.00‰ | -0.25 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -993.09 ‰ | -993.09 ‰ | 0.19 | 0.19 | -993 | 0.19 ‰ | -993 | 0.19 ‰ |  |  | 2015-10-23 | Jocelyn Turnbull | 14:02 | 0.0069146 | 0.0069146 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.1 |  |  | 0.00691 | 0.00019 | 10844 | 2015-10-23 | 10:48 | Albert Zondervan | 203172 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.69 | 0.02% | 36 | 5 | 6 | ... | 0.00665 | 0.00018 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | Water blank still poor - see if improves from change in carbasorb just after this wheel |  |  |  | 0.00665 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00078 | 0.00078 | 0.00691 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00665 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-19 | jenny dahl | 14:19 |  | Completed |  |  |  |  | 69113 | 2000000 |  |  | 68920 68949 69087 69095 69096 69097 | 68920~68949~69087~69095~69096~69097 | 3033 |  | 3033 3033 3033 3033 3033 3033 | 2015-10-23 | 13:46 | Jocelyn Turnbull |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99947 | 0.00188 | 68920 |  | XCAMS | ST | 0% |  | * straightforward measurement * only 1 error adjustment: for TP69103 only | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-11-30 | Water |  | JOB-203172 |  | water line nitrogen flush water vessel | 60145 | Normal | 14047/11 |  |  | 2015-10-13 | 2183 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0016 | 0.0019 | 68949 | [26 empty] | 2015-10-13 | 1326 | Gas Splitting | Completed
[168 empty] | 0.99991 | 0.00199 | 69087 | [26 empty] | 2015-10-13 | 2599 | Choose Gas Splitting | Completed
[168 empty] | 0.99915 | 0.00188 | 69095 | [26 empty] | 2015-10-16 | 9123 | Graphite | Completed
[168 empty] | 1.001 | 0.00187 | 69096 | [26 empty] | 2015-10-22 | 4027 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00191 | 69097 | [26 empty] | 2015-10-23 | 10844 | AMS Submission Results | Completed
[197 empty] | 2015-10-23 | 12871 | Close Job | Completed
[5 empty] | 5.6097e-05 | 0.0049975 | 98.1 | 90.0222 | 0.0684 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7168 | 7168 | 7158 | 42 | 41 | 42 | 7127.1 | 7210.1 | jenny dahl | -590.32 | 2.12 ‰ | 2015-10-22 | 2015-07-24 | 6 | 90 |  |  |  |  | -6.15‰ | -6.15‰ | 2.00‰ | 2.00‰ | -6.15 ‰ | 2.00 ‰ | AMS | -8.06‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -590.32 ‰ | -589.79 ‰ | 2.12 | 2.15 | -585.26 | 2.14 ‰ | -585.26 | 2.14 ‰ |  |  | 2015-10-23 | Jocelyn Turnbull | 14:03 | 0.40968 | 0.41021 | 0.00212 | 0.00215 | 0.00205 | 0.00205 | 2 |  |  | 0.41324 | 0.00137 | 10845 | 2015-10-23 | 10:48 | Albert Zondervan | 203174 | 0.00565 | 0.0048 | 0.00254 | 0.0022 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.97 | 0.21% | 37 | 5 | 6 | ... | 0.39742 | 0.00121 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 45.2707838085447669. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00565 | 0.00481 | 0.00247 | 0.394 | 0.39‰ | 0.00197 | 0.00‰ | Yes | 1 | 1 | 0.00078 | 0.00078 | 0.40968 | 0 | 0.00212 | 0 | 0 | 0.00565 | 0 | 0.00254 | 0 | 0.394 | 0 | 0.00197 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-19 | jenny dahl | 14:19 |  | Completed | 0.40968 | 0.41021 | 0.394 | 0.39451 | 69114 | 2000000 | 69113 69032 68963 68799 68609 68516 | 69113~69032~68963~68799~68609~68516 | 68920 68949 69087 69095 69096 69097 | 68920~68949~69087~69095~69096~69097 | 3033 | 3033 3031 3029 3026 3023 3020 | 3033 3033 3033 3033 3033 3033 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 0.96 mg | 0.13 | 0.96 | 2015-10-22 | 0.00665 | 69113 | 3033 | 0.86 | 0.99947 | 0.00188 | 68920 | ST | XCAMS | ST | 0% |  | * straightforward measurement * only 1 error adjustment: for TP69103 only | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-11-30 | Water |  | JOB-203174 |  |  | 60146 | Normal | 14047/12 |  |  | 2015-10-14 | 2187 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 1.0016 | 0.0019 | 68949 | ST | [25 empty] | 2015-10-14 | 1327 | Gas Splitting | Completed
[163 empty] | 2015-09-30 | 0.00941 | 68963 | 3029 | 0.84 | 0.99991 | 0.00199 | 69087 | EA | [25 empty] | 2015-10-14 | 2601 | Choose Gas Splitting | Completed
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 0.99915 | 0.00188 | 69095 | ST | [25 empty] | 2015-10-16 | 9124 | Graphite | Completed
[163 empty] | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 1.001 | 0.00187 | 69096 | EA | [25 empty] | 2015-10-22 | 4028 | C13 Measurement | Completed
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 0.9989 | 0.00191 | 69097 | ST | [25 empty] | 2015-10-23 | 10845 | AMS Submission Results | Completed
[197 empty] | 2015-10-23 | 12872 | Close Job | Completed
[7 empty] | 91.4 | 89.1644 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50455 | 50454 | 50455 | 312 | 312 | 312 |  |  | jenny dahl | -998.13 | 0.07 ‰ | 2015-11-10 | 2015-11-10 | 7 | 0 |  |  |  |  | 2.58‰ | 2.58‰ | 2.00‰ | 2.00‰ | 2.58 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0086 | 0 | 0.8165 | -998.13 ‰ | -998.13 ‰ | 0.07 | 0.07 | -998.11 | 0.07 ‰ | -998.11 | 0.07 ‰ |  |  | 2015-11-12 | Jocelyn Turnbull | 16:22 | 0.0018716 | 0.0018716 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3.4 |  |  | 0.00187 | 7e-05 | 10955 | 2015-11-12 | 13:56 | Albert Zondervan | 203125 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 27 | 5 | 7 | ... | 0.0018 | 7e-05 |  |  |  |  | 1.0045 | 1.0045 | Yes | Yes |  | nice blank |  |  |  | 0.0018 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00187 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0018 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-11-09 | jenny dahl | 10:09 |  | Completed |  |  |  |  | 69207 | 2000000 |  |  | 69251 69252 69253 69254 69255 69256 | 69251~69252~69253~69254~69255~69256 | 3035 |  | 3035 3035 3035 3035 3035 3035 | 2015-11-12 | 15:49 | Albert Zondervan |  | 0.62 mg |  | 0.62 | [5 empty] | 0.99878 | 0.00149 | 69251 |  | XCAMS | ST | 0% |  | - 1st wheel after source clean - blanks show clear sign of surface contamination; only later runs have been retained - mode 5 calibration once again vindicates over the other modes in that returns the best possible repeatability amongst the Ox-I cathodes | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-10-08 | Carbonate |  | JOB-203125 |  |  | 60251 |  | 14047/1 |  |  | 2015-10-23 | 3695 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00162 | 69252 | [26 empty] | 2015-10-30 | 2215 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0007 | 0.00178 | 69253 | [26 empty] | 2015-10-30 | 1331 | Gas Splitting | Completed
[168 empty] | 0.9992 | 0.00156 | 69254 | [26 empty] | 2015-10-30 | 2608 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00156 | 69255 | [26 empty] | 2015-11-03 | 9217 | Graphite | Completed
[168 empty] | 1.0003 | 0.00157 | 69256 | [26 empty] | 2015-11-10 | 4051 | C13 Measurement | Completed
[197 empty] | 2015-11-12 | 10955 | AMS Submission Results | Completed
[197 empty] | 2015-11-13 | 13020 | Close Job | Completed
[7 empty] | 94.9 | 89.0656 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47348 | 47347 | 47348 | 303 | 303 | 303 |  |  | jenny dahl | -997.24 | 0.10 ‰ | 2015-11-10 | 2015-11-10 | 19 | 0 |  |  |  |  | 3.69‰ | 3.69‰ | 2.00‰ | 2.00‰ | 3.69 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0086 | 0 | 0.8165 | -997.24 ‰ | -997.24 ‰ | 0.1 | 0.1 | -997.21 | 0.11 ‰ | -997.21 | 0.11 ‰ |  |  | 2015-11-12 | Jocelyn Turnbull | 16:22 | 0.0027555 | 0.0027555 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00276 | 0.0001 | 10958 | 2015-11-12 | 13:56 | Albert Zondervan | 203173 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 30 | 5 | 6 | ... | 0.00265 | 0.0001 |  |  |  |  | 1.0045 | 1.0045 | Yes | Yes |  | A really good water blank. Carbasorb on N2 tank was changed day defore this blank evolved. Not certain yet that was the problem, but certainly looks promising! |  |  |  | 0.00265 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00276 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00265 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-11-09 | jenny dahl | 10:09 |  | Completed |  |  |  |  | 69144 | 2000000 |  |  | 69251 69252 69253 69254 69255 69256 | 69251~69252~69253~69254~69255~69256 | 3035 |  | 3035 3035 3035 3035 3035 3035 | 2015-11-12 | 15:49 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99878 | 0.00149 | 69251 |  | XCAMS | ST | 0% |  | - 1st wheel after source clean - blanks show clear sign of surface contamination; only later runs have been retained - mode 5 calibration once again vindicates over the other modes in that returns the best possible repeatability amongst the Ox-I cathodes | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-11-30 | Water |  | JOB-203173 |  | A really good water blank. Carbasorb on N2 tank was changed day before this blank evolved. Not certain yet that was the problem, but certainly looks promising! | 60254 | Normal | 14047/11 |  |  | 2015-10-21 | 2205 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00162 | 69252 | [26 empty] | 2015-10-21 | 1329 | Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00178 | 69253 | [26 empty] | 2015-10-21 | 2604 | Choose Gas Splitting | Completed
[168 empty] | 0.9992 | 0.00156 | 69254 | [26 empty] | 2015-10-22 | 9154 | Graphite | Completed
[168 empty] | 1.0006 | 0.00156 | 69255 | [26 empty] | 2015-10-27 | 4045 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00157 | 69256 | [26 empty] | 2015-11-12 | 10958 | AMS Submission Results | Completed
[197 empty] | 2015-11-13 | 12998 | Close Job | Completed
[5 empty] | 5.8448e-05 | 0.0042257 | 96 | 89.9148 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7115 | 7115 | 7102 | 37 | 36 | 37 | 7078.6 | 7151.7 | jenny dahl | -587.59 | 1.88 ‰ | 2015-11-10 | 2015-05-14 | 19 | 180 |  |  |  |  | -5.79‰ | -5.79‰ | 2.00‰ | 2.00‰ | -5.79 ‰ | 2.00 ‰ | AMS | -8.07‰ | 0.2‰ | C13 | No | -19 | -19.0086 | 0 | 0.8165 | -587.59 ‰ | -586.91 ‰ | 1.88 | 1.89 | -582.49 | 1.90 ‰ | -582.49 | 1.90 ‰ |  |  | 2015-11-12 | Jocelyn Turnbull | 16:22 | 0.41241 | 0.41309 | 0.00188 | 0.00189 | 0.0018 | 0.0018 | 1.8 |  |  | 0.41564 | 0.00118 | 10959 | 2015-11-12 | 13:56 | Albert Zondervan | 203175 | 0.00514 | 0.0041 | 0.00231 | 0.0018 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.24 | 0.19% | 31 | 5 | 6 | ... | 0.39973 | 0.00101 |  |  |  |  | 1.0045 | 1.0045 | Yes | Yes |  | within 1-sigma from consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.790592573958361. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00514 | 0.00406 | 0.00289 | 0.39663 | 0.40‰ | 0.00173 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41241 | 0 | 0.00188 | 0 | 0 | 0.00514 | 0 | 0.00231 | 0 | 0.39663 | 0 | 0.00173 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-11-09 | jenny dahl | 10:09 |  | Completed | 0.41241 | 0.41309 | 0.39663 | 0.39728 | 69145 | 2000000 | 69144 69113 68963 68609 68516 | 69144~69113~68963~68609~68516 | 69251 69252 69253 69254 69255 69256 | 69251~69252~69253~69254~69255~69256 | 3035 | 3035 3033 3029 3023 3020 | 3035 3035 3035 3035 3035 3035 | 2015-11-30 | 15:31 | Jocelyn Turnbull |  | 1.00 mg | 0.13 | 1 | 2015-11-10 | 0.00265 | 69144 | 3035 | 0.95 | 0.99878 | 0.00149 | 69251 | ST | XCAMS | ST | 0% |  | - 1st wheel after source clean - blanks show clear sign of surface contamination; only later runs have been retained - mode 5 calibration once again vindicates over the other modes in that returns the best possible repeatability amongst the Ox-I cathodes | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-11-30 | Water |  | JOB-203175 |  |  | 60255 | Normal | 14047/12 |  |  | 2015-10-21 | 2206 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-10-22 | 0.00665 | 69113 | 3033 | 0.86 | 1.0008 | 0.00162 | 69252 | ST | [25 empty] | 2015-10-21 | 1330 | Gas Splitting | Completed
[163 empty] | 2015-09-30 | 0.00941 | 68963 | 3029 | 0.84 | 1.0007 | 0.00178 | 69253 | ST | [25 empty] | 2015-10-21 | 2606 | Choose Gas Splitting | Completed
[163 empty] | 2015-08-27 | 0.00273 | 68609 | 3023 | 1.21 | 0.9992 | 0.00156 | 69254 | ST | [25 empty] | 2015-10-22 | 9155 | Graphite | Completed
[163 empty] | 2015-08-11 | 0.00426 | 68516 | 3020 | 0.95 | 1.0006 | 0.00156 | 69255 | ST | [25 empty] | 2015-10-27 | 4046 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00157 | 69256 | [26 empty] | 2015-11-12 | 10959 | AMS Submission Results | Completed
[197 empty] | 2015-11-13 | 12999 | Close Job | Completed
[7 empty] | 114 | 89.1509 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37254 | 37254 | 37254 | 155 | 155 | 155 |  |  | jenny dahl | -990.32 | 0.19 ‰ | 2015-12-09 | 2015-12-09 | 28 | 0 |  |  |  |  | 3.63‰ | 3.63‰ | 2.00‰ | 2.00‰ | 3.63 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -990.32 ‰ | -990.32 ‰ | 0.19 | 0.19 | -990.2 | 0.19 ‰ | -990.2 | 0.19 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.0096805 | 0.0096805 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.7 |  |  | 0.00968 | 0.00019 | 11094 | 2015-12-12 | 14:02 | Albert Zondervan | 203558 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.97 | 0.02% | 5 | 5 | 6 | ... | 0.00931 | 0.00018 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes | [5 empty] | 0.00931 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00968 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00931 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:15 |  | Completed |  |  |  |  | 69274 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-12-22 | Water |  | JOB-203558 |  | freshly etched material water line nitrogen flush water vessel | 60429 | Normal | 14047/11 |  |  | 2015-11-09 | 2226 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [26 empty] | 2015-11-09 | 1337 | Gas Splitting | Completed
[168 empty] | 0.99975 | 0.00176 | 69370 | [26 empty] | 2015-11-09 | 2622 | Choose Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [26 empty] | 2015-11-11 | 9284 | Graphite | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [26 empty] | 2015-11-18 | 4072 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [26 empty] | 2015-12-17 | 11094 | AMS Submission Results | Completed
[197 empty] | 2015-12-21 | 13284 | Close Job | Completed
[5 empty] | 4.9813e-05 | 0.0031638 | 113.6 | 90.0239 | 0.0688 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7152 | 7151 | 7152 | 39 | 38 | 43 | 7113 | 7190.4 | jenny dahl | -589.46 | 1.98 ‰ | 2015-12-09 | 2015-06-12 | 27 | 180 |  |  |  |  | -6.10‰ | -6.10‰ | 2.00‰ | 2.00‰ | -6.10 ‰ | 2.00 ‰ | AMS | -8.08‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -589.46 ‰ | -589.46 ‰ | 1.98 | 2.18 | -584.39 | 2.00 ‰ | -584.39 | 2.00 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.41054 | 0.41054 | 0.00198 | 0.00218 | 0.0019 | 0.0019 | 2.3 |  |  | 0.41394 | 0.00125 | 11114 | 2015-12-12 | 14:02 | Albert Zondervan | 203598 | 0.00541 | 0.0054 | 0.00243 | 0.0024 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.05 | 0.2% | 24 | 5 | 6 | ... | 0.3981 | 0.00108 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 42.6902031838190491. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00541 | 0.00304 | 0.00345 | 0.39483 | 0.39‰ | 0.00183 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.41054 | 0 | 0.00198 | 0 | 0 | 0.00541 | 0 | 0.00243 | 0 | 0.39483 | 0 | 0.00183 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:17 |  | Completed |  |  |  |  | 69288 | 2000000 | 69274 69032 68799 68127 | 69274~69032~68799~68127 | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 | 3040 3031 3026 3009 | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.98 mg | 0.13 | 0.98 | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 0.99931 | 0.00169 | 69359 | EA | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-12-25 | Water |  | JOB-203598 |  |  | 60449 | Normal | 14047/12 |  |  | 2015-11-09 | 2227 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 0.99933 | 0.0018 | 69362 | EA | [25 empty] | 2015-11-09 | 1338 | Gas Splitting | Completed
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 0.99975 | 0.00176 | 69370 | EA | [25 empty] | 2015-11-09 | 2624 | Choose Gas Splitting | Completed
[163 empty] | 2015-06-15 | 0.00147 | 68127 | 3009 | 0.98 | 0.99761 | 0.00169 | 69375 | EA | [25 empty] | 2015-11-12 | 9298 | Graphite | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [26 empty] | 2015-11-18 | 4076 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [26 empty] | 2015-12-17 | 11114 | AMS Submission Results | Completed
[197 empty] | 2015-12-17 | 13260 | Close Job | Completed
[7 empty] | 112.6 | 89.2815 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34548 | 34547 | 34548 | 123 | 123 | 123 | 34425.3 | 34671.8 | jenny dahl | -986.44 | 0.21 ‰ | 2015-12-09 | 2015-12-09 | 14 | 0 |  |  |  |  | 2.16‰ | 2.16‰ | 2.00‰ | 2.00‰ | 2.16 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -986.44 ‰ | -986.44 ‰ | 0.21 | 0.21 | -986.27 | 0.21 ‰ | -986.27 | 0.21 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.013559 | 0.013559 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2.3 |  |  | 0.01356 | 0.00021 | 11119 | 2015-12-12 | 14:02 | Albert Zondervan | 203633 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.36 | 0.02% | 29 | 5 | 8 | ... | 0.01304 | 0.0002 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | seems way too high F |  |  |  | 0.01304 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.01356 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01304 | 0 | 0.0002 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:17 |  | Completed |  |  |  |  | 69371 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.80 mg |  | 0.8 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Test |  | 2016-01-06 | Water | Complete | JOB-203633 |  | Test of new small water vessel No job created for second c13 split failed | 60454 | Normal | 14047/11 |  |  | 2015-11-23 | 2230 | Graphite | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [26 empty] | 2015-11-23 | 2241 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99975 | 0.00176 | 69370 | [26 empty] | 2015-11-24 | 2245 | Water CO2 Evolution | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [26 empty] | 2015-11-24 | 1345 | Gas Splitting | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [26 empty] | 2015-11-24 | 2635 | Choose Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [26 empty] | 2015-11-25 | 9381 | Graphite | Completed
[197 empty] | 2015-11-27 | 4092 | C13 Measurement | Completed
[197 empty] | 2015-12-17 | 11119 | AMS Submission Results | Completed
[197 empty] | 2015-12-17 | 13262 | Close Job | Completed
[7 empty] | 112.5 | 89.2114 | 0.0015 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37753 | 37752 | 37753 | 156 | 156 | 156 | 37597.9 | 37910.1 | jenny dahl | -990.9 | 0.18 ‰ | 2015-12-09 | 2015-12-09 | 15 | 0 |  |  |  |  | 2.95‰ | 2.95‰ | 2.00‰ | 2.00‰ | 2.95 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -990.90 ‰ | -990.90 ‰ | 0.18 | 0.18 | -990.79 | 0.18 ‰ | -990.79 | 0.18 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.0090982 | 0.0090982 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.8 |  |  | 0.0091 | 0.00018 | 11127 | 2015-12-12 | 14:02 | Albert Zondervan | 203670 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.91 | 0.02% | 37 | 5 | 6 | ... | 0.00875 | 0.00017 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | seems too high F |  |  |  | 0.00875 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0091 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00875 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:17 |  | Completed |  |  |  |  | 69358 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.71 mg |  | 0.71 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Test |  | 2016-01-15 | Water | Complete | JOB-203670 |  | test of new small water vessel | 60462 | Normal | 14047/11 |  |  | 2015-11-20 | 2237 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [26 empty] | 2015-11-20 | 1341 | Gas Splitting | Completed
[168 empty] | 0.99975 | 0.00176 | 69370 | [26 empty] | 2015-11-20 | 2629 | Choose Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [26 empty] | 2015-11-24 | 9368 | Graphite | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [26 empty] | 2015-11-27 | 4090 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [26 empty] | 2015-12-17 | 11127 | AMS Submission Results | Completed
[197 empty] | 2015-12-17 | 13263 | Close Job | Completed
[5 empty] | 7.3019e-05 | 0.0066974 | 100.4 | 89.9067 | 0.0665 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7303 | 7303 | 7303 | 49 | 48 | 52 | 7255 | 7352.1 | jenny dahl | -597.14 | 2.44 ‰ | 2016-01-27 | 2015-07-31 | 8 | 180 |  |  |  |  | -6.58‰ | -6.58‰ | 2.00‰ | 2.00‰ | -6.58 ‰ | 2.00 ‰ | AMS | -8.03‰ | 0.2‰ | C13 | No | -19 | -19.006 | 0 | 0.8165 | -597.14 ‰ | -597.14 ‰ | 2.44 | 2.6 | -592.17 | 2.47 ‰ | -592.17 | 2.47 ‰ |  |  | 2016-02-01 | Jocelyn Turnbull | 9:47 | 0.40286 | 0.40286 | 0.00244 | 0.0026 | 0.00238 | 0.00238 | 2.5 |  |  | 0.40745 | 0.00126 | 11152 | 2016-01-28 | 13:18 | Albert Zondervan | 203022 | 0.00721 | 0.0072 | 0.00324 | 0.0032 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.29 | 0.24% | 1 | 5 | 6 | A.. | 0.39186 | 0.00127 |  |  |  |  | 1.0036 | 1.0036 | Yes | Yes |  | result disagrees from consensus value; does this indicate a problem with the blank subtraction? Flagged since it is a 4 sigma outlier wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 51.8228878765752279. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00721 | 0.00644 | 0.00244 | 0.38744 | 0.39‰ | 0.00237 | 0.00‰ | No | 1 | 1.0001 | 0.00098 | 0.00098 | 0.40286 | 0 | 0.00244 | 0 | 0 | 0.00721 | 0 | 0.00324 | 0 | 0.38744 | 0 | 0.00237 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-01-22 | jenny dahl | 12:24 |  | Completed |  |  |  |  | 69451 | 2000000 | 69454 69274 69032 68799 | 69454~69274~69032~68799 | 69441 69442 69465 69471 69472 69473 | 69441~69442~69465~69471~69472~69473 | 3042 | 3042 3040 3031 3026 | 3042 3042 3042 3042 3042 3042 | 2016-01-28 | 16:51 | Albert Zondervan |  | 0.90 mg | 0.13 | 0.9 | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 | 1.0038 | 0.00251 | 69441 | EA | XCAMS | EA | 0.16% |  | * good measurement * non-zero system error probably incurred by somewhat differently behaving (lower beam current) Ox-I in pos#0 | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2015-11-06 | Water | Complete | JOB-203022 |  |  | 60487 | Normal | 14047/12 |  |  | 2016-01-15 | 2279 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 0.99708 | 0.00233 | 69442 | EA | [25 empty] | 2016-01-15 | 1381 | Gas Splitting | Completed
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 0.99826 | 0.00234 | 69465 | EA | [25 empty] | 2016-01-15 | 2672 | Choose Gas Splitting | Completed
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 0.99991 | 0.00233 | 69471 | EA | [25 empty] | 2016-01-19 | 9461 | Graphite | Completed
[168 empty] | 1 | 0.00271 | 69472 | [26 empty] | 2016-01-28 | 4131 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00231 | 69473 | [26 empty] | 2016-02-01 | 11152 | AMS Submission Results | Completed
[197 empty] | 2016-02-04 | 13338 | Close Job | Completed
[7 empty] | 104.9 | 88.8485 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37826 | 37826 | 37826 | 138 | 138 | 138 |  |  | jenny dahl | -990.98 | 0.16 ‰ | 2016-01-27 | 2016-01-27 | 8 | 0 |  |  |  |  | 5.25‰ | 5.25‰ | 2.00‰ | 2.00‰ | 5.25 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.006 | 0 | 0.8165 | -990.98 ‰ | -990.98 ‰ | 0.16 | 0.16 | -990.87 | 0.16 ‰ | -990.87 | 0.16 ‰ |  |  | 2016-02-01 | Jocelyn Turnbull | 9:48 | 0.009015 | 0.009015 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00902 | 0.00016 | 11188 | 2016-01-28 | 13:18 | Albert Zondervan | 203968 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.9 | 0.02% | 36 | 5 | 7 | ... | 0.00867 | 0.00015 |  |  |  |  | 1.0036 | 1.0036 | Yes | Yes |  | clearly another poor result for the water line blank |  |  |  | 0.00867 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1.0001 | 0.00098 | 0.00098 | 0.00902 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00867 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-01-22 | jenny dahl | 12:24 |  | Completed |  |  |  |  | 69454 | 2000000 |  |  | 69441 69442 69465 69471 69472 69473 | 69441~69442~69465~69471~69472~69473 | 3042 |  | 3042 3042 3042 3042 3042 3042 | 2016-01-28 | 16:51 | Albert Zondervan |  | 1.00 mg |  | 1 | [5 empty] | 1.0038 | 0.00251 | 69441 |  | XCAMS | EA | 0.16% |  | * good measurement * non-zero system error probably incurred by somewhat differently behaving (lower beam current) Ox-I in pos#0 | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-03-10 | Water |  | JOB-203968 |  |  | 60523 | Normal | 14047/11 |  |  | 2016-01-15 | 2280 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99708 | 0.00233 | 69442 | [26 empty] | 2016-01-18 | 2284 | Water CO2 Evolution | Completed
[168 empty] | 0.99826 | 0.00234 | 69465 | [26 empty] | 2016-01-18 | 1382 | Gas Splitting | Completed
[168 empty] | 0.99991 | 0.00233 | 69471 | [26 empty] | 2016-01-18 | 2674 | Choose Gas Splitting | Completed
[168 empty] | 1 | 0.00271 | 69472 | [26 empty] | 2016-01-19 | 9464 | Graphite | Completed
[168 empty] | 1.0017 | 0.00231 | 69473 | [26 empty] | 2016-01-28 | 4135 | C13 Measurement | Completed
[197 empty] | 2016-02-01 | 11188 | AMS Submission Results | Completed
[197 empty] | 2016-02-04 | 13357 | Close Job | Completed
[7 empty] | 118 | 89.1422 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49936 | 49936 | 49936 | 293 | 292 | 293 |  |  | jenny dahl | -998 | 0.07 ‰ | 2016-02-02 | 2016-02-02 | 20 | 0 |  |  |  |  | 3.69‰ | 3.69‰ | 2.00‰ | 2.00‰ | 3.69 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -998.00 ‰ | -998.00 ‰ | 0.07 | 0.07 | -997.98 | 0.07 ‰ | -997.98 | 0.07 ‰ |  |  | 2016-02-04 | Jocelyn Turnbull | 14:03 | 0.0019964 | 0.0019964 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.002 | 7e-05 | 11220 | 2016-02-04 | 10:07 | Albert Zondervan | 203855 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 30 | 5 | 6 | ... | 0.00192 | 7e-05 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes | [5 empty] | 0.00192 | 0.00‰ | 7e-05 | 0.00‰ | No | 1 | 1.0001 | 0.00088 | 0.00088 | 0.002 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00192 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-01-28 | jenny dahl | 16:22 |  | Completed |  |  |  |  | 69404 | 2000000 |  |  | 69507 69508 69515 69516 69517 69529 | 69507~69508~69515~69516~69517~69529 | 3043 |  | 3043 3043 3043 3043 3043 3043 | 2016-02-04 | 11:32 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0036 | 0.00251 | 69507 |  | XCAMS | ST | 0.12% |  | - 1x failed cathode measurement: TP69444 (cat#31) - other than that, there were small error adjustments in CalAMS; these are not concerning | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-01-04 | Carbonate |  | JOB-203855 |  |  | 60527 |  | 14047/1 |  |  | 2016-01-12 | 2263 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0022 | 0.00219 | 69508 | [26 empty] | 2016-01-12 | 1367 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00208 | 69515 | [26 empty] | 2016-01-12 | 2658 | Choose Gas Splitting | Completed
[168 empty] | 0.9981 | 0.00208 | 69516 | [26 empty] | 2016-01-13 | 9414 | Graphite | Completed
[168 empty] | 0.99771 | 0.00208 | 69517 | [26 empty] | 2016-01-15 | 4113 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00211 | 69529 | [26 empty] | 2016-02-04 | 11220 | AMS Submission Results | Completed
[197 empty] | 2016-02-05 | 13381 | Close Job | Completed
[5 empty] | 6.9836e-05 | 0.007435 | 120.9 | 90.0891 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7092 | 7092 | 7092 | 53 | 53 | 56 | 7039 | 7145.7 | Margaret Norris | -586.41 | 2.75 ‰ | 2016-02-11 | 2015-08-15 | 13 | 180 |  |  |  |  | -8.04‰ | -8.04‰ | 2.00‰ | 2.00‰ | -8.04 ‰ | 2.00 ‰ | AMS | -8.16‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -586.41 ‰ | -586.41 ‰ | 2.75 | 2.9 | -581.31 | 2.78 ‰ | -581.31 | 2.78 ‰ |  |  | 2016-02-15 | Jocelyn Turnbull | 14:13 | 0.41359 | 0.41359 | 0.00275 | 0.0029 | 0.00269 | 0.00269 | 2.1 |  |  | 0.41905 | 0.00119 | 11294 | 2016-02-15 | 10:38 | Silvia Canessa | 203857 | 0.00872 | 0.0087 | 0.00392 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.36 | 0.27% | 23 | 5 | 6 | ... | 0.40301 | 0.00102 |  |  |  |  | 1.0042 | 1.0042 | Yes | Yes |  | agrees with consensus. JCT- interesting - the large blank correction looks correct - if use smaller blank correction, it comes out too young. wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 56.3222531453892288. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00872 | 0.00715 | 0.00398 | 0.39776 | 0.40‰ | 0.00259 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41359 | 0 | 0.00275 | 0 | 0 | 0.00872 | 0 | 0.00392 | 0 | 0.39776 | 0 | 0.00259 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-05 | Margaret Norris | 15:23 |  | Completed |  |  |  |  | 69548 | 2000000 | 69597 69454 69274 69032 68799 | 69597~69454~69274~69032~68799 | 69499 69500 69501 69527 69528 69545 | 69499~69500~69501~69527~69528~69545 | 3045 | 3045 3042 3040 3031 3026 | 3045 3045 3045 3045 3045 3045 | 2016-02-15 | 11:57 | Silvia Canessa |  | 0.93 mg | 0.13 | 0.93 | 2016-02-11 | 0.01476 | 69597 | 3045 | 0.88 | 1.0007 | 0.00157 | 69499 | EA | XCAMS | EA | 0% |  | * good measurement: - problem with extractor resolved beforehand - high beam currents achieved by pushing source - stable ratios - no error adjustments | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-02-15 | Water | Complete | JOB-203857 |  |  | 60607 | Normal | 14047/12 |  |  | 2016-01-28 | 2297 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 | 1.0007 | 0.00159 | 69500 | EA | [25 empty] | 2016-01-28 | 1386 | Gas Splitting | Completed
[163 empty] | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 0.99997 | 0.00159 | 69501 | EA | [25 empty] | 2016-01-28 | 2688 | Choose Gas Splitting | Completed
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 0.99838 | 0.00156 | 69527 | EA | [25 empty] | 2016-01-29 | 9558 | Graphite | Completed
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 | 1.0003 | 0.00155 | 69528 | EA | [25 empty] | 2016-02-01 | 4159 | C13 Measurement | Completed
[168 empty] | 0.99997 | 0.00159 | 69545 | [26 empty] | 2016-02-15 | 11294 | AMS Submission Results | Completed
[197 empty] | 2016-02-18 | 13466 | Close Job | Completed
[7 empty] | 120.2 | 89.5649 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33552 | 33552 | 33552 | 125 | 125 | 125 |  |  | Margaret Norris | -984.65 | 0.24 ‰ | 2016-02-11 | 2016-02-11 | 6 | 0 |  |  |  |  | -2.23‰ | -2.23‰ | 2.00‰ | 2.00‰ | -2.23 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -984.65 ‰ | -984.65 ‰ | 0.24 | 0.24 | -984.46 | 0.24 ‰ | -984.46 | 0.24 ‰ |  |  | 2016-02-15 | Jocelyn Turnbull | 14:13 | 0.015347 | 0.015347 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.5 |  |  | 0.01535 | 0.00024 | 11308 | 2016-02-15 | 10:38 | Silvia Canessa | 203944 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.53 | 0.02% | 36 | 5 | 7 | ... | 0.01476 | 0.00023 |  |  |  |  | 1.0042 | 1.0042 | Yes | Yes | [5 empty] | 0.01476 | 0.01‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01535 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01476 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-05 | Margaret Norris | 15:23 |  | Completed |  |  |  |  | 69597 | 2000000 |  |  | 69499 69500 69501 69527 69528 69545 | 69499~69500~69501~69527~69528~69545 | 3045 |  | 3045 3045 3045 3045 3045 3045 | 2016-02-15 | 11:57 | Silvia Canessa |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0007 | 0.00157 | 69499 |  | XCAMS | EA | 0% |  | * good measurement: - problem with extractor resolved beforehand - high beam currents achieved by pushing source - stable ratios - no error adjustments | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-03-07 | Water |  | JOB-203944 |  | memory effect test prepared on the 3/2/16 | 60606 | Normal | 14047/11 |  |  | 2016-02-03 | 2313 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00159 | 69500 | [26 empty] | 2016-02-04 | 2315 | Water CO2 Evolution | Completed
[168 empty] | 0.99997 | 0.00159 | 69501 | [26 empty] | 2016-02-04 | 1388 | Gas Splitting | Completed
[168 empty] | 0.99838 | 0.00156 | 69527 | [26 empty] | 2016-02-04 | 2690 | Choose Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00155 | 69528 | [26 empty] | 2016-02-05 | 9607 | Graphite | Completed
[168 empty] | 0.99997 | 0.00159 | 69545 | [26 empty] | 2016-02-10 | 4178 | C13 Measurement | Completed
[197 empty] | 2016-02-15 | 11308 | AMS Submission Results | Completed
[197 empty] | 2016-02-18 | 13472 | Close Job | Completed
[7 empty] | 159.3 | 90.142 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51377 | 51377 | 51377 | 294 | 293 | 294 |  |  | Margaret Norris | -998.33 | 0.06 ‰ | 2016-02-24 | 2015-08-28 | 12 | 180 | [8 empty] | 2.53 ‰ | 0.20 ‰ | C13 | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.455 | 0.08 | 0.08165 | -998.33 ‰ | -998.33 ‰ | 0.06 | 0.06 | -998.31 | 0.06 ‰ | -998.31 | 0.06 ‰ |  |  | 2016-03-07 | Jocelyn Turnbull | 11:46 | 0.0016686 | 0.0016686 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.2 |  |  | 0.00167 | 6e-05 | 11353 | 2016-03-05 | 13:26 | Albert Zondervan | 203018 |  |  |  |  | 45% | 1 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 1 |  | 0.17 | 0.01% | 36 | 5 | 6 | ... | 0.00164 | 6e-05 |  |  |  |  | 1.0181 | 1.0181 | Yes | Yes | [5 empty] | 0.00164 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00167 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00164 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-19 | Margaret Norris | 16:34 |  | Completed |  |  |  |  | 69640 | 2000000 |  |  | 69476 69629 69630 69631 69688 69700 | 69476~69629~69630~69631~69688~69700 | 3046 |  | 3046 3046 3046 3046 3046 3046 | 2016-03-07 | 9:26 | Jocelyn Turnbull |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99893 | 0.00156 | 69476 |  | XCAMS | EA | 0% |  | * CalAMS mode 1 was used to avoid 0.2% system error. * The full fractionation correction with mode 5 would further reduce run-to-run variability, but worsen inter-cathode variability. * The beam currents were good, but started to drop off early for the 0.5-0.6 mg C cathodes. This is another demonstration that something in the ion source blocks us to achieve high counting precision with 0.5 mg C samples. * The problems with too sharp sputter crater and lack of inter-cathode repeatability came to a head with next wheel TW3047. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-09-25 | Carbonate |  | JOB-203018 |  | Mode 1! Power surge during graphite, furnaces went off momentarily, turned back on and seemed okay | 60646 |  | 14047/1 |  |  | 2016-02-10 | 2319 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99907 | 0.00154 | 69629 | [26 empty] | 2016-02-10 | 1395 | Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00177 | 69630 | [26 empty] | 2016-02-10 | 2698 | Choose Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00185 | 69631 | [26 empty] | 2016-02-12 | 9650 | Graphite | Completed
[168 empty] | 1.0027 | 0.00175 | 69688 | [26 empty] | 2016-02-16 | 4179 | C13 Measurement | Completed
[168 empty] | 0.99866 | 0.00159 | 69700 | [26 empty] | 2016-03-07 | 11353 | AMS Submission Results | Completed
[197 empty] | 2016-03-08 | 13516 | Close Job | Completed
[5 empty] | 4.9073e-05 | 0.0066041 | 155.7 | 91.2478 | 0.0679 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7118 | 7117 | 7118 | 48 | 48 | 52 | 7069.6 | 7166.5 | Margaret Norris | -587.73 | 2.49 ‰ | 2016-02-24 | 2015-08-28 | 8 | 180 | [8 empty] | -8.04 ‰ | 0.20 ‰ | C13 | -8.04‰ | 0.2‰ | C13 | No | -19 | -19.455 | 0.08 | 0.08165 | -587.73 ‰ | -587.73 ‰ | 2.49 | 2.65 | -582.64 | 2.52 ‰ | -582.64 | 2.52 ‰ |  |  | 2016-03-07 | Jocelyn Turnbull | 11:46 | 0.41227 | 0.41227 | 0.00249 | 0.00265 | 0.00243 | 0.00243 | 2.4 |  |  | 0.41707 | 0.0012 | 11382 | 2016-03-05 | 13:26 | Albert Zondervan | 203945 | 0.00779 | 0.0078 | 0.00351 | 0.0035 | 45% | 1 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 1 |  | 41.23 | 0.25% | 34 | 5 | 6 | ... | 0.40559 | 0.00104 |  |  |  |  | 1.0181 | 1.0181 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 51.6931038333477483. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00779 | 0.00636 | 0.0035 | 0.40092 | 0.40‰ | 0.00236 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.41227 | 0 | 0.00249 | 0 | 0 | 0.00779 | 0 | 0.00351 | 0 | 0.40092 | 0 | 0.00236 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-19 | Margaret Norris | 16:34 |  | Completed |  |  |  |  | 69659 | 2000000 | 69662 69597 69454 69274 69144 69113 69032 68963 68799 | 69662~69597~69454~69274~69144~69113~69032~68963~68799 | 69476 69629 69630 69631 69688 69700 | 69476~69629~69630~69631~69688~69700 | 3046 | 3046 3045 3042 3040 3035 3033 3031 3029 3026 | 3046 3046 3046 3046 3046 3046 | 2016-03-07 | 11:37 | Jocelyn Turnbull |  | 0.90 mg | 0.13 | 0.9 | 2016-02-24 | 0.00778 | 69662 | 3046 | 0.91 | 0.99893 | 0.00156 | 69476 | EA | XCAMS | EA | 0% |  | * CalAMS mode 1 was used to avoid 0.2% system error. * The full fractionation correction with mode 5 would further reduce run-to-run variability, but worsen inter-cathode variability. * The beam currents were good, but started to drop off early for the 0.5-0.6 mg C cathodes. This is another demonstration that something in the ion source blocks us to achieve high counting precision with 0.5 mg C samples. * The problems with too sharp sputter crater and lack of inter-cathode repeatability came to a head with next wheel TW3047. | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-03-07 | Water | Complete | JOB-203945 |  | Mode 1! | 60648 | Normal | 14047/12 |  |  | 2016-02-15 | 2343 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-02-11 | 0.01476 | 69597 | 3045 | 0.88 | 0.99907 | 0.00154 | 69629 | ST | [25 empty] | 2016-02-15 | 1409 | Gas Splitting | Completed
[163 empty] | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 | 1.0005 | 0.00177 | 69630 | EA | [25 empty] | 2016-02-15 | 2719 | Choose Gas Splitting | Completed
[163 empty] | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 1.0011 | 0.00185 | 69631 | ST | [25 empty] | 2016-02-16 | 9669 | Graphite | Completed
[163 empty] | 2015-11-10 | 0.00265 | 69144 | 3035 | 0.95 | 1.0027 | 0.00175 | 69688 | ST | [25 empty] | 2016-02-23 | 4200 | C13 Measurement | Completed
[163 empty] | 2015-10-22 | 0.00665 | 69113 | 3033 | 0.86 | 0.99866 | 0.00159 | 69700 | EA | [25 empty] | 2016-03-07 | 11382 | AMS Submission Results | Completed
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 |  |  |  | EA | [25 empty] | 2016-03-08 | 13534 | Close Job | Completed
[163 empty] | 2015-09-30 | 0.00941 | 68963 | 3029 | 0.84 |  |  |  | EA
[163 empty] | 2015-09-17 | 0.00386 | 68799 | 3026 | 0.94 |  |  |  | EA
[7 empty] | 160.2 | 90.1592 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38871 | 38871 | 38871 | 134 | 134 | 134 |  |  | Margaret Norris | -992.08 | 0.13 ‰ | 2016-02-24 | 2015-08-28 | 8 | 180 | [8 empty] | 2.59 ‰ | 0.20 ‰ | C13 | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.455 | 0.08 | 0.08165 | -992.08 ‰ | -992.08 ‰ | 0.13 | 0.13 | -991.99 | 0.13 ‰ | -991.99 | 0.13 ‰ |  |  | 2016-03-07 | Jocelyn Turnbull | 11:46 | 0.0079154 | 0.0079154 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2.3 |  |  | 0.00792 | 0.00013 | 11389 | 2016-03-05 | 13:26 | Albert Zondervan | 204024 |  |  |  |  | 45% | 1 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 1 |  | 0.79 | 0.01% | 37 | 5 | 6 | ... | 0.00778 | 0.00013 |  |  |  |  | 1.0181 | 1.0181 | Yes | Yes |  | another poor result for the water-line blank |  |  |  | 0.00778 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00792 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00778 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-19 | Margaret Norris | 16:34 |  | Completed |  |  |  |  | 69662 | 2000000 |  |  | 69476 69629 69630 69631 69688 69700 | 69476~69629~69630~69631~69688~69700 | 3046 |  | 3046 3046 3046 3046 3046 3046 | 2016-03-07 | 11:37 | Jocelyn Turnbull |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99893 | 0.00156 | 69476 |  | XCAMS | EA | 0% |  | * CalAMS mode 1 was used to avoid 0.2% system error. * The full fractionation correction with mode 5 would further reduce run-to-run variability, but worsen inter-cathode variability. * The beam currents were good, but started to drop off early for the 0.5-0.6 mg C cathodes. This is another demonstration that something in the ion source blocks us to achieve high counting precision with 0.5 mg C samples. * The problems with too sharp sputter crater and lack of inter-cathode repeatability came to a head with next wheel TW3047. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-17 | Water | Complete | JOB-204024 |  | Mode 1! | 60647 | Normal | 14047/11 |  |  | 2016-02-15 | 2344 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99907 | 0.00154 | 69629 | [26 empty] | 2016-02-15 | 1407 | Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00177 | 69630 | [26 empty] | 2016-02-15 | 2712 | Choose Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00185 | 69631 | [26 empty] | 2016-02-16 | 9672 | Graphite | Completed
[168 empty] | 1.0027 | 0.00175 | 69688 | [26 empty] | 2016-02-23 | 4201 | C13 Measurement | Completed
[168 empty] | 0.99866 | 0.00159 | 69700 | [26 empty] | 2016-03-07 | 11389 | AMS Submission Results | Completed
[197 empty] | 2016-03-08 | 13541 | Close Job | Completed
[7 empty] | 100.6 | 89.0645 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39219 | 39218 | 39219 | 198 | 198 | 198 |  |  | jenny dahl | -992.42 | 0.19 ‰ | 2016-03-14 | 2016-03-14 | 17 | 0 |  |  |  |  | 2.42‰ | 2.42‰ | 2.00‰ | 2.00‰ | 2.42 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -992.42 ‰ | -992.42 ‰ | 0.19 | 0.19 | -992.33 | 0.19 ‰ | -992.33 | 0.19 ‰ |  |  | 2016-03-16 | Jocelyn Turnbull | 12:58 | 0.0075801 | 0.0075801 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.4 |  |  | 0.00758 | 0.00019 | 11420 | 2016-03-15 | 14:42 | Albert Zondervan | 204023 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.76 | 0.02% | 30 | 5 | 7 | ... | 0.00729 | 0.00018 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes | [5 empty] | 0.00729 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00758 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00729 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-29 | jenny dahl | 11:11 |  | Completed |  |  |  |  | 69760 | 2000000 |  |  | 69673 69674 69675 69676 69677 69754 | 69673~69674~69675~69676~69677~69754 | 3048 |  | 3048 3048 3048 3048 3048 3048 | 2016-03-15 | 16:38 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 1.0034 | 0.00214 | 69673 |  | XCAMS | ST | 0% |  | * 1st measurement after 2-week investigation into poor repeatability * good measurement * no system error * factor 1.09 multiplier for global run-to-run variability, which we think is due to long storage time of graphite | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-17 | Water |  | JOB-204023 |  | bad result although no obvious red flags | 60726 | Normal | 14047/11 |  |  | 2016-02-19 | 2349 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99701 | 0.00222 | 69674 | [26 empty] | 2016-02-23 | 2357 | Water CO2 Evolution | Completed
[168 empty] | 0.99923 | 0.00215 | 69675 | [26 empty] | 2016-02-23 | 1420 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00215 | 69676 | [26 empty] | 2016-02-23 | 2731 | Choose Gas Splitting | Completed
[168 empty] | 0.99875 | 0.00221 | 69677 | [26 empty] | 2016-02-26 | 9770 | Graphite | Completed
[168 empty] | 1.0004 | 0.00212 | 69754 | [26 empty] | 2016-03-03 | 4231 | C13 Measurement | Completed
[197 empty] | 2016-03-16 | 11420 | AMS Submission Results | Completed
[197 empty] | 2016-03-31 | 13742 | Close Job | Completed
[5 empty] | 7.4209e-05 | 0.0054066 | 112.5 | 89.7956 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7090 | 7090 | 7090 | 46 | 46 | 50 | 7044.3 | 7136.7 | jenny dahl | -586.32 | 2.38 ‰ | 2016-03-14 | 2015-09-16 | 17 | 180 |  |  |  |  | -5.74‰ | -5.74‰ | 2.00‰ | 2.00‰ | -5.74 ‰ | 2.00 ‰ | AMS | -8.08‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -586.32 ‰ | -586.32 ‰ | 2.38 | 2.56 | -581.21 | 2.41 ‰ | -581.21 | 2.41 ‰ |  |  | 2016-03-16 | Jocelyn Turnbull | 12:58 | 0.41368 | 0.41368 | 0.00238 | 0.00256 | 0.00232 | 0.00232 | 2.2 |  |  | 0.41775 | 0.0015 | 11421 | 2016-03-15 | 14:42 | Albert Zondervan | 204025 | 0.0065 | 0.0065 | 0.00293 | 0.0029 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.37 | 0.24% | 29 | 5 | 7 | ... | 0.40176 | 0.00134 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes |  | in agreement with consensus value within 1-sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 49.634225989925439. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0065 | 0.0052 | 0.00282 | 0.39785 | 0.40‰ | 0.00223 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.41368 | 0 | 0.00238 | 0 | 0 | 0.0065 | 0 | 0.00293 | 0 | 0.39785 | 0 | 0.00223 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-29 | jenny dahl | 11:11 |  | Completed |  |  |  |  | 69761 | 2000000 | 69760 69144 69113 68963 | 69760~69144~69113~68963 | 69673 69674 69675 69676 69677 69754 | 69673~69674~69675~69676~69677~69754 | 3048 | 3048 3035 3033 3029 | 3048 3048 3048 3048 3048 3048 | 2016-03-15 | 16:38 | Albert Zondervan |  | 0.90 mg | 0.13 | 0.9 | 2016-03-14 | 0.00729 | 69760 | 3048 | 0.85 | 1.0034 | 0.00214 | 69673 | ST | XCAMS | ST | 0% |  | * 1st measurement after 2-week investigation into poor repeatability * good measurement * no system error * factor 1.09 multiplier for global run-to-run variability, which we think is due to long storage time of graphite | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-03-17 | Water |  | JOB-204025 |  | Water line std | 60727 | Normal | 14047/12 |  |  | 2016-02-19 | 2350 | Carbonate CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2015-11-10 | 0.00265 | 69144 | 3035 | 0.95 | 0.99701 | 0.00222 | 69674 | ST | [25 empty] | 2016-02-24 | 2363 | Water CO2 Evolution | Completed
[163 empty] | 2015-10-22 | 0.00665 | 69113 | 3033 | 0.86 | 0.99923 | 0.00215 | 69675 | ST | [25 empty] | 2016-02-24 | 1421 | Gas Splitting | Completed
[163 empty] | 2015-09-30 | 0.00941 | 68963 | 3029 | 0.84 | 1.0009 | 0.00215 | 69676 | ST | [25 empty] | 2016-02-24 | 2733 | Choose Gas Splitting | Completed
[168 empty] | 0.99875 | 0.00221 | 69677 | [26 empty] | 2016-02-26 | 9771 | Graphite | Completed
[168 empty] | 1.0004 | 0.00212 | 69754 | [26 empty] | 2016-03-03 | 4230 | C13 Measurement | Completed
[197 empty] | 2016-03-16 | 11421 | AMS Submission Results | Completed
[197 empty] | 2016-03-16 | 13629 | Close Job | Completed
[5 empty] | 7.337e-05 | 0.0089907 | 135.5 | 89.8345 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7146 | 7146 | 7146 | 58 | 58 | 61 | 7088.1 | 7205 | jenny dahl | -589.19 | 2.99 ‰ | 2016-04-04 | 2015-10-07 | 13 | 180 |  |  |  |  | -6.43‰ | -6.43‰ | 2.00‰ | 2.00‰ | -6.43 ‰ | 2.00 ‰ | AMS | -8.09‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -589.19 ‰ | -589.19 ‰ | 2.99 | 3.13 | -584.12 | 3.03 ‰ | -584.12 | 3.03 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.41081 | 0.41081 | 0.00299 | 0.00313 | 0.00294 | 0.00294 | 1.5 |  |  | 0.41679 | 0.00124 | 11641 | 2016-04-05 | 15:51 | Albert Zondervan | 204167 | 0.0095 | 0.0095 | 0.00428 | 0.0043 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.08 | 0.3% | 32 | 5 | 6 | ... | 0.40084 | 0.00123 |  |  |  |  | 1.003 | 1.003 | Yes | Yes |  | result is spot on wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 61.1860161577472569. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0095 | 0.00871 | 0.00307 | 0.39509 | 0.40‰ | 0.00289 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.41081 | 0 | 0.00299 | 0 | 0 | 0.0095 | 0 | 0.00428 | 0 | 0.39509 | 0 | 0.00289 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69941 | 2000000 | 69662 69597 69454 69274 69032 | 69662~69597~69454~69274~69032 | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 | 3046 3045 3042 3040 3031 | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 1.22 mg | 0.13 | 1.22 | 2016-02-24 | 0.00778 | 69662 | 3046 | 0.91 | 1.0019 | 0.00219 | 69784 | EA | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-04-15 | Water |  | JOB-204167 |  |  | 60948 | Normal | 14047/12 |  |  | 2016-03-21 | 2415 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-02-11 | 0.01476 | 69597 | 3045 | 0.88 | 0.99913 | 0.0023 | 69812 | EA | [25 empty] | 2016-03-21 | 1475 | Gas Splitting | Completed
[163 empty] | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 | 1.0027 | 0.00228 | 69813 | EA | [25 empty] | 2016-03-21 | 2809 | Choose Gas Splitting | Completed
[163 empty] | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 0.99729 | 0.00233 | 69814 | EA | [25 empty] | 2016-03-22 | 9951 | Graphite | Completed
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 | 0.99855 | 0.00225 | 69928 | EA | [25 empty] | 2016-03-23 | 4333 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11641 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13845 | Close Job | Completed
[7 empty] | 131.2 | 88.9136 | 0.0022 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34811 | 34810 | 34811 | 133 | 133 | 133 | 34678.6 | 34944.8 | jenny dahl | -986.88 | 0.22 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 3.86‰ | 3.86‰ | 2.00‰ | 2.00‰ | 3.86 ‰ | 2.00 ‰ | AMS | 2.45‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.88 ‰ | -986.88 ‰ | 0.22 | 0.22 | -986.72 | 0.22 ‰ | -986.72 | 0.22 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013122 | 0.013122 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.3 |  |  | 0.01312 | 0.00022 | 11644 | 2016-04-05 | 15:51 | Albert Zondervan | 204312 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.31 | 0.02% | 8 | 5 | 7 | ... | 0.01262 | 0.00021 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01262 | 0.01‰ | 0.00021 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01312 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01262 | 0 | 0.00021 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69929 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.93 mg |  | 0.93 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.5 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  | JOB-204312 |  | new batch of standard prepared. test water blank | 60943 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [26 empty] | 2016-03-17 | 2404 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [26 empty] | 2016-03-17 | 1470 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [26 empty] | 2016-03-17 | 2799 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [26 empty] | 2016-03-18 | 9939 | Graphite | Completed
[197 empty] | 2016-03-22 | 4317 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11644 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13888 | Close Job | Completed
[7 empty] | 131.7 | 88.8554 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41645 | 41644 | 41645 | 193 | 193 | 193 | 41453.5 | 41840.3 | jenny dahl | -994.4 | 0.13 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.25‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -994.40 ‰ | -994.40 ‰ | 0.13 | 0.13 | -994.33 | 0.14 ‰ | -994.33 | 0.14 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.0056045 | 0.0056045 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2 |  |  | 0.0056 | 0.00013 | 11645 | 2016-04-05 | 15:51 | Albert Zondervan | 204314 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 22 | 5 | 7 | ... | 0.00539 | 0.00013 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.00539 | 0.01‰ | 0.00013 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.0056 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00539 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69925 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 1.00 mg |  | 1 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.3 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  | JOB-204314 |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel final weight incorrectly entered- back corrected to choose correct split. graphite yield will look odd. | 60944 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [26 empty] | 2016-03-17 | 2405 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [26 empty] | 2016-03-17 | 1469 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [26 empty] | 2016-03-17 | 2797 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [26 empty] | 2016-03-18 | 9935 | Graphite | Completed
[197 empty] | 2016-03-22 | 4318 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11645 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13889 | Close Job | Completed
[7 empty] | 126 | 89.0883 | 0.0027 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33232 | 33232 | 33232 | 114 | 114 | 114 | 33118.6 | 33347.5 | jenny dahl | -984.03 | 0.23 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 1.89‰ | 1.89‰ | 2.00‰ | 2.00‰ | 1.89 ‰ | 2.00 ‰ | AMS | 1.78‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -984.03 ‰ | -984.03 ‰ | 0.23 | 0.23 | -983.83 | 0.23 ‰ | -983.83 | 0.23 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.015971 | 0.015971 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 2.5 |  |  | 0.01597 | 0.00023 | 11646 | 2016-04-05 | 15:51 | Albert Zondervan | 204315 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.6 | 0.02% | 29 | 5 | 7 | ... | 0.01536 | 0.00022 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01536 | 0.02‰ | 0.00022 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01597 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01536 | 0 | 0.00022 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69926 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.90 mg |  | 0.9 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 1.8 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  | JOB-204315 |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60945 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [26 empty] | 2016-03-17 | 2406 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [26 empty] | 2016-03-17 | 1468 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [26 empty] | 2016-03-17 | 2795 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [26 empty] | 2016-03-18 | 9936 | Graphite | Completed
[197 empty] | 2016-03-22 | 4319 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11646 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13890 | Close Job | Completed
[7 empty] | 123.8 | 88.9962 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34311 | 34310 | 34311 | 131 | 131 | 131 | 34180.9 | 34443 | jenny dahl | -986.04 | 0.23 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.04 ‰ | -986.04 ‰ | 0.23 | 0.23 | -985.86 | 0.23 ‰ | -985.86 | 0.23 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013964 | 0.013964 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 2.6 |  |  | 0.01396 | 0.00023 | 11647 | 2016-04-05 | 15:51 | Albert Zondervan | 204316 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.4 | 0.02% | 15 | 5 | 7 | ... | 0.01343 | 0.00022 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01343 | 0.01‰ | 0.00022 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01396 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01343 | 0 | 0.00022 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69927 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.5 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  | JOB-204316 |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60946 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [26 empty] | 2016-03-17 | 2407 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [26 empty] | 2016-03-17 | 1467 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [26 empty] | 2016-03-17 | 2793 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [26 empty] | 2016-03-18 | 9937 | Graphite | Completed
[197 empty] | 2016-03-22 | 4320 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11647 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13891 | Close Job | Completed
[7 empty] | 132 | 88.9452 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34425 | 34425 | 34425 | 139 | 139 | 139 | 34287.5 | 34565.6 | jenny dahl | -986.23 | 0.24 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 3.50‰ | 3.50‰ | 2.00‰ | 2.00‰ | 3.50 ‰ | 2.00 ‰ | AMS | 2.17‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.23 ‰ | -986.23 ‰ | 0.24 | 0.24 | -986.06 | 0.24 ‰ | -986.06 | 0.24 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013767 | 0.013767 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.4 |  |  | 0.01377 | 0.00024 | 11648 | 2016-04-05 | 15:51 | Albert Zondervan | 204317 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.38 | 0.02% | 1 | 5 | 7 | ... | 0.01324 | 0.00023 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01324 | 0.01‰ | 0.00023 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01377 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01324 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69930 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.2 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water | Complete | JOB-204317 |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60947 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [26 empty] | 2016-03-17 | 2408 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [26 empty] | 2016-03-17 | 1471 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [26 empty] | 2016-03-17 | 2801 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [26 empty] | 2016-03-18 | 9940 | Graphite | Completed
[197 empty] | 2016-03-22 | 4321 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11648 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13892 | Close Job | Completed
[7 empty] | 135.2 | 88.8394 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 175 | 175 | 175 |  |  | jenny dahl | -994.28 | 0.12 ‰ | 2016-04-06 | 2016-04-06 | 14 | 0 |  |  |  |  | 5.01‰ | 5.01‰ | 2.00‰ | 2.00‰ | 5.01 ‰ | 2.00 ‰ | AMS | 2.70‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.89443 | -994.28 ‰ | -994.28 ‰ | 0.12 | 0.12 | -994.21 | 0.13 ‰ | -994.21 | 0.13 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:56 | 0.0057189 | 0.0057189 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00572 | 0.00012 | 11686 | 2016-04-07 | 13:39 | Albert Zondervan | 204238 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 4 | 5 | 6 | ... | 0.0055 | 0.00012 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00572 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-04 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 69954 | 2000000 |  |  | 69821 69916 69920 69945 69951 | 69821~69916~69920~69945~69951 | 3056 |  | 3056 3056 3056 3056 3056 | 2016-04-07 | 14:14 | Albert Zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99926 | 0.00157 | 69821 |  | XCAMS | EA | 0% |  | * good meas't, no surprises * no error adjustments * small Ox-I (TP69822) 3-sigma lower than large ones, so removed from list of calibration standards | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2016-04-26 | Water |  | JOB-204238 |  |  | 60981 | Normal | 14047/11 |  |  | 2016-03-22 | 2424 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00153 | 69916 | [26 empty] | 2016-03-22 | 1476 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00201 | 69920 | [26 empty] | 2016-03-22 | 2811 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00154 | 69945 | [26 empty] | 2016-03-23 | 9964 | Graphite | Completed
[168 empty] | 1.0001 | 0.0015 | 69951 | [26 empty] | 2016-03-24 | 4335 | C13 Measurement | Completed
[197 empty] | 2016-04-11 | 11686 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13847 | Close Job | Completed
[5 empty] | 6.3246e-05 | 0.00815 | 128.4 | 90.0096 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7190 | 7190 | 7190 | 55 | 54 | 58 | 7135.6 | 7245.1 | jenny dahl | -591.42 | 2.79 ‰ | 2016-04-06 | 2015-10-09 | 6 | 180 |  |  |  |  | -8.06‰ | -8.06‰ | 2.00‰ | 2.00‰ | -8.06 ‰ | 2.00 ‰ | AMS | -8.14‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.89443 | -591.42 ‰ | -591.42 ‰ | 2.79 | 2.93 | -586.38 | 2.82 ‰ | -586.38 | 2.82 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:56 | 0.40858 | 0.40858 | 0.00279 | 0.00293 | 0.00273 | 0.00273 | 2.1 |  |  | 0.41416 | 0.00116 | 11687 | 2016-04-07 | 13:39 | Albert Zondervan | 204239 | 0.00884 | 0.0088 | 0.00398 | 0.004 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 40.86 | 0.28% | 11 | 5 | 6 | ... | 0.39831 | 0.00099 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes |  | agreement with expectation value within 1-sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.6800950603804393. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00884 | 0.00788 | 0.00319 | 0.39294 | 0.39‰ | 0.00263 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.40858 | 0 | 0.00279 | 0 | 0 | 0.00884 | 0 | 0.00398 | 0 | 0.39294 | 0 | 0.00263 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-04 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 69976 | 2000000 | 69954 69662 69597 69454 69274 69032 | 69954~69662~69597~69454~69274~69032 | 69821 69916 69920 69945 69951 | 69821~69916~69920~69945~69951 | 3056 | 3056 3046 3045 3042 3040 3031 | 3056 3056 3056 3056 3056 | 2016-04-07 | 14:14 | Albert Zondervan |  | 0.88 mg | 0.13 | 0.88 | 2016-04-06 | 0.0055 | 69954 | 3056 | 1.09 | 0.99926 | 0.00157 | 69821 | EA | XCAMS | EA | 0% |  | * good meas't, no surprises * no error adjustments * small Ox-I (TP69822) 3-sigma lower than large ones, so removed from list of calibration standards | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-04-26 | Water |  | JOB-204239 |  |  | 60982 | Normal | 14047/12 |  |  | 2016-03-30 | 2438 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-02-24 | 0.00778 | 69662 | 3046 | 0.91 | 0.9993 | 0.00153 | 69916 | EA | [25 empty] | 2016-03-30 | 1479 | Gas Splitting | Completed
[163 empty] | 2016-02-11 | 0.01476 | 69597 | 3045 | 0.88 | 1.0013 | 0.00201 | 69920 | EA | [25 empty] | 2016-03-30 | 2815 | Choose Gas Splitting | Completed
[163 empty] | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 | 1.0006 | 0.00154 | 69945 | EA | [25 empty] | 2016-03-31 | 9986 | Graphite | Completed
[163 empty] | 2015-12-09 | 0.00931 | 69274 | 3040 | 0.88 | 1.0001 | 0.0015 | 69951 | EA | [25 empty] | 2016-04-04 | 4351 | C13 Measurement | Completed
[163 empty] | 2015-10-16 | 0.007 | 69032 | 3031 | 0.87 |  |  |  | EA | [25 empty] | 2016-04-11 | 11687 | AMS Submission Results | Completed
[197 empty] | 2016-04-11 | 13848 | Close Job | Completed
[7 empty] | 102.2 | 88.8807 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50961 | 50961 | 50961 | 237 | 237 | 237 |  |  | jenny dahl | -998.24 | 0.05 ‰ | 2016-04-11 | 2016-04-11 | 34 | 0 |  |  |  |  | 6.74‰ | 6.74‰ | 2.00‰ | 2.00‰ | 6.74 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0041 | 0 | 0.8165 | -998.24 ‰ | -998.24 ‰ | 0.05 | 0.05 | -998.22 | 0.05 ‰ | -998.22 | 0.05 ‰ |  |  | 2016-04-14 | Jocelyn Turnbull | 14:04 | 0.0017573 | 0.0017573 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2 |  |  | 0.00176 | 5e-05 | 11714 | 2016-04-14 | 12:48 | Albert Zondervan | 204225 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.18 | 0.01% | 26 | 5 | 6 | ... | 0.00169 | 5e-05 |  |  |  |  | 1.0055 | 1.0055 | Yes | Yes | [5 empty] | 0.00169 | 0.00‰ | 5e-05 | 0.00‰ | No | 1 | 1.0001 | 0.00102 | 0.00102 | 0.00176 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.00169 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-08 | jenny dahl | 8:44 |  | Completed |  |  |  |  | 69833 | 2000000 |  |  | 69905 69906 69935 69936 70000 70031 | 69905~69906~69935~69936~70000~70031 | 3057 |  | 3057 3057 3057 3057 3057 3057 | 2016-04-14 | 13:10 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99798 | 0.00252 | 69905 |  | XCAMS | ST | 0.18% |  | * good end result, after initial troubles with tuning * non-zero but acceptable system error | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-14 | Carbonate | Complete | JOB-204225 |  |  | 61021 |  | 14047/1 |  |  | 2016-03-04 | 2379 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99895 | 0.0026 | 69906 | [26 empty] | 2016-03-04 | 1422 | Gas Splitting | Completed
[168 empty] | 0.99754 | 0.00253 | 69935 | [26 empty] | 2016-03-04 | 2735 | Choose Gas Splitting | Completed
[168 empty] | 0.99965 | 0.00225 | 69936 | [26 empty] | 2016-03-08 | 9843 | Graphite | Completed
[168 empty] | 1.0021 | 0.0025 | 70000 | [26 empty] | 2016-03-11 | 4255 | C13 Measurement | Completed
[168 empty] | 1.0043 | 0.00261 | 70031 | [26 empty] | 2016-04-14 | 11714 | AMS Submission Results | Completed
[197 empty] | 2016-04-15 | 13915 | Close Job | Completed
[7 empty] | 89.9 | 88.822 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40987 | 40986 | 40987 | 205 | 205 | 205 |  |  | jenny dahl | -993.92 | 0.16 ‰ | 2016-04-11 | 2016-04-11 | 4 | 0 |  |  |  |  | 7.40‰ | 7.40‰ | 2.00‰ | 2.00‰ | 7.40 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0041 | 0 | 0.8165 | -993.92 ‰ | -993.92 ‰ | 0.16 | 0.16 | -993.84 | 0.16 ‰ | -993.84 | 0.16 ‰ |  |  | 2016-04-14 | Jocelyn Turnbull | 14:04 | 0.0060828 | 0.0060828 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.4 |  |  | 0.00608 | 0.00016 | 11715 | 2016-04-14 | 12:48 | Albert Zondervan | 204230 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.61 | 0.02% | 1 | 5 | 6 | ... | 0.00585 | 0.00015 |  |  |  |  | 1.0055 | 1.0055 | Yes | Yes | [5 empty] | 0.00585 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1.0001 | 0.00102 | 0.00102 | 0.00608 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00585 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-08 | jenny dahl | 8:44 |  | Completed |  |  |  |  | 70029 | 2000000 |  |  | 69905 69906 69935 69936 70000 70031 | 69905~69906~69935~69936~70000~70031 | 3057 |  | 3057 3057 3057 3057 3057 3057 | 2016-04-14 | 13:10 | Albert Zondervan |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99798 | 0.00252 | 69905 |  | XCAMS | ST | 0.18% |  | * good end result, after initial troubles with tuning * non-zero but acceptable system error |  | RRL | [5 empty] | Background Inorganic |  | 2016-04-25 | Water | Not Complete | JOB-204230 |  |  | 61022 | Normal | 14047/11 |  |  | 2016-04-04 | 2452 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99895 | 0.0026 | 69906 | [6 empty] | 2.5 | [19 empty] | 2016-04-04 | 1484 | Gas Splitting | Completed
[168 empty] | 0.99754 | 0.00253 | 69935 | [26 empty] | 2016-04-04 | 2822 | Choose Gas Splitting | Completed
[168 empty] | 0.99965 | 0.00225 | 69936 | [26 empty] | 2016-04-05 | 4362 | C13 Measurement | Cancelled
[168 empty] | 1.0021 | 0.0025 | 70000 | [26 empty] | 2016-04-07 | 10039 | Graphite | Completed
[168 empty] | 1.0043 | 0.00261 | 70031 | [26 empty] | 2016-04-14 | 11715 | AMS Submission Results | Completed
[197 empty] | 2016-04-13 | 4378 | C13 Measurement | Completed
[197 empty] | 2016-04-15 | 13916 | Close Job | Completed
[5 empty] | 7.2377e-05 | 0.004819 | 86.8 | 90.4805 | 0.0686 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7120 | 7120 | 7120 | 41 | 41 | 45 | 7079.3 | 7161.6 | jenny dahl | -587.86 | 2.11 ‰ | 2016-04-11 | 2015-10-14 | 4 | 180 |  |  |  |  | -11.06‰ | -11.06‰ | 2.00‰ | 2.00‰ | -11.06 ‰ | 2.00 ‰ | AMS | -8.05‰ | 0.2‰ | C13 | No | -19 | -19.0041 | 0 | 0.8165 | -587.86 ‰ | -587.86 ‰ | 2.11 | 2.31 | -582.77 | 2.14 ‰ | -582.77 | 2.14 ‰ |  |  | 2016-04-14 | Jocelyn Turnbull | 14:04 | 0.41214 | 0.41214 | 0.00211 | 0.00231 | 0.00204 | 0.00204 | 2 |  |  | 0.41567 | 0.00137 | 11716 | 2016-04-14 | 12:48 | Albert Zondervan | 204231 | 0.00561 | 0.0056 | 0.00252 | 0.0025 | 45% | 5 | 5/07/22 | 14:06:06 | Jocelyn Turnbull | 0 |  | 41.21 | 0.21% | 6 | 5 | 6 | ... | 0.39976 | 0.00141 |  |  |  |  | 1.0055 | 1.0055 | Yes | Yes |  | in agreement with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.9582251369313337. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00561 | 0.00463 | 0.00206 | 0.39637 | 0.40‰ | 0.00209 | 0.00‰ | No | 1 | 1.0001 | 0.00102 | 0.00102 | 0.41214 | 0 | 0.00211 | 0 | 0 | 0.00561 | 0 | 0.00252 | 0 | 0.39637 | 0 | 0.00209 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-08 | jenny dahl | 8:44 |  | Completed |  |  |  |  | 70030 | 2000000 | 70029 69760 69144 69113 | 70029~69760~69144~69113 | 69905 69906 69935 69936 70000 70031 | 69905~69906~69935~69936~70000~70031 | 3057 | 3057 3048 3035 3033 | 3057 3057 3057 3057 3057 3057 | 2016-04-14 | 13:10 | Albert Zondervan |  | 0.98 mg | 0.13 | 0.98 | 2016-04-11 | 0.00585 | 70029 | 3057 | 0.81 | 0.99798 | 0.00252 | 69905 | ST | XCAMS | ST | 0.18% |  | * good end result, after initial troubles with tuning * non-zero but acceptable system error | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-04-25 | Water |  | JOB-204231 |  |  | 61023 | Normal | 14047/12 |  |  | 2016-04-04 | 2453 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-03-14 | 0.00729 | 69760 | 3048 | 0.85 | 0.99895 | 0.0026 | 69906 | ST | [25 empty] | 2016-04-04 | 1482 | Gas Splitting | Completed
[163 empty] | 2015-11-10 | 0.00265 | 69144 | 3035 | 0.95 | 0.99754 | 0.00253 | 69935 | ST | [25 empty] | 2016-04-04 | 2818 | Choose Gas Splitting | Completed
[163 empty] | 2015-10-22 | 0.00665 | 69113 | 3033 | 0.86 | 0.99965 | 0.00225 | 69936 | ST | [25 empty] | 2016-04-07 | 10040 | Graphite | Completed
[168 empty] | 1.0021 | 0.0025 | 70000 | [26 empty] | 2016-04-13 | 4367 | C13 Measurement | Completed
[168 empty] | 1.0043 | 0.00261 | 70031 | [26 empty] | 2016-04-14 | 11716 | AMS Submission Results | Completed
[197 empty] | 2016-04-15 | 13917 | Close Job | Completed
[7 empty] | 117.4 | 89.0196 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47946 | 47945 | 47946 | 294 | 293 | 294 |  |  | jenny dahl | -997.44 | 0.09 ‰ | 2016-04-13 | 2015-10-16 | 22 | 180 |  |  |  |  | 3.27‰ | 3.27‰ | 2.00‰ | 2.00‰ | 3.27 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -997.44 ‰ | -997.44 ‰ | 0.09 | 0.09 | -997.41 | 0.09 ‰ | -997.41 | 0.09 ‰ |  |  | 2016-04-15 | Jocelyn Turnbull | 14:36 | 0.0025579 | 0.0025579 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.1 |  |  | 0.00256 | 9e-05 | 11788 | 2016-04-14 | 14:16 | Albert Zondervan | 204243 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.26 | 0.01% | 1 | 5 | 4 | ... | 0.00246 | 9e-05 |  |  |  |  | 1.0036 | 1.0036 | Yes | Yes | [5 empty] | 0.00246 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00256 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00246 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-12 | jenny dahl | 15:49 |  | Completed |  |  |  |  | 69931 | 2000000 |  |  | 70033 70039 70040 70044 70072 70074 | 70033~70039~70040~70044~70072~70074 | 3059 |  | 3059 3059 3059 3059 3059 3059 | 2016-04-14 | 15:01 | Albert Zondervan |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0008 | 0.0018 | 70033 |  | XCAMS | EA | 0% |  | * good measurement; no tune needed * no error adjustments | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-15 | Carbonate |  | JOB-204243 |  |  | 61062 |  | 14047/1 |  |  | 2016-03-18 | 2414 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99921 | 0.00178 | 70039 | [26 empty] | 2016-03-22 | 9941 | Graphite | Completed
[168 empty] | 0.99827 | 0.00182 | 70040 | [26 empty] | 2016-03-23 | 4332 | C13 Measurement | Completed
[168 empty] | 0.99962 | 0.00185 | 70044 | [26 empty] | 2016-04-15 | 11788 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00191 | 70072 | [26 empty] | 2016-04-21 | 14007 | Close Job | Completed
[168 empty] | 1.0011 | 0.00199 | 70074
[7 empty] | 118.8 | 89.4771 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44368 | 44368 | 44368 | 251 | 251 | 251 |  |  | Margaret Norris | -996.01 | 0.12 ‰ | 2016-05-12 | 2016-05-12 | 42 | 0 |  |  |  |  | 3.67‰ | 3.67‰ | 2.00‰ | 2.00‰ | 3.67 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -996.01 ‰ | -996.01 ‰ | 0.12 | 0.12 | -995.96 | 0.13 ‰ | -995.96 | 0.13 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0039928 | 0.0039928 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.7 |  |  | 0.00399 | 0.00012 | 11809 | 2016-05-20 | 14:21 | Albert Zondervan | 203019 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 8 | 5 | 4 | ..T | 0.00384 | 0.00012 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00384 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00399 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00384 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:11 |  | Completed |  |  |  |  | 69975 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-09-25 | Carbonate | Complete | JOB-203019 | standard, no charge | XCAMS test | 62085 |  | 14047/1 |  |  | 2016-03-29 | 2427 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-03-31 | 9985 | Graphite | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-04-04 | 4347 | C13 Measurement | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-11-02 | 11809 | AMS Submission Results | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-11-02 | 15127 | Close Job | Completed
[168 empty] | 0.99708 | 0.00202 | 70176
[168 empty] | 0.99827 | 0.00189 | 70178
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 106.7 | 89.4952 | 0.0018 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36568 | 36568 | 36568 | 151 | 150 | 151 | 36419 | 36720.1 | Margaret Norris | -989.46 | 0.20 ‰ | 2016-05-12 | 2016-05-12 | 6 | 0 |  |  |  |  | 3.47‰ | 3.47‰ | 2.00‰ | 2.00‰ | 3.47 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0039 | 0 | 0.5 | -989.46 ‰ | -989.46 ‰ | 0.2 | 0.2 | -989.33 | 0.20 ‰ | -989.33 | 0.20 ‰ |  |  | 2016-10-12 | Margaret Norris | 10:47 | 0.010544 | 0.010544 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.5 |  |  | 0.01054 | 0.0002 | 11811 | 2016-05-20 | 14:21 | Albert Zondervan | 204423 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 22 | 5 | 6 | ..T | 0.01014 | 0.00019 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.01014 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.01054 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01014 | 0 | 0.00019 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70214 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-03 | Water | Complete | JOB-204423 | Charged in timesheet June 2016 Jeremy 2 hours to GCT. | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61888 | Normal | 14047/11 |  |  | 2016-05-05 | 2492 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-05-05 | 2498 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-05-05 | 1513 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-05-05 | 2873 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-05-06 | 10224 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [26 empty] | 2016-10-12 | 11811 | AMS Submission Results | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [26 empty] | 2016-10-12 | 14975 | Close Job | Completed
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 100.5 | 89.1645 | 0.0011 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40315 | 40315 | 40315 | 202 | 202 | 202 | 40115.7 | 40519.9 | Margaret Norris | -993.39 | 0.17 ‰ | 2016-05-12 | 2016-05-12 | 6 | 0 |  |  |  |  | 7.19‰ | 7.19‰ | 2.00‰ | 2.00‰ | 7.19 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0039 | 0 | 0.5 | -993.39 ‰ | -993.39 ‰ | 0.17 | 0.17 | -993.31 | 0.17 ‰ | -993.31 | 0.17 ‰ |  |  | 2016-10-12 | Jocelyn Turnbull | 11:10 | 0.0066131 | 0.0066131 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.8 |  |  | 0.00661 | 0.00017 | 11812 | 2016-05-20 | 14:21 | Albert Zondervan | 204424 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.66 | 0.02% | 6 | 5 | 6 | ..T | 0.00636 | 0.00016 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00636 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00661 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00636 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70216 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 1.07 mg |  | 1.07 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-03 | Water | Complete | JOB-204424 | Charged in timesheet June 2016 Jeremy 2 hours to GCT. | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61889 | Normal | 14047/11 |  |  | 2016-05-05 | 2493 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-05-05 | 2499 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-05-05 | 1515 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-05-05 | 2877 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-05-06 | 10226 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [26 empty] | 2016-10-12 | 11812 | AMS Submission Results | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [26 empty] | 2016-10-12 | 14976 | Close Job | Completed
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 110.7 | 89.244 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50106 | 50105 | 50106 | 385 | 384 | 385 |  |  | Margaret Norris | -998.05 | 0.09 ‰ | 2016-05-12 | 2016-05-12 | 22 | 0 |  |  |  |  | 6.29‰ | 6.29‰ | 2.00‰ | 2.00‰ | 6.29 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -998.05 ‰ | -998.05 ‰ | 0.09 | 0.09 | -998.02 | 0.09 ‰ | -998.02 | 0.09 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0019548 | 0.0019548 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00195 | 9e-05 | 11813 | 2016-05-20 | 14:21 | Albert Zondervan | 204426 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 13 | 5 | 4 | ..T | 0.00188 | 9e-05 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00188 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00195 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00188 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70111 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-04-22 | Carbonate |  | JOB-204426 | XCAMS test - no charge | XCAMS test | 62081 |  | 14047/1 |  |  | 2016-04-18 | 2479 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-04-20 | 10121 | Graphite | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-04-29 | 4394 | C13 Measurement | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-11-02 | 11813 | AMS Submission Results | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-11-02 | 15124 | Close Job | Completed
[168 empty] | 0.99708 | 0.00202 | 70176
[168 empty] | 0.99827 | 0.00189 | 70178
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 92.5 | 89.0489 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48382 | 48381 | 48382 | 310 | 310 | 310 |  |  | Margaret Norris | -997.58 | 0.09 ‰ | 2016-05-12 | 2016-05-12 | 22 | 0 |  |  |  |  | 8.50‰ | 8.50‰ | 2.00‰ | 2.00‰ | 8.50 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -997.58 ‰ | -997.58 ‰ | 0.09 | 0.09 | -997.55 | 0.09 ‰ | -997.55 | 0.09 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0024227 | 0.0024227 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.7 |  |  | 0.00242 | 9e-05 | 11814 | 2016-05-20 | 14:21 | Albert Zondervan | 204427 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.24 | 0.01% | 4 | 5 | 4 | ..T | 0.00233 | 9e-05 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00233 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00242 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00233 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70112 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-04-22 | Carbonate |  | JOB-204427 | XCAMS test - no charge | XCAMS test | 62082 |  | 14047/1 |  |  | 2016-04-18 | 2480 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-04-20 | 10122 | Graphite | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-04-29 | 4398 | C13 Measurement | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-11-02 | 11814 | AMS Submission Results | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-11-02 | 15126 | Close Job | Completed
[168 empty] | 0.99708 | 0.00202 | 70176
[168 empty] | 0.99827 | 0.00189 | 70178
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 113.6 | 89.6097 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43385 | 43385 | 43385 | 241 | 240 | 241 |  |  | Margaret Norris | -995.49 | 0.14 ‰ | 2016-05-12 | 2016-05-12 | 21 | 0 |  |  |  |  | 2.18‰ | 2.18‰ | 2.00‰ | 2.00‰ | 2.18 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -995.49 ‰ | -995.49 ‰ | 0.14 | 0.14 | -995.43 | 0.14 ‰ | -995.43 | 0.14 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0045127 | 0.0045127 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00451 | 0.00014 | 11815 | 2016-05-20 | 14:21 | Albert Zondervan | 204428 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.45 | 0.01% | 16 | 5 | 4 | ..T | 0.00434 | 0.00013 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00434 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00451 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00434 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70126 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 1.07 mg |  | 1.07 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-04-22 | Carbonate |  | JOB-204428 | XCAMS test - no charge | XCAMS test | 62086 |  | 14047/1 |  |  | 2016-04-18 | 2481 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-04-21 | 10136 | Graphite | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-04-29 | 4400 | C13 Measurement | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-11-02 | 11815 | AMS Submission Results | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-11-02 | 15125 | Close Job | Completed
[168 empty] | 0.99708 | 0.00202 | 70176
[168 empty] | 0.99827 | 0.00189 | 70178
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 102.9 | 89.1757 | 0.0012 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39802 | 39801 | 39802 | 178 | 177 | 178 | 39625.8 | 39981.3 | Margaret Norris | -992.95 | 0.16 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 7.06‰ | 7.06‰ | 2.00‰ | 2.00‰ | 7.06 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -992.95 ‰ | -992.95 ‰ | 0.16 | 0.16 | -992.86 | 0.16 ‰ | -992.86 | 0.16 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0070498 | 0.0070498 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00705 | 0.00016 | 11841 | 2016-05-20 | 14:21 | Albert Zondervan | 204507 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.7 | 0.02% | 37 | 5 | 7 | ..T | 0.00678 | 0.00015 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00678 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00705 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00678 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70222 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water | Complete | JOB-204507 | blank, no charge | Water blank test cutoff at higher pressure than normal cut off at 0.1 | 62087 | Normal | 14047/11 |  |  | 2016-05-05 | 2494 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-05-06 | 2503 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-05-06 | 1518 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-05-06 | 2882 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-05-10 | 10232 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [26 empty] | 2016-05-16 | 4421 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [26 empty] | 2016-11-02 | 11841 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [26 empty] | 2016-11-02 | 15128 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 106.5 | 89.3498 | 0.0012 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39557 | 39556 | 39557 | 184 | 183 | 184 | 39374.7 | 39742.5 | Margaret Norris | -992.73 | 0.17 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 5.10‰ | 5.10‰ | 2.00‰ | 2.00‰ | 5.10 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -992.73 ‰ | -992.73 ‰ | 0.17 | 0.17 | -992.64 | 0.17 ‰ | -992.64 | 0.17 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0072682 | 0.0072682 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2.2 |  |  | 0.00727 | 0.00017 | 11842 | 2016-05-20 | 14:21 | Albert Zondervan | 204508 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.73 | 0.02% | 29 | 5 | 7 | ..T | 0.00699 | 0.00016 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00699 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00727 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00699 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70223 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.4 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water |  | JOB-204508 | blank, no charge | Water blank test cutoff at higher pressure than normal cut off at 1mBar | 62088 | Normal | 14047/11 |  |  | 2016-05-05 | 2495 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-05-06 | 2502 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-05-06 | 1519 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-05-06 | 2884 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-05-10 | 10233 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [26 empty] | 2016-05-16 | 4422 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [26 empty] | 2016-11-02 | 11842 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [26 empty] | 2016-11-02 | 15130 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 105.2 | 89.3472 | 0.0016 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37263 | 37262 | 37263 | 164 | 164 | 164 | 37100.4 | 37428.6 | Margaret Norris | -990.33 | 0.20 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 5.13‰ | 5.13‰ | 2.00‰ | 2.00‰ | 5.13 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -990.33 ‰ | -990.33 ‰ | 0.2 | 0.2 | -990.21 | 0.20 ‰ | -990.21 | 0.20 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0096701 | 0.0096701 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2 |  |  | 0.00967 | 0.0002 | 11843 | 2016-05-20 | 14:21 | Albert Zondervan | 204509 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.97 | 0.02% | 31 | 5 | 7 | ..T | 0.0093 | 0.00019 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.0093 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00967 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.0093 | 0 | 0.00019 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70224 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.93 mg |  | 0.93 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water |  | JOB-204509 | blank, no charge | Water blank test cutoff at higher pressure than normal cut off point was 0.01 | 62089 | Normal | 14047/11 |  |  | 2016-05-05 | 2496 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [26 empty] | 2016-05-06 | 2501 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [26 empty] | 2016-05-06 | 1517 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [26 empty] | 2016-05-06 | 2880 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [26 empty] | 2016-05-10 | 10234 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [26 empty] | 2016-05-16 | 4423 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [26 empty] | 2016-11-02 | 11843 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [26 empty] | 2016-11-02 | 15129 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 126.5 | 89.0588 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50366 | 50366 | 50366 | 353 | 353 | 353 |  |  | Margaret Norris | -998.11 | 0.08 ‰ | 2016-05-20 | 2015-11-22 | 17 | 180 |  |  |  |  | 3.36‰ | 3.36‰ | 2.00‰ | 2.00‰ | 3.36 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.003 | 0 | 0.70711 | -998.11 ‰ | -998.11 ‰ | 0.08 | 0.08 | -998.08 | 0.08 ‰ | -998.08 | 0.08 ‰ |  |  | 2016-05-25 | Jocelyn Turnbull | 16:01 | 0.0018924 | 0.0018924 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.4 |  |  | 0.00189 | 8e-05 | 11867 | 2016-05-25 | 14:48 | Albert Zondervan | 204369 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 11 | 5 | 4 | ... | 0.00182 | 8e-05 |  |  |  |  | 1.0041 | 1.0041 | Yes | Yes | Uncertainty on this result is larger than normally reported. The primary standards used to normalise this result exhibited a larger scatter than normal, due to a component failure during analysis. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. |  |  |  |  | 0.00182 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 1.0002 | 0.00169 | 0.00169 | 0.00189 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00182 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-17 | Margaret Norris | 13:01 |  | Completed |  |  |  |  | 70172 | 2000000 |  |  | 70135 70136 70137 70138 70149 70174 70179 70199 | 70135~70136~70137~70138~70149~70174~70179~70199 | 3061 |  | 3061 3061 3061 3061 3061 3061 3061 3061 |  |  |  |  | 0.87 mg | 0.44 | 0.87 | [5 empty] | 0.99477 | 0.00471 | 70135 |  | XCAMS | ST | 0.44% | Uncertainties on these results are larger than normally reported. The primary standards used to normalise your results exhibited a larger scatter than normal, due to a component failure during the analysis, and implying a larger uncertainty on the authentic samples. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. | * lots to add here still JCT - despite the clear problems with scatter in OxI, the secondary standards look very good. R14047/2 water secondary was 7087±35 (vs 7134 consensus); R24889/4 FIRI-D wood was 4485±26 (vs 4508 ± 3 consensus); R24889/9 FIRI-I cellulose was 4463±26 (vs 4485±5 consensus). R14047/2 is slightly low, but we don't do very well with this material at any time. The others are within 1 sigma. Thus I chose not to increase the wheel-to-wheel error over the usual values we apply. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-04-15 | Carbonate |  | JOB-204369 |  | do with Yoganandan foram samples | 61139 |  | 14047/1 |  |  | 2016-05-02 | 2489 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0044 | 0.00509 | 70136 | [26 empty] | 2016-05-03 | 10182 | Graphite | Completed
[168 empty] | 1.0005 | 0.00473 | 70137 | [26 empty] | 2016-05-10 | 4413 | C13 Measurement | Completed
[168 empty] | 0.99594 | 0.00473 | 70138 | [26 empty] | 2016-05-25 | 11867 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00474 | 70149 | [26 empty] | 2016-05-30 | 14089 | Close Job | Completed
[168 empty] | 1.0059 | 0.00476 | 70174
[168 empty] | 1.0034 | 0.00478 | 70179
[168 empty] | 0.99424 | 0.00473 | 70199
[7 empty] | 131 | 88.9489 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37393 | 37393 | 37393 | 140 | 140 | 140 |  |  | Margaret Norris | -990.49 | 0.17 ‰ | 2016-05-20 | 2015-11-22 | 14 | 180 |  |  |  |  | 4.60‰ | 4.60‰ | 2.00‰ | 2.00‰ | 4.60 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.003 | 0 | 0.70711 | -990.49 ‰ | -990.49 ‰ | 0.17 | 0.17 | -990.37 | 0.17 ‰ | -990.37 | 0.17 ‰ |  |  | 2016-05-25 | Jocelyn Turnbull | 16:01 | 0.0095141 | 0.0095141 | 0.00017 | 0.00017 | 0.00016 | 0.00016 | 1.7 |  |  | 0.00951 | 0.00017 | 11870 | 2016-05-25 | 14:48 | Albert Zondervan | 204422 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 6 | 5 | 6 | ... | 0.00915 | 0.00016 |  |  |  |  | 1.0041 | 1.0041 | Yes | Yes | Uncertainty on this result is larger than normally reported. The primary standards used to normalise this result exhibited a larger scatter than normal, due to a component failure during analysis. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. |  |  |  |  | 0.00915 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 1.0002 | 0.00169 | 0.00169 | 0.00951 | 0 | 0.00017 | 0 | 0 |  |  |  |  | 0.00915 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-17 | Margaret Norris | 13:01 |  | Completed |  |  |  |  | 70215 | 2000000 |  |  | 70135 70136 70137 70138 70149 70174 70179 70199 | 70135~70136~70137~70138~70149~70174~70179~70199 | 3061 | 3057 3056 3048 3046 3045 3042 3040 3061 | 3061 3061 3061 3061 3061 3061 3061 3061 | 2016-05-25 | 15:52 | Jocelyn Turnbull | Yes | 1.10 mg | 0.44 | 1.1 | [5 empty] | 0.99477 | 0.00471 | 70135 |  | XCAMS | ST | 0.44% | Uncertainties on these results are larger than normally reported. The primary standards used to normalise your results exhibited a larger scatter than normal, due to a component failure during the analysis, and implying a larger uncertainty on the authentic samples. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. | * lots to add here still JCT - despite the clear problems with scatter in OxI, the secondary standards look very good. R14047/2 water secondary was 7087±35 (vs 7134 consensus); R24889/4 FIRI-D wood was 4485±26 (vs 4508 ± 3 consensus); R24889/9 FIRI-I cellulose was 4463±26 (vs 4485±5 consensus). R14047/2 is slightly low, but we don't do very well with this material at any time. The others are within 1 sigma. Thus I chose not to increase the wheel-to-wheel error over the usual values we apply. |  | RRL | [5 empty] | Background Inorganic |  | 2016-06-03 | Water |  | JOB-204422 | really to be billed to GCT- if so at what price? | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61140 | Normal | 14047/11 |  |  | 2016-05-05 | 2491 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0044 | 0.00509 | 70136 | [26 empty] | 2016-05-05 | 2497 | Water CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00473 | 70137 | [26 empty] | 2016-05-05 | 1514 | Gas Splitting | Completed
[168 empty] | 0.99594 | 0.00473 | 70138 | [26 empty] | 2016-05-05 | 2875 | Choose Gas Splitting | Completed
[168 empty] | 1.003 | 0.00474 | 70149 | [26 empty] | 2016-05-06 | 10225 | Graphite | Completed
[168 empty] | 1.0059 | 0.00476 | 70174 | [26 empty] | 2016-05-25 | 11870 | AMS Submission Results | Completed
[168 empty] | 1.0034 | 0.00478 | 70179 | [26 empty] | 2016-05-30 | 14101 | Close Job | Completed
[168 empty] | 0.99424 | 0.00473 | 70199
[7 empty] | 76.8 | 92.029 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37571 | 37570 | 37571 | 278 | 278 | 278 |  |  | jenny dahl | -990.69 | 0.32 ‰ | 2016-06-09 | 2016-05-26 | 14 | 14 |  |  |  |  | 3.09‰ | 3.09‰ | 2.00‰ | 2.00‰ | 3.09 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -990.69 ‰ | -990.69 ‰ | 0.32 | 0.32 | -990.58 | 0.33 ‰ | -990.58 | 0.33 ‰ |  |  | 2016-06-10 | Jocelyn Turnbull | 16:52 | 0.0093062 | 0.0093062 | 0.00032 | 0.00032 | 0.00032 | 0.00032 | 2.1 |  |  | 0.00931 | 0.00032 | 11987 | 2016-06-10 | 15:56 | Albert Zondervan | 204398 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.93 | 0.03% | 36 | 5 | 6 | ..S | 0.00895 | 0.00031 |  |  |  |  | 1.0373 | 1.0373 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00895 | 0.01‰ | 0.00031 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00931 | 0 | 0.00032 | 0 | 0 |  |  | 0 | 0 | 0.00895 | 0 | 0.00031 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-31 | jenny dahl | 15:12 |  | Completed |  |  |  |  | 70291 | 2000000 |  |  | 70244 70246 70259 70262 70293 70309 | 70244~70246~70259~70262~70293~70309 | 3064 |  | 3064 3064 3064 3064 3064 3064 | 2016-06-10 | 16:25 | Jocelyn Turnbull |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0011 | 0.00167 | 70244 |  | SSAMS | EA | 0% |  | * 2nd wheel of unknowns measured on SSAMS at ANU * good results and FIRI-D secondaries appear OK * system error is zero * only blanks showed some excess run-to-run variance, which has been accounted for in the reported uncertainties Decided to use only blanks from this wheel for correction -this is only the second wheel from SSAMS and the pervious wheel (3061) only had air materials. Organic blanks are quite poor, but both are consistent. Water blank is also poor but not unusual. Therefore seems most appropriate to use these and not blanks run on XCAMS. FIRI-D all come out just on the low side of 1-sigma with this larger blank, but ok I think | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-06-03 | Water |  | JOB-204398 |  |  | 61197 | Normal | 14047/11 |  |  | 2016-05-24 | 2510 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0012 | 0.00166 | 70246 | [26 empty] | 2016-05-24 | 1526 | Gas Splitting | Completed
[168 empty] | 0.9984 | 0.00166 | 70259 | [26 empty] | 2016-05-24 | 2891 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00165 | 70262 | [26 empty] | 2016-05-26 | 10301 | Graphite | Completed
[168 empty] | 0.99805 | 0.00164 | 70293 | [26 empty] | 2016-05-30 | 4448 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00165 | 70309 | [26 empty] | 2016-06-10 | 11987 | AMS Submission Results | Completed
[197 empty] | 2016-06-13 | 14157 | Close Job | Completed
[5 empty] | 0.00032234 | 0.0093062 | 78.3 | 92.9927 | 0.0708 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7085 | 7084 | 7085 | 55 | 54 | 58 | 7030 | 7139.8 | jenny dahl | -586.03 | 2.83 ‰ | 2016-06-09 | 2016-05-26 | 14 | 14 |  |  |  |  | -7.30‰ | -7.30‰ | 2.00‰ | 2.00‰ | -7.30 ‰ | 2.00 ‰ | AMS | -8.18‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -586.03 ‰ | -586.03 ‰ | 2.83 | 2.98 | -580.92 | 2.86 ‰ | -580.92 | 2.86 ‰ |  |  | 2016-06-10 | Jocelyn Turnbull | 16:52 | 0.41397 | 0.41397 | 0.00283 | 0.00298 | 0.00278 | 0.00278 | 2.1 |  |  | 0.41958 | 0.00123 | 11988 | 2016-06-10 | 15:56 | Albert Zondervan | 204399 | 0.00895 | 0.009 | 0.00403 | 0.004 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.4 | 0.28% | 25 | 5 | 6 | ..S | 0.40352 | 0.00106 |  |  |  |  | 1.0373 | 1.0373 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.7790855770054752. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00895 | 0.00895 | ? | 0.39813 | 0.40‰ | 0.00267 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.41397 | 0 | 0.00283 | 0 | 0 | 0.00895 | 0 | 0.00403 | 0 | 0.39813 | 0 | 0.00267 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-31 | jenny dahl | 15:12 |  | Completed |  |  |  |  | 70292 | 2000000 | 70291 | 70291 | 70244 70246 70259 70262 70293 70309 | 70244~70246~70259~70262~70293~70309 | 3064 | 3064 | 3064 3064 3064 3064 3064 3064 | 2016-06-10 | 16:25 | Jocelyn Turnbull |  | 1.03 mg | 0.13 | 1.03 | 2016-06-09 | 0.00895 | 70291 | 3064 | 0.95 | 1.0011 | 0.00167 | 70244 | EA | SSAMS | EA | 0% |  | * 2nd wheel of unknowns measured on SSAMS at ANU * good results and FIRI-D secondaries appear OK * system error is zero * only blanks showed some excess run-to-run variance, which has been accounted for in the reported uncertainties Decided to use only blanks from this wheel for correction -this is only the second wheel from SSAMS and the pervious wheel (3061) only had air materials. Organic blanks are quite poor, but both are consistent. Water blank is also poor but not unusual. Therefore seems most appropriate to use these and not blanks run on XCAMS. FIRI-D all come out just on the low side of 1-sigma with this larger blank, but ok I think | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-03 | Water |  | JOB-204399 |  | hold off until we have results from water blanks in TST2 new carbosorb | 61198 | Normal | 14047/12 |  |  | 2016-05-24 | 2511 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0012 | 0.00166 | 70246 | [26 empty] | 2016-05-24 | 1527 | Gas Splitting | Completed
[168 empty] | 0.9984 | 0.00166 | 70259 | [26 empty] | 2016-05-24 | 2893 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00165 | 70262 | [26 empty] | 2016-05-26 | 10302 | Graphite | Completed
[168 empty] | 0.99805 | 0.00164 | 70293 | [26 empty] | 2016-05-30 | 4449 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00165 | 70309 | [26 empty] | 2016-06-10 | 11988 | AMS Submission Results | Completed
[197 empty] | 2016-06-13 | 14158 | Close Job | Completed
[7 empty] | 80.9 | 91.8653 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45883 | 45883 | 45883 | 227 | 227 | 227 |  |  | Jocelyn Turnbull | -996.69 | 0.09 ‰ | 2016-07-02 | 2016-04-03 | 39 | 90 |  |  |  |  | 3.24‰ | 3.24‰ | 2.00‰ | 2.00‰ | 3.24 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -996.69 ‰ | -996.69 ‰ | 0.09 | 0.09 | -996.65 | 0.09 ‰ | -996.65 | 0.09 ‰ |  |  | 2016-07-04 | Jocelyn Turnbull | 9:54 | 0.0033065 | 0.0033065 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.4 |  |  | 0.00331 | 9e-05 | 12037 | 2016-07-04 | 8:52 | Albert Zondervan | 204584 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 3 | 5 | 4 | ..S | 0.00318 | 9e-05 |  |  |  |  | 1.0357 | 1.0357 | Yes | Yes | [5 empty] | 0.00318 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00331 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00318 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-10 | Jocelyn Turnbull | 12:00 |  | Completed |  |  |  |  | 70285 | 2000000 |  |  | 70264 70269 70346 70348 70368 70370 | 70264~70269~70346~70348~70368~70370 | 3066 |  | 3066 3066 3066 3066 3066 3066 | 2016-07-04 | 9:48 | Jocelyn Turnbull |  | 0.79 mg |  | 0.79 | [5 empty] | 0.99725 | 0.00163 | 70264 |  | SSAMS | EA | 0% |  | * 4th wheel with unknowns run by S. Fallon on his SSAMS at ANU * physical wheel had been sent over before removing the ANU-sucrose in cathode pos#39 and moving the Ox-I from pos#0 there; this was corrected before measurement and now also in RLIMS * no error adjustments needed; good repeatability all around * again, all blanks seem a bit high Used only blanks run at ANU for blank corrections for this wheel | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-05-24 | Carbonate | Complete | JOB-204584 | blank, no charge | to go with Bahk samples | 61352 |  | 14047/1 |  |  | 2016-05-23 | 2504 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9998 | 0.00155 | 70269 | [26 empty] | 2016-05-24 | 10295 | Graphite | Completed
[168 empty] | 0.99948 | 0.00151 | 70346 | [26 empty] | 2016-05-30 | 4442 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00149 | 70348 | [26 empty] | 2016-07-04 | 12037 | AMS Submission Results | Completed
[168 empty] | 0.99984 | 0.00148 | 70368 | [26 empty] | 2016-07-04 | 14360 | Close Job | Completed
[168 empty] | 1.0022 | 0.00148 | 70370
[7 empty] | 83.7 | 91.8977 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36017 | 36017 | 36017 | 111 | 110 | 111 |  |  | Jocelyn Turnbull | -988.71 | 0.16 ‰ | 2016-07-02 | 2016-04-03 | 31 | 90 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -988.71 ‰ | -988.71 ‰ | 0.16 | 0.16 | -988.57 | 0.16 ‰ | -988.57 | 0.16 ‰ |  |  | 2016-07-04 | Jocelyn Turnbull | 9:54 | 0.011292 | 0.011292 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.5 |  |  | 0.01129 | 0.00016 | 12038 | 2016-07-04 | 8:52 | Albert Zondervan | 204491 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.13 | 0.02% | 23 | 5 | 6 | ..S | 0.01086 | 0.00015 |  |  |  |  | 1.0357 | 1.0357 | Yes | Yes | [5 empty] | 0.01086 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.01129 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.01086 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-10 | Jocelyn Turnbull | 12:00 |  | Completed |  |  |  |  | 70334 | 2000000 |  |  | 70264 70269 70346 70348 70368 70370 | 70264~70269~70346~70348~70368~70370 | 3066 |  | 3066 3066 3066 3066 3066 3066 | 2016-07-04 | 9:46 | Jocelyn Turnbull |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99725 | 0.00163 | 70264 |  | SSAMS | EA | 0% |  | * 4th wheel with unknowns run by S. Fallon on his SSAMS at ANU * physical wheel had been sent over before removing the ANU-sucrose in cathode pos#39 and moving the Ox-I from pos#0 there; this was corrected before measurement and now also in RLIMS * no error adjustments needed; good repeatability all around * again, all blanks seem a bit high Used only blanks run at ANU for blank corrections for this wheel | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-06-24 | Water |  | JOB-204491 | blank, no charge |  | 61353 | Normal | 14047/11 |  |  | 2016-05-31 | 2532 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00155 | 70269 | [26 empty] | 2016-05-31 | 1533 | Gas Splitting | Completed
[168 empty] | 0.99948 | 0.00151 | 70346 | [26 empty] | 2016-05-31 | 2900 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00149 | 70348 | [26 empty] | 2016-06-01 | 10344 | Graphite | Completed
[168 empty] | 0.99984 | 0.00148 | 70368 | [26 empty] | 2016-06-08 | 4467 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00148 | 70370 | [26 empty] | 2016-07-04 | 12038 | AMS Submission Results | Completed
[197 empty] | 2016-07-04 | 14362 | Close Job | Completed
[5 empty] | 0.0001404 | 0.010915 | 82.5 | 92.8423 | 0.0702 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7103 | 7103 | 7103 | 59 | 59 | 62 | 7044.2 | 7162.9 | Jocelyn Turnbull | -586.98 | 3.05 ‰ | 2016-07-02 | 2016-04-03 | 31 | 90 |  |  |  |  | -7.32‰ | -7.32‰ | 2.00‰ | 2.00‰ | -7.32 ‰ | 2.00 ‰ | AMS | -8.22‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -586.98 ‰ | -586.98 ‰ | 3.05 | 3.19 | -581.88 | 3.09 ‰ | -581.88 | 3.09 ‰ |  |  | 2016-07-04 | Jocelyn Turnbull | 9:54 | 0.41302 | 0.41302 | 0.00305 | 0.00319 | 0.00301 | 0.00301 | 2.2 |  |  | 0.41922 | 0.00115 | 12039 | 2016-07-04 | 8:52 | Albert Zondervan | 204492 | 0.00991 | 0.0099 | 0.00446 | 0.0045 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.3 | 0.31% | 2 | 5 | 6 | ..S | 0.40318 | 0.00097 |  |  |  |  | 1.0357 | 1.0357 | Yes | Yes |  | wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 62.065784964928025. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00991 | 0.0105 | 0.00135 | 0.39721 | 0.40‰ | 0.00289 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.41302 | 0 | 0.00305 | 0 | 0 | 0.00991 | 0 | 0.00446 | 0 | 0.39721 | 0 | 0.00289 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-10 | Jocelyn Turnbull | 12:00 |  | Completed |  |  |  |  | 70336 | 2000000 | 70334 70291 | 70334~70291 | 70264 70269 70346 70348 70368 70370 | 70264~70269~70346~70348~70368~70370 | 3066 | 3066 3064 | 3066 3066 3066 3066 3066 3066 | 2016-07-04 | 9:46 | Jocelyn Turnbull |  | 0.96 mg | 0.13 | 0.96 | 2016-07-02 | 0.01086 | 70334 | 3066 | 0.89 | 0.99725 | 0.00163 | 70264 | EA | SSAMS | EA | 0% |  | * 4th wheel with unknowns run by S. Fallon on his SSAMS at ANU * physical wheel had been sent over before removing the ANU-sucrose in cathode pos#39 and moving the Ox-I from pos#0 there; this was corrected before measurement and now also in RLIMS * no error adjustments needed; good repeatability all around * again, all blanks seem a bit high Used only blanks run at ANU for blank corrections for this wheel | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-24 | Water |  | JOB-204492 | secondary, no charge |  | 61354 | Normal | 14047/12 |  |  | 2016-05-31 | 2533 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-06-09 | 0.00895 | 70291 | 3064 | 0.95 | 0.9998 | 0.00155 | 70269 | EA | [25 empty] | 2016-05-31 | 1534 | Gas Splitting | Completed
[168 empty] | 0.99948 | 0.00151 | 70346 | [26 empty] | 2016-05-31 | 2902 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00149 | 70348 | [26 empty] | 2016-06-01 | 10346 | Graphite | Completed
[168 empty] | 0.99984 | 0.00148 | 70368 | [26 empty] | 2016-06-08 | 4468 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00148 | 70370 | [26 empty] | 2016-07-04 | 12039 | AMS Submission Results | Completed
[197 empty] | 2016-07-04 | 14363 | Close Job | Completed
[7 empty] | 76.9 | 88.738 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37938 | 37938 | 37938 | 141 | 140 | 141 |  |  | jenny dahl | -991.11 | 0.16 ‰ | 2016-06-27 | 2016-03-29 | 13 | 90 |  |  |  |  | 1.03‰ | 1.03‰ | 2.00‰ | 2.00‰ | 1.03 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | -991.11 ‰ | -991.11 ‰ | 0.16 | 0.16 | -991 | 0.16 ‰ | -991 | 0.16 ‰ |  |  | 2016-06-29 | Jocelyn Turnbull | 11:17 | 0.0088902 | 0.0088902 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00889 | 0.00016 | 12140 | 2016-06-29 | 10:23 | Albert Zondervan | 204641 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.89 | 0.02% | 10 | 5 | 6 | ... | 0.00855 | 0.00015 |  |  |  |  | 0.9982 | 0.9982 | Yes | Yes | [5 empty] | 0.00855 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00889 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00855 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-27 | jenny dahl | 8:53 |  | Completed |  |  |  |  | 70408 | 2000000 |  |  | 70376 70431 70437 70440 70441 | 70376~70431~70437~70440~70441 | 3070 |  | 3070 3070 3070 3070 3070 | 2016-06-29 | 11:07 | Jocelyn Turnbull |  | 1.07 mg |  | 1.07 | [5 empty] | 1.0006 | 0.00162 | 70376 |  | XCAMS | EA | 0% |  | * OxI in pos#7 (TP70373) is a 5-sigma high outlier and was therefore removed from the list of calibration standards. No reason for this deviation has been found yet. The lowering of Cs-oven temperature halfway during the measurement did not influence the degree of this outlying behaviour. * Excess global run-to-run variance required a 1.06 multiplication of all Poisson uncertainties. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-07-19 | Water | Complete | JOB-204641 | Blank, no charge |  | 61312 | Normal | 14047/11 |  |  | 2016-06-10 | 2562 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99896 | 0.00163 | 70431 | [26 empty] | 2016-06-10 | 1544 | Gas Splitting | Completed
[168 empty] | 0.99727 | 0.00175 | 70437 | [26 empty] | 2016-06-10 | 2913 | Choose Gas Splitting | Completed
[168 empty] | 1.0021 | 0.00164 | 70440 | [26 empty] | 2016-06-14 | 10418 | Graphite | Completed
[168 empty] | 1.0007 | 0.00157 | 70441 | [26 empty] | 2016-06-22 | 4506 | C13 Measurement | Completed
[197 empty] | 2016-06-29 | 12140 | AMS Submission Results | Completed
[197 empty] | 2016-06-29 | 14338 | Close Job | Completed
[5 empty] | 7.2093e-05 | 0.0073667 | 86 | 89.3298 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7140 | 7139 | 7140 | 48 | 48 | 52 | 7091.5 | 7188.4 | jenny dahl | -588.85 | 2.48 ‰ | 2016-06-27 | 2016-03-29 | 13 | 90 |  |  |  |  | -5.60‰ | -5.60‰ | 2.00‰ | 2.00‰ | -5.60 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | -588.85 ‰ | -588.85 ‰ | 2.48 | 2.65 | -583.78 | 2.51 ‰ | -583.78 | 2.51 ‰ |  |  | 2016-06-29 | Jocelyn Turnbull | 11:17 | 0.41115 | 0.41115 | 0.00248 | 0.00265 | 0.00242 | 0.00242 | 2.1 |  |  | 0.41592 | 0.0012 | 12141 | 2016-06-29 | 10:23 | Albert Zondervan | 204643 | 0.0076 | 0.0076 | 0.00342 | 0.0034 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.11 | 0.25% | 18 | 5 | 6 | ... | 0.4 | 0.00103 |  |  |  |  | 0.9982 | 0.9982 | Yes | Yes |  | Result is bang on. wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 51.7383475937902615. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0076 | 0.00709 | 0.00177 | 0.39541 | 0.40‰ | 0.00233 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.41115 | 0 | 0.00248 | 0 | 0 | 0.0076 | 0 | 0.00342 | 0 | 0.39541 | 0 | 0.00233 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-27 | jenny dahl | 8:53 |  | Completed |  |  |  |  | 70409 | 2000000 | 70408 70291 70215 70029 69954 | 70408~70291~70215~70029~69954 | 70376 70431 70437 70440 70441 | 70376~70431~70437~70440~70441 | 3070 | 3070 3064 3061 3057 3056 | 3070 3070 3070 3070 3070 | 2016-06-29 | 11:07 | Jocelyn Turnbull |  | 0.96 mg | 0.13 | 0.96 | 2016-06-27 | 0.00855 | 70408 | 3070 | 1.07 | 1.0006 | 0.00162 | 70376 | EA | XCAMS | EA | 0% |  | * OxI in pos#7 (TP70373) is a 5-sigma high outlier and was therefore removed from the list of calibration standards. No reason for this deviation has been found yet. The lowering of Cs-oven temperature halfway during the measurement did not influence the degree of this outlying behaviour. * Excess global run-to-run variance required a 1.06 multiplication of all Poisson uncertainties. | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-07-19 | Water |  | JOB-204643 | Blank, no charge |  | 61313 | Normal | 14047/12 |  |  | 2016-06-10 | 2563 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-06-09 | 0.00895 | 70291 | 3064 | 0.95 | 0.99896 | 0.00163 | 70431 | ST | [25 empty] | 2016-06-10 | 1545 | Gas Splitting | Completed
[163 empty] | 2016-05-20 | 0.00915 | 70215 | 3061 | 1.1 | 0.99727 | 0.00175 | 70437 | ST | [25 empty] | 2016-06-10 | 2915 | Choose Gas Splitting | Completed
[163 empty] | 2016-04-11 | 0.00585 | 70029 | 3057 | 0.81 | 1.0021 | 0.00164 | 70440 | EA | [25 empty] | 2016-06-14 | 10419 | Graphite | Completed
[163 empty] | 2016-04-06 | 0.0055 | 69954 | 3056 | 1.09 | 1.0007 | 0.00157 | 70441 | EA | [25 empty] | 2016-06-22 | 4507 | C13 Measurement | Completed
[197 empty] | 2016-06-29 | 12141 | AMS Submission Results | Completed
[197 empty] | 2016-06-29 | 14339 | Close Job | Completed
[7 empty] | 118.4 | 88.423 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52531 | 52531 | 52531 | 289 | 288 | 289 |  |  | jenny dahl | -998.55 | 0.05 ‰ | 2016-07-04 | 2016-07-04 | 6 | 0 |  |  |  |  | 1.42‰ | 1.42‰ | 2.00‰ | 2.00‰ | 1.42 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -998.55 ‰ | -998.55 ‰ | 0.05 | 0.05 | -998.54 | 0.05 ‰ | -998.54 | 0.05 ‰ |  |  | 2016-07-11 | Jocelyn Turnbull | 11:33 | 0.0014453 | 0.0014453 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2.8 |  |  | 0.00145 | 5e-05 | 12180 | 2016-07-08 | 14:51 | Albert Zondervan | 204952 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.14 | 0.01% | 15 | 5 | 4 | ... | 0.00139 | 5e-05 |  |  |  |  | 0.99504 | 0.99504 | Yes | Yes | [5 empty] | 0.00139 | 0.00‰ | 5e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00145 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.00139 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-29 | jenny dahl | 16:06 |  | Completed |  |  |  |  | 70517 | 2000000 |  |  | 70462 70465 70529 70530 70538 70539 | 70462~70465~70529~70530~70538~70539 | 3071 |  | 3071 3071 3071 3071 3071 3071 | 2016-07-08 | 16:03 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99876 | 0.00135 | 70462 |  | XCAMS | ST | 0% |  | * good data despite having to re-tune XCAMS because of power mains problems * measurement was extended beyond time needed for regular precision to get an impression how good the repeatability is now: With on average 550 k counts per OxI, we can't see a system error. * 1 cathode failed to give a reliable result: TP70453 (cat#4) * all other variances agree with their Poisson expectation: GOOD !! | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-07-01 | Carbonate | Not Complete | JOB-204952 | blank, no charge | loaded into tiny microcentrifuge tube to match with all other 3 samples. | 61392 |  | 14047/1 |  |  | 2016-06-24 | 2583 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99909 | 0.00136 | 70465 | [26 empty] | 2016-06-28 | 10527 | Graphite | Completed
[168 empty] | 0.99922 | 0.00138 | 70529 | [26 empty] | 2016-06-30 | 4534 | C13 Measurement | Completed
[168 empty] | 0.99974 | 0.00138 | 70530 | [26 empty] | 2016-07-11 | 12180 | AMS Submission Results | Completed
[168 empty] | 1.0029 | 0.00139 | 70538 | [26 empty] | 2016-07-12 | 14417 | Close Job | Completed
[168 empty] | 1.0004 | 0.00131 | 70539
[7 empty] | 115.2 | 88.4851 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39580 | 39579 | 39580 | 127 | 126 | 127 |  |  | jenny dahl | -992.75 | 0.11 ‰ | 2016-07-04 | 2016-07-04 | 17 | 0 |  |  |  |  | 0.72‰ | 0.72‰ | 2.00‰ | 2.00‰ | 0.72 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -992.75 ‰ | -992.75 ‰ | 0.11 | 0.11 | -992.66 | 0.12 ‰ | -992.66 | 0.12 ‰ |  |  | 2016-07-11 | Jocelyn Turnbull | 11:33 | 0.0072474 | 0.0072474 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.4 |  |  | 0.00725 | 0.00011 | 12182 | 2016-07-08 | 14:51 | Albert Zondervan | 204977 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.72 | 0.01% | 30 | 5 | 6 | ... | 0.00697 | 0.00011 |  |  |  |  | 0.99504 | 0.99504 | Yes | Yes | [5 empty] | 0.00697 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00725 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00697 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-29 | jenny dahl | 16:06 |  | Completed |  |  |  |  | 70455 | 2000000 |  |  | 70462 70465 70529 70530 70538 70539 | 70462~70465~70529~70530~70538~70539 | 3071 |  | 3071 3071 3071 3071 3071 3071 | 2016-07-08 | 16:03 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 0.99876 | 0.00135 | 70462 |  | XCAMS | ST | 0% |  | * good data despite having to re-tune XCAMS because of power mains problems * measurement was extended beyond time needed for regular precision to get an impression how good the repeatability is now: With on average 550 k counts per OxI, we can't see a system error. * 1 cathode failed to give a reliable result: TP70453 (cat#4) * all other variances agree with their Poisson expectation: GOOD !! | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-08-03 | Water |  | JOB-204977 | blank, no charge |  | 61393 | Normal | 14047/11 |  |  | 2016-06-15 | 2569 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99909 | 0.00136 | 70465 | [26 empty] | 2016-06-15 | 1550 | Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00138 | 70529 | [26 empty] | 2016-06-15 | 2927 | Choose Gas Splitting | Completed
[168 empty] | 0.99974 | 0.00138 | 70530 | [26 empty] | 2016-06-17 | 10465 | Graphite | Completed
[168 empty] | 1.0029 | 0.00139 | 70538 | [26 empty] | 2016-06-22 | 4509 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00131 | 70539 | [26 empty] | 2016-07-11 | 12182 | AMS Submission Results | Completed
[197 empty] | 2016-07-12 | 14420 | Close Job | Completed
[5 empty] | 7.4038e-05 | 0.0074846 | 120.5 | 89.4926 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7116 | 7116 | 7116 | 46 | 45 | 49 | 7071 | 7162.2 | jenny dahl | -587.66 | 2.34 ‰ | 2016-07-04 | 2016-01-06 | 17 | 180 |  |  |  |  | -10.55‰ | -10.55‰ | 2.00‰ | 2.00‰ | -10.55 ‰ | 2.00 ‰ | AMS | -8.30‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -587.66 ‰ | -587.66 ‰ | 2.34 | 2.52 | -582.57 | 2.37 ‰ | -582.57 | 2.37 ‰ |  |  | 2016-07-11 | Jocelyn Turnbull | 11:33 | 0.41234 | 0.41234 | 0.00234 | 0.00252 | 0.00228 | 0.00228 | 2.4 |  |  | 0.41694 | 0.00106 | 12183 | 2016-07-08 | 14:51 | Albert Zondervan | 204978 | 0.00732 | 0.0073 | 0.00329 | 0.0033 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.23 | 0.23% | 37 | 5 | 6 | ... | 0.40098 | 0.00088 |  |  |  |  | 0.99504 | 0.99504 | Yes | Yes |  | result consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 49.0328515220162487. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00732 | 0.0072 | 0.00137 | 0.39656 | 0.40‰ | 0.00219 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.41234 | 0 | 0.00234 | 0 | 0 | 0.00732 | 0 | 0.00329 | 0 | 0.39656 | 0 | 0.00219 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-29 | jenny dahl | 16:06 |  | Completed |  |  |  |  | 70456 | 2000000 | 70455 70215 70029 69760 | 70455~70215~70029~69760 | 70462 70465 70529 70530 70538 70539 | 70462~70465~70529~70530~70538~70539 | 3071 | 3071 3061 3057 3048 | 3071 3071 3071 3071 3071 3071 | 2016-07-08 | 16:03 | Albert Zondervan |  | 0.84 mg | 0.13 | 0.84 | 2016-07-04 | 0.00697 | 70455 | 3071 | 0.84 | 0.99876 | 0.00135 | 70462 | ST | XCAMS | ST | 0% |  | * good data despite having to re-tune XCAMS because of power mains problems * measurement was extended beyond time needed for regular precision to get an impression how good the repeatability is now: With on average 550 k counts per OxI, we can't see a system error. * 1 cathode failed to give a reliable result: TP70453 (cat#4) * all other variances agree with their Poisson expectation: GOOD !! | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-08-03 | Water | Not Complete | JOB-204978 | secondary, no charge |  | 61394 | Normal | 14047/12 |  |  | 2016-06-15 | 2570 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-05-20 | 0.00915 | 70215 | 3061 | 1.1 | 0.99909 | 0.00136 | 70465 | ST | [25 empty] | 2016-06-15 | 1552 | Gas Splitting | Completed
[163 empty] | 2016-04-11 | 0.00585 | 70029 | 3057 | 0.81 | 0.99922 | 0.00138 | 70529 | ST | [25 empty] | 2016-06-16 | 2929 | Choose Gas Splitting | Completed
[163 empty] | 2016-03-14 | 0.00729 | 69760 | 3048 | 0.85 | 0.99974 | 0.00138 | 70530 | ST | [25 empty] | 2016-06-17 | 10466 | Graphite | Completed
[168 empty] | 1.0029 | 0.00139 | 70538 | [26 empty] | 2016-06-22 | 4510 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00131 | 70539 | [26 empty] | 2016-07-11 | 12183 | AMS Submission Results | Completed
[197 empty] | 2016-07-12 | 14421 | Close Job | Completed
[7 empty] | 103.6 | 88.4727 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38254 | 38254 | 38254 | 147 | 146 | 147 |  |  | jenny dahl | -991.45 | 0.16 ‰ | 2016-07-18 | 2016-07-18 | 18 | 0 |  |  |  |  | 3.03‰ | 3.03‰ | 2.00‰ | 2.00‰ | 3.03 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -991.45 ‰ | -991.45 ‰ | 0.16 | 0.16 | -991.35 | 0.16 ‰ | -991.35 | 0.16 ‰ |  |  | 2016-07-20 | jenny dahl | 11:15 | 0.0085471 | 0.0085471 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.4 |  |  | 0.00855 | 0.00016 | 12261 | 2016-07-19 | 13:01 | Albert Zondervan | 205068 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.85 | 0.02% | 10 | 5 | 6 | ... | 0.00822 | 0.00015 |  |  |  |  | 0.9972 | 0.9972 | Yes | Yes | [5 empty] | 0.00822 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00855 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00822 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-14 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 70550 | 2000000 |  |  | 70419 70421 70428 70505 70506 70507 | 70419~70421~70428~70505~70506~70507 | 3073 |  | 3073 3073 3073 3073 3073 3073 | 2016-07-19 | 14:44 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 1.0003 | 0.00171 | 70419 |  | XCAMS | EA | 0% |  | * straightforward problem-free measurement | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-08-16 | Water |  | JOB-205068 | blank, no charge |  | 61491 | Normal | 14047/11 |  |  | 2016-06-29 | 2601 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9984 | 0.00173 | 70421 | [26 empty] | 2016-06-29 | 1562 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00172 | 70428 | [26 empty] | 2016-06-29 | 2959 | Choose Gas Splitting | Completed
[168 empty] | 0.99782 | 0.00173 | 70505 | [26 empty] | 2016-06-30 | 10560 | Graphite | Completed
[168 empty] | 1.0023 | 0.00173 | 70506 | [26 empty] | 2016-07-04 | 4555 | C13 Measurement | Completed
[168 empty] | 0.99914 | 0.00172 | 70507 | [26 empty] | 2016-07-20 | 12261 | AMS Submission Results | Completed
[197 empty] | 2016-07-20 | 14499 | Close Job | Completed
[5 empty] | 5.6497e-05 | 0.008722 | 100.8 | 89.4629 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7156 | 7156 | 7156 | 57 | 57 | 60 | 7099.5 | 7213.8 | jenny dahl | -589.71 | 2.92 ‰ | 2016-07-18 | 2016-01-20 | 18 | 180 |  |  |  |  | -8.07‰ | -8.07‰ | 2.00‰ | 2.00‰ | -8.07 ‰ | 2.00 ‰ | AMS | -8.14‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -589.71 ‰ | -589.71 ‰ | 2.92 | 3.06 | -584.64 | 2.96 ‰ | -584.64 | 2.96 ‰ |  |  | 2016-07-20 | jenny dahl | 11:15 | 0.41029 | 0.41029 | 0.00292 | 0.00306 | 0.00287 | 0.00287 | 2.3 |  |  | 0.41606 | 0.00127 | 12262 | 2016-07-19 | 13:01 | Albert Zondervan | 205069 | 0.00916 | 0.0092 | 0.00412 | 0.0041 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.03 | 0.29% | 30 | 5 | 6 | ... | 0.40014 | 0.00111 |  |  |  |  | 0.9972 | 0.9972 | Yes | Yes |  | result is spot on consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 59.9438677708221368. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00916 | 0.00842 | 0.0027 | 0.39459 | 0.39‰ | 0.00276 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.41029 | 0 | 0.00292 | 0 | 0 | 0.00916 | 0 | 0.00412 | 0 | 0.39459 | 0 | 0.00276 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-14 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 70551 | 2000000 | 70550 70334 70408 70291 69954 69662 69597 69454 | 70550~70334~70408~70291~69954~69662~69597~69454 | 70419 70421 70428 70505 70506 70507 | 70419~70421~70428~70505~70506~70507 | 3073 | 3073 3066 3070 3064 3056 3046 3045 3042 | 3073 3073 3073 3073 3073 3073 | 2016-07-19 | 14:44 | Albert Zondervan |  | 0.89 mg | 0.13 | 0.89 | 2016-07-18 | 0.00822 | 70550 | 3073 | 0.78 | 1.0003 | 0.00171 | 70419 | EA | XCAMS | EA | 0% |  | * straightforward problem-free measurement | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-08-16 | Water | Complete | JOB-205069 | secondary, no charge |  | 61492 | Normal | 14047/12 |  |  | 2016-06-29 | 2602 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-07-02 | 0.01086 | 70334 | 3066 | 0.89 | 0.9984 | 0.00173 | 70421 | EA | [25 empty] | 2016-06-29 | 1563 | Gas Splitting | Completed
[163 empty] | 2016-06-27 | 0.00855 | 70408 | 3070 | 1.07 | 1.002 | 0.00172 | 70428 | EA | [25 empty] | 2016-06-29 | 2961 | Choose Gas Splitting | Completed
[163 empty] | 2016-06-09 | 0.00895 | 70291 | 3064 | 0.95 | 0.99782 | 0.00173 | 70505 | EA | [25 empty] | 2016-06-30 | 10561 | Graphite | Completed
[163 empty] | 2016-04-06 | 0.0055 | 69954 | 3056 | 1.09 | 1.0023 | 0.00173 | 70506 | EA | [25 empty] | 2016-07-04 | 4556 | C13 Measurement | Completed
[163 empty] | 2016-02-24 | 0.00778 | 69662 | 3046 | 0.91 | 0.99914 | 0.00172 | 70507 | EA | [25 empty] | 2016-07-20 | 12262 | AMS Submission Results | Completed
[163 empty] | 2016-02-11 | 0.01476 | 69597 | 3045 | 0.88 |  |  |  | EA | [25 empty] | 2016-07-20 | 14500 | Close Job | Completed
[163 empty] | 2016-01-27 | 0.00867 | 69454 | 3042 | 1 |  |  |  | EA
[7 empty] | 97.1 | 88.2885 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39330 | 39329 | 39330 | 167 | 167 | 167 |  |  | Margaret Norris | -992.52 | 0.16 ‰ | 2016-07-20 | 2016-01-22 | 5 | 180 |  |  |  |  | 3.25‰ | 3.25‰ | 2.00‰ | 2.00‰ | 3.25 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -992.52 ‰ | -992.52 ‰ | 0.16 | 0.16 | -992.43 | 0.16 ‰ | -992.43 | 0.16 ‰ |  |  | 2016-07-25 | jenny dahl | 12:27 | 0.0074761 | 0.0074761 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.00748 | 0.00016 | 12293 | 2016-07-21 | 12:13 | Albert Zondervan | 205070 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.75 | 0.02% | 16 | 5 | 6 | ... | 0.00719 | 0.00015 |  |  |  |  | 0.99534 | 0.99534 | Yes | Yes | [5 empty] | 0.00719 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 0.99999 | 0.00094 | 0.00094 | 0.00748 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00719 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-18 | Margaret Norris | 11:31 |  | Completed |  |  |  |  | 70636 | 2000000 |  |  | 70463 70471 70473 70531 70536 70537 | 70463~70471~70473~70531~70536~70537 | 3074 |  | 3074 3074 3074 3074 3074 3074 | 2016-07-21 | 13:21 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99951 | 0.00227 | 70463 |  | XCAMS | ST | 0.13% |  | * straightforward measurement | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-08-16 | Water |  | JOB-205070 |  |  | 61523 | Normal | 14047/11 |  |  | 2016-07-12 | 2642 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00228 | 70471 | [26 empty] | 2016-07-12 | 1580 | Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00231 | 70473 | [26 empty] | 2016-07-12 | 2984 | Choose Gas Splitting | Completed
[168 empty] | 0.99564 | 0.00233 | 70531 | [26 empty] | 2016-07-15 | 10646 | Graphite | Completed
[168 empty] | 0.99985 | 0.00225 | 70536 | [26 empty] | 2016-07-19 | 4605 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00235 | 70537 | [26 empty] | 2016-07-25 | 12293 | AMS Submission Results | Completed
[197 empty] | 2016-07-26 | 14558 | Close Job | Completed
[5 empty] | 6.6861e-05 | 0.0074831 | 95.4 | 89.298 | 0.0687 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7157 | 7157 | 7157 | 49 | 49 | 52 | 7108.6 | 7206.7 | Margaret Norris | -589.76 | 2.51 ‰ | 2016-07-20 | 2016-01-22 | 5 | 180 |  |  |  |  | -8.09‰ | -8.09‰ | 2.00‰ | 2.00‰ | -8.09 ‰ | 2.00 ‰ | AMS | -8.72‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -589.76 ‰ | -589.76 ‰ | 2.51 | 2.67 | -584.69 | 2.54 ‰ | -584.69 | 2.54 ‰ |  |  | 2016-07-25 | jenny dahl | 12:27 | 0.41024 | 0.41024 | 0.00251 | 0.00267 | 0.00245 | 0.00245 | 2.2 |  |  | 0.41483 | 0.00138 | 12294 | 2016-07-21 | 12:13 | Albert Zondervan | 205071 | 0.00729 | 0.0073 | 0.00328 | 0.0033 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.02 | 0.25% | 30 | 5 | 6 | ... | 0.39895 | 0.00133 |  |  |  |  | 0.99534 | 0.99534 | Yes | Yes |  | agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 52.2961864808680143. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00729 | 0.0072 | 0.00119 | 0.39454 | 0.39‰ | 0.00241 | 0.00‰ | Yes | 1 | 0.99999 | 0.00094 | 0.00094 | 0.41024 | 0 | 0.00251 | 0 | 0 | 0.00729 | 0 | 0.00328 | 0 | 0.39454 | 0 | 0.00241 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-18 | Margaret Norris | 11:31 |  | Completed |  |  |  |  | 70637 | 2000000 | 70636 70455 70215 70029 69760 | 70636~70455~70215~70029~69760 | 70463 70471 70473 70531 70536 70537 | 70463~70471~70473~70531~70536~70537 | 3074 | 3074 3071 3061 3057 3048 | 3074 3074 3074 3074 3074 3074 | 2016-07-21 | 13:21 | Albert Zondervan |  | 0.88 mg | 0.13 | 0.88 | 2016-07-20 | 0.00719 | 70636 | 3074 | 0.95 | 0.99951 | 0.00227 | 70463 | ST | XCAMS | ST | 0.13% |  | * straightforward measurement | -8.7 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-08-16 | Water |  | JOB-205071 |  |  | 61524 | Normal | 14047/12 |  |  | 2016-07-12 | 2643 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-07-04 | 0.00697 | 70455 | 3071 | 0.84 | 1.0015 | 0.00228 | 70471 | ST | [25 empty] | 2016-07-12 | 1579 | Gas Splitting | Completed
[163 empty] | 2016-05-20 | 0.00915 | 70215 | 3061 | 1.1 | 1.0017 | 0.00231 | 70473 | ST | [25 empty] | 2016-07-12 | 2982 | Choose Gas Splitting | Completed
[163 empty] | 2016-04-11 | 0.00585 | 70029 | 3057 | 0.81 | 0.99564 | 0.00233 | 70531 | ST | [25 empty] | 2016-07-15 | 10647 | Graphite | Completed
[163 empty] | 2016-03-14 | 0.00729 | 69760 | 3048 | 0.85 | 0.99985 | 0.00225 | 70536 | ST | [25 empty] | 2016-07-19 | 4606 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00235 | 70537 | [26 empty] | 2016-07-25 | 12294 | AMS Submission Results | Completed
[197 empty] | 2016-07-26 | 14559 | Close Job | Completed
[7 empty] | 96.1 | 88.2736 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48077 | 48077 | 48077 | 299 | 298 | 299 |  |  | Margaret Norris | -997.48 | 0.09 ‰ | 2016-07-20 | 2016-01-22 | 12 | 180 |  |  |  |  | 3.42‰ | 3.42‰ | 2.00‰ | 2.00‰ | 3.42 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -997.48 ‰ | -997.48 ‰ | 0.09 | 0.09 | -997.45 | 0.09 ‰ | -997.45 | 0.09 ‰ |  |  | 2016-07-25 | jenny dahl | 12:27 | 0.0025163 | 0.0025163 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00252 | 9e-05 | 12318 | 2016-07-21 | 12:13 | Albert Zondervan | 205185 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.25 | 0.01% | 36 | 5 | 4 | ... | 0.00242 | 9e-05 |  |  |  |  | 0.99534 | 0.99534 | Yes | Yes | [5 empty] | 0.00242 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 0.99999 | 0.00094 | 0.00094 | 0.00252 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00242 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-18 | Margaret Norris | 11:31 |  | Completed |  |  |  |  | 70597 | 2000000 |  |  | 70463 70471 70473 70531 70536 70537 | 70463~70471~70473~70531~70536~70537 | 3074 |  | 3074 3074 3074 3074 3074 3074 | 2016-07-21 | 13:21 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99951 | 0.00227 | 70463 |  | XCAMS | ST | 0.13% |  | * straightforward measurement | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2016-07-29 | Carbonate |  | JOB-205185 |  |  | 61548 |  | 14047/1 |  |  | 2016-07-06 | 2626 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00228 | 70471 | [26 empty] | 2016-07-08 | 10607 | Graphite | Completed
[168 empty] | 1.0017 | 0.00231 | 70473 | [26 empty] | 2016-07-11 | 4586 | C13 Measurement | Completed
[168 empty] | 0.99564 | 0.00233 | 70531 | [26 empty] | 2016-07-25 | 12318 | AMS Submission Results | Completed
[168 empty] | 0.99985 | 0.00225 | 70536 | [26 empty] | 2016-07-26 | 14547 | Close Job | Completed
[168 empty] | 1.0018 | 0.00235 | 70537
[7 empty] | 95.7 | 88.1196 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44000 | 44000 | 44000 | 240 | 239 | 240 |  |  | Margaret Norris | -995.82 | 0.12 ‰ | 2016-08-24 | 2016-08-24 | 22 | 0 |  |  |  |  | 3.61‰ | 3.61‰ | 2.00‰ | 2.00‰ | 3.61 ‰ | 2.00 ‰ | AMS | 2.69‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -995.82 ‰ | -995.82 ‰ | 0.12 | 0.12 | -995.77 | 0.13 ‰ | -995.77 | 0.13 ‰ |  |  | 2016-08-26 | Jocelyn Turnbull | 13:44 | 0.00418 | 0.00418 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.7 |  |  | 0.00418 | 0.00012 | 12440 | 2016-08-25 | 16:16 | Albert Zondervan | 205186 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.42 | 0.01% | 26 | 5 | 6 | ... | 0.00402 | 0.00012 |  |  |  |  | 0.99379 | 0.99379 | Yes | Yes |  | result for this IAEA-C1 processing blank appears to have come out with higher F (younger) relative to previous determinations. Longish graphite time - maybe that's why? |  |  |  | 0.00402 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 0.99999 | 0.00074 | 0.00074 | 0.00418 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00402 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-08-24 | Margaret Norris | 8:57 |  | Completed |  |  |  |  | 70688 | 2000000 |  |  | 70716 70717 70718 70719 70732 70733 | 70716~70717~70718~70719~70732~70733 | 3078 |  | 3078 3078 3078 3078 3078 3078 | 2016-08-26 | 11:32 | Albert Zondervan |  | 1.11 mg |  | 1.11 | [5 empty] | 0.99806 | 0.00181 | 70716 |  | XCAMS | ST | 0% |  | * good measurement overall * a few minor error adjustments | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2016-07-29 | Carbonate |  | JOB-205186 |  |  | 61627 |  | 14047/1 |  |  | 2016-07-29 | 2658 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99835 | 0.00179 | 70717 | [26 empty] | 2016-07-29 | 1584 | Gas Splitting | Completed
[168 empty] | 0.99898 | 0.0018 | 70718 | [26 empty] | 2016-07-29 | 2992 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00178 | 70719 | [26 empty] | 2016-08-02 | 10698 | Graphite | Completed
[168 empty] | 1.0019 | 0.00183 | 70732 | [26 empty] | 2016-08-04 | 4635 | C13 Measurement | Completed
[168 empty] | 1.0015 | 0.00187 | 70733 | [26 empty] | 2016-08-26 | 12440 | AMS Submission Results | Completed
[197 empty] | 2016-08-31 | 14677 | Close Job | Completed
[7 empty] | 155.8 | 87.7847 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52192 | 52191 | 52192 | 332 | 332 | 332 |  |  | jeremy thompson | -998.49 | 0.06 ‰ | 2016-08-26 | 2016-08-26 | 15 | 0 |  |  |  |  | 5.72‰ | 5.72‰ | 2.00‰ | 2.00‰ | 5.72 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -998.49 ‰ | -998.49 ‰ | 0.06 | 0.06 | -998.47 | 0.06 ‰ | -998.47 | 0.06 ‰ |  |  | 2016-08-30 | Jocelyn Turnbull | 9:51 | 0.0015077 | 0.0015077 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.6 |  |  | 0.00151 | 6e-05 | 12488 | 2016-08-29 | 11:25 | Albert Zondervan | 205189 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 30 | 5 | 4 | ... | 0.00145 | 6e-05 |  |  |  |  | 0.9921 | 0.9921 | Yes | Yes | [5 empty] | 0.00145 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00151 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00145 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-08-25 | jeremy thompson | 10:51 |  | Completed |  |  |  |  | 70776 | 2000000 |  |  | 70698 70699 70701 70721 70724 70758 | 70698~70699~70701~70721~70724~70758 | 3079 |  | 3079 3079 3079 3079 3079 3079 | 2016-08-29 | 11:46 | Albert Zondervan |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0026 | 0.00184 | 70698 |  | XCAMS | EA | 0.11% |  | * good measurement, with small system error and no other uncertainty adjustments * high beam currents generally * exception: cathode pos#15 (TP70745) had lower output than all others, for unknown reason | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-07-29 | Carbonate |  | JOB-205189 |  |  | 61667 |  | 14047/1 |  |  | 2016-08-10 | 2711 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00183 | 70699 | [26 empty] | 2016-08-11 | 10786 | Graphite | Completed
[168 empty] | 0.99965 | 0.00183 | 70701 | [26 empty] | 2016-08-15 | 4692 | C13 Measurement | Completed
[168 empty] | 0.99817 | 0.00184 | 70721 | [26 empty] | 2016-08-30 | 12488 | AMS Submission Results | Completed
[168 empty] | 0.99783 | 0.00182 | 70724 | [26 empty] | 2016-09-01 | 14745 | Close Job | Completed
[168 empty] | 1.0011 | 0.00183 | 70758
[7 empty] | 144.4 | 88.1432 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39230 | 39229 | 39230 | 132 | 132 | 132 |  |  | jeremy thompson | -992.43 | 0.12 ‰ | 2016-08-26 | 2016-08-26 | 17 | 0 |  |  |  |  | 1.63‰ | 1.63‰ | 2.00‰ | 2.00‰ | 1.63 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -992.43 ‰ | -992.43 ‰ | 0.12 | 0.12 | -992.34 | 0.13 ‰ | -992.34 | 0.13 ‰ |  |  | 2016-08-30 | Jocelyn Turnbull | 9:51 | 0.0075697 | 0.0075697 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 3.4 |  |  | 0.00757 | 0.00012 | 12490 | 2016-08-29 | 11:25 | Albert Zondervan | 205193 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.76 | 0.01% | 10 | 5 | 7 | ... | 0.00728 | 0.00012 |  |  |  |  | 0.9921 | 0.9921 | Yes | Yes | [5 empty] | 0.00728 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00757 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00728 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-08-25 | jeremy thompson | 10:51 |  | Completed |  |  |  |  | 70748 | 2000000 |  |  | 70698 70699 70701 70721 70724 70758 | 70698~70699~70701~70721~70724~70758 | 3079 |  | 3079 3079 3079 3079 3079 3079 | 2016-08-29 | 11:46 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0026 | 0.00184 | 70698 |  | XCAMS | EA | 0.11% |  | * good measurement, with small system error and no other uncertainty adjustments * high beam currents generally * exception: cathode pos#15 (TP70745) had lower output than all others, for unknown reason | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-08-26 | Water | Complete | JOB-205193 |  |  | 61668 | Normal | 14047/11 |  |  | 2016-08-05 | 2682 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00183 | 70699 | [26 empty] | 2016-08-08 | 2695 | Water CO2 Evolution | Completed
[168 empty] | 0.99965 | 0.00183 | 70701 | [26 empty] | 2016-08-08 | 1599 | Gas Splitting | Completed
[168 empty] | 0.99817 | 0.00184 | 70721 | [26 empty] | 2016-08-08 | 3017 | Choose Gas Splitting | Completed
[168 empty] | 0.99783 | 0.00182 | 70724 | [26 empty] | 2016-08-09 | 10758 | Graphite | Completed
[168 empty] | 1.0011 | 0.00183 | 70758 | [26 empty] | 2016-08-11 | 4674 | C13 Measurement | Completed
[197 empty] | 2016-08-30 | 12490 | AMS Submission Results | Completed
[197 empty] | 2016-09-01 | 14748 | Close Job | Completed
[5 empty] | 6.1816e-05 | 0.008125 | 145.2 | 88.9227 | 0.0662 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 |  |  |  | 7548 | 7547 |  | 54 | 54 |  | 7493.6 | 7602.3 | jeremy thompson | -609.21 | 2.64 ‰ | 2016-08-26 | 2016-02-28 | 15 | 180 |  |  |  |  | -7.15‰ |  | 2.00‰ |  | -7.15 ‰ | 2.00 ‰ | AMS | -9.69‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -609.21 ‰ |  | 2.64 |  | -604.39 | 2.68 ‰ | -604.39 | 2.68 ‰ |  |  | 2016-08-30 | Jocelyn Turnbull | 9:50 | 0.39079 |  | 0.00264 |  | 0.0026 | 0.0026 | 1.8 |  |  | 0.39613 | 0.00107 | 12491 | 2016-08-29 | 11:25 | Albert Zondervan | 205194 | 0.00823 |  | 0.0037 |  | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 39.08 | 0.26% | 37 | 5 | 7 | A.. | 0.38097 | 0.00099 |  |  |  |  | 0.9921 | 0.9921 | Yes | Yes |  | This result is vastly different than the consensus value. Note that this can't be an issue of poor blank, because the deviation from the consensus value is towards LOWER fraction-modern. If contamination is the issue, then the contaminant must have a much lower F than IAEA-C2. wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.2920674110203. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00823 | 0.00781 | 0.00178 | 0.37583 |  | 0.00253 |  | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.39079 | 0 | 0.00264 | 0 | 0 | 0.00823 | 0 | 0.0037 | 0 | 0.37583 | 0 | 0.00253 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-08-25 | jeremy thompson | 10:51 |  | Cancelled |  |  |  |  | 70778 | 2000000 | 70748 70550 70334 70408 70291 69954 | 70748~70550~70334~70408~70291~69954 | 70698 70699 70701 70721 70724 70758 | 70698~70699~70701~70721~70724~70758 | 3079 | 3079 3073 3066 3070 3064 3056 | 3079 3079 3079 3079 3079 3079 | 2016-08-29 | 11:46 | Albert Zondervan |  | 1.09 mg | 0.13 | 1.09 | 2016-08-26 | 0.00728 | 70748 | 3079 | 0.87 | 1.0026 | 0.00184 | 70698 | EA | XCAMS | EA | 0.11% |  | * good measurement, with small system error and no other uncertainty adjustments * high beam currents generally * exception: cathode pos#15 (TP70745) had lower output than all others, for unknown reason | -9.7 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-08-26 | Water | Complete | JOB-205194 |  | Appears to be swapped with JOB 205265. Cancelled! |  | Normal | 14047/12 |  |  | 2016-08-05 | 2690 | Water CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-07-18 | 0.00822 | 70550 | 3073 | 0.78 | 1.0007 | 0.00183 | 70699 | EA | [25 empty] | 2016-08-10 | 2712 | Water CO2 Evolution | Completed
[163 empty] | 2016-07-02 | 0.01086 | 70334 | 3066 | 0.89 | 0.99965 | 0.00183 | 70701 | EA | [25 empty] | 2016-08-10 | 1604 | Gas Splitting | Completed
[163 empty] | 2016-06-27 | 0.00855 | 70408 | 3070 | 1.07 | 0.99817 | 0.00184 | 70721 | EA | [25 empty] | 2016-08-10 | 3023 | Choose Gas Splitting | Completed
[163 empty] | 2016-06-09 | 0.00895 | 70291 | 3064 | 0.95 | 0.99783 | 0.00182 | 70724 | EA | [25 empty] | 2016-08-11 | 10788 | Graphite | Completed
[163 empty] | 2016-04-06 | 0.0055 | 69954 | 3056 | 1.09 | 1.0011 | 0.00183 | 70758 | EA | [25 empty] | 2016-08-15 | 4693 | C13 Measurement | Completed
[197 empty] | 2016-08-30 | 12491 | AMS Submission Results | Cancelled
[197 empty] | 2016-09-01 | 14749 | Close Job | Completed
[5 empty] | 6.2092e-05 | 0.0072621 | 124.2 | 89.118 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7123 | 7122 | 7123 | 46 | 45 | 49 | 7077.1 | 7168.3 | Margaret Norris | -587.97 | 2.34 ‰ | 2016-09-20 | 2016-03-24 | 7 | 180 |  |  |  |  | -7.73‰ | -7.73‰ | 2.00‰ | 2.00‰ | -7.73 ‰ | 2.00 ‰ | AMS | -8.20‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -587.97 ‰ | -587.97 ‰ | 2.34 | 2.52 | -582.88 | 2.37 ‰ | -582.88 | 2.37 ‰ |  |  | 2016-09-21 | Jocelyn Turnbull | 12:53 | 0.41203 | 0.41203 | 0.00234 | 0.00252 | 0.00228 | 0.00228 | 2.1 |  |  | 0.41649 | 0.00117 | 12586 | 2016-09-21 | 11:00 | Albert Zondervan | 205321 | 0.00711 | 0.0071 | 0.0032 | 0.0032 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.2 | 0.23% | 6 | 5 | 6 | ... | 0.40055 | 0.001 |  |  |  |  | 0.9937 | 0.9937 | Yes | Yes |  | result consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 49.0632400476090188. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00711 | 0.00698 | 0.00126 | 0.39626 | 0.40‰ | 0.00219 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41203 | 0 | 0.00234 | 0 | 0 | 0.00711 | 0 | 0.0032 | 0 | 0.39626 | 0 | 0.00219 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-19 | Margaret Norris | 10:09 |  | Completed |  |  |  |  | 70934 | 2000000 | 70636 70455 70215 70029 70935 | 70636~70455~70215~70029~70935 | 70820 70821 70832 70833 70834 70883 | 70820~70821~70832~70833~70834~70883 | 3082 | 3074 3071 3061 3057 3082 | 3082 3082 3082 3082 3082 3082 | 2016-09-21 | 11:44 | Albert Zondervan |  | 0.97 mg | 0.13 | 0.97 | 2016-09-20 | 0.00638 | 70935 | 3082 | 0.94 | 1.0001 | 0.00156 | 70820 | ST | XCAMS | ST | 0% |  | * good meas't * minor error adjustments for 3 cathodes * on basis of wildly different 14C concentration, it was found that two samples (TP70934,5) had been permutated; this was found to be caused by a swap on the water-line | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-09-21 | Water | Complete | JOB-205321 |  | JCT 21/9/16. Changed R number from 14047/11 to 14047/12 since this was obviously switched with Travertine at evolution d13C should be + 2.5 | 61810 | Normal | 14047/12 |  |  | 2016-09-12 | 2745 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-07-20 | 0.00719 | 70636 | 3074 | 0.95 | 0.99934 | 0.00157 | 70821 | ST | [25 empty] | 2016-09-12 | 1620 | Gas Splitting | Completed
[163 empty] | 2016-07-04 | 0.00697 | 70455 | 3071 | 0.84 | 1.0006 | 0.00157 | 70832 | ST | [25 empty] | 2016-09-12 | 3048 | Choose Gas Splitting | Completed
[163 empty] | 2016-05-20 | 0.00915 | 70215 | 3061 | 1.1 | 1.0002 | 0.00157 | 70833 | ST | [25 empty] | 2016-09-13 | 10944 | Graphite | Completed
[163 empty] | 2016-04-11 | 0.00585 | 70029 | 3057 | 0.81 | 0.99901 | 0.00157 | 70834 | ST | [25 empty] | 2016-09-20 | 4736 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00157 | 70883 | [26 empty] | 2016-09-21 | 12586 | AMS Submission Results | Completed
[197 empty] | 2016-09-23 | 14895 | Close Job | Completed
[7 empty] | 124.4 | 88.2242 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40290 | 40290 | 40290 | 151 | 151 | 151 |  |  | Margaret Norris | -993.37 | 0.12 ‰ | 2016-09-20 | 2016-09-20 | 7 | 0 |  |  |  |  | 2.32‰ | 2.32‰ | 2.00‰ | 2.00‰ | 2.32 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -993.37 ‰ | -993.37 ‰ | 0.12 | 0.12 | -993.28 | 0.13 ‰ | -993.28 | 0.13 ‰ |  |  | 2016-09-21 | Jocelyn Turnbull | 12:53 | 0.0066339 | 0.0066339 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00663 | 0.00012 | 12587 | 2016-09-21 | 11:00 | Albert Zondervan | 205322 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 26 | 5 | 6 | ... | 0.00638 | 0.00012 |  |  |  |  | 0.9937 | 0.9937 | Yes | Yes | [5 empty] | 0.00638 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00663 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00638 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-19 | Margaret Norris | 10:09 |  | Completed |  |  |  |  | 70935 | 2000000 |  |  | 70820 70821 70832 70833 70834 70883 | 70820~70821~70832~70833~70834~70883 | 3082 |  | 3082 3082 3082 3082 3082 3082 | 2016-09-21 | 11:44 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0001 | 0.00156 | 70820 |  | XCAMS | ST | 0% |  | * good meas't * minor error adjustments for 3 cathodes * on basis of wildly different 14C concentration, it was found that two samples (TP70934,5) had been permutated; this was found to be caused by a swap on the water-line | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-09-21 | Water |  | JOB-205322 |  | JCT 21/9/16. Changed R number from 14047/12 to 14047/11 since this was obviously switched with c1 marble at evolution d13C should be - 8.2 | 61809 | Normal | 14047/11 |  |  | 2016-09-12 | 2746 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99934 | 0.00157 | 70821 | [26 empty] | 2016-09-12 | 1619 | Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00157 | 70832 | [26 empty] | 2016-09-12 | 3046 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 70833 | [26 empty] | 2016-09-13 | 10945 | Graphite | Completed
[168 empty] | 0.99901 | 0.00157 | 70834 | [26 empty] | 2016-09-20 | 4737 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00157 | 70883 | [26 empty] | 2016-09-21 | 12587 | AMS Submission Results | Completed
[197 empty] | 2016-09-23 | 14896 | Close Job | Completed
[7 empty] | 122.9 | 88.3784 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48662 | 48662 | 48662 | 250 | 249 | 250 |  |  | Margaret Norris | -997.66 | 0.07 ‰ | 2016-09-20 | 2016-09-20 | 33 | 0 |  |  |  |  | 0.57‰ | 0.57‰ | 2.00‰ | 2.00‰ | 0.57 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0 | 0.8165 | -997.66 ‰ | -997.66 ‰ | 0.07 | 0.07 | -997.63 | 0.07 ‰ | -997.63 | 0.07 ‰ |  |  | 2016-09-21 | Jocelyn Turnbull | 12:53 | 0.0023395 | 0.0023395 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.7 |  |  | 0.00234 | 7e-05 | 12588 | 2016-09-21 | 11:00 | Albert Zondervan | 205323 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.23 | 0.01% | 8 | 5 | 5 | ... | 0.00225 | 7e-05 |  |  |  |  | 0.9937 | 0.9937 | Yes | Yes | [5 empty] | 0.00225 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00234 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00225 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-19 | Margaret Norris | 10:09 |  | Completed |  |  |  |  | 70810 | 2000000 |  |  | 70820 70821 70832 70833 70834 70883 | 70820~70821~70832~70833~70834~70883 | 3082 |  | 3082 3082 3082 3082 3082 3082 | 2016-09-21 | 11:44 | Albert Zondervan |  | 0.79 mg |  | 0.79 | [5 empty] | 1.0001 | 0.00156 | 70820 |  | XCAMS | ST | 0% |  | * good meas't * minor error adjustments for 3 cathodes * on basis of wildly different 14C concentration, it was found that two samples (TP70934,5) had been permutated; this was found to be caused by a swap on the water-line |  | RRL | [5 empty] | Background Inorganic |  | 2016-08-24 | Carbonate |  | JOB-205323 |  |  | 61808 |  | 14047/1 |  |  | 2016-08-17 | 2722 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99934 | 0.00157 | 70821 | [26 empty] | 2016-08-17 | 1610 | Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00157 | 70832 | [26 empty] | 2016-08-17 | 3030 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 70833 | [26 empty] | 2016-08-18 | 10820 | Graphite | Completed
[168 empty] | 0.99901 | 0.00157 | 70834 | [26 empty] | 2016-09-21 | 12588 | AMS Submission Results | Completed
[168 empty] | 1.0007 | 0.00157 | 70883 | [26 empty] | 2016-09-23 | 14897 | Close Job | Completed
[7 empty] | 126.5 | 88.1565 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51973 | 51973 | 51973 | 323 | 323 | 323 |  |  | jenny dahl | -998.45 | 0.06 ‰ | 2016-09-30 | 2016-09-30 | 15 | 0 |  |  |  |  | 4.48‰ | 4.48‰ | 2.00‰ | 2.00‰ | 4.48 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -998.45 ‰ | -998.45 ‰ | 0.06 | 0.06 | -998.43 | 0.06 ‰ | -998.43 | 0.06 ‰ |  |  | 2016-10-03 | Jocelyn Turnbull | 9:05 | 0.0015493 | 0.0015493 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.1 |  |  | 0.00155 | 6e-05 | 12649 | 2016-10-01 | 14:01 | Albert Zondervan | 205434 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 5 | 5 | 4 | ... | 0.00149 | 6e-05 |  |  |  |  | 0.99507 | 0.99507 | Yes | Yes | [5 empty] | 0.00149 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00155 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00149 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-27 | jenny dahl | 14:56 |  | Completed |  |  |  |  | 70966 | 2000000 |  |  | 70889 70890 70952 70953 70959 70960 71033 | 70889~70890~70952~70953~70959~70960~71033 | 3083 |  | 3083 3083 3083 3083 3083 3083 3083 | 2016-10-01 | 14:50 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0003 | 0.00156 | 70889 |  | XCAMS | EA | 0% |  | * solid meas't without issues | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-09-21 | Carbonate | Complete | JOB-205434 |  |  | 61848 |  | 14047/1 |  |  | 2016-09-14 | 2752 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99956 | 0.00156 | 70890 | [26 empty] | 2016-09-15 | 10976 | Graphite | Completed
[168 empty] | 0.99996 | 0.00158 | 70952 | [26 empty] | 2016-09-20 | 4739 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00155 | 70953 | [26 empty] | 2016-10-03 | 12649 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00156 | 70959 | [26 empty] | 2016-10-06 | 14961 | Close Job | Completed
[168 empty] | 0.99895 | 0.00155 | 70960
[168 empty] | 0.99894 | 0.00153 | 71033
[7 empty] | 103 | 88.0137 | 0.0094 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23244 | 23243 | 23244 | 57 | 57 | 57 | 23186.7 | 23301.4 | jenny dahl | -944.62 | 0.40 ‰ | 2016-09-30 | 2016-09-30 | 3 | 0 |  |  |  |  | 6.11‰ | 6.11‰ | 2.00‰ | 2.00‰ | 6.11 ‰ | 2.00 ‰ | AMS | 1.95‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -944.62 ‰ | -944.62 ‰ | 0.4 | 0.4 | -943.94 | 0.40 ‰ | -943.94 | 0.40 ‰ |  |  | 2016-10-03 | Jocelyn Turnbull | 9:05 | 0.055379 | 0.055379 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 2.2 |  |  | 0.05538 | 0.0004 | 12656 | 2016-10-01 | 14:01 | Albert Zondervan | 205517 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 5.54 | 0.04% | 11 | 5 | 6 | .X. | 0.05326 | 0.00038 |  |  |  |  | 0.99507 | 0.99507 | Yes | Yes | [5 empty] | 0.05326 | 0.05‰ | 0.00038 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.05538 | 0 | 0.0004 | 0 | 0 |  |  | 0 | 0 | 0.05326 | 0 | 0.00038 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-27 | jenny dahl | 14:56 |  | Completed |  |  |  |  | 71042 | 2000000 |  |  | 70889 70890 70952 70953 70959 70960 71033 | 70889~70890~70952~70953~70959~70960~71033 | 3083 |  | 3083 3083 3083 3083 3083 3083 3083 | 2016-10-01 | 14:50 | Albert Zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0003 | 0.00156 | 70889 |  | XCAMS | EA | 0% |  | * solid meas't without issues | 2 | RRL | [5 empty] | Background Test |  | 2016-11-10 | Water | Complete | JOB-205517 | JCT ADDED INVOICE 540GCT21 | Water blank test. No flush with nitrogen, large vessel | 61849 | Normal | 14047/11 |  |  | 2016-09-26 | 2762 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00156 | 70890 | [26 empty] | 2016-09-26 | 1622 | Gas Splitting | Completed
[168 empty] | 0.99996 | 0.00158 | 70952 | [26 empty] | 2016-09-26 | 3054 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00155 | 70953 | [26 empty] | 2016-09-27 | 11052 | Graphite | Completed
[168 empty] | 1.0015 | 0.00156 | 70959 | [26 empty] | 2016-10-03 | 4749 | C13 Measurement | Completed
[168 empty] | 0.99895 | 0.00155 | 70960 | [26 empty] | 2016-10-03 | 12656 | AMS Submission Results | Completed
[168 empty] | 0.99894 | 0.00153 | 71033 | [26 empty] | 2016-10-06 | 14970 | Close Job | Completed
[197 empty] | 2017-03-27 | 4996 | Invoice | Completed
[7 empty] | 121.9 | 88.4284 | 0.0033 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31558 | 31557 | 31558 | 89 | 89 | 89 | 31469.1 | 31647.4 | jenny dahl | -980.33 | 0.22 ‰ | 2016-09-30 | 2016-09-30 | 3 | 0 |  |  |  |  | 1.40‰ | 1.40‰ | 2.00‰ | 2.00‰ | 1.40 ‰ | 2.00 ‰ | AMS | 2.21‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -980.33 ‰ | -980.33 ‰ | 0.22 | 0.22 | -980.08 | 0.22 ‰ | -980.08 | 0.22 ‰ |  |  | 2016-10-03 | Jocelyn Turnbull | 9:05 | 0.019673 | 0.019673 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.4 |  |  | 0.01967 | 0.00022 | 12657 | 2016-10-01 | 14:01 | Albert Zondervan | 205518 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 1.97 | 0.02% | 31 | 5 | 6 | .X. | 0.01892 | 0.00021 |  |  |  |  | 0.99507 | 0.99507 | Yes | Yes | [5 empty] | 0.01892 | 0.02‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.01967 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01892 | 0 | 0.00021 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-27 | jenny dahl | 14:56 |  | Completed |  |  |  |  | 71043 | 2000000 |  |  | 70889 70890 70952 70953 70959 70960 71033 | 70889~70890~70952~70953~70959~70960~71033 | 3083 |  | 3083 3083 3083 3083 3083 3083 3083 | 2016-10-01 | 14:50 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0003 | 0.00156 | 70889 |  | XCAMS | EA | 0% |  | * solid meas't without issues | 2.2 | RRL | [5 empty] | Background Test |  | 2016-11-10 | Water | Not Complete | JOB-205518 | JCT ADDED INVOICE 540GCT21 | Water blank test. No flush with nitrogen, small vessel. C13 readings | 61850 | Normal | 14047/11 |  |  | 2016-09-26 | 2763 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00156 | 70890 | [26 empty] | 2016-09-26 | 1626 | Gas Splitting | Completed
[168 empty] | 0.99996 | 0.00158 | 70952 | [26 empty] | 2016-09-26 | 3062 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00155 | 70953 | [26 empty] | 2016-09-27 | 11053 | Graphite | Completed
[168 empty] | 1.0015 | 0.00156 | 70959 | [26 empty] | 2016-10-03 | 4750 | C13 Measurement | Completed
[168 empty] | 0.99895 | 0.00155 | 70960 | [26 empty] | 2016-10-03 | 12657 | AMS Submission Results | Completed
[168 empty] | 0.99894 | 0.00153 | 71033 | [26 empty] | 2016-10-06 | 14971 | Close Job | Completed
[197 empty] | 2017-03-27 | 4997 | Invoice | Completed
[7 empty] | 106 | 88.4189 | 0.0007 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43784 | 43783 | 43784 | 214 | 213 | 214 |  |  | Cathy Ginnane | -995.71 | 0.11 ‰ | 2016-10-05 | 2016-10-05 | 15 | 0 |  |  |  |  | 1.64‰ | 1.64‰ | 2.00‰ | 2.00‰ | 1.64 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -995.71 ‰ | -995.71 ‰ | 0.11 | 0.11 | -995.65 | 0.12 ‰ | -995.65 | 0.12 ‰ |  |  | 2016-10-18 | Jocelyn Turnbull | 10:14 | 0.0042944 | 0.0042944 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 7.9 |  |  | 0.00429 | 0.00011 | 12700 | 2016-10-06 | 17:18 | Albert Zondervan | 202251 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.43 | 0.01% | 20 | 5 | 3 | ... | 0.00413 | 0.00011 |  |  |  |  | 0.99521 | 0.99521 | Yes | Yes | [5 empty] | 0.00413 | 0.00‰ | 0.00011 | 0.00‰ | No | 1 | 1 | 0.00095 | 0.00095 | 0.00429 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00413 | 0 | 0.00011 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2016-10-04 | Cathy Ginnane | 15:19 |  | Completed |  |  |  |  | 70986 | 2000000 |  |  | 71018 71019 71020 71039 71040 71041 | 71018~71019~71020~71039~71040~71041 | 3085 |  | 3085 3085 3085 3085 3085 3085 | 2016-10-18 | 9:33 | Albert Zondervan |  | 0.30 mg |  | 0.3 | [5 empty] | 0.99835 | 0.00229 | 71018 |  | XCAMS | ST | 0.16% |  | * wheel with also small (SM+TY) samples * good measurement with expected response towards smaller sample sizes * small system error hi all Jocelyn and I have refined my earlier blank-subtraction of TW3085 (ACE): when sample weight <= 0.3 mgC, then the sample is a small and must receive the Santos-method blank correction using small blanks and OxIs in the same wheel; when sample weight > 0.3 mgC, then we apply the normal procedure, spanning several months of recent measurements. Two groups were identified to blank-correct: inorganics: 6x unknowns (shell, foram) + 4x secondaries (C2-travertine) + 5x blanks (C1-marble) organics: 8x unknowns (soil, sediment, plant) + 3x secondaries (FIRI-I) + 3x blanks (KauriRRd-wood) Since we have a mix of smalls and larges in this wheel, both methods are applied. So, apart from sample-type grouping there is mass grouping, thus making it a 2 x 2 = 4 groups exercise. Instead of the regular blank-subtraction procedure for large samples, the mass-dependency in the OxIs and blanks in this wheel have been used to derive the modern and dead contaminants, m and d, resp. (We discussed this in the last team meeting.) The method is described in Guaciara Santos et al. [NIM B 259 (2007) 293-302]. No information from other wheels has been used. While doing this, the 'smalls correction procedure' and the associated Excel spreadsheet was revamped. However, it is not complete yet. One of the issues to sort is that the correction formulae should IMO apply to the parameter ratio-to-standard, not fraction-modern. The other major challenge is to make the new procedure more transparent and simple to use. I will try to continue with that until I have it ready such that any operator can do this task. With the newly developed Excel spreadsheet, I determined the parameters for the contaminants: dead-C contaminant: d = 0.00167 +- 0.00015 mgC, derived from the OxIs modern-C contaminant: m = 0.00084 +- 0.000014 mgC, derived from the KauriRRd modern-C contaminant: m = 0.00092 +- 0.000006 mgC, derived from the C1-marble The values for m and d were used for the smalls, but not their uncertainties put out by the regresssion. Because of lack of statistics on repeatability amongst small blanks, we must err on the side of caution. For that reason, we set the undertainty on m and d equal to the m and d values themselves. The small OxIs were nicely corrected to the average value for the large OxIs. With the corrections made, the secondaries look a lot more consistent with their expectation values. This set is ready for a final check prior reporting. cheers Albert |  | RRL | [5 empty] | Background Inorganic |  | 2015-06-01 | Carbonate | Not Complete | JOB-202251 |  | JCT - added back to ACE since there was space REMOVE FROM ACE 0.30mg small std to match levy sample, aim for 0.2mg C sample data swapped with job 202251 which was loaded incorrectly at graphite to allow graphite data to follow on correctly | 61909 |  | 14047/1 |  |  | 2016-09-19 | 2755 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00231 | 71019 | [26 empty] | 2016-09-20 | 10996 | Graphite | Completed
[168 empty] | 1.0012 | 0.00233 | 71020 | [26 empty] | 2016-10-18 | 12700 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00234 | 71039 | [26 empty] | 2016-10-18 | 14988 | Close Job | Completed
[168 empty] | 1.0032 | 0.00234 | 71040
[168 empty] | 0.99656 | 0.0023 | 71041
[7 empty] | 105.2 | 88.4853 | 0.0003 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50961 | 50961 | 50961 | 332 | 332 | 332 |  |  | Cathy Ginnane | -998.24 | 0.07 ‰ | 2016-10-05 | 2016-10-05 | 15 | 0 |  |  |  |  | 0.88‰ | 0.88‰ | 2.00‰ | 2.00‰ | 0.88 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -998.24 ‰ | -998.24 ‰ | 0.07 | 0.07 | -998.22 | 0.07 ‰ | -998.22 | 0.07 ‰ |  |  | 2016-10-18 | Jocelyn Turnbull | 10:14 | 0.0017573 | 0.0017573 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 4.7 |  |  | 0.00176 | 7e-05 | 12717 | 2016-10-06 | 17:18 | Albert Zondervan | 205473 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.18 | 0.01% | 37 | 5 | 3 | ... | 0.00169 | 7e-05 |  |  |  |  | 0.99521 | 0.99521 | Yes | Yes | [5 empty] | 0.00169 | 0.00‰ | 7e-05 | 0.00‰ | No | 1 | 1 | 0.00095 | 0.00095 | 0.00176 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00169 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-10-04 | Cathy Ginnane | 15:19 |  | Completed |  |  |  |  | 70992 | 2000000 |  |  | 71018 71019 71020 71039 71040 71041 | 71018~71019~71020~71039~71040~71041 | 3085 |  | 3085 3085 3085 3085 3085 3085 | 2016-10-18 | 9:33 | Albert Zondervan |  | 0.46 mg |  | 0.46 | [5 empty] | 0.99835 | 0.00229 | 71018 |  | XCAMS | ST | 0.16% |  | * wheel with also small (SM+TY) samples * good measurement with expected response towards smaller sample sizes * small system error hi all Jocelyn and I have refined my earlier blank-subtraction of TW3085 (ACE): when sample weight <= 0.3 mgC, then the sample is a small and must receive the Santos-method blank correction using small blanks and OxIs in the same wheel; when sample weight > 0.3 mgC, then we apply the normal procedure, spanning several months of recent measurements. Two groups were identified to blank-correct: inorganics: 6x unknowns (shell, foram) + 4x secondaries (C2-travertine) + 5x blanks (C1-marble) organics: 8x unknowns (soil, sediment, plant) + 3x secondaries (FIRI-I) + 3x blanks (KauriRRd-wood) Since we have a mix of smalls and larges in this wheel, both methods are applied. So, apart from sample-type grouping there is mass grouping, thus making it a 2 x 2 = 4 groups exercise. Instead of the regular blank-subtraction procedure for large samples, the mass-dependency in the OxIs and blanks in this wheel have been used to derive the modern and dead contaminants, m and d, resp. (We discussed this in the last team meeting.) The method is described in Guaciara Santos et al. [NIM B 259 (2007) 293-302]. No information from other wheels has been used. While doing this, the 'smalls correction procedure' and the associated Excel spreadsheet was revamped. However, it is not complete yet. One of the issues to sort is that the correction formulae should IMO apply to the parameter ratio-to-standard, not fraction-modern. The other major challenge is to make the new procedure more transparent and simple to use. I will try to continue with that until I have it ready such that any operator can do this task. With the newly developed Excel spreadsheet, I determined the parameters for the contaminants: dead-C contaminant: d = 0.00167 +- 0.00015 mgC, derived from the OxIs modern-C contaminant: m = 0.00084 +- 0.000014 mgC, derived from the KauriRRd modern-C contaminant: m = 0.00092 +- 0.000006 mgC, derived from the C1-marble The values for m and d were used for the smalls, but not their uncertainties put out by the regresssion. Because of lack of statistics on repeatability amongst small blanks, we must err on the side of caution. For that reason, we set the undertainty on m and d equal to the m and d values themselves. The small OxIs were nicely corrected to the average value for the large OxIs. With the corrections made, the secondaries look a lot more consistent with their expectation values. This set is ready for a final check prior reporting. cheers Albert |  | RRL | [5 empty] | Background Inorganic |  | 2016-10-05 | Carbonate |  | JOB-205473 |  | 0.40mg small to match the size of Levy samples, make when we know what size to aim for. loaded into tiny centrifuge tube to match small levy sample | 61918 |  | 14047/1 |  |  | 2016-09-08 | 2743 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00231 | 71019 | [26 empty] | 2016-09-20 | 11002 | Graphite | Completed
[168 empty] | 1.0012 | 0.00233 | 71020 | [26 empty] | 2016-10-18 | 12717 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00234 | 71039 | [26 empty] | 2016-10-18 | 14995 | Close Job | Completed
[168 empty] | 1.0032 | 0.00234 | 71040
[168 empty] | 0.99656 | 0.0023 | 71041
[7 empty] | 79 | 88.6306 | 0.001 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40810 | 40810 | 40810 | 201 | 201 | 201 |  |  | Cathy Ginnane | -993.78 | 0.16 ‰ | 2016-10-05 | 2016-10-05 | 15 | 0 |  |  |  |  | -0.76‰ | -0.76‰ | 2.00‰ | 2.00‰ | -0.76 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -993.78 ‰ | -993.78 ‰ | 0.16 | 0.16 | -993.71 | 0.16 ‰ | -993.71 | 0.16 ‰ |  |  | 2016-10-18 | Jocelyn Turnbull | 10:14 | 0.006218 | 0.006218 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 13.4 |  |  | 0.00622 | 0.00016 | 12728 | 2016-10-06 | 17:18 | Albert Zondervan | 205520 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.62 | 0.02% | 2 | 5 | 3 | ... | 0.00598 | 0.00015 |  |  |  |  | 0.99521 | 0.99521 | Yes | Yes | [5 empty] | 0.00598 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1 | 0.00095 | 0.00095 | 0.00622 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00598 | 0 | 0.00015 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2016-10-04 | Cathy Ginnane | 15:19 |  | Completed |  |  |  |  | 70984 | 2000000 |  |  | 71018 71019 71020 71039 71040 71041 | 71018~71019~71020~71039~71040~71041 | 3085 |  | 3085 3085 3085 3085 3085 3085 | 2016-10-18 | 9:33 | Albert Zondervan |  | 0.17 mg |  | 0.17 | [5 empty] | 0.99835 | 0.00229 | 71018 |  | XCAMS | ST | 0.16% |  | * wheel with also small (SM+TY) samples * good measurement with expected response towards smaller sample sizes * small system error hi all Jocelyn and I have refined my earlier blank-subtraction of TW3085 (ACE): when sample weight <= 0.3 mgC, then the sample is a small and must receive the Santos-method blank correction using small blanks and OxIs in the same wheel; when sample weight > 0.3 mgC, then we apply the normal procedure, spanning several months of recent measurements. Two groups were identified to blank-correct: inorganics: 6x unknowns (shell, foram) + 4x secondaries (C2-travertine) + 5x blanks (C1-marble) organics: 8x unknowns (soil, sediment, plant) + 3x secondaries (FIRI-I) + 3x blanks (KauriRRd-wood) Since we have a mix of smalls and larges in this wheel, both methods are applied. So, apart from sample-type grouping there is mass grouping, thus making it a 2 x 2 = 4 groups exercise. Instead of the regular blank-subtraction procedure for large samples, the mass-dependency in the OxIs and blanks in this wheel have been used to derive the modern and dead contaminants, m and d, resp. (We discussed this in the last team meeting.) The method is described in Guaciara Santos et al. [NIM B 259 (2007) 293-302]. No information from other wheels has been used. While doing this, the 'smalls correction procedure' and the associated Excel spreadsheet was revamped. However, it is not complete yet. One of the issues to sort is that the correction formulae should IMO apply to the parameter ratio-to-standard, not fraction-modern. The other major challenge is to make the new procedure more transparent and simple to use. I will try to continue with that until I have it ready such that any operator can do this task. With the newly developed Excel spreadsheet, I determined the parameters for the contaminants: dead-C contaminant: d = 0.00167 +- 0.00015 mgC, derived from the OxIs modern-C contaminant: m = 0.00084 +- 0.000014 mgC, derived from the KauriRRd modern-C contaminant: m = 0.00092 +- 0.000006 mgC, derived from the C1-marble The values for m and d were used for the smalls, but not their uncertainties put out by the regresssion. Because of lack of statistics on repeatability amongst small blanks, we must err on the side of caution. For that reason, we set the undertainty on m and d equal to the m and d values themselves. The small OxIs were nicely corrected to the average value for the large OxIs. With the corrections made, the secondaries look a lot more consistent with their expectation values. This set is ready for a final check prior reporting. cheers Albert |  | RRL | [5 empty] | Background Inorganic |  | 2016-10-07 | Carbonate |  | JOB-205520 | JCT ADDED INVOICE 540GCT21 | used in blank correction within RLIMS- change size category? 0.17mg small std to match levy sample, aim for 0.15mg C | 61899 |  | 14047/1 |  |  | 2016-09-19 | 2756 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00231 | 71019 | [26 empty] | 2016-09-20 | 10994 | Graphite | Completed
[168 empty] | 1.0012 | 0.00233 | 71020 | [26 empty] | 2016-10-18 | 12728 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00234 | 71039 | [26 empty] | 2016-10-18 | 14996 | Close Job | Completed
[168 empty] | 1.0032 | 0.00234 | 71040 | [26 empty] | 2017-03-27 | 16037 | Close Job | Completed
[168 empty] | 0.99656 | 0.0023 | 71041 | [26 empty] | 2017-03-27 | 4998 | Invoice | Completed
[7 empty] | 71.5 | 88.7821 | 0.0013 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38863 | 38863 | 38863 | 189 | 189 | 189 |  |  | Cathy Ginnane | -992.08 | 0.19 ‰ | 2016-10-05 | 2016-10-05 | 13 | 0 |  |  |  |  | -2.46‰ | -2.46‰ | 2.00‰ | 2.00‰ | -2.46 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -992.08 ‰ | -992.08 ‰ | 0.19 | 0.19 | -991.98 | 0.19 ‰ | -991.98 | 0.19 ‰ |  |  | 2016-10-18 | Jocelyn Turnbull | 10:14 | 0.0079232 | 0.0079232 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 25.1 |  |  | 0.00792 | 0.00019 | 12729 | 2016-10-06 | 17:18 | Albert Zondervan | 205521 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.79 | 0.02% | 12 | 5 | 3 | ... | 0.00762 | 0.00018 |  |  |  |  | 0.99521 | 0.99521 | Yes | Yes | [5 empty] | 0.00762 | 0.01‰ | 0.00018 | 0.00‰ | No | 1 | 1 | 0.00095 | 0.00095 | 0.00792 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00762 | 0 | 0.00018 | 0 | Small |  | 0.9878 | 0.95 | 1 | 2016-10-04 | Cathy Ginnane | 15:19 |  | Completed |  |  |  |  | 71015 | 2000000 |  |  | 71018 71019 71020 71039 71040 71041 | 71018~71019~71020~71039~71040~71041 | 3085 |  | 3085 3085 3085 3085 3085 3085 | 2016-10-18 | 9:33 | Albert Zondervan |  | 0.11 mg |  | 0.11 | [5 empty] | 0.99835 | 0.00229 | 71018 |  | XCAMS | ST | 0.16% |  | * wheel with also small (SM+TY) samples * good measurement with expected response towards smaller sample sizes * small system error hi all Jocelyn and I have refined my earlier blank-subtraction of TW3085 (ACE): when sample weight <= 0.3 mgC, then the sample is a small and must receive the Santos-method blank correction using small blanks and OxIs in the same wheel; when sample weight > 0.3 mgC, then we apply the normal procedure, spanning several months of recent measurements. Two groups were identified to blank-correct: inorganics: 6x unknowns (shell, foram) + 4x secondaries (C2-travertine) + 5x blanks (C1-marble) organics: 8x unknowns (soil, sediment, plant) + 3x secondaries (FIRI-I) + 3x blanks (KauriRRd-wood) Since we have a mix of smalls and larges in this wheel, both methods are applied. So, apart from sample-type grouping there is mass grouping, thus making it a 2 x 2 = 4 groups exercise. Instead of the regular blank-subtraction procedure for large samples, the mass-dependency in the OxIs and blanks in this wheel have been used to derive the modern and dead contaminants, m and d, resp. (We discussed this in the last team meeting.) The method is described in Guaciara Santos et al. [NIM B 259 (2007) 293-302]. No information from other wheels has been used. While doing this, the 'smalls correction procedure' and the associated Excel spreadsheet was revamped. However, it is not complete yet. One of the issues to sort is that the correction formulae should IMO apply to the parameter ratio-to-standard, not fraction-modern. The other major challenge is to make the new procedure more transparent and simple to use. I will try to continue with that until I have it ready such that any operator can do this task. With the newly developed Excel spreadsheet, I determined the parameters for the contaminants: dead-C contaminant: d = 0.00167 +- 0.00015 mgC, derived from the OxIs modern-C contaminant: m = 0.00084 +- 0.000014 mgC, derived from the KauriRRd modern-C contaminant: m = 0.00092 +- 0.000006 mgC, derived from the C1-marble The values for m and d were used for the smalls, but not their uncertainties put out by the regresssion. Because of lack of statistics on repeatability amongst small blanks, we must err on the side of caution. For that reason, we set the undertainty on m and d equal to the m and d values themselves. The small OxIs were nicely corrected to the average value for the large OxIs. With the corrections made, the secondaries look a lot more consistent with their expectation values. This set is ready for a final check prior reporting. cheers Albert |  | RRL | [5 empty] | Background Inorganic |  | 2016-10-07 | Carbonate |  | JOB-205521 | JCT ADDED INVOICE 540GCT21 | used in blank correction within RLIMS- change size category? 0.13mg SLUSHIE small std to match levy sample, aim for 0.1mg C sample data swapped with job 202251 after graphite for that sample to make data follow correctly | 61893 |  | 14047/1 |  |  | 2016-09-19 | 2757 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00231 | 71019 | [26 empty] | 2016-09-22 | 11025 | Graphite | Completed
[168 empty] | 1.0012 | 0.00233 | 71020 | [26 empty] | 2016-10-18 | 12729 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00234 | 71039 | [26 empty] | 2016-10-18 | 14997 | Close Job | Completed
[168 empty] | 1.0032 | 0.00234 | 71040 | [26 empty] | 2017-03-27 | 16038 | Close Job | Completed
[168 empty] | 0.99656 | 0.0023 | 71041 | [26 empty] | 2017-03-27 | 4999 | Invoice | Completed
[7 empty] | 105.2 | 88.6383 | 0.0005 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46459 | 46459 | 46459 | 244 | 243 | 244 |  |  | Cathy Ginnane | -996.92 | 0.09 ‰ | 2016-10-05 | 2016-10-05 | 14 | 0 |  |  |  |  | -0.84‰ | -0.84‰ | 2.00‰ | 2.00‰ | -0.84 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -996.92 ‰ | -996.92 ‰ | 0.09 | 0.09 | -996.88 | 0.09 ‰ | -996.88 | 0.09 ‰ |  |  | 2016-10-18 | Jocelyn Turnbull | 10:14 | 0.0030778 | 0.0030778 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 7.1 |  |  | 0.00308 | 9e-05 | 12730 | 2016-10-06 | 17:18 | Albert Zondervan | 205522 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.31 | 0.01% | 28 | 5 | 3 | ... | 0.00296 | 9e-05 |  |  |  |  | 0.99521 | 0.99521 | Yes | Yes | [5 empty] | 0.00296 | 0.00‰ | 9e-05 | 0.00‰ | No | 1 | 1 | 0.00095 | 0.00095 | 0.00308 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00296 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-10-04 | Cathy Ginnane | 15:19 |  | Completed |  |  |  |  | 71006 | 2000000 |  |  | 71018 71019 71020 71039 71040 71041 | 71018~71019~71020~71039~71040~71041 | 3085 |  | 3085 3085 3085 3085 3085 3085 | 2016-10-18 | 9:33 | Albert Zondervan |  | 0.32 mg |  | 0.32 | [5 empty] | 0.99835 | 0.00229 | 71018 |  | XCAMS | ST | 0.16% |  | * wheel with also small (SM+TY) samples * good measurement with expected response towards smaller sample sizes * small system error hi all Jocelyn and I have refined my earlier blank-subtraction of TW3085 (ACE): when sample weight <= 0.3 mgC, then the sample is a small and must receive the Santos-method blank correction using small blanks and OxIs in the same wheel; when sample weight > 0.3 mgC, then we apply the normal procedure, spanning several months of recent measurements. Two groups were identified to blank-correct: inorganics: 6x unknowns (shell, foram) + 4x secondaries (C2-travertine) + 5x blanks (C1-marble) organics: 8x unknowns (soil, sediment, plant) + 3x secondaries (FIRI-I) + 3x blanks (KauriRRd-wood) Since we have a mix of smalls and larges in this wheel, both methods are applied. So, apart from sample-type grouping there is mass grouping, thus making it a 2 x 2 = 4 groups exercise. Instead of the regular blank-subtraction procedure for large samples, the mass-dependency in the OxIs and blanks in this wheel have been used to derive the modern and dead contaminants, m and d, resp. (We discussed this in the last team meeting.) The method is described in Guaciara Santos et al. [NIM B 259 (2007) 293-302]. No information from other wheels has been used. While doing this, the 'smalls correction procedure' and the associated Excel spreadsheet was revamped. However, it is not complete yet. One of the issues to sort is that the correction formulae should IMO apply to the parameter ratio-to-standard, not fraction-modern. The other major challenge is to make the new procedure more transparent and simple to use. I will try to continue with that until I have it ready such that any operator can do this task. With the newly developed Excel spreadsheet, I determined the parameters for the contaminants: dead-C contaminant: d = 0.00167 +- 0.00015 mgC, derived from the OxIs modern-C contaminant: m = 0.00084 +- 0.000014 mgC, derived from the KauriRRd modern-C contaminant: m = 0.00092 +- 0.000006 mgC, derived from the C1-marble The values for m and d were used for the smalls, but not their uncertainties put out by the regresssion. Because of lack of statistics on repeatability amongst small blanks, we must err on the side of caution. For that reason, we set the undertainty on m and d equal to the m and d values themselves. The small OxIs were nicely corrected to the average value for the large OxIs. With the corrections made, the secondaries look a lot more consistent with their expectation values. This set is ready for a final check prior reporting. cheers Albert |  | RRL | [5 empty] | Background Inorganic |  | 2016-10-07 | Carbonate |  | JOB-205522 | JCT ADDED INVOICE 540GCT21 | 0.27mg SLUSHIE small std to match levy sample, aim for 0.25mg C | 61911 |  | 14047/1 |  |  | 2016-09-19 | 2758 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0001 | 0.00231 | 71019 | [26 empty] | 2016-09-21 | 11016 | Graphite | Completed
[168 empty] | 1.0012 | 0.00233 | 71020 | [26 empty] | 2016-10-18 | 12730 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00234 | 71039 | [26 empty] | 2016-10-18 | 14998 | Close Job | Completed
[168 empty] | 1.0032 | 0.00234 | 71040 | [26 empty] | 2017-03-27 | 5000 | Invoice | Completed
[168 empty] | 0.99656 | 0.0023 | 71041
[7 empty] | 97.8 | 88.4872 | 0.0017 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36761 | 36760 | 36761 | 146 | 146 | 148 | 36615.5 | 36908.4 | margaret norris | -989.71 | 0.19 ‰ | 2016-11-25 | 2016-11-25 | 35 | 0 |  |  |  |  | 3.16‰ | 3.16‰ | 2.00‰ | 2.00‰ | 3.16 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.75593 | -989.71 ‰ | -989.71 ‰ | 0.19 | 0.19 | -989.58 | 0.19 ‰ | -989.58 | 0.19 ‰ |  |  | 2016-11-28 | Jocelyn Turnbull | 13:03 | 0.010294 | 0.010294 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.7 |  |  | 0.01029 | 0.00019 | 12943 | 2016-11-28 | 8:51 | albert zondervan | 205635 |  |  |  |  | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 1.03 | 0.02% | 11 | 5 | 6 | A.. | 0.0099 | 0.00018 |  |  |  |  | 0.99749 | 0.99749 | Yes | Yes |  | 35 days old wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 147.5403229275772136. If Saved values show old error bar then the revised result has not been re-reported to the client. |  |  |  | 0.0099 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01029 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0099 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-11-21 | margaret norris | 8:31 |  | Completed |  |  |  |  | 71187 | 2000000 |  |  | 71234 71246 71253 71254 71269 71314 71320 | 71234~71246~71253~71254~71269~71314~71320 | 3091 |  | 3091 3091 3091 3091 3091 3091 3091 | 2016-11-28 | 9:32 | albert zondervan |  | 0.74 mg | 0.13 | 0.74 | [5 empty] | 1.0003 | 0.00197 | 71234 |  | XCAMS | EA | 0% |  | TW3091 measurement: No tuning required as we've come off back of TW3090. error adjustments: * factor 2.7 for run-to-run variability of plant TP71277 * factor 1.3 for run-to-run variability of the small sediment TP71313 * no other adjustments warranted unknowns: * TP71277 generated an average rts that has been derived from a far too variable set of individual runs. I do not trust this average to be correct and therefore I call it a fail. * None of the other cathodes in the wheel show such a behaviour. * Small sediment TP71313 was blank-corrected along with all larges. primaries: The wheel contained 7x cathodes with OxI: all via EA. All of them performed equally well, so I changed the category for TP71269 to primary standard. blanks: Two groups have been identified in this wheel to blank-correct: * organics via AAA: 26x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied * organics via CellExtr: 0x unknowns + 1x secondaries + 0x blanks; a 360-days window was applied to find a decent number of blanks to go with it secondaries: * Both secondaries (FIRI-I, D) agree with their expectation value. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-12-09 | Water |  | JOB-205635 |  | TEST Water blank test. Flush with nitrogen before added sample as powder, C13 readings | 62189 | Normal | 14047/11 |  |  | 2016-10-20 | 2764 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.00184 | 71246 | [26 empty] | 2016-10-20 | 1637 | Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00188 | 71253 | [26 empty] | 2016-10-20 | 3077 | Choose Gas Splitting | Completed
[168 empty] | 0.99791 | 0.00192 | 71254 | [26 empty] | 2016-10-21 | 11197 | Graphite | Completed
[168 empty] | 1.0015 | 0.00193 | 71269 | [26 empty] | 2016-11-01 | 4777 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00185 | 71314 | [26 empty] | 2016-11-28 | 12943 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00188 | 71320 | [26 empty] | 2016-12-05 | 15339 | Close Job | Completed
[7 empty] | 98.1 | 88.1866 | 0.001 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41453 | 41453 | 41453 | 189 | 189 | 190 | 41265.9 | 41644.9 | margaret norris | -994.26 | 0.14 ‰ | 2016-11-25 | 2016-11-25 | 35 | 0 |  |  |  |  | 6.58‰ | 6.58‰ | 2.00‰ | 2.00‰ | 6.58 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0 | 0.75593 | -994.26 ‰ | -994.26 ‰ | 0.14 | 0.14 | -994.19 | 0.14 ‰ | -994.19 | 0.14 ‰ |  |  | 2016-11-28 | Jocelyn Turnbull | 13:03 | 0.0057397 | 0.0057397 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.7 |  |  | 0.00574 | 0.00014 | 12944 | 2016-11-28 | 8:51 | albert zondervan | 205636 |  |  |  |  | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 31 | 5 | 5 | A.. | 0.00552 | 0.00013 |  |  |  |  | 0.99749 | 0.99749 | Yes | Yes |  | 35 days old wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 190.3323755000963192. If Saved values show old error bar then the revised result has not been re-reported to the client. |  |  |  | 0.00552 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00574 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00552 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-11-21 | margaret norris | 8:31 |  | Completed |  |  |  |  | 71188 | 2000000 |  |  | 71234 71246 71253 71254 71269 71314 71320 | 71234~71246~71253~71254~71269~71314~71320 | 3091 |  | 3091 3091 3091 3091 3091 3091 3091 | 2016-11-28 | 9:32 | albert zondervan |  | 1.24 mg | 0.13 | 1.24 | [5 empty] | 1.0003 | 0.00197 | 71234 |  | XCAMS | EA | 0% |  | TW3091 measurement: No tuning required as we've come off back of TW3090. error adjustments: * factor 2.7 for run-to-run variability of plant TP71277 * factor 1.3 for run-to-run variability of the small sediment TP71313 * no other adjustments warranted unknowns: * TP71277 generated an average rts that has been derived from a far too variable set of individual runs. I do not trust this average to be correct and therefore I call it a fail. * None of the other cathodes in the wheel show such a behaviour. * Small sediment TP71313 was blank-corrected along with all larges. primaries: The wheel contained 7x cathodes with OxI: all via EA. All of them performed equally well, so I changed the category for TP71269 to primary standard. blanks: Two groups have been identified in this wheel to blank-correct: * organics via AAA: 26x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied * organics via CellExtr: 0x unknowns + 1x secondaries + 0x blanks; a 360-days window was applied to find a decent number of blanks to go with it secondaries: * Both secondaries (FIRI-I, D) agree with their expectation value. |  | RRL | [5 empty] | Unknown Inorganic |  | 2016-12-09 | Water | Complete | JOB-205636 |  | TEST Water blank test. Flush with nitrogen before added sample as pieces , C13 readings | 62190 | Normal | 14047/11 |  |  | 2016-10-20 | 2765 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.00184 | 71246 | [26 empty] | 2016-10-20 | 1638 | Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00188 | 71253 | [26 empty] | 2016-10-20 | 3079 | Choose Gas Splitting | Completed
[168 empty] | 0.99791 | 0.00192 | 71254 | [26 empty] | 2016-10-21 | 11198 | Graphite | Completed
[168 empty] | 1.0015 | 0.00193 | 71269 | [26 empty] | 2016-11-28 | 12944 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.00185 | 71314 | [26 empty] | 2016-12-05 | 15340 | Close Job | Completed
[168 empty] | 1.0019 | 0.00188 | 71320
[7 empty] | 99.8 | 88.6789 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37836 | 37835 | 37836 | 111 | 111 | 111 |  |  | Margaret Norris | -991 | 0.12 ‰ | 2016-12-06 | 2016-12-06 | 11 | 0 |  |  |  |  | 0.91‰ | 0.91‰ | 2.00‰ | 2.00‰ | 0.91 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -991.00 ‰ | -991.00 ‰ | 0.12 | 0.12 | -990.88 | 0.13 ‰ | -990.88 | 0.13 ‰ |  |  | 2016-12-08 | Jocelyn Turnbull | 11:30 | 0.0090046 | 0.0090046 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.7 |  |  | 0.009 | 0.00012 | 13036 | 2016-12-08 | 8:33 | Albert Zondervan | 205820 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 37 | 5 | 6 | ... | 0.00866 | 0.00012 |  |  |  |  | 0.99741 | 0.99741 | Yes | Yes | [5 empty] | 0.00866 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.0005 | 0.0005 | 0.009 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00866 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-12-05 | Margaret Norris | 9:05 |  | Completed |  |  |  |  | 71386 | 2000000 |  |  | 71321 71359 71360 71368 71369 71455 | 71321~71359~71360~71368~71369~71455 | 3093 |  | 3093 3093 3093 3093 3093 3093 | 2016-12-08 | 9:48 | Albert Zondervan |  | 0.79 mg |  | 0.79 | [5 empty] | 0.99847 | 0.00121 | 71321 |  | XCAMS | EA | 0% |  | TW3093 measurement: * no issues CalAMS error adjustments: * none primaries: * The wheel contained 6x cathodes with OxI: all via EA. Three groups have been identified in this wheel to blank-correct: * ANSTO cellulose extraction: 2x unknowns + 1x secondaries + 1x blanks a 30-days window was applied leaving only the one blank to blank-subtract with 1x FIRI-D: result is 2.5-sigma below consensus value * other organics: a 180-days window was applied AAA pretreated: 7x unknowns + 0x secondaries + 1x blanks untreated: 17x unknowns + 1x secondaries + 0x blanks 1x FIRI-I: result is consistent with consensus value * waters: 1x unknowns + 0x secondaries + 1x blanks unknowns: * no issues | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-01-05 | Water |  | JOB-205820 |  | replaces sample lost in graphite | 62268 | Normal | 14047/11 |  |  | 2016-11-24 | 2768 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00127 | 71359 | [26 empty] | 2016-11-24 | 1645 | Gas Splitting | Completed
[168 empty] | 0.99883 | 0.0012 | 71360 | [26 empty] | 2016-11-24 | 3096 | Choose Gas Splitting | Completed
[168 empty] | 0.99959 | 0.00121 | 71368 | [26 empty] | 2016-11-25 | 11396 | Graphite | Completed
[168 empty] | 1.0013 | 0.00123 | 71369 | [26 empty] | 2016-12-09 | 4808 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00123 | 71455 | [26 empty] | 2016-12-08 | 13036 | AMS Submission Results | Completed
[197 empty] | 2016-12-16 | 15541 | Close Job | Completed
[7 empty] | 113.3 | 88.3371 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47563 | 47562 | 47563 | 218 | 217 | 218 |  |  | Margaret Norris | -997.32 | 0.07 ‰ | 2016-12-12 | 2016-12-12 | 11 | 0 |  |  |  |  | 4.00‰ | 4.00‰ | 2.00‰ | 2.00‰ | 4.00 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -997.32 ‰ | -997.32 ‰ | 0.07 | 0.07 | -997.28 | 0.07 ‰ | -997.28 | 0.07 ‰ |  |  | 2016-12-20 | Jocelyn Turnbull | 11:55 | 0.0026827 | 0.0026827 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.00268 | 7e-05 | 13080 | 2016-12-14 | 8:33 | Albert Zondervan | 205887 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.27 | 0.01% | 32 | 5 | 5 | ... | 0.00258 | 7e-05 |  |  |  |  | 0.99663 | 0.99663 | Yes | Yes | [5 empty] | 0.00258 | 0.00‰ | 7e-05 | 0.00‰ | No | 1 | 1 | 0.00081 | 0.00081 | 0.00268 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00258 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-12-08 | Margaret Norris | 15:49 |  | Completed |  |  |  |  | 71443 | 2000000 |  |  | 71477 71478 71479 71491 71492 71493 | 71477~71478~71479~71491~71492~71493 | 3094 |  | 3094 3094 3094 3094 3094 3094 | 2016-12-14 | 16:48 | Albert Zondervan |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0007 | 0.00191 | 71477 |  | XCAMS | EA | 0.1% |  | TW3094 measurement: * no issues with XCAMS CalAMS error adjustments: * factor 1.03 global uncertainty multiplier, due to excess run-to-run variability * 0.10% system error primaries: * The wheel contained 6x cathodes with OxI: all via EA. Four groups have been identified in this wheel to blank-correct: * carbonates on carbonate line: 8x unknowns + 1x secondaries + 1x blanks a 180-days window was applied 1x IAEA-C2 (travertine): result is almost 3-sigma below consensus value * woods via cellulose extraction & ANSTO: 8x unknowns + 2x secondaries + 2x blanks a 180-days window was applied FIRI-D: 1x result is 2-sigma below consensus value 1x result is almost 3-sigma below consensus value * other organics: a 180-days window was applied AAA pretreated: 2x unknowns + 0x secondaries + 0x blanks * woods via cellulose extraction & Rafter a 180-days window was applied 2x unknowns + 1x secondaries + 0x blanks FIRI-D: result is almost 3-sigma below consensus value furthermore in this wheel * 1x test of the waterline: result is similar to that of any processing blank, much better than the waterline blanks we have been getting up to now * 3x ramped-pyrolysis samples: Their measurement has failed because of too low carbon beam intensity and 14C countrate. One can't call these smalls anymore, as they are too small to generate any meaningful result. |  | RRL | [5 empty] | Background Inorganic |  | 2016-12-12 | Carbonate |  | JOB-205887 |  |  | 62308 |  | 14047/1 |  |  | 2016-11-28 | 2769 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00197 | 71478 | [26 empty] | 2016-11-28 | 1647 | Gas Splitting | Completed
[168 empty] | 0.99848 | 0.00206 | 71479 | [26 empty] | 2016-11-28 | 3102 | Choose Gas Splitting | Completed
[168 empty] | 0.99716 | 0.0019 | 71491 | [26 empty] | 2016-12-02 | 11453 | Graphite | Completed
[168 empty] | 1.0007 | 0.00201 | 71492 | [26 empty] | 2016-12-20 | 13080 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00203 | 71493 | [26 empty] | 2016-12-20 | 15584 | Close Job | Completed
[7 empty] | 95.3 | 88.3985 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42010 | 42010 | 42010 | 187 | 187 | 187 |  |  | margaret norris | -994.65 | 0.12 ‰ | 2017-02-10 | 2017-02-10 | 22 | 0 |  |  |  |  | 3.59‰ | 3.59‰ | 2.00‰ | 2.00‰ | 3.59 ‰ | 2.00 ‰ | AMS | 2.27‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -994.65 ‰ | -994.65 ‰ | 0.12 | 0.12 | -994.58 | 0.13 ‰ | -994.58 | 0.13 ‰ |  |  | 2017-02-13 | Jocelyn Turnbull | 11:07 | 0.0053549 | 0.0053549 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.7 |  |  | 0.00535 | 0.00012 | 13133 | 2017-02-13 | 8:17 | Albert Zondervan | 206002 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 10 | 5 | 6 | ... | 0.00515 | 0.00012 |  |  |  |  | 0.99692 | 0.99692 | Yes | Yes | [5 empty] | 0.00515 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00535 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00515 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-02 | margaret norris | 10:52 |  | Completed |  |  |  |  | 71554 | 2000000 |  |  | 71511 71512 71513 71595 71596 71597 | 71511~71512~71513~71595~71596~71597 | 3096 |  | 3096 3096 3096 3096 3096 3096 | 2017-02-13 | 10:49 | Albert Zondervan |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99691 | 0.00182 | 71511 |  | XCAMS | ST | 0% |  | TW3096 measurement: * first wheel of year 2017 AD * unknowns are all waters * large samples only * XCAMS worked fine, except for anomalously low Cs and ionizer settings to obtain normal beam intensity CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: cat#14: factor 1.3 * 0% system error primaries: * The wheel contained 6x cathodes with large OxI, combusted via ST. One group of unknowns is identified in this wheel to blank-correct: * waters: 28x unknowns + 1x secondaries (IAEA-C2, R14047/12) + 1x blanks (IAEA-C1, R14047/11) 4x blanks were found for last 180 days IAEA-C2 comes out consistent with consensus value | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-02-10 | Water | Complete | JOB-206002 |  |  | 62376 | Normal | 14047/11 |  |  | 2017-01-17 | 2789 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99979 | 0.00243 | 71512 | [26 empty] | 2017-01-17 | 1659 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00187 | 71513 | [26 empty] | 2017-01-17 | 3121 | Choose Gas Splitting | Completed
[168 empty] | 1.0025 | 0.00172 | 71595 | [26 empty] | 2017-01-19 | 11564 | Graphite | Completed
[168 empty] | 1 | 0.00172 | 71596 | [26 empty] | 2017-01-26 | 4830 | C13 Measurement | Completed
[168 empty] | 0.99951 | 0.00171 | 71597 | [26 empty] | 2017-02-13 | 13133 | AMS Submission Results | Completed
[197 empty] | 2017-02-13 | 15677 | Close Job | Completed
[5 empty] | 6.2353e-05 | 0.0071413 | 94.5 | 89.4886 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7129 | 7129 | 7129 | 46 | 46 | 49 | 7083.3 | 7175.4 | margaret norris | -588.31 | 2.36 ‰ | 2017-02-10 | 2016-08-14 | 23 | 180 |  |  |  |  | -8.63‰ | -8.63‰ | 2.00‰ | 2.00‰ | -8.63 ‰ | 2.00 ‰ | AMS | -8.32‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -588.31 ‰ | -588.31 ‰ | 2.36 | 2.54 | -583.23 | 2.39 ‰ | -583.23 | 2.39 ‰ |  |  | 2017-02-13 | Jocelyn Turnbull | 11:07 | 0.41169 | 0.41169 | 0.00236 | 0.00254 | 0.0023 | 0.0023 | 2.1 |  |  | 0.416 | 0.00131 | 13134 | 2017-02-13 | 8:17 | Albert Zondervan | 206003 | 0.00687 | 0.0069 | 0.00309 | 0.0031 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.17 | 0.24% | 30 | 5 | 6 | ... | 0.40008 | 0.00115 |  |  |  |  | 0.99692 | 0.99692 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 49.4642584523931916. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00687 | 0.00687 | 0.00148 | 0.39593 | 0.40‰ | 0.00221 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.41169 | 0 | 0.00236 | 0 | 0 | 0.00687 | 0 | 0.00309 | 0 | 0.39593 | 0 | 0.00221 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-02 | margaret norris | 10:52 |  | Completed |  |  |  |  | 71544 | 2000000 | 71554 71386 70935 70748 | 71554~71386~70935~70748 | 71511 71512 71513 71595 71596 71597 | 71511~71512~71513~71595~71596~71597 | 3096 | 3096 3093 3082 3079 | 3096 3096 3096 3096 3096 3096 | 2017-02-13 | 10:49 | Albert Zondervan |  | 0.98 mg | 0.13 | 0.98 | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 0.99691 | 0.00182 | 71511 | EA | XCAMS | ST | 0% |  | TW3096 measurement: * first wheel of year 2017 AD * unknowns are all waters * large samples only * XCAMS worked fine, except for anomalously low Cs and ionizer settings to obtain normal beam intensity CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: cat#14: factor 1.3 * 0% system error primaries: * The wheel contained 6x cathodes with large OxI, combusted via ST. One group of unknowns is identified in this wheel to blank-correct: * waters: 28x unknowns + 1x secondaries (IAEA-C2, R14047/12) + 1x blanks (IAEA-C1, R14047/11) 4x blanks were found for last 180 days IAEA-C2 comes out consistent with consensus value | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-02-10 | Water | Complete | JOB-206003 |  |  | 62377 | Normal | 14047/12 |  |  | 2017-01-17 | 2790 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 | 0.99979 | 0.00243 | 71512 | ST | [25 empty] | 2017-01-17 | 1657 | Gas Splitting | Completed
[163 empty] | 2016-09-20 | 0.00638 | 70935 | 3082 | 0.94 | 1.0009 | 0.00187 | 71513 | EA | [25 empty] | 2017-01-17 | 3118 | Choose Gas Splitting | Completed
[163 empty] | 2016-08-26 | 0.00728 | 70748 | 3079 | 0.87 | 1.0025 | 0.00172 | 71595 | ST | [25 empty] | 2017-01-18 | 11554 | Graphite | Completed
[168 empty] | 1 | 0.00172 | 71596 | [26 empty] | 2017-01-26 | 4831 | C13 Measurement | Completed
[168 empty] | 0.99951 | 0.00171 | 71597 | [26 empty] | 2017-02-13 | 13134 | AMS Submission Results | Completed
[197 empty] | 2017-02-13 | 15678 | Close Job | Completed
[7 empty] | 95.1 | 88.5976 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41780 | 41779 | 41780 | 197 | 197 | 197 |  |  | Margaret Norris | -994.49 | 0.14 ‰ | 2017-02-21 | 2017-02-21 | 6 | 0 |  |  |  |  | 2.58‰ | 2.58‰ | 2.00‰ | 2.00‰ | 2.58 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -994.49 ‰ | -994.49 ‰ | 0.14 | 0.14 | -994.42 | 0.14 ‰ | -994.42 | 0.14 ‰ |  |  | 2017-02-27 | Jocelyn Turnbull | 15:31 | 0.0055109 | 0.0055109 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 3 |  |  | 0.00551 | 0.00014 | 13220 | 2017-02-22 | 15:26 | Albert Zondervan | 206056 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 13 | 5 | 6 | ... | 0.0053 | 0.00013 |  |  |  |  | 0.99815 | 0.99815 | Yes | Yes | [5 empty] | 0.0053 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00551 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0053 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-20 | Margaret Norris | 8:38 |  | Completed |  |  |  |  | 71702 | 2000000 |  |  | 71580 71581 71601 71602 71615 71680 | 71580~71581~71601~71602~71615~71680 | 3098 |  | 3098 3098 3098 3098 3098 3098 | 2017-02-23 | 13:10 | Albert Zondervan |  | 0.69 mg |  | 0.69 | [5 empty] | 1.0026 | 0.00164 | 71580 |  | XCAMS | EA | 0% |  | TW3098 measurement: * no issues CalAMS error adjustments: * run-to-run variability of whole dataset is ok * factor 1.3 to account for run-to-run variability of cathodes TP71601 and 71599 * 0% system error primaries: * The wheel contained 6x cathodes with OxI: all via EA. secondaries: * 1x IAEA-C2: result consistent with expectation value * 1x FIRI-I: result consistent with expectation value blanks: * Four groups have been identified to blank-correct: - waters: 6x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - carbonate: 2x unknowns + 0x secondaries + 0x blanks; a 180-days window was applied - organics via AAA and other non-CE: 14x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - organics via CE-GNSmethod: 5x unknowns + 0x secondaries + 1x blanks; a 180-days window was applied unknowns: * no issues | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-02-15 | Water |  | JOB-206056 |  |  | 62483 | Normal | 14047/11 | HAT |  | 2017-02-13 | 2831 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99774 | 0.00166 | 71581 | [26 empty] | 2017-02-13 | 1678 | Gas Splitting | Completed
[168 empty] | 0.9998 | 0.00217 | 71601 | [26 empty] | 2017-02-13 | 3148 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00169 | 71602 | [26 empty] | 2017-02-15 | 11712 | Graphite | Completed
[168 empty] | 1.0005 | 0.00169 | 71615 | [26 empty] | 2017-02-21 | 4871 | C13 Measurement | Completed
[168 empty] | 0.99868 | 0.00169 | 71680 | [26 empty] | 2017-02-27 | 13220 | AMS Submission Results | Completed
[197 empty] | 2017-03-01 | 15839 | Close Job | Completed
[5 empty] | 5.6619e-05 | 0.0068552 | 100.9 | 89.5144 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7162 | 7162 | 7162 | 44 | 44 | 48 | 7118.3 | 7206.3 | Margaret Norris | -590 | 2.25 ‰ | 2017-02-21 | 2016-08-25 | 6 | 180 |  |  |  |  | -7.69‰ | -7.69‰ | 2.00‰ | 2.00‰ | -7.69 ‰ | 2.00 ‰ | AMS | -8.20‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -590.00 ‰ | -590.00 ‰ | 2.25 | 2.43 | -584.94 | 2.28 ‰ | -584.94 | 2.28 ‰ |  |  | 2017-02-27 | Jocelyn Turnbull | 15:31 | 0.41 | 0.41 | 0.00225 | 0.00243 | 0.00218 | 0.00218 | 2.3 |  |  | 0.41413 | 0.00124 | 13221 | 2017-02-22 | 15:26 | Albert Zondervan | 206057 | 0.00655 | 0.0066 | 0.00295 | 0.003 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41 | 0.22% | 33 | 5 | 6 | ... | 0.39828 | 0.00107 |  |  |  |  | 0.99815 | 0.99815 | Yes | Yes |  | results is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 47.6077817668362208. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00655 | 0.00659 | 0.00146 | 0.39431 | 0.39‰ | 0.0021 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.41 | 0 | 0.00225 | 0 | 0 | 0.00655 | 0 | 0.00295 | 0 | 0.39431 | 0 | 0.0021 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-20 | Margaret Norris | 8:38 |  | Completed |  |  |  |  | 71703 | 2000000 | 71702 71554 71386 70935 70748 | 71702~71554~71386~70935~70748 | 71580 71581 71601 71602 71615 71680 | 71580~71581~71601~71602~71615~71680 | 3098 | 3098 3096 3093 3082 3079 | 3098 3098 3098 3098 3098 3098 | 2017-02-23 | 13:10 | Albert Zondervan |  | 0.90 mg | 0.13 | 0.9 | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0026 | 0.00164 | 71580 | EA | XCAMS | EA | 0% |  | TW3098 measurement: * no issues CalAMS error adjustments: * run-to-run variability of whole dataset is ok * factor 1.3 to account for run-to-run variability of cathodes TP71601 and 71599 * 0% system error primaries: * The wheel contained 6x cathodes with OxI: all via EA. secondaries: * 1x IAEA-C2: result consistent with expectation value * 1x FIRI-I: result consistent with expectation value blanks: * Four groups have been identified to blank-correct: - waters: 6x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - carbonate: 2x unknowns + 0x secondaries + 0x blanks; a 180-days window was applied - organics via AAA and other non-CE: 14x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - organics via CE-GNSmethod: 5x unknowns + 0x secondaries + 1x blanks; a 180-days window was applied unknowns: * no issues | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-02-15 | Water |  | JOB-206057 |  |  | 62484 | Normal | 14047/12 | HAT |  | 2017-02-13 | 2832 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 0.99774 | 0.00166 | 71581 | ST | [25 empty] | 2017-02-13 | 1679 | Gas Splitting | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 | 0.9998 | 0.00217 | 71601 | EA | [25 empty] | 2017-02-13 | 3150 | Choose Gas Splitting | Completed
[163 empty] | 2016-09-20 | 0.00638 | 70935 | 3082 | 0.94 | 1.0006 | 0.00169 | 71602 | ST | [25 empty] | 2017-02-15 | 11713 | Graphite | Completed
[163 empty] | 2016-08-26 | 0.00728 | 70748 | 3079 | 0.87 | 1.0005 | 0.00169 | 71615 | EA | [25 empty] | 2017-02-21 | 4872 | C13 Measurement | Completed
[168 empty] | 0.99868 | 0.00169 | 71680 | [26 empty] | 2017-02-27 | 13221 | AMS Submission Results | Completed
[197 empty] | 2017-03-01 | 15840 | Close Job | Completed
[7 empty] | 91.3 | 88.5952 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39352 | 39352 | 39352 | 156 | 156 | 156 |  |  | jenny dahl | -992.54 | 0.15 ‰ | 2017-03-17 | 2017-03-17 | 21 | 0 |  |  |  |  | 2.71‰ | 2.71‰ | 2.00‰ | 2.00‰ | 2.71 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -992.54 ‰ | -992.54 ‰ | 0.15 | 0.15 | -992.45 | 0.15 ‰ | -992.45 | 0.15 ‰ |  |  | 2017-03-20 | Jocelyn Turnbull | 13:02 | 0.0074553 | 0.0074553 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.2 |  |  | 0.00746 | 0.00015 | 13411 | 2017-03-20 | 8:36 | Albert Zondervan | 206559 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.75 | 0.01% | 15 | 5 | 6 | .X. | 0.00717 | 0.00014 |  |  |  |  | 0.99826 | 0.99826 | Yes | Yes | [5 empty] | 0.00717 | 0.01‰ | 0.00014 | 0.00‰ | No | 1 | 1 | 0.00087 | 0.00087 | 0.00746 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00717 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-14 | jenny dahl | 17:44 |  | Completed |  |  |  |  | 71759 | 2000000 |  |  | 71747 71762 71802 71840 71843 71844 | 71747~71762~71802~71840~71843~71844 | 3103 |  | 3103 3103 3103 3103 3103 3103 | 2017-03-20 | 9:37 | Albert Zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99774 | 0.00213 | 71747 |  | XCAMS | EA | 0.14% |  | TW3103 measurement: * virTW=ASE1 * no ST * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.14% system error primaries: * 6x large OxI groups identified to blank-correct: * water: 3x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 5x blanks IAEA-C2 (TP71760) is deviating >3-sigma from consensus value - TEST SAMPLE USING POWDER MIXED IN WITH WATER - NOT A WORRY RIGHT NOW * organic via ANSTO-method cellulose extraction: 12x unknowns + 2x secondaries + 2x blanks a 180-day window was applied, capturing 5x blanks both FIRI-D (TP71749,71798) are deviating from consensus value, towards lower F NOTE WE HAVE CHECKED THE CONSENSUS VALUE AND IT IS 4532 FOR AMS LABS WHICH MAKES THESE MUCH CLOSER TO THE EXPECTED VALUE * organic via AAA: 9x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 11x blanks FIRI-I (TP71838) is 1-sigma away from consensus value | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-04-18 | Water |  | JOB-206559 | JCT ADDED INVOICE 540GCT21 | test with marble in DI water | 62622 | Normal | 14047/11 |  |  | 2017-02-22 | 2838 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00212 | 71762 | [26 empty] | 2017-02-22 | 1686 | Gas Splitting | Completed
[168 empty] | 0.99878 | 0.00213 | 71802 | [26 empty] | 2017-02-22 | 3161 | Choose Gas Splitting | Completed
[168 empty] | 0.99881 | 0.00213 | 71840 | [26 empty] | 2017-02-24 | 11769 | Graphite | Completed
[168 empty] | 1.0038 | 0.00217 | 71843 | [26 empty] | 2017-02-27 | 4897 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00216 | 71844 | [26 empty] | 2017-03-20 | 13411 | AMS Submission Results | Completed
[197 empty] | 2017-03-27 | 16035 | Close Job | Completed
[197 empty] | 2017-03-27 | 5024 | Invoice | Completed
[5 empty] | 5.8237e-05 | 0.0067954 | 80 | 89.7392 | 0.0698 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 6984 | 6984 | 6984 | 43 | 42 | 46 | 6941.6 | 7027.2 | jenny dahl | -580.82 | 2.24 ‰ | 2017-03-17 | 2016-09-18 | 21 | 180 |  |  |  |  | -10.08‰ | -10.08‰ | 2.00‰ | 2.00‰ | -10.08 ‰ | 2.00 ‰ | AMS | -8.26‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -580.82 ‰ | -580.82 ‰ | 2.24 | 2.43 | -575.64 | 2.26 ‰ | -575.64 | 2.26 ‰ |  |  | 2017-03-20 | Jocelyn Turnbull | 13:02 | 0.41918 | 0.41918 | 0.00224 | 0.00243 | 0.00217 | 0.00217 | 4 |  |  | 0.42324 | 0.00126 | 13412 | 2017-03-20 | 8:36 | Albert Zondervan | 206560 | 0.00653 | 0.0065 | 0.00294 | 0.0029 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.92 | 0.22% | 29 | 5 | 4 | T.. | 0.40704 | 0.00123 |  |  |  |  | 0.99826 | 0.99826 | Yes | Yes |  | result is deviating by >3-sigma from expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 46.4995690831103443. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00653 | 0.00653 | 0.00145 | 0.40314 | 0.40‰ | 0.00216 | 0.00‰ | No | 1 | 1 | 0.00087 | 0.00087 | 0.41918 | 0 | 0.00224 | 0 | 0 | 0.00653 | 0 | 0.00294 | 0 | 0.40314 | 0 | 0.00216 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-14 | jenny dahl | 17:44 |  | Completed |  |  |  |  | 71760 | 2000000 | 71759 71702 71554 71386 70935 | 71759~71702~71554~71386~70935 | 71747 71762 71802 71840 71843 71844 | 71747~71762~71802~71840~71843~71844 | 3103 | 3103 3098 3096 3093 3082 | 3103 3103 3103 3103 3103 3103 | 2017-03-20 | 9:37 | Albert Zondervan |  | 0.51 mg | 0.13 | 0.51 | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 0.99774 | 0.00213 | 71747 | ST | XCAMS | EA | 0.14% |  | TW3103 measurement: * virTW=ASE1 * no ST * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.14% system error primaries: * 6x large OxI groups identified to blank-correct: * water: 3x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 5x blanks IAEA-C2 (TP71760) is deviating >3-sigma from consensus value - TEST SAMPLE USING POWDER MIXED IN WITH WATER - NOT A WORRY RIGHT NOW * organic via ANSTO-method cellulose extraction: 12x unknowns + 2x secondaries + 2x blanks a 180-day window was applied, capturing 5x blanks both FIRI-D (TP71749,71798) are deviating from consensus value, towards lower F NOTE WE HAVE CHECKED THE CONSENSUS VALUE AND IT IS 4532 FOR AMS LABS WHICH MAKES THESE MUCH CLOSER TO THE EXPECTED VALUE * organic via AAA: 9x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 11x blanks FIRI-I (TP71838) is 1-sigma away from consensus value | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-04-18 | Water |  | JOB-206560 | JCT ADDED INVOICE 540GCT21 | test with powder suspended in D.I. water. JCT - this is a test that looks like not such a good idea! Result is quite odd. | 62623 | Normal | 14047/12 |  |  | 2017-02-22 | 2839 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0003 | 0.00212 | 71762 | EA | [25 empty] | 2017-02-24 | 11770 | Graphite | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 0.99878 | 0.00213 | 71802 | ST | [25 empty] | 2017-02-27 | 4898 | C13 Measurement | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 | 0.99881 | 0.00213 | 71840 | EA | [25 empty] | 2017-03-20 | 13412 | AMS Submission Results | Completed
[163 empty] | 2016-09-20 | 0.00638 | 70935 | 3082 | 0.94 | 1.0038 | 0.00217 | 71843 | EA | [25 empty] | 2017-03-27 | 16034 | Close Job | Completed
[168 empty] | 1.0007 | 0.00216 | 71844 | [26 empty] | 2017-03-27 | 5025 | Invoice | Completed
[5 empty] | 5.6896e-05 | 0.0072282 | 93.6 | 89.2833 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7089 | 7088 | 7089 | 43 | 43 | 47 | 7045.7 | 7131.7 | jenny dahl | -586.22 | 2.22 ‰ | 2017-03-24 | 2016-09-25 | 9 | 180 |  |  |  |  | -5.10‰ | -5.10‰ | 2.00‰ | 2.00‰ | -5.10 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.89 | 0.89443 | -586.22 ‰ | -586.22 ‰ | 2.22 | 2.4 | -581.12 | 2.24 ‰ | -581.12 | 2.24 ‰ |  |  | 2017-03-27 | Jocelyn Turnbull | 10:06 | 0.41378 | 0.41378 | 0.00222 | 0.0024 | 0.00215 | 0.00215 | 2.2 |  |  | 0.41808 | 0.00104 | 13437 | 2017-03-27 | 8:02 | Albert Zondervan | 206597 | 0.00687 | 0.0069 | 0.00309 | 0.0031 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.38 | 0.22% | 6 | 5 | 6 | ... | 0.40208 | 0.0012 |  |  |  |  | 0.99818 | 0.99818 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 46.6697566445353733. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00687 | 0.00694 | 0.00159 | 0.39794 | 0.40‰ | 0.00223 | 0.00‰ | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.41378 | 0 | 0.00222 | 0 | 0 | 0.00687 | 0 | 0.00309 | 0 | 0.39794 | 0 | 0.00223 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-23 | jenny dahl | 16:30 |  | Completed |  |  |  |  | 71869 | 2000000 | 71907 71759 71702 71554 71386 | 71907~71759~71702~71554~71386 | 71885 71886 71935 71937 71938 | 71885~71886~71935~71937~71938 | 3104 | 3104 3103 3098 3096 3093 | 3104 3104 3104 3104 3104 | 2017-03-27 | 8:31 | Albert Zondervan |  | 0.92 mg | 0.13 | 0.92 | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 1.0024 | 0.00251 | 71885 | EA | XCAMS | ST | 0.21% |  | TW3104 measurement: * virTW=KID * no EA * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.21% system error primaries: * 6x large OxI, of which 1x was rejected as calibration standard: TP71936 is an outlier in 13/12 and rts. Cathy is querying this. groups identified to blank-correct: * bone + organic : 3x EXPRESS bone unknowns + 1x organic unknown + 2x secondaries + 1x blanks a 180-day window was applied FIRI-L whalebone and IAEA-C6 sucrose secondaries come out good * carbonates: 8x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good * water: 5x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good Some IRMS del13C values are still to be reported. | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-04-25 | Water |  | JOB-206597 |  |  | 62704 | Normal | 14047/12 |  |  | 2017-03-14 | 2848 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0028 | 0.00251 | 71886 | ST | [25 empty] | 2017-03-14 | 1696 | Gas Splitting | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 0.99735 | 0.0025 | 71935 | EA | [25 empty] | 2017-03-14 | 3179 | Choose Gas Splitting | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 0.99799 | 0.0025 | 71937 | EA | [25 empty] | 2017-03-15 | 11879 | Graphite | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 | 0.99948 | 0.00253 | 71938 | ST | [25 empty] | 2017-03-23 | 4919 | C13 Measurement | Completed
[197 empty] | 2017-03-27 | 13437 | AMS Submission Results | Completed
[197 empty] | 2017-03-29 | 16096 | Close Job | Completed
[7 empty] | 97.6 | 88.5292 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52474 | 52473 | 52474 | 286 | 286 | 286 |  |  | jenny dahl | -998.54 | 0.05 ‰ | 2017-03-24 | 2017-03-24 | 7 | 0 |  |  |  |  | 3.37‰ | 3.37‰ | 2.00‰ | 2.00‰ | 3.37 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0.89 | 0.89443 | -998.54 ‰ | -998.54 ‰ | 0.05 | 0.05 | -998.53 | 0.05 ‰ | -998.53 | 0.05 ‰ |  |  | 2017-03-27 | Jocelyn Turnbull | 10:05 | 0.0014557 | 0.0014557 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2 |  |  | 0.00146 | 5e-05 | 13438 | 2017-03-27 | 8:02 | Albert Zondervan | 206598 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 26 | 5 | 5 | ... | 0.0014 | 5e-05 |  |  |  |  | 0.99818 | 0.99818 | Yes | Yes |  | C13 is off between AMS and IRMS- confused because there is no C13 process for this sample. Not only that the C13 ought to be +2.4. The gas split from the parent job had a d13C of +3.37 so -8.00 is wrong. MN |  |  |  | 0.0014 | 0.00‰ | 5e-05 | 0.00‰ | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.00146 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.0014 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-23 | jenny dahl | 16:30 |  | Completed |  |  |  |  | 71898 | 2000000 |  |  | 71885 71886 71935 71937 71938 | 71885~71886~71935~71937~71938 | 3104 |  | 3104 3104 3104 3104 3104 | 2017-03-27 | 8:31 | Albert Zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0024 | 0.00251 | 71885 |  | XCAMS | ST | 0.21% |  | TW3104 measurement: * virTW=KID * no EA * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.21% system error primaries: * 6x large OxI, of which 1x was rejected as calibration standard: TP71936 is an outlier in 13/12 and rts. Cathy is querying this. groups identified to blank-correct: * bone + organic : 3x EXPRESS bone unknowns + 1x organic unknown + 2x secondaries + 1x blanks a 180-day window was applied FIRI-L whalebone and IAEA-C6 sucrose secondaries come out good * carbonates: 8x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good * water: 5x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good Some IRMS del13C values are still to be reported. |  | RRL | [5 empty] | Background Inorganic |  | 2017-03-21 | Carbonate | Not Complete | JOB-206598 |  | martial sourced from 206734 | 62702 |  | 14047/1 |  |  | 2017-03-16 | 2852 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0028 | 0.00251 | 71886 | [26 empty] | 2017-03-16 | 1700 | Gas Splitting | Completed
[168 empty] | 0.99735 | 0.0025 | 71935 | [26 empty] | 2017-03-16 | 3185 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.0025 | 71937 | [26 empty] | 2017-03-17 | 11908 | Graphite | Completed
[168 empty] | 0.99948 | 0.00253 | 71938 | [26 empty] | 2017-03-27 | 13438 | AMS Submission Results | Completed
[197 empty] | 2017-03-29 | 16088 | Close Job | Completed
[7 empty] | 98.9 | 88.5038 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38422 | 38422 | 38422 | 108 | 108 | 108 |  |  | jenny dahl | -991.63 | 0.11 ‰ | 2017-03-24 | 2017-03-24 | 3 | 0 |  |  |  |  | 3.66‰ | 3.66‰ | 2.00‰ | 2.00‰ | 3.66 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.89 | 0.89443 | -991.63 ‰ | -991.63 ‰ | 0.11 | 0.11 | -991.53 | 0.11 ‰ | -991.53 | 0.11 ‰ |  |  | 2017-03-27 | Jocelyn Turnbull | 10:06 | 0.0083703 | 0.0083703 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.5 |  |  | 0.00837 | 0.00011 | 13442 | 2017-03-27 | 8:02 | Albert Zondervan | 206721 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.84 | 0.01% | 17 | 5 | 6 | ... | 0.00805 | 0.00011 |  |  |  |  | 0.99818 | 0.99818 | Yes | Yes | [5 empty] | 0.00805 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.00837 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00805 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-23 | jenny dahl | 16:30 |  | Completed |  |  |  |  | 71907 | 2000000 |  |  | 71885 71886 71935 71937 71938 | 71885~71886~71935~71937~71938 | 3104 |  | 3104 3104 3104 3104 3104 | 2017-03-27 | 8:31 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0024 | 0.00251 | 71885 |  | XCAMS | ST | 0.21% |  | TW3104 measurement: * virTW=KID * no EA * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.21% system error primaries: * 6x large OxI, of which 1x was rejected as calibration standard: TP71936 is an outlier in 13/12 and rts. Cathy is querying this. groups identified to blank-correct: * bone + organic : 3x EXPRESS bone unknowns + 1x organic unknown + 2x secondaries + 1x blanks a 180-day window was applied FIRI-L whalebone and IAEA-C6 sucrose secondaries come out good * carbonates: 8x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good * water: 5x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good Some IRMS del13C values are still to be reported. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2017-05-09 | Water | Not Complete | JOB-206721 |  | to replace sample that was lost into a melted tube in graphite | 62703 | Normal | 14047/11 |  |  | 2017-03-17 | 2863 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0028 | 0.00251 | 71886 | [26 empty] | 2017-03-17 | 1701 | Gas Splitting | Completed
[168 empty] | 0.99735 | 0.0025 | 71935 | [26 empty] | 2017-03-17 | 3187 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.0025 | 71937 | [26 empty] | 2017-03-21 | 11917 | Graphite | Completed
[168 empty] | 0.99948 | 0.00253 | 71938 | [26 empty] | 2017-03-27 | 4943 | C13 Measurement | Completed
[197 empty] | 2017-03-27 | 13442 | AMS Submission Results | Completed
[197 empty] | 2017-03-29 | 16101 | Close Job | Completed
[7 empty] | 95.3 | 88.571 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52027 | 52027 | 52027 | 326 | 325 | 326 |  |  | margaret norris | -998.46 | 0.06 ‰ | 2017-04-06 | 2017-04-06 | 6 | 0 |  |  |  |  | 2.50‰ | 2.50‰ | 2.00‰ | 2.00‰ | 2.50 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -998.46 ‰ | -998.46 ‰ | 0.06 | 0.06 | -998.44 | 0.06 ‰ | -998.44 | 0.06 ‰ |  |  | 2017-04-11 | Cathy Ginnane | 9:14 | 0.0015389 | 0.0015389 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.7 |  |  | 0.00154 | 6e-05 | 13489 | 2017-04-10 | 11:27 | Albert Zondervan | 206819 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 10 | 5 | 7 | ... | 0.00148 | 6e-05 |  |  |  |  | 0.99778 | 0.99778 | Yes | Yes | [5 empty] | 0.00148 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00154 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00148 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-04-06 | margaret norris | 8:50 |  | Completed |  |  |  |  | 72000 | 2000000 |  |  | 71940 71941 71977 71978 72022 72024 | 71940~71941~71977~71978~72022~72024 | 3106 |  | 3106 3106 3106 3106 3106 3106 | 2017-04-10 | 12:02 | Albert Zondervan |  | 0.76 mg |  | 0.76 | [5 empty] | 1.002 | 0.0015 | 71940 |  | XCAMS | EA | 0% |  | TW3106 measurement: * no problems with XCAMS * regular precision * calibrated with 6x cathodes with OxI via EA CalAMS error adjustments: * 0% system error * uncertainty multipliers, due to excess run-to-run variability: cat#11: factor 1.3 cat#30: factor 1.3 Three groups have been identified to blank-correct: * organics via AAA or without prep-chem: 21x unknowns + 0x secondaries + 1x blanks; a 180-days window was applied * organics via CE: 0x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied * carbonates: 6x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2017-05-22 | Carbonate |  | JOB-206819 |  |  | 62773 | Normal | 14047/1 |  |  | 2017-03-16 | 4302 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00151 | 71941 | [26 empty] | 2017-03-27 | 2872 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9976 | 0.0015 | 71977 | [26 empty] | 2017-03-27 | 1709 | Gas Splitting | Completed
[168 empty] | 0.99924 | 0.00151 | 71978 | [26 empty] | 2017-03-27 | 3200 | Choose Gas Splitting | Completed
[168 empty] | 0.99967 | 0.0015 | 72022 | [26 empty] | 2017-03-31 | 12010 | Graphite | Completed
[168 empty] | 0.99985 | 0.00149 | 72024 | [26 empty] | 2017-04-04 | 4967 | C13 Measurement | Completed
[197 empty] | 2017-04-11 | 13489 | AMS Submission Results | Completed
[197 empty] | 2017-04-11 | 16152 | Close Job | Completed
[7 empty] | 130.3 | 88.2694 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37725 | 37725 | 37725 | 137 | 137 | 137 |  |  | margaret norris | -990.87 | 0.16 ‰ | 2017-05-05 | 2017-05-05 | 24 | 0 |  |  |  |  | 5.18‰ | 5.18‰ | 2.00‰ | 2.00‰ | 5.18 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -990.87 ‰ | -990.87 ‰ | 0.16 | 0.16 | -990.76 | 0.16 ‰ | -990.76 | 0.16 ‰ |  |  | 2017-05-08 | Jocelyn Turnbull | 6:59 | 0.0091294 | 0.0091294 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00913 | 0.00016 | 13673 | 2017-05-06 | 13:53 | Albert Zondervan | 206979 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.91 | 0.02% | 20 | 5 | 6 | ... | 0.00878 | 0.00015 |  |  |  |  | 0.99704 | 0.99704 | Yes | Yes | [5 empty] | 0.00878 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00913 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00878 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-04 | margaret norris | 11:18 |  | Completed |  |  |  |  | 72071 | 2000000 |  |  | 72120 72122 72125 72171 72175 72198 | 72120~72122~72125~72171~72175~72198 | 3110 |  | 3110 3110 3110 3110 3110 3110 | 2017-05-06 | 14:12 | Albert Zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 0.9985 | 0.00158 | 72120 |  | XCAMS | EA | 0% |  | TW3110 measurement: * no problems with XCAMS * normal precision * EA + waterline CalAMS error adjustments: * 0% system error * factor 1.4 for TP72112 because of excess run-to-run variability primaries: * The wheel contained 6x cathodes with OxI secondaries: * 1x IAEA-C2: results consistent with expectation value * 1x FIRI-I: results consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 13x unknowns + 1x secondaries + 1x blanks waters: 15x unknowns + 1x secondaries + 1x blanks * Waterline blank is rather high (again). unknowns: * no failures * partial loss of graphite in loading TP72124 has not affected its precision | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-05-30 | Water | Complete | JOB-206979 |  |  | 62930 | Normal | 14047/11 |  |  | 2017-04-10 | 2881 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00155 | 72122 | [26 empty] | 2017-04-10 | 1733 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00155 | 72125 | [26 empty] | 2017-04-10 | 3228 | Choose Gas Splitting | Completed
[168 empty] | 0.99954 | 0.00158 | 72171 | [26 empty] | 2017-04-11 | 12081 | Graphite | Completed
[168 empty] | 0.99833 | 0.00157 | 72175 | [26 empty] | 2017-04-19 | 4993 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00159 | 72198 | [26 empty] | 2017-05-08 | 13673 | AMS Submission Results | Completed
[197 empty] | 2017-05-17 | 16426 | Close Job | Completed
[5 empty] | 5.3451e-05 | 0.0074514 | 127.1 | 89.3145 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7153 | 7153 | 7153 | 50 | 49 | 53 | 7103.4 | 7202.9 | margaret norris | -589.53 | 2.54 ‰ | 2017-05-05 | 2016-11-06 | 16 | 180 |  |  |  |  | -6.58‰ | -6.58‰ | 2.00‰ | 2.00‰ | -6.58 ‰ | 2.00 ‰ | AMS | -8.26‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -589.53 ‰ | -589.53 ‰ | 2.54 | 2.7 | -584.46 | 2.57 ‰ | -584.46 | 2.57 ‰ |  |  | 2017-05-08 | Jocelyn Turnbull | 7:00 | 0.41047 | 0.41047 | 0.00254 | 0.0027 | 0.00249 | 0.00249 | 2.1 |  |  | 0.41499 | 0.00148 | 13674 | 2017-05-06 | 13:53 | Albert Zondervan | 206980 | 0.00719 | 0.0072 | 0.00324 | 0.0032 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.05 | 0.25% | 34 | 5 | 6 | ... | 0.39911 | 0.00133 |  |  |  |  | 0.99704 | 0.99704 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 52.9292946646679144. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00719 | 0.00716 | 0.00162 | 0.39476 | 0.39‰ | 0.00239 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41047 | 0 | 0.00254 | 0 | 0 | 0.00719 | 0 | 0.00324 | 0 | 0.39476 | 0 | 0.00239 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-04 | margaret norris | 11:18 |  | Completed |  |  |  |  | 72112 | 2000000 | 72071 71907 71759 71702 71554 71386 | 72071~71907~71759~71702~71554~71386 | 72120 72122 72125 72171 72175 72198 | 72120~72122~72125~72171~72175~72198 | 3110 | 3110 3104 3103 3098 3096 3093 | 3110 3110 3110 3110 3110 3110 | 2017-05-06 | 14:12 | Albert Zondervan |  | 0.90 mg | 0.13 | 0.9 | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 0.9985 | 0.00158 | 72120 | EA | XCAMS | EA | 0% |  | TW3110 measurement: * no problems with XCAMS * normal precision * EA + waterline CalAMS error adjustments: * 0% system error * factor 1.4 for TP72112 because of excess run-to-run variability primaries: * The wheel contained 6x cathodes with OxI secondaries: * 1x IAEA-C2: results consistent with expectation value * 1x FIRI-I: results consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 13x unknowns + 1x secondaries + 1x blanks waters: 15x unknowns + 1x secondaries + 1x blanks * Waterline blank is rather high (again). unknowns: * no failures * partial loss of graphite in loading TP72124 has not affected its precision | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-05-30 | Water |  | JOB-206980 |  |  | 62931 | Normal | 14047/12 |  |  | 2017-04-18 | 2884 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 1.0019 | 0.00155 | 72122 | ST | [25 empty] | 2017-04-18 | 1734 | Gas Splitting | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0003 | 0.00155 | 72125 | EA | [25 empty] | 2017-04-18 | 3230 | Choose Gas Splitting | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 0.99954 | 0.00158 | 72171 | EA | [25 empty] | 2017-04-19 | 12122 | Graphite | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 0.99833 | 0.00157 | 72175 | ST | [25 empty] | 2017-04-26 | 5003 | C13 Measurement | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 | 1.0013 | 0.00159 | 72198 | EA | [25 empty] | 2017-05-08 | 13674 | AMS Submission Results | Completed
[197 empty] | 2017-05-17 | 16427 | Close Job | Completed
|  |  | 0 | 0 | 5.1163e-05 | 0.0075136 | 120.1 | 88.5552 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37341 | 37340 | 37341 | 122 | 122 | 122 | 37219.7 | 37464 | margaret Norris | -990.42 | 0.15 ‰ | 2017-05-12 | 2016-11-13 | 15 | 180 |  |  |  |  | 1.55‰ | 1.55‰ | 2.00‰ | 2.00‰ | 1.55 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -990.42 ‰ | -990.42 ‰ | 0.15 | 0.15 | -990.31 | 0.15 ‰ | -990.31 | 0.15 ‰ |  |  | 2017-05-15 | Jocelyn Turnbull | 11:02 | 0.0095765 | 0.0095765 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.8 |  |  | 0.00958 | 0.00015 | 13758 | 2017-05-15 | 9:14 | Albert Zondervan | 207067 | 0 | 0 | 0 | 0 | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.96 | 0.01% | 22 | 5 | 7 | ... | 0.00921 | 0.00014 |  |  |  |  | 0.99666 | 0.99666 | Yes | Yes |  |  | 0.00728 | 0.00722 | 0.0015 | 0.00921 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 0.99999 | 0.00062 | 0.00062 | 0.00958 | 0 | 0.00015 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00921 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-11 | margaret Norris | 11:39 |  | Completed |  |  |  |  | 72164 | 2000000 | 72237 72071 71907 71759 71702 71554 71386 | 72237~72071~71907~71759~71702~71554~71386 | 72069 72088 72089 72090 72153 72163 | 72069~72088~72089~72090~72153~72163 | 3112 | 3112 3110 3104 3103 3098 3096 3093 | 3112 3112 3112 3112 3112 3112 | 2017-05-15 | 10:20 | Albert Zondervan | Yes | 0.70 mg |  | 0.7 | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 0.99735 | 0.00145 | 72069 | EA | XCAMS | ST | 0% |  | TW3112 measurement: * no problems with XCAMS * normal precision * ST + waterline CalAMS error adjustments: * 0% system error * large factor 3.2 for TP72253 because of excess run-to-run variability (see comment below) primaries: * The wheel contained 6x cathodes with OxI, with 450k counts each on average. secondaries: * 1x FIRI-I: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 25x unknowns + 1x secondaries + 0x blanks (+ 1x EA blank) waters: 2x unknowns + 0x secondaries + 1x blanks + 2x tests * The waterline test blanks were forced to category 'Test Background' and to receive no blank correction. unknowns: * no failures, but: * The result for TP72253 (water unknown) is suspect because of trends in the raw data suggesting ratio inhomogeneity. | 2.3 | RRL | [5 empty] | Background Test |  | 2017-06-21 | Water | Complete | JOB-207067 | 540GCT23 (test sample) | test job to compare with Kapuni bubble -flushed before addition -chunk rather than powder acid heated in water bath to melt marble chip | 63010 | Normal | 14047/11 |  |  | 2017-04-26 | 2895 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 1.0001 | 0.00149 | 72088 | ST | [25 empty] | 2017-04-26 | 1738 | Gas Splitting | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 0.99931 | 0.00147 | 72089 | EA | [25 empty] | 2017-04-26 | 3238 | Choose Gas Splitting | Cancelled
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0013 | 0.00169 | 72090 | EA | [25 empty] | 2017-04-26 | 2897 | Water CO2 Evolution | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0014 | 0.00151 | 72153 | ST | [25 empty] | 2017-04-27 | 12174 | Graphite | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 1.0009 | 0.0015 | 72163 | EA | [25 empty] | 2017-05-01 | 5025 | C13 Measurement | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 |  |  |  | ST | [25 empty] | 2017-05-15 | 13758 | AMS Submission Results | Completed
[197 empty] | 2017-05-17 | 16433 | Close Job | Completed
[197 empty] | 2017-05-24 | 5180 | Invoice | Completed
|  |  | 0 | 0 | 2.927e-05 | 0.0034707 | 123.8 | 88.4731 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38594 | 38593 | 38594 | 173 | 173 | 173 |  |  | margaret Norris | -991.81 | 0.18 ‰ | 2017-05-12 | 2016-11-13 | 7 | 180 |  |  |  |  | 2.48‰ | 2.48‰ | 2.00‰ | 2.00‰ | 2.48 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -991.81 ‰ | -991.81 ‰ | 0.18 | 0.18 | -991.71 | 0.18 ‰ | -991.71 | 0.18 ‰ |  |  | 2017-05-15 | Jocelyn Turnbull | 11:02 | 0.0081936 | 0.0081936 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.2 |  |  | 0.00819 | 0.00018 | 13764 | 2017-05-15 | 9:14 | Albert Zondervan | 207150 | 0 | 0 | 0 | 0 | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.82 | 0.02% | 3 | 5 | 6 | ... | 0.00788 | 0.00017 |  |  |  |  | 0.99666 | 0.99666 | Yes | Yes |  |  | 0.0054 | 0.00333 | 0.00299 | 0.00788 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 0.99999 | 0.00062 | 0.00062 | 0.00819 | 0 | 0.00018 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00788 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-11 | margaret Norris | 11:39 |  | Completed |  |  |  |  | 72237 | 2000000 | 72071 72000 71898 71907 71759 71702 71554 71443 71386 | 72071~72000~71898~71907~71759~71702~71554~71443~71386 | 72069 72088 72089 72090 72153 72163 | 72069~72088~72089~72090~72153~72163 | 3112 | 3110 3106 3104 3104 3103 3098 3096 3094 3093 | 3112 3112 3112 3112 3112 3112 | 2017-05-15 | 10:20 | Albert Zondervan | Yes | 0.89 mg |  | 0.89 | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 0.99735 | 0.00145 | 72069 | EA | XCAMS | ST | 0% |  | TW3112 measurement: * no problems with XCAMS * normal precision * ST + waterline CalAMS error adjustments: * 0% system error * large factor 3.2 for TP72253 because of excess run-to-run variability (see comment below) primaries: * The wheel contained 6x cathodes with OxI, with 450k counts each on average. secondaries: * 1x FIRI-I: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 25x unknowns + 1x secondaries + 0x blanks (+ 1x EA blank) waters: 2x unknowns + 0x secondaries + 1x blanks + 2x tests * The waterline test blanks were forced to category 'Test Background' and to receive no blank correction. unknowns: * no failures, but: * The result for TP72253 (water unknown) is suspect because of trends in the raw data suggesting ratio inhomogeneity. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-06-28 | Water | Complete | JOB-207150 |  |  | 63011 | Normal | 14047/11 |  |  | 2017-05-03 | 2900 | Water CO2 Evolution | Completed | 14047/1 | IAEA-C1: Carrara Marble | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-04-06 | 0.00148 | 72000 | 3106 | 0.76 | 1.0001 | 0.00149 | 72088 | EA | [25 empty] | 2017-05-03 | 1743 | Gas Splitting | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE
[163 empty] | 2017-03-24 | 0.0014 | 71898 | 3104 | 1.05 | 0.99931 | 0.00147 | 72089 | ST | [25 empty] | 2017-05-03 | 3245 | Choose Gas Splitting | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 1.0013 | 0.00169 | 72090 | EA | [25 empty] | 2017-05-05 | 12247 | Graphite | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0014 | 0.00151 | 72153 | EA | [25 empty] | 2017-05-08 | 5057 | C13 Measurement | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0009 | 0.0015 | 72163 | ST | [25 empty] | 2017-05-15 | 13764 | AMS Submission Results | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 |  |  |  | EA | [25 empty] | 2017-05-17 | 16436 | Close Job | Completed
[163 empty] | 2016-12-12 | 0.00258 | 71443 | 3094 | 0.99 |  |  |  | EA
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79
[7 empty] | 119.9 | 88.5589 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41948 | 41948 | 41948 | 186 | 185 | 186 |  |  | jenny dahl | -994.6 | 0.12 ‰ | 2017-05-24 | 2017-05-24 | 2 | 0 |  |  |  |  | 1.63‰ | 1.63‰ | 2.00‰ | 2.00‰ | 1.63 ‰ | 2.00 ‰ | AMS | 2.11‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -994.60 ‰ | -994.60 ‰ | 0.12 | 0.12 | -994.54 | 0.13 ‰ | -994.54 | 0.13 ‰ |  |  | 2017-05-26 | Jocelyn Turnbull | 12:24 | 0.0053965 | 0.0053965 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.0054 | 0.00012 | 13835 | 2017-05-25 | 14:45 | Albert Zondervan | 207250 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 16 | 5 | 6 | ... | 0.00519 | 0.00012 |  |  |  |  | 0.99677 | 0.99677 | Yes | Yes | [5 empty] | 0.00519 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.0054 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00519 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-05-24 | jenny dahl | 8:19 |  | Completed |  |  |  |  | 72341 | 2000000 |  |  | 72240 72250 72317 72319 72326 72336 | 72240~72250~72317~72319~72326~72336 | 3114 |  | 3114 3114 3114 3114 3114 3114 | 2017-05-26 | 12:04 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 1.0007 | 0.00164 | 72240 |  | XCAMS | ST | 0% |  | TW3114 measurement: * no problems with XCAMS * normal precision * large samples, except 2 organic unknowns just under 0.3 mg C CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes secondaries: * 1x FIRI-I: result is 2-sigma deviant from expectation value * 1x FIRI-L: result consistent with expectation value * 1x IAEA-C2: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: organics via AAA: 17x unknowns + 2x secondaries + 1x blanks waters: 7x unknowns + 1x secondaries + 1x blanks shells: 3x unknowns + 0x secondaries + 0x blanks * Two organic unknowns are below the 0.3 mgC threshold: TP72206,72207. It was checked if blank-correcting them with the Santos method (using the spreadsheet for smalls) makes a difference. On the basis of one size-matching Kauri Renton Rd blank (TP71784, TW3100), MCC and DCC for these two unknowns were calculated. Changing to these MCC values makes no significant difference, but changing to these DCC values has a significant effect. This is all explained qualitatively by the fact that the unknowns are near-modern, for which we expect the blank correction to be dominated by the dead carbon component. Hence, the Santos correction has been applied for those two unknowns. unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-07-10 | Water |  | JOB-207250 |  | to be run with Troy's samples | 63089 | Normal | 14047/11 |  |  | 2017-05-18 | 2920 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99918 | 0.00164 | 72250 | [26 empty] | 2017-05-18 | 1758 | Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00166 | 72317 | [26 empty] | 2017-05-18 | 3263 | Choose Gas Splitting | Completed
[168 empty] | 0.99904 | 0.00164 | 72319 | [26 empty] | 2017-05-22 | 12351 | Graphite | Completed
[168 empty] | 0.99838 | 0.00165 | 72326 | [26 empty] | 2017-05-24 | 5102 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00165 | 72336 | [26 empty] | 2017-05-26 | 13835 | AMS Submission Results | Completed
[197 empty] | 2017-05-29 | 16537 | Close Job | Completed
[5 empty] | 4.7338e-05 | 0.0072089 | 121.5 | 89.3502 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7174 | 7173 | 7174 | 46 | 46 | 50 | 7127.6 | 7219.8 | jenny dahl | -590.58 | 2.35 ‰ | 2017-05-24 | 2016-11-25 | 2 | 180 |  |  |  |  | -7.24‰ | -7.24‰ | 2.00‰ | 2.00‰ | -7.24 ‰ | 2.00 ‰ | AMS | -8.29‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -590.58 ‰ | -590.58 ‰ | 2.35 | 2.52 | -585.53 | 2.38 ‰ | -585.53 | 2.38 ‰ |  |  | 2017-05-26 | Jocelyn Turnbull | 12:24 | 0.40942 | 0.40942 | 0.00235 | 0.00252 | 0.00229 | 0.00229 | 1.7 |  |  | 0.41385 | 0.00121 | 13839 | 2017-05-25 | 14:45 | Albert Zondervan | 207271 | 0.00702 | 0.007 | 0.00316 | 0.0032 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 40.94 | 0.23% | 15 | 5 | 6 | ... | 0.39801 | 0.00104 |  |  |  |  | 0.99677 | 0.99677 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 49.5043615001721049. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00702 | 0.00693 | 0.00158 | 0.39375 | 0.39‰ | 0.0022 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.40942 | 0 | 0.00235 | 0 | 0 | 0.00702 | 0 | 0.00316 | 0 | 0.39375 | 0 | 0.0022 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-05-24 | jenny dahl | 8:19 |  | Completed |  |  |  |  | 72342 | 2000000 | 72341 72237 72071 71907 71759 71702 71554 71386 | 72341~72237~72071~71907~71759~71702~71554~71386 | 72240 72250 72317 72319 72326 72336 | 72240~72250~72317~72319~72326~72336 | 3114 | 3114 3112 3110 3104 3103 3098 3096 3093 | 3114 3114 3114 3114 3114 3114 | 2017-05-26 | 12:04 | Albert Zondervan |  | 1.09 mg | 0.13 | 1.09 | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 1.0007 | 0.00164 | 72240 | EA | XCAMS | ST | 0% |  | TW3114 measurement: * no problems with XCAMS * normal precision * large samples, except 2 organic unknowns just under 0.3 mg C CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes secondaries: * 1x FIRI-I: result is 2-sigma deviant from expectation value * 1x FIRI-L: result consistent with expectation value * 1x IAEA-C2: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: organics via AAA: 17x unknowns + 2x secondaries + 1x blanks waters: 7x unknowns + 1x secondaries + 1x blanks shells: 3x unknowns + 0x secondaries + 0x blanks * Two organic unknowns are below the 0.3 mgC threshold: TP72206,72207. It was checked if blank-correcting them with the Santos method (using the spreadsheet for smalls) makes a difference. On the basis of one size-matching Kauri Renton Rd blank (TP71784, TW3100), MCC and DCC for these two unknowns were calculated. Changing to these MCC values makes no significant difference, but changing to these DCC values has a significant effect. This is all explained qualitatively by the fact that the unknowns are near-modern, for which we expect the blank correction to be dominated by the dead carbon component. Hence, the Santos correction has been applied for those two unknowns. unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-07-13 | Water |  | JOB-207271 |  |  | 63090 | Normal | 14047/12 |  |  | 2017-05-18 | 2921 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 0.99918 | 0.00164 | 72250 | ST | [25 empty] | 2017-05-18 | 1757 | Gas Splitting | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 1.0006 | 0.00166 | 72317 | EA | [25 empty] | 2017-05-18 | 3262 | Choose Gas Splitting | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 0.99904 | 0.00164 | 72319 | EA | [25 empty] | 2017-05-22 | 12352 | Graphite | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 0.99838 | 0.00165 | 72326 | ST | [25 empty] | 2017-05-24 | 5103 | C13 Measurement | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0021 | 0.00165 | 72336 | EA | [25 empty] | 2017-05-26 | 13839 | AMS Submission Results | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 |  |  |  | ST | [25 empty] | 2017-05-29 | 16541 | Close Job | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 |  |  |  | ST
[7 empty] | 101.1 | 89.0039 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49217 | 49216 | 49217 | 230 | 229 | 230 |  |  | julia collins | -997.82 | 0.06 ‰ | 2017-06-15 | 2017-06-15 | 34 | 0 |  |  |  |  | 4.26‰ | 4.26‰ | 2.00‰ | 2.00‰ | 4.26 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -997.82 ‰ | -997.82 ‰ | 0.06 | 0.06 | -997.79 | 0.06 ‰ | -997.79 | 0.06 ‰ |  |  | 2017-06-20 | Jocelyn Turnbull | 7:39 | 0.0021836 | 0.0021836 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.7 |  |  | 0.00218 | 6e-05 | 13903 | 2017-06-19 | 14:49 | rachel corran | 207095 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.22 | 0.01% | 6 | 5 | 6 | ... | 0.0021 | 6e-05 |  |  |  |  | 1.0044 | 1.0044 | Yes | Yes | [5 empty] | 0.0021 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00218 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0021 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-06-08 | julia collins | 8:47 |  | Completed |  |  |  |  | 72285 | 2000000 |  |  | 72409 72410 72411 72422 72427 72430 | 72409~72410~72411~72422~72427~72430 | 3116 |  | 3116 3116 3116 3116 3116 3116 | 2017-06-19 | 15:43 | rachel corran |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99903 | 0.00137 | 72409 |  | XCAMS | EA | 0% |  | TW3116 measurement: * no problems with XCAMS * during some runs high cathode current was generated; runs rejected * large samples CalAMS error adjustments: * 0% system error * factor 1.2 adjustment on cathodes 11 and 32 primaries: * 6x OxI cathodes secondaries: * 2x IAEA-C2: result consistent with expectation value * 1x FIRI-I:cellulose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: carbonate samples: 17x unknowns + 2x secondaries + 2x blanks organic samples (AAA): 8x unknowns + 1x secondaries + 1x blanks * The Kapuni result is consistent with previous analyses (ignoring the last in TW3115) unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-06-21 | Carbonate | Not Complete | JOB-207095 |  |  | 63169 | Normal | 14047/1 |  |  | 2017-05-09 | 2908 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99919 | 0.00137 | 72410 | [26 empty] | 2017-05-09 | 1747 | Gas Splitting | Completed
[168 empty] | 0.9998 | 0.00137 | 72411 | [26 empty] | 2017-05-09 | 3249 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.0014 | 72422 | [26 empty] | 2017-05-12 | 12295 | Graphite | Completed
[168 empty] | 1.0003 | 0.00139 | 72427 | [26 empty] | 2017-05-16 | 5074 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00141 | 72430 | [26 empty] | 2017-06-20 | 13903 | AMS Submission Results | Completed
[197 empty] | 2017-06-20 | 16667 | Close Job | Completed
[7 empty] | 89.4 | 89.218 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 54558 | 54558 | 54558 | 372 | 371 | 372 |  |  | julia collins | -998.88 | 0.05 ‰ | 2017-06-15 | 2017-06-15 | 8 | 0 |  |  |  |  | 1.85‰ | 1.85‰ | 2.00‰ | 2.00‰ | 1.85 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -998.88 ‰ | -998.88 ‰ | 0.05 | 0.05 | -998.86 | 0.05 ‰ | -998.86 | 0.05 ‰ |  |  | 2017-06-20 | Jocelyn Turnbull | 7:39 | 0.001123 | 0.001123 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2 |  |  | 0.00112 | 5e-05 | 13920 | 2017-06-19 | 14:49 | rachel corran | 207197 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.11 | 0.01% | 26 | 5 | 6 | ... | 0.00108 | 5e-05 |  |  |  |  | 1.0044 | 1.0044 | Yes | Yes | [5 empty] | 0.00108 | 0.00‰ | 5e-05 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00112 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.00108 | 0 | 5e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-06-08 | julia collins | 8:47 |  | Completed |  |  |  |  | 72434 | 2000000 |  |  | 72409 72410 72411 72422 72427 72430 | 72409~72410~72411~72422~72427~72430 | 3116 |  | 3116 3116 3116 3116 3116 3116 | 2017-06-19 | 15:43 | rachel corran |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99903 | 0.00137 | 72409 |  | XCAMS | EA | 0% |  | TW3116 measurement: * no problems with XCAMS * during some runs high cathode current was generated; runs rejected * large samples CalAMS error adjustments: * 0% system error * factor 1.2 adjustment on cathodes 11 and 32 primaries: * 6x OxI cathodes secondaries: * 2x IAEA-C2: result consistent with expectation value * 1x FIRI-I:cellulose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: carbonate samples: 17x unknowns + 2x secondaries + 2x blanks organic samples (AAA): 8x unknowns + 1x secondaries + 1x blanks * The Kapuni result is consistent with previous analyses (ignoring the last in TW3115) unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2017-06-29 | Carbonate |  | JOB-207197 |  | to be run with Meltzner repeat | 63170 | Normal | 14047/1 |  |  | 2017-06-06 | 2944 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99919 | 0.00137 | 72410 | [26 empty] | 2017-06-06 | 1764 | Gas Splitting | Completed
[168 empty] | 0.9998 | 0.00137 | 72411 | [26 empty] | 2017-06-06 | 3273 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.0014 | 72422 | [26 empty] | 2017-06-07 | 12444 | Graphite | Completed
[168 empty] | 1.0003 | 0.00139 | 72427 | [26 empty] | 2017-06-09 | 5136 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00141 | 72430 | [26 empty] | 2017-06-20 | 13920 | AMS Submission Results | Completed
[197 empty] | 2017-06-20 | 16669 | Close Job | Completed
[5 empty] | 4.6703e-05 | 0.0066725 | 139.9 | 89.1738 | 0.0677 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7312 | 7312 | 7312 | 44 | 44 | 48 | 7268.2 | 7356.2 | julia collins | -597.58 | 2.20 ‰ | 2017-06-28 | 2016-12-30 | 6 | 180 |  |  |  |  | -6.73‰ | -6.73‰ | 2.00‰ | 2.00‰ | -6.73 ‰ | 2.00 ‰ | AMS | -8.09‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -597.58 ‰ | -597.58 ‰ | 2.2 | 2.38 | -592.61 | 2.23 ‰ | -592.61 | 2.23 ‰ |  |  | 2017-06-29 | Jocelyn Turnbull | 10:55 | 0.40242 | 0.40242 | 0.0022 | 0.00238 | 0.00214 | 0.00214 | 1.6 |  |  | 0.40663 | 0.0011 | 14021 | 2017-06-29 | 9:51 | Albert Zondervan | 207317 | 0.0066 | 0.0066 | 0.00297 | 0.003 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 40.24 | 0.22% | 17 | 5 | 6 | ... | 0.39107 | 0.00093 |  |  |  |  | 0.99532 | 0.99532 | Yes | Yes |  | result NOT consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 47.5858627999766374. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0066 | 0.00641 | 0.00153 | 0.38702 | 0.39‰ | 0.00206 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.40242 | 0 | 0.0022 | 0 | 0 | 0.0066 | 0 | 0.00297 | 0 | 0.38702 | 0 | 0.00206 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-06-27 | julia collins | 13:26 |  | Completed |  |  |  |  | 72524 | 2000000 | 72525 72341 72237 72071 71907 71759 71702 71554 | 72525~72341~72237~72071~71907~71759~71702~71554 | 72487 72488 72489 72515 72529 72547 | 72487~72488~72489~72515~72529~72547 | 3119 | 3119 3114 3112 3110 3104 3103 3098 3096 | 3119 3119 3119 3119 3119 3119 | 2017-06-29 | 10:10 | Albert Zondervan |  | 1.17 mg | 0.13 | 1.17 | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 0.99902 | 0.0015 | 72487 | ST | XCAMS | ST | 0% |  | TW3119 measurement: * no problems with XCAMS * large samples CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes, via ST secondaries: * IAEA-C6 (tune cathode): result consistent with expectation value * IAEA-C2 via water-line: result 3-sigma older than expectation CRA * FIRI-I cellulose: result consistent with expectation value * FIRI-L whalebone: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)water: 14x unknowns + 1x secondaries + 1x blanks bone + other organics: 7x unknowns + 3x secondaries + 1x blanks unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-07-26 | Water | Complete | JOB-207317 |  |  | 63290 | Normal | 14047/12 |  |  | 2017-06-21 | 2956 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 1.0007 | 0.00148 | 72488 | EA | [25 empty] | 2017-06-21 | 1782 | Gas Splitting | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 1.0007 | 0.00149 | 72489 | EA | [25 empty] | 2017-06-21 | 3301 | Choose Gas Splitting | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 1.0009 | 0.00154 | 72515 | ST | [25 empty] | 2017-06-22 | 12534 | Graphite | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 0.99954 | 0.0015 | 72529 | EA | [25 empty] | 2017-06-23 | 5180 | C13 Measurement | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 0.99924 | 0.00152 | 72547 | ST | [25 empty] | 2017-06-29 | 14021 | AMS Submission Results | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 |  |  |  | ST | [25 empty] | 2017-06-29 | 16815 | Close Job | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 |  |  |  | ST
[7 empty] | 133.5 | 88.3038 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41795 | 41795 | 41795 | 167 | 167 | 167 |  |  | julia collins | -994.5 | 0.11 ‰ | 2017-06-28 | 2017-06-28 | 6 | 0 |  |  |  |  | 3.06‰ | 3.06‰ | 2.00‰ | 2.00‰ | 3.06 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0008 | 0 | 0.8165 | -994.50 ‰ | -994.50 ‰ | 0.11 | 0.11 | -994.43 | 0.12 ‰ | -994.43 | 0.12 ‰ |  |  | 2017-06-29 | Jocelyn Turnbull | 10:55 | 0.0055005 | 0.0055005 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.9 |  |  | 0.0055 | 0.00011 | 14038 | 2017-06-29 | 9:51 | Albert Zondervan | 207397 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 31 | 5 | 6 | ... | 0.00529 | 0.00011 |  |  |  |  | 0.99532 | 0.99532 | Yes | Yes | [5 empty] | 0.00529 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.0055 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00529 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-06-27 | julia collins | 13:26 |  | Completed |  |  |  |  | 72525 | 2000000 |  |  | 72487 72488 72489 72515 72529 72547 | 72487~72488~72489~72515~72529~72547 | 3119 |  | 3119 3119 3119 3119 3119 3119 | 2017-06-29 | 10:10 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99902 | 0.0015 | 72487 |  | XCAMS | ST | 0% |  | TW3119 measurement: * no problems with XCAMS * large samples CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes, via ST secondaries: * IAEA-C6 (tune cathode): result consistent with expectation value * IAEA-C2 via water-line: result 3-sigma older than expectation CRA * FIRI-I cellulose: result consistent with expectation value * FIRI-L whalebone: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)water: 14x unknowns + 1x secondaries + 1x blanks bone + other organics: 7x unknowns + 3x secondaries + 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-08-16 | Water | Not Complete | JOB-207397 |  | gas too old to use for c13 | 63289 | Normal | 14047/11 |  |  | 2017-06-21 | 2955 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00148 | 72488 | [26 empty] | 2017-06-21 | 2963 | Water CO2 Evolution | Completed
[168 empty] | 1.0007 | 0.00149 | 72489 | [26 empty] | 2017-06-21 | 1787 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00154 | 72515 | [26 empty] | 2017-06-21 | 3305 | Choose Gas Splitting | Completed
[168 empty] | 0.99954 | 0.0015 | 72529 | [26 empty] | 2017-06-22 | 12535 | Graphite | Completed
[168 empty] | 0.99924 | 0.00152 | 72547 | [26 empty] | 2017-06-29 | 14038 | AMS Submission Results | Completed
[197 empty] | 2017-07-07 | 16879 | Close Job | Cancelled
[7 empty] | 139.2 | 88.4115 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49489 | 49488 | 49489 | 277 | 276 | 277 |  |  | julia collins | -997.89 | 0.07 ‰ | 2017-08-18 | 2017-08-18 | 3 | 0 |  |  |  |  | 0.66‰ | 0.66‰ | 2.00‰ | 2.00‰ | 0.66 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.75593 | -997.89 ‰ | -997.89 ‰ | 0.07 | 0.07 | -997.86 | 0.07 ‰ | -997.86 | 0.07 ‰ |  |  | 2017-08-21 | Margaret Norris | 14:14 | 0.0021108 | 0.0021108 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.6 |  |  | 0.00211 | 7e-05 | 14237 | 2017-08-20 | 21:14 | Albert Zondervan | 207478 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.21 | 0.01% | 13 | 5 | 6 | ... | 0.00203 | 7e-05 |  |  |  |  | 0.99415 | 0.99415 | Yes | Yes | [5 empty] | 0.00203 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00211 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00203 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-08-16 | julia collins | 8:53 |  | Completed |  |  |  |  | 72778 | 2000000 |  |  | 72689 72690 72737 72754 72793 72794 72795 | 72689~72690~72737~72754~72793~72794~72795 | 3125 |  | 3125 3125 3125 3125 3125 3125 3125 | 2017-08-20 | 22:02 | Albert Zondervan |  | 0.75 mg |  | 0.75 | [5 empty] | 1.0005 | 0.0015 | 72689 |  | XCAMS | ST | 0% |  | TW3125 measurement: * no problems with XCAMS * regular precision * large samples CalAMS error adjustments: * 1.03 uncertainty multiplier to account for global run-to-run variability * 0% system error primaries: * 7x OxI cathodes: no failures secondaries: * 1x IAEA-C2: result consistent with consensus value * 1x IAEA-C6: result consistent with consensus value * 1x FIRI-I: result consistent with consensus value * 1x FIRI-L: result consistent with consensus value blanks: * Two groups have been identified to blank-correct: 26x organic samples: wood, sediment, bone 22x unknowns + 3x secondaries + 1x blanks window = 180 days 6x inorganic samples: shell, carbonate 4x unknowns + 1x secondaries + 1x blanks window = 180 days unknowns: * no failures * 3x smaller samples have larger AMS-IRMS differences for del13C, which is to be expected | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-08-31 | Carbonate | Complete | JOB-207478 |  |  | 63443 | Normal | 14047/1 |  |  | 2017-08-14 | 2974 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0013 | 0.00152 | 72690 | [26 empty] | 2017-08-14 | 1812 | Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00149 | 72737 | [26 empty] | 2017-08-14 | 3357 | Choose Gas Splitting | Completed
[168 empty] | 0.9964 | 0.00147 | 72754 | [26 empty] | 2017-08-15 | 12788 | Graphite | Completed
[168 empty] | 1.0009 | 0.00147 | 72793 | [26 empty] | 2017-08-21 | 5201 | C13 Measurement | Completed
[168 empty] | 0.99992 | 0.00149 | 72794 | [26 empty] | 2017-08-21 | 14237 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00147 | 72795 | [26 empty] | 2017-08-22 | 17066 | Close Job | Completed
[7 empty] | 121.1 | 88.4253 | 0.0025 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33790 | 33789 | 33790 | 106 | 106 | 106 |  |  | margaret norris | -985.1 | 0.20 ‰ | 2017-08-24 | 2017-08-24 | 8 | 0 |  |  |  |  | 2.58‰ | 2.58‰ | 2.00‰ | 2.00‰ | 2.58 ‰ | 2.00 ‰ | AMS | 2.14‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -985.10 ‰ | -985.10 ‰ | 0.2 | 0.2 | -984.92 | 0.20 ‰ | -984.92 | 0.20 ‰ |  |  | 2017-08-28 | Cathy Ginnane | 15:44 | 0.0149 | 0.0149 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.3 |  |  | 0.0149 | 0.0002 | 14277 | 2017-08-28 | 9:01 | Albert Zondervan | 207533 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 1.49 | 0.02% | 8 | 5 | 6 | ... | 0.01433 | 0.00019 |  |  |  |  | 0.99622 | 0.99622 | Yes | Yes |  | very high blank- no unusual observations were made during processing evolution or graphite of this blank material. |  |  |  | 0.01433 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.0149 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01433 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-08-23 | margaret norris | 15:47 |  | Completed |  |  |  |  | 72804 | 2000000 |  |  | 72767 72768 72846 72849 72857 72858 | 72767~72768~72846~72849~72857~72858 | 3126 |  | 3126 3126 3126 3126 3126 3126 |  |  |  |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0011 | 0.00204 | 72767 |  | XCAMS | EA | 0.12% |  | TW3126 measurement: * no problems with XCAMS * regular precision * large samples CalAMS error adjustments: * 0.12% system error category line (method) unkn/blan/seco/prim window [d] ----------¦--------------¦--------------------¦---------- inorganic carbonate 8/1/1/0 180 water water 3/1/1/0 180 organic EA (AAA) 15/1/2/6 180 unknowns: * no failures blanks: * Both IAEA-C1 blanks are uncharacteristically high relative to previous C1 blanks in previous wheels. The amounts by which they are high are about the same, suggesting there is a common contamination factor at play. Querying RLIMS does not reveal an obvious connection between these two preparations and no special comments have been made for them. These C1 blanks have not been ignored for the blank-subtraction calculation. The gas for these two blanks were put into spare M/S bottles not the usual M/s bottles. * The AAA-treated Kauri Renton Rd and Kapuni CO2 blanks come out consistent with previous analyses. secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x EA IAEA-C6: result consistent with consensus value * 1x EA FIRI-I: result consistent with consensus value primaries: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-09-07 | Water |  | JOB-207533 |  |  | 63499 | Normal | 14047/11 |  |  | 2017-08-15 | 2977 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00203 | 72768 | [26 empty] | 2017-08-15 | 1829 | Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00202 | 72846 | [26 empty] | 2017-08-15 | 3381 | Choose Gas Splitting | Completed
[168 empty] | 0.99693 | 0.00203 | 72849 | [26 empty] | 2017-08-16 | 12814 | Graphite | Completed
[168 empty] | 0.99913 | 0.00203 | 72857 | [26 empty] | 2017-08-21 | 5225 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00202 | 72858 | [26 empty] | 2017-08-28 | 14277 | AMS Submission Results | Completed
[197 empty] | 2017-08-30 | 17102 | Close Job | Completed
[5 empty] | 5.2318e-05 | 0.0076599 | 124 | 89.4187 | 0.069 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7157 | 7157 | 7157 | 51 | 51 | 54 | 7106.4 | 7208.9 | margaret norris | -589.76 | 2.62 ‰ | 2017-08-24 | 2017-02-25 | 8 | 180 |  |  |  |  | -8.56‰ | -8.56‰ | 2.00‰ | 2.00‰ | -8.56 ‰ | 2.00 ‰ | AMS | -8.12‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -589.76 ‰ | -589.76 ‰ | 2.62 | 2.78 | -584.69 | 2.65 ‰ | -584.69 | 2.65 ‰ |  |  | 2017-08-28 | Cathy Ginnane | 15:44 | 0.41024 | 0.41024 | 0.00262 | 0.00278 | 0.00256 | 0.00256 | 2.3 |  |  | 0.41534 | 0.0012 | 14278 | 2017-08-28 | 9:01 | Albert Zondervan | 207534 | 0.0081 | 0.0081 | 0.00365 | 0.0037 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.02 | 0.26% | 15 | 5 | 6 | ... | 0.39944 | 0.00114 |  |  |  |  | 0.99622 | 0.99622 | Yes | Yes |  | result consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 54.3597310969805252. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0081 | 0.00736 | 0.00307 | 0.39454 | 0.39‰ | 0.00251 | 0.00‰ | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.41024 | 0 | 0.00262 | 0 | 0 | 0.0081 | 0 | 0.00365 | 0 | 0.39454 | 0 | 0.00251 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-08-23 | margaret norris | 15:47 |  | Completed |  |  |  |  | 72805 | 2000000 | 72804 72525 72341 72237 72071 71907 71759 | 72804~72525~72341~72237~72071~71907~71759 | 72767 72768 72846 72849 72857 72858 | 72767~72768~72846~72849~72857~72858 | 3126 | 3126 3119 3114 3112 3110 3104 3103 | 3126 3126 3126 3126 3126 3126 | 2017-08-28 | 9:33 | Albert Zondervan |  | 0.82 mg | 0.13 | 0.82 | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 1.0011 | 0.00204 | 72767 | EA | XCAMS | EA | 0.12% |  | TW3126 measurement: * no problems with XCAMS * regular precision * large samples CalAMS error adjustments: * 0.12% system error category line (method) unkn/blan/seco/prim window [d] ----------¦--------------¦--------------------¦---------- inorganic carbonate 8/1/1/0 180 water water 3/1/1/0 180 organic EA (AAA) 15/1/2/6 180 unknowns: * no failures blanks: * Both IAEA-C1 blanks are uncharacteristically high relative to previous C1 blanks in previous wheels. The amounts by which they are high are about the same, suggesting there is a common contamination factor at play. Querying RLIMS does not reveal an obvious connection between these two preparations and no special comments have been made for them. These C1 blanks have not been ignored for the blank-subtraction calculation. The gas for these two blanks were put into spare M/S bottles not the usual M/s bottles. * The AAA-treated Kauri Renton Rd and Kapuni CO2 blanks come out consistent with previous analyses. secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x EA IAEA-C6: result consistent with consensus value * 1x EA FIRI-I: result consistent with consensus value primaries: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-09-07 | Water |  | JOB-207534 |  |  | 63500 | Normal | 14047/12 |  |  | 2017-08-15 | 2978 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 1.0017 | 0.00203 | 72768 | ST | [25 empty] | 2017-08-15 | 1827 | Gas Splitting | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 1.0023 | 0.00202 | 72846 | EA | [25 empty] | 2017-08-15 | 3375 | Choose Gas Splitting | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 0.99693 | 0.00203 | 72849 | ST | [25 empty] | 2017-08-16 | 12815 | Graphite | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 0.99913 | 0.00203 | 72857 | ST | [25 empty] | 2017-08-21 | 5226 | C13 Measurement | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 0.9989 | 0.00202 | 72858 | ST | [25 empty] | 2017-08-28 | 14278 | AMS Submission Results | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 |  |  |  | EA | [25 empty] | 2017-08-30 | 17103 | Close Job | Completed
[7 empty] | 120.9 | 88.4514 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38553 | 38553 | 38553 | 142 | 141 | 142 |  |  | margaret norris | -991.76 | 0.15 ‰ | 2017-08-24 | 2017-08-24 | 8 | 0 |  |  |  |  | 2.28‰ | 2.28‰ | 2.00‰ | 2.00‰ | 2.28 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -991.76 ‰ | -991.76 ‰ | 0.15 | 0.15 | -991.66 | 0.15 ‰ | -991.66 | 0.15 ‰ |  |  | 2017-08-28 | Cathy Ginnane | 15:44 | 0.0082352 | 0.0082352 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.8 |  |  | 0.00824 | 0.00015 | 14300 | 2017-08-28 | 9:01 | Albert Zondervan | 207751 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.82 | 0.01% | 29 | 5 | 6 | ... | 0.00792 | 0.00014 |  |  |  |  | 0.99622 | 0.99622 | Yes | Yes |  | very high blank- no unusual observations in evolution or graphite were made during processing of this blank material. Note that the last of the prepped material appears to have been used- so subsequent blanks will be made from freshly etched C1 marble |  |  |  | 0.00792 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00824 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00792 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-08-23 | margaret norris | 15:47 |  | Completed |  |  |  |  | 72811 | 2000000 |  |  | 72767 72768 72846 72849 72857 72858 | 72767~72768~72846~72849~72857~72858 | 3126 |  | 3126 3126 3126 3126 3126 3126 |  |  |  |  | 0.71 mg |  | 0.71 | [5 empty] | 1.0011 | 0.00204 | 72767 |  | XCAMS | EA | 0.12% |  | TW3126 measurement: * no problems with XCAMS * regular precision * large samples CalAMS error adjustments: * 0.12% system error category line (method) unkn/blan/seco/prim window [d] ----------¦--------------¦--------------------¦---------- inorganic carbonate 8/1/1/0 180 water water 3/1/1/0 180 organic EA (AAA) 15/1/2/6 180 unknowns: * no failures blanks: * Both IAEA-C1 blanks are uncharacteristically high relative to previous C1 blanks in previous wheels. The amounts by which they are high are about the same, suggesting there is a common contamination factor at play. Querying RLIMS does not reveal an obvious connection between these two preparations and no special comments have been made for them. These C1 blanks have not been ignored for the blank-subtraction calculation. The gas for these two blanks were put into spare M/S bottles not the usual M/s bottles. * The AAA-treated Kauri Renton Rd and Kapuni CO2 blanks come out consistent with previous analyses. secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x EA IAEA-C6: result consistent with consensus value * 1x EA FIRI-I: result consistent with consensus value primaries: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-10-10 | Carbonate | Complete | JOB-207751 |  |  | 63522 | Normal | 14047/1 |  |  | 2017-08-15 | 2985 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00203 | 72768 | [26 empty] | 2017-08-15 | 1826 | Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00202 | 72846 | [26 empty] | 2017-08-15 | 3374 | Choose Gas Splitting | Completed
[168 empty] | 0.99693 | 0.00203 | 72849 | [26 empty] | 2017-08-16 | 12821 | Graphite | Completed
[168 empty] | 0.99913 | 0.00203 | 72857 | [26 empty] | 2017-08-21 | 5232 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00202 | 72858 | [26 empty] | 2017-08-28 | 14300 | AMS Submission Results | Completed
[197 empty] | 2017-08-30 | 17115 | Close Job | Completed
[7 empty] | 123.9 | 88.5148 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36616 | 36615 | 36616 | 128 | 127 | 128 |  |  | Julia Collins | -989.52 | 0.17 ‰ | 2017-09-06 | 2017-09-06 | 12 | 0 |  |  |  |  | 0.20‰ | 0.20‰ | 2.00‰ | 2.00‰ | 0.20 ‰ | 2.00 ‰ | AMS | 2.07‰ | 0.2‰ | C13 | No | -19 | -19.0019 | 0 | 0.8165 | -989.52 ‰ | -989.52 ‰ | 0.17 | 0.17 | -989.39 | 0.17 ‰ | -989.39 | 0.17 ‰ |  |  | 2017-09-08 | margaret norris | 8:35 | 0.010481 | 0.010481 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.8 |  |  | 0.01048 | 0.00017 | 14357 | 2017-09-07 | 11:40 | Albert Zondervan | 207703 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 16 | 5 | 8 | ... | 0.01008 | 0.00016 |  |  |  |  | 0.99486 | 0.99486 | Yes | Yes | [5 empty] | 0.01008 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.01048 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01008 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-05 | Julia Collins | 14:57 |  | Completed |  |  |  |  | 72897 | 2000000 |  |  | 72942 72958 72959 72960 72961 72971 | 72942~72958~72959~72960~72961~72971 | 3128 |  | 3128 3128 3128 3128 3128 3128 | 2017-09-07 | 12:08 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99943 | 0.00155 | 72942 |  | XCAMS | ST | 0% |  | TW3128 measurement: * no problems with XCAMS * normal precision * large samples * no EA (except IAEA-C6) CalAMS error adjustments: * 1.3 uncertainty multiplier for cat#2 (TP72937) to account for run-to-run variability * 0% system error primaries: * OxI cathodes: 6x no failures blanks: * Two groups have been identified to blank-correct: organics: 17x unknowns + 2x secondaries + 1x blanks window = 180 days waters: 10x unknowns + 1x secondaries + 1x blanks window = 180 days water-line blank is still pretty bad * Kapuni has low rts: good secondaries: * 1x IAEA-C2: consistent with expectation value * 1x TIRI-I: consistent with expectation value * 1x TIRI-L: consistent with expectation value unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-10-05 | Water |  | JOB-207703 |  |  | 63572 | Normal | 14047/11 |  |  | 2017-08-21 | 2988 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.002 | 0.00147 | 72958 | [26 empty] | 2017-08-21 | 1831 | Gas Splitting | Completed
[168 empty] | 0.99835 | 0.00153 | 72959 | [26 empty] | 2017-08-21 | 3385 | Choose Gas Splitting | Cancelled
[168 empty] | 1.0005 | 0.00155 | 72960 | [26 empty] | 2017-08-21 | 1832 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00154 | 72961 | [26 empty] | 2017-08-21 | 3386 | Choose Gas Splitting | Completed
[168 empty] | 0.99829 | 0.00158 | 72971 | [26 empty] | 2017-08-25 | 12907 | Graphite | Completed
[197 empty] | 2017-08-30 | 5241 | C13 Measurement | Completed
[197 empty] | 2017-09-08 | 14357 | AMS Submission Results | Completed
[197 empty] | 2017-09-08 | 17146 | Close Job | Completed
[5 empty] | 4.9909e-05 | 0.0079138 | 112.3 | 89.5875 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7100 | 7100 | 7100 | 52 | 52 | 55 | 7048.4 | 7152.5 | Julia Collins | -586.83 | 2.67 ‰ | 2017-09-06 | 2017-03-10 | 12 | 180 |  |  |  |  | -11.77‰ | -11.77‰ | 2.00‰ | 2.00‰ | -11.77 ‰ | 2.00 ‰ | AMS | -8.37‰ | 0.2‰ | C13 | No | -19 | -19.0019 | 0 | 0.8165 | -586.83 ‰ | -586.83 ‰ | 2.67 | 2.83 | -581.73 | 2.71 ‰ | -581.73 | 2.71 ‰ |  |  | 2017-09-08 | margaret norris | 8:35 | 0.41317 | 0.41317 | 0.00267 | 0.00283 | 0.00262 | 0.00262 | 1.8 |  |  | 0.4184 | 0.00122 | 14358 | 2017-09-07 | 11:40 | Albert Zondervan | 207704 | 0.00835 | 0.0084 | 0.00376 | 0.0038 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.32 | 0.27% | 17 | 5 | 6 | ... | 0.40239 | 0.00105 |  |  |  |  | 0.99486 | 0.99486 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 55.0592512855370147. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00835 | 0.00761 | 0.00293 | 0.39736 | 0.40‰ | 0.00252 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.41317 | 0 | 0.00267 | 0 | 0 | 0.00835 | 0 | 0.00376 | 0 | 0.39736 | 0 | 0.00252 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-05 | Julia Collins | 14:57 |  | Completed |  |  |  |  | 72898 | 2000000 | 72897 72804 72525 72341 72237 72071 71907 71759 | 72897~72804~72525~72341~72237~72071~71907~71759 | 72942 72958 72959 72960 72961 72971 | 72942~72958~72959~72960~72961~72971 | 3128 | 3128 3126 3119 3114 3112 3110 3104 3103 | 3128 3128 3128 3128 3128 3128 | 2017-09-07 | 12:08 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 0.99943 | 0.00155 | 72942 | EA | XCAMS | ST | 0% |  | TW3128 measurement: * no problems with XCAMS * normal precision * large samples * no EA (except IAEA-C6) CalAMS error adjustments: * 1.3 uncertainty multiplier for cat#2 (TP72937) to account for run-to-run variability * 0% system error primaries: * OxI cathodes: 6x no failures blanks: * Two groups have been identified to blank-correct: organics: 17x unknowns + 2x secondaries + 1x blanks window = 180 days waters: 10x unknowns + 1x secondaries + 1x blanks window = 180 days water-line blank is still pretty bad * Kapuni has low rts: good secondaries: * 1x IAEA-C2: consistent with expectation value * 1x TIRI-I: consistent with expectation value * 1x TIRI-L: consistent with expectation value unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-10-05 | Water | Complete | JOB-207704 |  |  | 63575 | Normal | 14047/12 |  |  | 2017-08-21 | 2989 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 1.002 | 0.00147 | 72958 | ST | [25 empty] | 2017-08-21 | 1830 | Gas Splitting | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 0.99835 | 0.00153 | 72959 | EA | [25 empty] | 2017-08-21 | 3383 | Choose Gas Splitting | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 1.0005 | 0.00155 | 72960 | ST | [25 empty] | 2017-08-25 | 12908 | Graphite | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 1.0012 | 0.00154 | 72961 | ST | [25 empty] | 2017-08-30 | 5242 | C13 Measurement | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 0.99829 | 0.00158 | 72971 | ST | [25 empty] | 2017-09-08 | 14358 | AMS Submission Results | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 |  |  |  | EA | [25 empty] | 2017-09-08 | 17147 | Close Job | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 |  |  |  | ST
[7 empty] | 111.8 | 88.5151 | 0.001 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS |  | 0.2‰ | Estimate | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 1.9 |  |  |  |  | 14422 | 2017-09-11 | 8:48 | Albert Zondervan | 207917 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 1 | 5 | 5 | ... | 0.00399 | 6e-05 | [7 empty] | Yes | [5 empty] | 0.00399 | 0.00‰ | 6e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.00399 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73015 | 2000000 |  |  |  |  | 3130 | [6 empty] | 0.99 mg |  | 0.99 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Unknown Inorganic |  | 2017-11-02 | Water |  | JOB-207917 |  | carbonate wheel test | 64957 | Normal | 14047/11 |  |  | 2017-09-07 | 3027 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[197 empty] | 2017-09-07 | 1855 | Gas Splitting | Completed
[197 empty] | 2017-09-07 | 3424 | Choose Gas Splitting | Completed
[197 empty] | 2017-09-08 | 13025 | Graphite | Completed
[197 empty] | 2018-05-28 | 14422 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18924 | Close Job | Completed
[197 empty] | 2018-05-28 | 7021 | Invoice | Completed
[7 empty] | 115.8 | 88.571 | 0.0015 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS |  | 0.2‰ | Estimate | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 2.6 |  |  |  |  | 14423 | 2017-09-11 | 8:48 | Albert Zondervan | 207918 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 36 | 5 | 5 | ... | 0.00608 | 8e-05 | [7 empty] | Yes | [5 empty] | 0.00608 | 0.01‰ | 8e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.00608 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73016 | 2000000 |  |  |  |  | 3130 | [6 empty] | 0.77 mg |  | 0.77 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Unknown Inorganic |  | 2017-11-02 | Water |  | JOB-207918 |  | carbonate wheel test | 64958 | Normal | 14047/11 |  |  | 2017-09-07 | 3028 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[197 empty] | 2017-09-07 | 1856 | Gas Splitting | Completed
[197 empty] | 2017-09-07 | 3426 | Choose Gas Splitting | Completed
[197 empty] | 2017-09-08 | 13026 | Graphite | Completed
[197 empty] | 2018-05-28 | 14423 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18925 | Close Job | Completed
[197 empty] | 2018-05-28 | 7022 | Invoice | Completed
[7 empty] | 115.5 | 88.5167 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS | -8.00‰ | 0.2‰ | Submitted | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 1.9 |  |  |  |  | 14424 | 2017-09-11 | 8:48 | Albert Zondervan | 207919 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 0 | 5 | 5 | ... | 0.00105 | 3e-05 | [7 empty] | Yes | [5 empty] | 0.00105 | 0.00‰ | 3e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.00105 | 0 | 3e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73017 | 2000000 |  |  |  |  | 3130 | [6 empty] | 1.13 mg |  | 1.13 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Background Inorganic |  | 2017-11-02 | Carbonate |  | JOB-207919 |  | carbonate wheel test | 64959 | Normal | 14047/1 |  |  | 2017-09-07 | 3029 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[197 empty] | 2017-09-07 | 1854 | Gas Splitting | Completed
[197 empty] | 2017-09-07 | 3422 | Choose Gas Splitting | Completed
[197 empty] | 2017-09-08 | 13027 | Graphite | Completed
[197 empty] | 2018-05-28 | 14424 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18926 | Close Job | Completed
[197 empty] | 2018-05-28 | 7023 | Invoice | Completed
[7 empty] | 111.6 | 88.5509 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS | -8.00‰ | 0.2‰ | Submitted | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 2.1 |  |  |  |  | 14425 | 2017-09-11 | 8:48 | Albert Zondervan | 207920 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 37 | 5 | 5 | ... | 0.0009 | 3e-05 | [7 empty] | Yes | [5 empty] | 0.0009 | 0.00‰ | 3e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.0009 | 0 | 3e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73018 | 2000000 |  |  |  |  | 3130 | [6 empty] | 1.03 mg |  | 1.03 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Background Inorganic |  | 2017-11-02 | Carbonate | Complete | JOB-207920 |  | carbonate wheel test | 64960 | Normal | 14047/1 |  |  | 2017-09-07 | 3030 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[197 empty] | 2017-09-07 | 1853 | Gas Splitting | Completed
[197 empty] | 2017-09-07 | 3420 | Choose Gas Splitting | Completed
[197 empty] | 2017-09-08 | 13028 | Graphite | Completed
[197 empty] | 2018-05-28 | 14425 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18927 | Close Job | Completed
[197 empty] | 2018-05-28 | 7024 | Invoice | Completed
[7 empty] | 109.8 | 88.5635 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37580 | 37579 | 37580 | 135 | 134 | 135 |  |  | helen zhang | -990.7 | 0.16 ‰ | 2017-09-15 | 2017-09-15 | 8 | 0 |  |  |  |  | 3.25‰ | 3.25‰ | 2.00‰ | 2.00‰ | 3.25 ‰ | 2.00 ‰ | AMS | 2.30‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -990.70 ‰ | -990.70 ‰ | 0.16 | 0.16 | -990.59 | 0.16 ‰ | -990.59 | 0.16 ‰ |  |  | 2017-09-18 | Jocelyn Turnbull | 11:00 | 0.0092958 | 0.0092958 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.7 |  |  | 0.0093 | 0.00016 | 14497 | 2017-09-18 | 9:14 | Albert Zondervan | 207819 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.93 | 0.02% | 22 | 5 | 6 | ... | 0.00894 | 0.00015 |  |  |  |  | 0.99844 | 0.99844 | Yes | Yes | [5 empty] | 0.00894 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.0093 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00894 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-14 | helen zhang | 15:10 |  | Completed |  |  |  |  | 73012 | 2000000 |  |  | 73003 73007 73040 73041 73056 73057 | 73003~73007~73040~73041~73056~73057 | 3132 |  | 3132 3132 3132 3132 3132 3132 | 2017-09-18 | 10:31 | Albert Zondervan |  | 0.74 mg |  | 0.74 | [5 empty] | 1.0002 | 0.00173 | 73003 |  | XCAMS | EA | 0% |  | TW3132 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Four groups have been identified to blank-correct: organics via AAA: 7x unknowns + 2x secondaries + 1x blanks window = 180 days: captures 10x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks water: 8x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks carbonate: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 2.5 sigma * 1x IAEA-C2 on carbonate-line: result differs from consensus value by 2.0 sigma unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-10-16 | Water |  | JOB-207819 |  | New etched material carbosorb not changed | 63724 | Normal | 14047/11 |  |  | 2017-09-04 | 3014 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00169 | 73007 | [26 empty] | 2017-09-04 | 1852 | Gas Splitting | Completed
[168 empty] | 1.0027 | 0.0017 | 73040 | [26 empty] | 2017-09-04 | 3418 | Choose Gas Splitting | Completed
[168 empty] | 0.99938 | 0.00168 | 73041 | [26 empty] | 2017-09-07 | 13022 | Graphite | Completed
[168 empty] | 0.99918 | 0.00171 | 73056 | [26 empty] | 2017-09-11 | 5290 | C13 Measurement | Completed
[168 empty] | 0.99859 | 0.0016 | 73057 | [26 empty] | 2017-09-18 | 14497 | AMS Submission Results | Completed
[197 empty] | 2017-09-19 | 17286 | Close Job | Completed
[5 empty] | 5.0306e-05 | 0.0081126 | 104.1 | 89.4896 | 0.0684 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7274 | 7273 | 7274 | 55 | 54 | 58 | 7219 | 7328.5 | helen zhang | -595.64 | 2.75 ‰ | 2017-09-15 | 2017-03-19 | 8 | 180 |  |  |  |  | -7.13‰ | -7.13‰ | 2.00‰ | 2.00‰ | -7.13 ‰ | 2.00 ‰ | AMS | -8.22‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -595.64 ‰ | -595.64 ‰ | 2.75 | 2.9 | -590.65 | 2.79 ‰ | -590.65 | 2.79 ‰ |  |  | 2017-09-18 | Jocelyn Turnbull | 11:00 | 0.40436 | 0.40436 | 0.00275 | 0.0029 | 0.0027 | 0.0027 | 2 |  |  | 0.4098 | 0.00121 | 14498 | 2017-09-18 | 9:14 | Albert Zondervan | 207820 | 0.00857 | 0.0086 | 0.00386 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 40.44 | 0.28% | 11 | 5 | 6 | ... | 0.39412 | 0.00105 |  |  |  |  | 0.99844 | 0.99844 | Yes | Yes |  | result is 2.5 sigma different from consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.6256866876391512. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00857 | 0.0078 | 0.00289 | 0.38888 | 0.39‰ | 0.0026 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.40436 | 0 | 0.00275 | 0 | 0 | 0.00857 | 0 | 0.00386 | 0 | 0.38888 | 0 | 0.0026 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-14 | helen zhang | 15:10 |  | Completed |  |  |  |  | 73013 | 2000000 | 73012 72897 72804 72525 72341 72237 72071 71907 | 73012~72897~72804~72525~72341~72237~72071~71907 | 73003 73007 73040 73041 73056 73057 | 73003~73007~73040~73041~73056~73057 | 3132 | 3132 3128 3126 3119 3114 3112 3110 3104 | 3132 3132 3132 3132 3132 3132 | 2017-09-18 | 10:31 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 1.0002 | 0.00173 | 73003 | ST | XCAMS | EA | 0% |  | TW3132 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Four groups have been identified to blank-correct: organics via AAA: 7x unknowns + 2x secondaries + 1x blanks window = 180 days: captures 10x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks water: 8x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks carbonate: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 2.5 sigma * 1x IAEA-C2 on carbonate-line: result differs from consensus value by 2.0 sigma unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-10-16 | Water |  | JOB-207820 |  |  | 63725 | Normal | 14047/12 |  |  | 2017-09-04 | 3015 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 1.0002 | 0.00169 | 73007 | EA | [25 empty] | 2017-09-04 | 1851 | Gas Splitting | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 1.0027 | 0.0017 | 73040 | ST | [25 empty] | 2017-09-04 | 3416 | Choose Gas Splitting | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 0.99938 | 0.00168 | 73041 | ST | [25 empty] | 2017-09-07 | 13023 | Graphite | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 0.99918 | 0.00171 | 73056 | ST | [25 empty] | 2017-09-11 | 5291 | C13 Measurement | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 0.99859 | 0.0016 | 73057 | EA | [25 empty] | 2017-09-18 | 14498 | AMS Submission Results | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 |  |  |  | ST | [25 empty] | 2017-09-19 | 17287 | Close Job | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 |  |  |  | EA
[7 empty] | 95.2 | 88.5184 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49812 | 49811 | 49812 | 330 | 329 | 330 |  |  | helen zhang | -997.97 | 0.08 ‰ | 2017-09-15 | 2017-09-15 | 11 | 0 |  |  |  |  | 3.76‰ | 3.76‰ | 2.00‰ | 2.00‰ | 3.76 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -997.97 ‰ | -997.97 ‰ | 0.08 | 0.08 | -997.95 | 0.08 ‰ | -997.95 | 0.08 ‰ |  |  | 2017-09-18 | Jocelyn Turnbull | 11:00 | 0.0020276 | 0.0020276 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00203 | 8e-05 | 14505 | 2017-09-18 | 9:14 | Albert Zondervan | 207884 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 1 | 5 | 5 | ... | 0.00195 | 8e-05 |  |  |  |  | 0.99844 | 0.99844 | Yes | Yes | [5 empty] | 0.00195 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00203 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00195 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-14 | helen zhang | 15:10 |  | Completed |  |  |  |  | 72967 | 2000000 |  |  | 73003 73007 73040 73041 73056 73057 | 73003~73007~73040~73041~73056~73057 | 3132 |  | 3132 3132 3132 3132 3132 3132 | 2017-09-18 | 10:31 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 1.0002 | 0.00173 | 73003 |  | XCAMS | EA | 0% |  | TW3132 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Four groups have been identified to blank-correct: organics via AAA: 7x unknowns + 2x secondaries + 1x blanks window = 180 days: captures 10x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks water: 8x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks carbonate: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 2.5 sigma * 1x IAEA-C2 on carbonate-line: result differs from consensus value by 2.0 sigma unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-10-27 | Carbonate |  | JOB-207884 |  | new acid etched and ground marble. | 63723 | Normal | 14047/1 |  |  | 2017-09-01 | 3009 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0002 | 0.00169 | 73007 | [26 empty] | 2017-09-01 | 1843 | Gas Splitting | Completed
[168 empty] | 1.0027 | 0.0017 | 73040 | [26 empty] | 2017-09-01 | 3404 | Choose Gas Splitting | Completed
[168 empty] | 0.99938 | 0.00168 | 73041 | [26 empty] | 2017-09-04 | 12977 | Graphite | Completed
[168 empty] | 0.99918 | 0.00171 | 73056 | [26 empty] | 2017-09-18 | 14505 | AMS Submission Results | Completed
[168 empty] | 0.99859 | 0.0016 | 73057 | [26 empty] | 2017-09-19 | 17288 | Close Job | Completed
[7 empty] | 92.1 | 89.2754 | 0.0004 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48044 | 48044 | 48044 | 231 | 231 | 231 |  |  | jenny dahl | -997.47 | 0.07 ‰ | 2017-09-27 | 2017-09-27 | 9 | 0 |  |  |  |  | -2.35‰ | -2.35‰ | 2.00‰ | 2.00‰ | -2.35 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -997.47 ‰ | -997.47 ‰ | 0.07 | 0.07 | -997.44 | 0.07 ‰ | -997.44 | 0.07 ‰ |  |  | 2017-09-29 | Jocelyn Turnbull | 13:53 | 0.0025267 | 0.0025267 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 6.3 |  |  | 0.00253 | 7e-05 | 14569 | 2017-09-29 | 10:48 | Albert Zondervan | 207894 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.25 | 0.01% | 8 | 5 | 3 | ... | 0.00243 | 7e-05 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes | [5 empty] | 0.00243 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00253 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00243 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-27 | jenny dahl | 8:22 |  | Completed |  |  |  |  | 73104 | 2000000 |  |  | 73082 73088 73089 73101 73102 73103 | 73082~73088~73089~73101~73102~73103 | 3134 |  | 3134 3134 3134 3134 3134 3134 | 2017-09-29 | 13:38 | Albert Zondervan |  | 0.32 mg |  | 0.32 | [5 empty] | 1.0015 | 0.00135 | 73082 |  | XCAMS | ST | 0% |  | TW3134 measurement: * no problems with XCAMS * mostly small samples * none processed via EA, except for 1x IAEA-C6 CalAMS error adjustments: * 0% system error * 1.2 uncertainty multiplier for cat#14 and #28, both a large OxI, to account for run-to-run variability primaries: * 6x large OxI cathodes * 4x small OxI cathodes: d = 0.00078 mgC secondaries: * 2x small IAEA-C2 travertine: 1x result consistent with expectation value 1x result 2-sigma away from expectation value * 4x small FIRI-I cellulose: results consistent with expectation value blanks: * Three groups have been identified to blank-correct: small carbonates (CarbLine): 4x unknowns + 2x secondaries + 3x blanks m = 0.00049 mgC small organics (AAA): 9x unknowns + 4x secondaries + 4x blanks m = 0.00137 mgC large organics (AAA): 2x unknowns + 0x secondaries + 0x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-09-26 | Carbonate |  | JOB-207894 |  | samples were pre-cracked on combustion line and put into MS bottles for graphite. try to get 0.25mg C to match sample size. transferred gas into RRL-42 | 63803 |  | 14047/1 |  |  | 2017-09-05 | 3023 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99926 | 0.00133 | 73088 | [26 empty] | 2017-09-18 | 13114 | Graphite | Completed
[168 empty] | 0.99741 | 0.00169 | 73089 | [26 empty] | 2017-09-29 | 14569 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00138 | 73101 | [26 empty] | 2017-09-29 | 17459 | Close Job | Completed
[168 empty] | 1.0023 | 0.00178 | 73102
[168 empty] | 0.99955 | 0.00138 | 73103
[7 empty] | 81.6 | 89.2775 | 0.0004 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47563 | 47562 | 47563 | 249 | 249 | 249 |  |  | jenny dahl | -997.32 | 0.08 ‰ | 2017-09-27 | 2017-09-27 | 9 | 0 |  |  |  |  | -2.38‰ | -2.38‰ | 2.00‰ | 2.00‰ | -2.38 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -997.32 ‰ | -997.32 ‰ | 0.08 | 0.08 | -997.28 | 0.08 ‰ | -997.28 | 0.08 ‰ |  |  | 2017-09-29 | Jocelyn Turnbull | 13:53 | 0.0026827 | 0.0026827 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 8.9 |  |  | 0.00268 | 8e-05 | 14570 | 2017-09-29 | 10:48 | Albert Zondervan | 207895 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.27 | 0.01% | 9 | 5 | 3 | ... | 0.00258 | 8e-05 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes | [5 empty] | 0.00258 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00268 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00258 | 0 | 8e-05 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2017-09-27 | jenny dahl | 8:22 |  | Completed |  |  |  |  | 73105 | 2000000 |  |  | 73082 73088 73089 73101 73102 73103 | 73082~73088~73089~73101~73102~73103 | 3134 |  | 3134 3134 3134 3134 3134 3134 | 2017-09-29 | 13:38 | Albert Zondervan |  | 0.21 mg |  | 0.21 | [5 empty] | 1.0015 | 0.00135 | 73082 |  | XCAMS | ST | 0% |  | TW3134 measurement: * no problems with XCAMS * mostly small samples * none processed via EA, except for 1x IAEA-C6 CalAMS error adjustments: * 0% system error * 1.2 uncertainty multiplier for cat#14 and #28, both a large OxI, to account for run-to-run variability primaries: * 6x large OxI cathodes * 4x small OxI cathodes: d = 0.00078 mgC secondaries: * 2x small IAEA-C2 travertine: 1x result consistent with expectation value 1x result 2-sigma away from expectation value * 4x small FIRI-I cellulose: results consistent with expectation value blanks: * Three groups have been identified to blank-correct: small carbonates (CarbLine): 4x unknowns + 2x secondaries + 3x blanks m = 0.00049 mgC small organics (AAA): 9x unknowns + 4x secondaries + 4x blanks m = 0.00137 mgC large organics (AAA): 2x unknowns + 0x secondaries + 0x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-09-26 | Carbonate |  | JOB-207895 |  | samples were pre-cracked on combustion line and put into MS bottles for graphite. try to get 0.2mg C to match sample size. transferred gas into RRl-43 | 63804 |  | 14047/1 |  |  | 2017-09-05 | 3024 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99926 | 0.00133 | 73088 | [26 empty] | 2017-09-18 | 13115 | Graphite | Completed
[168 empty] | 0.99741 | 0.00169 | 73089 | [26 empty] | 2017-09-29 | 14570 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00138 | 73101 | [26 empty] | 2017-09-29 | 17460 | Close Job | Completed
[168 empty] | 1.0023 | 0.00178 | 73102
[168 empty] | 0.99955 | 0.00138 | 73103
[7 empty] | 47.4 | 89.5083 | 0.0007 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43960 | 43960 | 43960 | 258 | 258 | 258 |  |  | jenny dahl | -995.8 | 0.14 ‰ | 2017-09-27 | 2017-09-27 | 7 | 0 |  |  |  |  | -4.95‰ | -4.95‰ | 2.00‰ | 2.00‰ | -4.95 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -995.80 ‰ | -995.80 ‰ | 0.14 | 0.14 | -995.75 | 0.14 ‰ | -995.75 | 0.14 ‰ |  |  | 2017-09-29 | Jocelyn Turnbull | 13:53 | 0.0042008 | 0.0042008 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 22.4 |  |  | 0.0042 | 0.00014 | 14585 | 2017-09-29 | 10:48 | Albert Zondervan | 207957 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.42 | 0.01% | 10 | 5 | 3 | ... | 0.00404 | 0.00013 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes | [5 empty] | 0.00404 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.0042 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00404 | 0 | 0.00013 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2017-09-27 | jenny dahl | 8:22 |  | Completed |  |  |  |  | 73129 | 2000000 |  |  | 73082 73088 73089 73101 73102 73103 | 73082~73088~73089~73101~73102~73103 | 3134 |  | 3134 3134 3134 3134 3134 3134 | 2017-09-29 | 13:38 | Albert Zondervan |  | 0.11 mg |  | 0.11 | [5 empty] | 1.0015 | 0.00135 | 73082 |  | XCAMS | ST | 0% |  | TW3134 measurement: * no problems with XCAMS * mostly small samples * none processed via EA, except for 1x IAEA-C6 CalAMS error adjustments: * 0% system error * 1.2 uncertainty multiplier for cat#14 and #28, both a large OxI, to account for run-to-run variability primaries: * 6x large OxI cathodes * 4x small OxI cathodes: d = 0.00078 mgC secondaries: * 2x small IAEA-C2 travertine: 1x result consistent with expectation value 1x result 2-sigma away from expectation value * 4x small FIRI-I cellulose: results consistent with expectation value blanks: * Three groups have been identified to blank-correct: small carbonates (CarbLine): 4x unknowns + 2x secondaries + 3x blanks m = 0.00049 mgC small organics (AAA): 9x unknowns + 4x secondaries + 4x blanks m = 0.00137 mgC large organics (AAA): 2x unknowns + 0x secondaries + 0x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-10-10 | Carbonate |  | JOB-207957 |  | aiming for 0.1mg C to match small sample | 63805 |  | 14047/1 |  |  | 2017-09-19 | 3059 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99926 | 0.00133 | 73088 | [26 empty] | 2017-09-20 | 13139 | Graphite | Completed
[168 empty] | 0.99741 | 0.00169 | 73089 | [26 empty] | 2017-09-29 | 14585 | AMS Submission Results | Completed
[168 empty] | 0.99995 | 0.00138 | 73101 | [26 empty] | 2017-09-29 | 17464 | Close Job | Completed
[168 empty] | 1.0023 | 0.00178 | 73102
[168 empty] | 0.99955 | 0.00138 | 73103
[7 empty] | 113.3 | 88.7738 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34110 | 34110 | 34110 | 111 | 110 | 111 |  |  | Helen Zhang | -985.68 | 0.20 ‰ | 2017-09-29 | 2017-09-29 | 17 | 0 |  |  |  |  | 0.89‰ | 0.89‰ | 2.00‰ | 2.00‰ | 0.89 ‰ | 2.00 ‰ | AMS | 2.28‰ | 0.2‰ | C13 | No | -19 | -19.0035 | 0 | 0.8165 | -985.68 ‰ | -985.68 ‰ | 0.2 | 0.2 | -985.51 | 0.20 ‰ | -985.51 | 0.20 ‰ |  |  | 2017-10-03 | Cathy Ginnane | 9:31 | 0.014318 | 0.014318 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 3 |  |  | 0.01432 | 0.0002 | 14617 | 2017-10-02 | 9:26 | Albert Zondervan | 207874 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.43 | 0.02% | 31 | 5 | 4 | ... | 0.01377 | 0.00019 |  |  |  |  | 0.99846 | 0.99846 | Yes | Yes | [5 empty] | 0.01377 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.01432 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01377 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-28 | Helen Zhang | 10:27 |  | Completed |  |  |  |  | 73029 | 2000000 |  |  | 73087 73121 73122 73131 73132 73161 | 73087~73121~73122~73131~73132~73161 | 3135 |  | 3135 3135 3135 3135 3135 3135 | 2017-10-02 | 9:40 | Albert Zondervan |  | 0.67 mg |  | 0.67 | [5 empty] | 0.99957 | 0.00154 | 73087 |  | XCAMS | ST | 0% |  | TW3135 measurement: * no problems with XCAMS * normal precision * large, (ground)water samples CalAMS error adjustments: * 0% system error * uncertainty of 3x unknowns (TP73109, TP73111, TP73099) adjusted by factor 1.3x to account for excess run-to-run variability primaries: * 6x OxI cathodes blanks: * One groups have been identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x IAEA-C2 via waterline: result differs from consensus value by 2.0 sigma Note that this disagreement resembles that of TP73013 in TW3132. unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-10-26 | Water | Complete | JOB-207874 |  | first stds after water line move | 63850 | Normal | 14047/11 |  |  | 2017-09-11 | 3035 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00154 | 73121 | [26 empty] | 2017-09-12 | 13039 | Graphite | Completed
[168 empty] | 0.9982 | 0.00157 | 73122 | [26 empty] | 2017-09-18 | 5296 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00168 | 73131 | [26 empty] | 2017-10-03 | 14617 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00167 | 73132 | [26 empty] | 2017-10-03 | 17489 | Close Job | Completed
[168 empty] | 1.0006 | 0.00158 | 73161
[5 empty] | 5.4036e-05 | 0.0085179 | 110.5 | 89.8171 | 0.0682 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7249 | 7249 | 7249 | 58 | 58 | 61 | 7191 | 7307.4 | Helen Zhang | -594.41 | 2.94 ‰ | 2017-09-29 | 2017-04-02 | 17 | 180 |  |  |  |  | -10.73‰ | -10.73‰ | 2.00‰ | 2.00‰ | -10.73 ‰ | 2.00 ‰ | AMS | -8.36‰ | 0.2‰ | C13 | No | -19 | -19.0035 | 0 | 0.8165 | -594.41 ‰ | -594.41 ‰ | 2.94 | 3.08 | -589.4 | 2.97 ‰ | -589.4 | 2.97 ‰ |  |  | 2017-10-03 | Cathy Ginnane | 9:31 | 0.40559 | 0.40559 | 0.00294 | 0.00308 | 0.00289 | 0.00289 | 2.2 |  |  | 0.41148 | 0.00122 | 14619 | 2017-10-02 | 9:26 | Albert Zondervan | 207925 | 0.00928 | 0.0093 | 0.00418 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 40.56 | 0.29% | 33 | 5 | 6 | ... | 0.39573 | 0.00105 |  |  |  |  | 0.99846 | 0.99846 | Yes | Yes |  | result is 2.0 sigma different from consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 60.9413438427569189. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00928 | 0.00819 | 0.00341 | 0.39007 | 0.39‰ | 0.00278 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.40559 | 0 | 0.00294 | 0 | 0 | 0.00928 | 0 | 0.00418 | 0 | 0.39007 | 0 | 0.00278 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-28 | Helen Zhang | 10:27 |  | Completed |  |  |  |  | 73030 | 2000000 | 73029 73012 72897 72804 72525 72341 72237 72071 | 73029~73012~72897~72804~72525~72341~72237~72071 | 73087 73121 73122 73131 73132 73161 | 73087~73121~73122~73131~73132~73161 | 3135 | 3135 3132 3128 3126 3119 3114 3112 3110 | 3135 3135 3135 3135 3135 3135 | 2017-10-02 | 9:40 | Albert Zondervan |  | 0.87 mg | 0.13 | 0.87 | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 0.99957 | 0.00154 | 73087 | EA | XCAMS | ST | 0% |  | TW3135 measurement: * no problems with XCAMS * normal precision * large, (ground)water samples CalAMS error adjustments: * 0% system error * uncertainty of 3x unknowns (TP73109, TP73111, TP73099) adjusted by factor 1.3x to account for excess run-to-run variability primaries: * 6x OxI cathodes blanks: * One groups have been identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x IAEA-C2 via waterline: result differs from consensus value by 2.0 sigma Note that this disagreement resembles that of TP73013 in TW3132. unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-11-06 | Water | Complete | JOB-207925 |  | split from gas in J 207875 first stds after water line move. | 63852 | Normal | 14047/12 |  |  | 2017-09-11 | 3036 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 0.9998 | 0.00154 | 73121 | ST | [25 empty] | 2017-09-11 | 1857 | Gas Splitting | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 0.9982 | 0.00157 | 73122 | ST | [25 empty] | 2017-09-11 | 3429 | Choose Gas Splitting | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 1.0017 | 0.00168 | 73131 | ST | [25 empty] | 2017-09-12 | 13040 | Graphite | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 1.0005 | 0.00167 | 73132 | EA | [25 empty] | 2017-09-18 | 5297 | C13 Measurement | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 | 1.0006 | 0.00158 | 73161 | ST | [25 empty] | 2017-10-03 | 14619 | AMS Submission Results | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 |  |  |  | EA | [25 empty] | 2017-10-03 | 17491 | Close Job | Completed
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 |  |  |  | ST
[7 empty] | 117.3 | 88.4817 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49255 | 49254 | 49255 | 269 | 269 | 269 |  |  | Julia Collins | -997.83 | 0.07 ‰ | 2017-10-20 | 2017-10-20 | 7 | 0 |  |  |  |  | 3.27‰ | 3.27‰ | 2.00‰ | 2.00‰ | 3.27 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -997.83 ‰ | -997.83 ‰ | 0.07 | 0.07 | -997.8 | 0.07 ‰ | -997.8 | 0.07 ‰ |  |  | 2017-10-24 | Cathy Ginnane | 15:48 | 0.0021732 | 0.0021732 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.8 |  |  | 0.00217 | 7e-05 | 14741 | 2017-10-24 | 8:19 | Albert Zondervan | 208296 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.22 | 0.01% | 34 | 5 | 3 | ... | 0.00209 | 7e-05 |  |  |  |  | 0.99753 | 0.99753 | Yes | Yes | [5 empty] | 0.00209 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00217 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00209 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | Julia Collins | 8:25 |  | Completed |  |  |  |  | 73297 | 2000000 |  |  | 73207 73208 73209 73244 73245 73246 | 73207~73208~73209~73244~73245~73246 | 3138 |  | 3138 3138 3138 3138 3138 3138 | 2017-10-24 | 9:25 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99989 | 0.00153 | 73207 |  | XCAMS | ST | 0% |  | TW3138 measurement: * no problems with XCAMS * normal precision * large samples * no EA (apart from 1x IAEA-C6) CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Three groups have been identified to blank-correct: organics via AAA: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 10x blanks of the 10x KRR blanks, the one in this wheel (TP73241) is the worst water: 3x unknowns + 0x secondaries + 0x blanks window = 180 days: captures 8x blanks carbonate: 14x unknowns + 1x secondaries + 3x blanks of the 3x blanks, 2x are tests window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 on carbonate-line: result differs from consensus value unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-10-24 | Carbonate |  | JOB-208296 |  |  | 63923 |  | 14047/1 |  |  | 2017-10-11 | 3088 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00158 | 73208 | [26 empty] | 2017-10-13 | 13307 | Graphite | Completed
[168 empty] | 1.0026 | 0.00161 | 73209 | [26 empty] | 2017-10-24 | 14741 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00163 | 73244 | [26 empty] | 2017-10-25 | 17599 | Close Job | Completed
[168 empty] | 0.9978 | 0.0016 | 73245
[168 empty] | 0.9985 | 0.00162 | 73246
[7 empty] | 116.1 | 88.576 | 0.0003 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 51919 | 51919 | 51919 | 321 | 321 | 321 | 51604.4 | 52247.3 | Julia Collins | -998.44 | 0.06 ‰ | 2017-10-20 | 2017-10-20 | 1 | 0 |  |  |  |  | 2.20‰ | 2.20‰ | 2.00‰ | 2.00‰ | 2.20 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -998.44 ‰ | -998.44 ‰ | 0.06 | 0.06 | -998.42 | 0.06 ‰ | -998.42 | 0.06 ‰ |  |  | 2017-10-24 | Cathy Ginnane | 15:48 | 0.0015597 | 0.0015597 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2 |  |  | 0.00156 | 6e-05 | 14745 | 2017-10-24 | 8:19 | Albert Zondervan | 208432 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.16 | 0.01% | 8 | 5 | 5 | ... | 0.0015 | 6e-05 |  |  |  |  | 0.99753 | 0.99753 | Yes | Yes | [5 empty] | 0.0015 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00156 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0015 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | Julia Collins | 8:25 |  | Completed |  |  |  |  | 73337 | 2000000 |  |  | 73207 73208 73209 73244 73245 73246 | 73207~73208~73209~73244~73245~73246 | 3138 |  | 3138 3138 3138 3138 3138 3138 | 2017-10-24 | 9:25 | Albert Zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99989 | 0.00153 | 73207 |  | XCAMS | ST | 0% |  | TW3138 measurement: * no problems with XCAMS * normal precision * large samples * no EA (apart from 1x IAEA-C6) CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Three groups have been identified to blank-correct: organics via AAA: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 10x blanks of the 10x KRR blanks, the one in this wheel (TP73241) is the worst water: 3x unknowns + 0x secondaries + 0x blanks window = 180 days: captures 8x blanks carbonate: 14x unknowns + 1x secondaries + 3x blanks of the 3x blanks, 2x are tests window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 on carbonate-line: result differs from consensus value unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2017-11-07 | Carbonate | Complete | JOB-208432 |  | powdered etched C1 marble for comparison | 63924 |  | 14047/1 |  |  | 2017-10-17 | 3106 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00158 | 73208 | [26 empty] | 2017-10-17 | 1885 | Gas Splitting | Completed
[168 empty] | 1.0026 | 0.00161 | 73209 | [26 empty] | 2017-10-17 | 3473 | Choose Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00163 | 73244 | [26 empty] | 2017-10-19 | 13347 | Graphite | Completed
[168 empty] | 0.9978 | 0.0016 | 73245 | [26 empty] | 2017-10-24 | 14745 | AMS Submission Results | Completed
[168 empty] | 0.9985 | 0.00162 | 73246 | [26 empty] | 2017-10-25 | 17601 | Close Job | Completed
[197 empty] | 2018-05-11 | 6857 | Invoice | Completed
[7 empty] | 115 | 88.5102 | 0.0003 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 51604 | 51604 | 51604 | 309 | 308 | 309 | 51301.2 | 51919.4 | Julia Collins | -998.38 | 0.06 ‰ | 2017-10-20 | 2017-10-20 | 1 | 0 |  |  |  |  | 2.94‰ | 2.94‰ | 2.00‰ | 2.00‰ | 2.94 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0009 | 0 | 0.8165 | -998.38 ‰ | -998.38 ‰ | 0.06 | 0.06 | -998.36 | 0.06 ‰ | -998.36 | 0.06 ‰ |  |  | 2017-10-24 | Cathy Ginnane | 15:48 | 0.0016221 | 0.0016221 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2 |  |  | 0.00162 | 6e-05 | 14746 | 2017-10-24 | 8:19 | Albert Zondervan | 208435 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.16 | 0.01% | 18 | 5 | 5 | ... | 0.00156 | 6e-05 |  |  |  |  | 0.99753 | 0.99753 | Yes | Yes | [5 empty] | 0.00156 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00162 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00156 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | Julia Collins | 8:25 |  | Completed |  |  |  |  | 73338 | 2000000 |  |  | 73207 73208 73209 73244 73245 73246 | 73207~73208~73209~73244~73245~73246 | 3138 |  | 3138 3138 3138 3138 3138 3138 | 2017-10-24 | 9:25 | Albert Zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99989 | 0.00153 | 73207 |  | XCAMS | ST | 0% |  | TW3138 measurement: * no problems with XCAMS * normal precision * large samples * no EA (apart from 1x IAEA-C6) CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Three groups have been identified to blank-correct: organics via AAA: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 10x blanks of the 10x KRR blanks, the one in this wheel (TP73241) is the worst water: 3x unknowns + 0x secondaries + 0x blanks window = 180 days: captures 8x blanks carbonate: 14x unknowns + 1x secondaries + 3x blanks of the 3x blanks, 2x are tests window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 on carbonate-line: result differs from consensus value unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2017-11-07 | Carbonate |  | JOB-208435 |  | etched not ground chips of C1 marble for comparison | 63925 |  | 14047/1 |  |  | 2017-10-17 | 3107 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00158 | 73208 | [26 empty] | 2017-10-17 | 1886 | Gas Splitting | Completed
[168 empty] | 1.0026 | 0.00161 | 73209 | [26 empty] | 2017-10-17 | 3475 | Choose Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00163 | 73244 | [26 empty] | 2017-10-19 | 13348 | Graphite | Completed
[168 empty] | 0.9978 | 0.0016 | 73245 | [26 empty] | 2017-10-24 | 14746 | AMS Submission Results | Completed
[168 empty] | 0.9985 | 0.00162 | 73246 | [26 empty] | 2017-10-25 | 17602 | Close Job | Completed
[197 empty] | 2018-05-11 | 6858 | Invoice | Completed
[7 empty] | 117.9 | 88.4905 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39308 | 39307 | 39308 | 134 | 133 | 134 |  |  | margaret norris | -992.5 | 0.12 ‰ | 2017-10-24 | 2017-10-24 | 26 | 0 |  |  |  |  | 3.93‰ | 3.93‰ | 2.00‰ | 2.00‰ | 3.93 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.50 ‰ | -992.50 ‰ | 0.12 | 0.12 | -992.41 | 0.13 ‰ | -992.41 | 0.13 ‰ |  |  | 2017-10-30 | Cathy Ginnane | 13:48 | 0.0074969 | 0.0074969 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.5 |  |  | 0.0075 | 0.00012 | 14760 | 2017-10-26 | 15:31 | Albert Zondervan | 207963 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.75 | 0.01% | 9 | 5 | 7 | ... | 0.00721 | 0.00012 |  |  |  |  | 0.99829 | 0.99829 | Yes | Yes | [5 empty] | 0.00721 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.0075 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00721 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | margaret norris | 15:11 |  | Completed |  |  |  |  | 73196 | 2000000 |  |  | 73047 73061 73272 73273 73321 73324 | 73047~73061~73272~73273~73321~73324 | 3139 |  | 3139 3139 3139 3139 3139 3139 | 2017-10-26 | 16:13 | Albert Zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99809 | 0.00144 | 73047 |  | XCAMS | EA | 0% |  | TW3139 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * 3x groups have been identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 9x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks organics not via CE: 3x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 5x blanks via EA secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 3 sigma unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-11-17 | Water | Complete | JOB-207963 |  |  | 63971 | Normal | 14047/11 |  |  | 2017-09-25 | 3076 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99996 | 0.00151 | 73061 | [26 empty] | 2017-09-27 | 3078 | Water CO2 Evolution | Completed
[168 empty] | 1.0015 | 0.00143 | 73272 | [26 empty] | 2017-09-27 | 1876 | Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00144 | 73273 | [26 empty] | 2017-09-27 | 3458 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00146 | 73321 | [26 empty] | 2017-09-28 | 13206 | Graphite | Completed
[168 empty] | 0.99995 | 0.00143 | 73324 | [26 empty] | 2017-09-29 | 5348 | C13 Measurement | Completed
[197 empty] | 2017-10-30 | 14760 | AMS Submission Results | Completed
[197 empty] | 2017-10-31 | 17643 | Close Job | Completed
[5 empty] | 4.9586e-05 | 0.0083567 | 115.7 | 89.4377 | 0.0679 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7313 | 7313 | 7313 | 56 | 56 | 59 | 7257.1 | 7369.9 | margaret norris | -597.64 | 2.83 ‰ | 2017-10-24 | 2017-04-27 | 28 | 180 |  |  |  |  | -6.71‰ | -6.71‰ | 2.00‰ | 2.00‰ | -6.71 ‰ | 2.00 ‰ | AMS | -8.18‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -597.64 ‰ | -597.64 ‰ | 2.83 | 2.97 | -592.67 | 2.86 ‰ | -592.67 | 2.86 ‰ |  |  | 2017-10-30 | Cathy Ginnane | 13:48 | 0.40236 | 0.40236 | 0.00283 | 0.00297 | 0.00278 | 0.00278 | 1.8 |  |  | 0.40813 | 0.00107 | 14761 | 2017-10-26 | 15:31 | Albert Zondervan | 207964 | 0.00905 | 0.0091 | 0.00407 | 0.0041 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 40.24 | 0.28% | 22 | 5 | 6 | ... | 0.39251 | 0.00089 |  |  |  |  | 0.99829 | 0.99829 | Yes | Yes |  | result is too old by 3-sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 59.2316751288941697. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00905 | 0.00803 | 0.00326 | 0.38696 | 0.39‰ | 0.00267 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.40236 | 0 | 0.00283 | 0 | 0 | 0.00905 | 0 | 0.00407 | 0 | 0.38696 | 0 | 0.00267 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | margaret norris | 15:11 |  | Completed |  |  |  |  | 73175 | 2000000 | 73196 73029 73012 72897 72804 72525 72341 72237 72071 | 73196~73029~73012~72897~72804~72525~72341~72237~72071 | 73047 73061 73272 73273 73321 73324 | 73047~73061~73272~73273~73321~73324 | 3139 | 3139 3135 3132 3128 3126 3119 3114 3112 3110 | 3139 3139 3139 3139 3139 3139 | 2017-10-26 | 16:13 | Albert Zondervan |  | 1.00 mg | 0.13 | 1 | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 0.99809 | 0.00144 | 73047 | EA | XCAMS | EA | 0% |  | TW3139 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * 3x groups have been identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 9x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks organics not via CE: 3x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 5x blanks via EA secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 3 sigma unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-11-17 | Water |  | JOB-207964 |  | water line used for extraction, preparation method not documented but likely flushed with N2 gas, MN. | 63976 | Normal | 14047/12 |  |  | 2017-09-25 | 3077 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 0.99996 | 0.00151 | 73061 | ST | [25 empty] | 2017-09-25 | 1875 | Gas Splitting | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 1.0015 | 0.00143 | 73272 | ST | [25 empty] | 2017-09-25 | 3456 | Choose Gas Splitting | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 1.0011 | 0.00144 | 73273 | ST | [25 empty] | 2017-09-26 | 13185 | Graphite | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 0.99937 | 0.00146 | 73321 | EA | [25 empty] | 2017-09-29 | 5346 | C13 Measurement | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 | 0.99995 | 0.00143 | 73324 | ST | [25 empty] | 2017-10-30 | 14761 | AMS Submission Results | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 |  |  |  | EA | [25 empty] | 2017-10-31 | 17644 | Close Job | Completed
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 |  |  |  | ST
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 |  |  |  | EA
[7 empty] | 139.5 | 88.2422 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40987 | 40986 | 40987 | 165 | 164 | 165 |  |  | julia collins | -993.92 | 0.12 ‰ | 2017-11-08 | 2017-11-08 | 8 | 0 |  |  |  |  | 4.63‰ | 4.63‰ | 2.00‰ | 2.00‰ | 4.63 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0005 | 0 | 0.8165 | -993.92 ‰ | -993.92 ‰ | 0.12 | 0.12 | -993.84 | 0.13 ‰ | -993.84 | 0.13 ‰ |  |  | 2017-11-13 | Jocelyn Turnbull | 16:57 | 0.0060828 | 0.0060828 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00608 | 0.00012 | 14876 | 2017-11-10 | 8:32 | Albert Zondervan | 208365 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 9 | 5 | 5 | ... | 0.00585 | 0.00012 |  |  |  |  | 0.99619 | 0.99619 | Yes | Yes | [5 empty] | 0.00585 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00608 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00585 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-08 | julia collins | 8:12 |  | Completed |  |  |  |  | 73420 | 2000000 |  |  | 73363 73364 73365 73366 73367 73384 | 73363~73364~73365~73366~73367~73384 | 3142 |  | 3142 3142 3142 3142 3142 3142 | 2017-11-10 | 10:37 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99968 | 0.00153 | 73363 |  | XCAMS | ST | 0% |  | TW3142 measurement: * no problems with XCAMS * few small samples * none processed via EA CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes * 3x small OxI cathodes: d = 0.00075 mgC secondaries: * 1x small IAEA-C2 travertine: result consistent with expectation value * 1x small FIRI-I cellulose: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: small organics (AAA): blank-corrected using data from this wheel and TW3134 3x unknowns + 1x secondaries + 2x blanks m = 0.00127 mgC large organics (AAA): 7x unknowns + 1x secondaries + 1x blanks (ground)waters: 12x unknowns + 1x secondaries + 1x blanks * machine blank: An artefact has reappeared in the 2D spectrum of the particle detector. This is most likely due to Si-28 from the stubby behind the graphite in the cathode and is visible only when the Cs sputter beam reaches the stubby. Recently we reverted back to using stubbies cut from Al welding wire, which contains some Si. Any Si-28 ions that 'leak through' fall outside the C-14 counting gate, thus not affecting the machine blank. unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-12-11 | Water | Complete | JOB-208365 |  | new carbosorb, 113% CO2 yield on graphite seems high; split issue??? | 64123 | Normal | 14047/11 |  |  | 2017-10-30 | 3118 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99883 | 0.00153 | 73364 | [26 empty] | 2017-10-30 | 1900 | Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00158 | 73365 | [26 empty] | 2017-10-30 | 3496 | Choose Gas Splitting | Completed
[168 empty] | 1.002 | 0.0016 | 73366 | [26 empty] | 2017-10-31 | 13430 | Graphite | Completed
[168 empty] | 1.0001 | 0.00157 | 73367 | [26 empty] | 2017-11-13 | 14876 | AMS Submission Results | Completed
[168 empty] | 0.99926 | 0.00153 | 73384 | [26 empty] | 2017-11-14 | 17808 | Close Job | Completed
[5 empty] | 4.8234e-05 | 0.007943 | 126.5 | 89.2992 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7073 | 7073 | 7073 | 53 | 53 | 56 | 7020.2 | 7127 | julia collins | -585.44 | 2.76 ‰ | 2017-11-08 | 2017-05-12 | 12 | 180 |  |  |  |  | -7.26‰ | -7.26‰ | 2.00‰ | 2.00‰ | -7.26 ‰ | 2.00 ‰ | AMS | -8.33‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -585.44 ‰ | -585.44 ‰ | 2.76 | 2.91 | -580.33 | 2.79 ‰ | -580.33 | 2.79 ‰ |  |  | 2017-11-13 | Jocelyn Turnbull | 16:57 | 0.41456 | 0.41456 | 0.00276 | 0.00291 | 0.0027 | 0.0027 | 1.6 |  |  | 0.42001 | 0.0012 | 14877 | 2017-11-10 | 8:32 | Albert Zondervan | 208366 | 0.00873 | 0.0087 | 0.00393 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.46 | 0.28% | 27 | 5 | 6 | ... | 0.40394 | 0.00103 |  |  |  |  | 0.99619 | 0.99619 | Yes | Yes |  | result consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 56.3960670690070797. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00873 | 0.00764 | 0.00343 | 0.39869 | 0.40‰ | 0.0026 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41456 | 0 | 0.00276 | 0 | 0 | 0.00873 | 0 | 0.00393 | 0 | 0.39869 | 0 | 0.0026 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-08 | julia collins | 8:12 |  | Completed |  |  |  |  | 73394 | 2000000 | 73420 73196 73029 73012 72897 72804 72525 72341 72237 | 73420~73196~73029~73012~72897~72804~72525~72341~72237 | 73363 73364 73365 73366 73367 73384 | 73363~73364~73365~73366~73367~73384 | 3142 | 3142 3139 3135 3132 3128 3126 3119 3114 3112 | 3142 3142 3142 3142 3142 3142 | 2017-11-10 | 10:37 | Albert Zondervan |  | 1.13 mg | 0.13 | 1.13 | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 0.99968 | 0.00153 | 73363 | ST | XCAMS | ST | 0% |  | TW3142 measurement: * no problems with XCAMS * few small samples * none processed via EA CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes * 3x small OxI cathodes: d = 0.00075 mgC secondaries: * 1x small IAEA-C2 travertine: result consistent with expectation value * 1x small FIRI-I cellulose: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: small organics (AAA): blank-corrected using data from this wheel and TW3134 3x unknowns + 1x secondaries + 2x blanks m = 0.00127 mgC large organics (AAA): 7x unknowns + 1x secondaries + 1x blanks (ground)waters: 12x unknowns + 1x secondaries + 1x blanks * machine blank: An artefact has reappeared in the 2D spectrum of the particle detector. This is most likely due to Si-28 from the stubby behind the graphite in the cathode and is visible only when the Cs sputter beam reaches the stubby. Recently we reverted back to using stubbies cut from Al welding wire, which contains some Si. Any Si-28 ions that 'leak through' fall outside the C-14 counting gate, thus not affecting the machine blank. unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-12-11 | Water |  | JOB-208366 |  |  | 64124 | Normal | 14047/12 |  |  | 2017-10-25 | 3117 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 0.99883 | 0.00153 | 73364 | ST | [25 empty] | 2017-10-25 | 1897 | Gas Splitting | Completed
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 1.0004 | 0.00158 | 73365 | ST | [25 empty] | 2017-10-25 | 3492 | Choose Gas Splitting | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 1.002 | 0.0016 | 73366 | EA | [25 empty] | 2017-10-27 | 13404 | Graphite | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 1.0001 | 0.00157 | 73367 | ST | [25 empty] | 2017-10-30 | 5398 | C13 Measurement | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 | 0.99926 | 0.00153 | 73384 | EA | [25 empty] | 2017-11-13 | 14877 | AMS Submission Results | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 |  |  |  | ST | [25 empty] | 2017-11-14 | 17809 | Close Job | Completed
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 |  |  |  | EA
[163 empty] | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 |  |  |  | ST
[5 empty] | 4.7995e-05 | 0.0082159 | 122.9 | 89.3431 | 0.0688 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7156 | 7155 | 7156 | 56 | 56 | 59 | 7099.9 | 7212.2 | julia collins | -589.68 | 2.87 ‰ | 2017-11-14 | 2017-05-18 | 8 | 180 |  |  |  |  | -9.96‰ | -9.96‰ | 2.00‰ | 2.00‰ | -9.96 ‰ | 2.00 ‰ | AMS | -8.34‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -589.68 ‰ | -589.68 ‰ | 2.87 | 3.01 | -584.61 | 2.90 ‰ | -584.61 | 2.90 ‰ |  |  | 2017-11-16 | Jocelyn Turnbull | 11:58 | 0.41032 | 0.41032 | 0.00287 | 0.00301 | 0.00282 | 0.00282 | 2.4 |  |  | 0.41603 | 0.00121 | 14941 | 2017-11-15 | 10:16 | Albert Zondervan | 208517 | 0.00907 | 0.0091 | 0.00408 | 0.0041 | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 41.03 | 0.29% | 29 | 5 | 6 | ... | 0.40011 | 0.00112 |  |  |  |  | 0.99396 | 0.99396 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 59.0142063368895746. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00907 | 0.0079 | 0.00349 | 0.39462 | 0.39‰ | 0.00274 | 0.00‰ | No | 1 | 1.0001 | 0.00079 | 0.00079 | 0.41032 | 0 | 0.00287 | 0 | 0 | 0.00907 | 0 | 0.00408 | 0 | 0.39462 | 0 | 0.00274 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-13 | julia collins | 11:40 |  | Completed |  |  |  |  | 73464 | 2000000 | 73485 73420 73196 73029 73012 72897 72804 72525 72341 | 73485~73420~73196~73029~73012~72897~72804~72525~72341 | 73480 73483 73484 73497 73498 73499 | 73480~73483~73484~73497~73498~73499 | 3143 | 3143 3142 3139 3135 3132 3128 3126 3119 3114 | 3143 3143 3143 3143 3143 3143 | 2017-11-15 | 10:28 | Albert Zondervan |  | 0.86 mg | 0.13 | 0.86 | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 0.99775 | 0.00184 | 73480 | ST | XCAMS | ST | 0.1% |  | TW3143 measurement: * no problems with XCAMS, apart from drifting stripper-gas pressure * large samples * none processed via EA CalAMS error adjustments: * 0.10% system error * error adjustments for cat#1,22,35 (TP73496,735009, 73484) factor 1.3 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 1x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics (bone): 20x unknowns + 1x secondaries + 1x blanks (ground)waters: 8x unknowns + 1x secondaries + 1x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-12-28 | Water | Complete | JOB-208517 |  |  | 64197 | Normal | 14047/12 |  |  | 2017-11-02 | 3133 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 0.99839 | 0.00189 | 73483 | ST | [25 empty] | 2017-11-02 | 1907 | Gas Splitting | Completed
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 1.0032 | 0.00232 | 73484 | EA | [25 empty] | 2017-11-02 | 3505 | Choose Gas Splitting | Completed
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 1.0011 | 0.0019 | 73497 | ST | [25 empty] | 2017-11-06 | 13474 | Graphite | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 1.0013 | 0.00189 | 73498 | EA | [25 empty] | 2017-11-07 | 5430 | C13 Measurement | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 | 0.99965 | 0.0019 | 73499 | ST | [25 empty] | 2017-11-16 | 14941 | AMS Submission Results | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 |  |  |  | EA | [25 empty] | 2017-11-28 | 17919 | Close Job | Completed
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 |  |  |  | ST
[163 empty] | 2017-05-24 | 0.00519 | 72341 | 3114 | 0.85 |  |  |  | ST
[7 empty] | 126.4 | 88.1154 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35922 | 35921 | 35922 | 117 | 116 | 117 |  |  | julia collins | -988.57 | 0.17 ‰ | 2017-11-14 | 2017-11-14 | 6 | 0 |  |  |  |  | 3.83‰ | 3.83‰ | 2.00‰ | 2.00‰ | 3.83 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0 | 0.8165 | -988.57 ‰ | -988.57 ‰ | 0.17 | 0.17 | -988.43 | 0.17 ‰ | -988.43 | 0.17 ‰ |  |  | 2017-11-16 | Jocelyn Turnbull | 11:58 | 0.011427 | 0.011427 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.01143 | 0.00017 | 14945 | 2017-11-15 | 10:16 | Albert Zondervan | 208538 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.14 | 0.02% | 8 | 5 | 5 | ... | 0.01099 | 0.00016 |  |  |  |  | 0.99396 | 0.99396 | Yes | Yes |  | not a great water blank |  |  |  | 0.01099 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 1.0001 | 0.00079 | 0.00079 | 0.01143 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01099 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-13 | julia collins | 11:40 |  | Completed |  |  |  |  | 73485 | 2000000 |  |  | 73480 73483 73484 73497 73498 73499 | 73480~73483~73484~73497~73498~73499 | 3143 |  | 3143 3143 3143 3143 3143 3143 | 2017-11-15 | 10:28 | Albert Zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 0.99775 | 0.00184 | 73480 |  | XCAMS | ST | 0.1% |  | TW3143 measurement: * no problems with XCAMS, apart from drifting stripper-gas pressure * large samples * none processed via EA CalAMS error adjustments: * 0.10% system error * error adjustments for cat#1,22,35 (TP73496,735009, 73484) factor 1.3 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 1x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics (bone): 20x unknowns + 1x secondaries + 1x blanks (ground)waters: 8x unknowns + 1x secondaries + 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-02 | Water | Complete | JOB-208538 |  | loaded marble chips | 64201 | Normal | 14047/11 |  |  | 2017-11-07 | 3140 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99839 | 0.00189 | 73483 | [26 empty] | 2017-11-07 | 1918 | Gas Splitting | Completed
[168 empty] | 1.0032 | 0.00232 | 73484 | [26 empty] | 2017-11-07 | 3524 | Choose Gas Splitting | Completed
[168 empty] | 1.0011 | 0.0019 | 73497 | [26 empty] | 2017-11-08 | 13495 | Graphite | Completed
[168 empty] | 1.0013 | 0.00189 | 73498 | [26 empty] | 2017-11-16 | 14945 | AMS Submission Results | Completed
[168 empty] | 0.99965 | 0.0019 | 73499 | [26 empty] | 2017-11-28 | 17922 | Close Job | Completed
[5 empty] | 4.8529e-05 | 0.0086073 | 97.8 | 89.4048 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7153 | 7152 | 7153 | 58 | 57 | 60 | 7095.5 | 7210.5 | Julia Collins | -589.52 | 2.94 ‰ | 2017-12-04 | 2017-06-07 | 13 | 180 |  |  |  |  | -8.96‰ | -8.96‰ | 2.00‰ | 2.00‰ | -8.96 ‰ | 2.00 ‰ | AMS | -8.23‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -589.52 ‰ | -589.52 ‰ | 2.94 | 3.08 | -584.45 | 2.98 ‰ | -584.45 | 2.98 ‰ |  |  | 2017-12-06 | Jocelyn Turnbull | 16:44 | 0.41048 | 0.41048 | 0.00294 | 0.00308 | 0.00289 | 0.00289 | 2.8 |  |  | 0.41635 | 0.00122 | 15061 | 2017-12-05 | 16:51 | Albert Zondervan | 208512 | 0.00934 | 0.0093 | 0.0042 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:35 | Jocelyn Turnbull | 0 |  | 41.05 | 0.29% | 34 | 5 | 6 | ... | 0.40042 | 0.00105 |  |  |  |  | 0.99565 | 0.99565 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 60.3014694598261703. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00934 | 0.00827 | 0.00324 | 0.39477 | 0.39‰ | 0.00278 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41048 | 0 | 0.00294 | 0 | 0 | 0.00934 | 0 | 0.0042 | 0 | 0.39477 | 0 | 0.00278 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-01 | Julia Collins | 12:37 |  | Completed |  |  |  |  | 73569 | 2000000 | 73570 73485 73420 73196 73029 73012 72897 72804 72525 | 73570~73485~73420~73196~73029~73012~72897~72804~72525 | 73525 73526 73527 73637 73644 73646 | 73525~73526~73527~73637~73644~73646 | 3146 | 3146 3143 3142 3139 3135 3132 3128 3126 3119 | 3146 3146 3146 3146 3146 3146 | 2017-12-06 | 12:27 | Albert Zondervan |  | 0.77 mg | 0.13 | 0.77 | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 1.0008 | 0.00159 | 73525 | ST | XCAMS | EA | 0% |  | TW3146 measurement: * no problems with XCAMS * large samples * none processed via ST CalAMS error adjustments: * 0% system error * error adjustments: none necessary primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 2x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)waters: 11x unknowns + 1x secondaries + 1x blanks The sample swap between cat#13, 34 has been resolved. organics: CE: 3x unknowns + 0x secondaries + 1x blanks non-CE: 12x unknowns + 4x secondaries + 0x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-12-26 | Water |  | JOB-208512 |  | This job was originally labelled as R14047/11, but it is actually R14047/12. Changed - JCT. | 64284 | Normal | 14047/12 |  |  | 2017-11-20 | 3170 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 1.0003 | 0.00158 | 73526 | EA | [25 empty] | 2017-11-20 | 1937 | Gas Splitting | Completed
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 0.99854 | 0.00156 | 73527 | ST | [25 empty] | 2017-11-20 | 3545 | Choose Gas Splitting | Completed
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 0.99977 | 0.0016 | 73637 | EA | [25 empty] | 2017-11-21 | 13579 | Graphite | Completed
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 1.0006 | 0.00163 | 73644 | ST | [25 empty] | 2017-11-29 | 5498 | C13 Measurement | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 | 1 | 0.00164 | 73646 | EA | [25 empty] | 2017-12-06 | 15061 | AMS Submission Results | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 |  |  |  | ST | [25 empty] | 2017-12-07 | 18018 | Close Job | Completed
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 |  |  |  | ST
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 |  |  |  | EA
[7 empty] | 98.2 | 88.4688 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38842 | 38842 | 38842 | 137 | 136 | 137 |  |  | Julia Collins | -992.06 | 0.14 ‰ | 2017-12-04 | 2017-12-04 | 13 | 0 |  |  |  |  | 1.52‰ | 1.52‰ | 2.00‰ | 2.00‰ | 1.52 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0001 | 0 | 0.8165 | -992.06 ‰ | -992.06 ‰ | 0.14 | 0.14 | -991.96 | 0.14 ‰ | -991.96 | 0.14 ‰ |  |  | 2017-12-06 | Jocelyn Turnbull | 16:44 | 0.007944 | 0.007944 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.2 |  |  | 0.00794 | 0.00014 | 15062 | 2017-12-05 | 16:51 | Albert Zondervan | 208513 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.79 | 0.01% | 13 | 5 | 5 | ... | 0.00764 | 0.00013 |  |  |  |  | 0.99565 | 0.99565 | Yes | Yes | [5 empty] | 0.00764 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00794 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00764 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-01 | Julia Collins | 12:37 |  | Completed |  |  |  |  | 73570 | 2000000 |  |  | 73525 73526 73527 73637 73644 73646 | 73525~73526~73527~73637~73644~73646 | 3146 |  | 3146 3146 3146 3146 3146 3146 | 2017-12-06 | 12:27 | Albert Zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0008 | 0.00159 | 73525 |  | XCAMS | EA | 0% |  | TW3146 measurement: * no problems with XCAMS * large samples * none processed via ST CalAMS error adjustments: * 0% system error * error adjustments: none necessary primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 2x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)waters: 11x unknowns + 1x secondaries + 1x blanks The sample swap between cat#13, 34 has been resolved. organics: CE: 3x unknowns + 0x secondaries + 1x blanks non-CE: 12x unknowns + 4x secondaries + 0x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-12-26 | Water |  | JOB-208513 |  | This job was originally labelled as R14047/12, but it is actually R14047/11. Changed - JCT. | 64283 | Normal | 14047/11 |  |  | 2017-11-20 | 3171 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00158 | 73526 | [26 empty] | 2017-11-20 | 1938 | Gas Splitting | Completed
[168 empty] | 0.99854 | 0.00156 | 73527 | [26 empty] | 2017-11-20 | 3547 | Choose Gas Splitting | Completed
[168 empty] | 0.99977 | 0.0016 | 73637 | [26 empty] | 2017-11-21 | 13580 | Graphite | Completed
[168 empty] | 1.0006 | 0.00163 | 73644 | [26 empty] | 2017-12-06 | 15062 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00164 | 73646 | [26 empty] | 2017-12-07 | 18019 | Close Job | Completed
[5 empty] | 4.5228e-05 | 0.008665 | 102.2 | 89.2012 | 0.0675 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7381 | 7381 | 7381 | 58 | 58 | 61 | 7323.2 | 7439.8 | Julia Collins | -601.03 | 2.90 ‰ | 2017-12-06 | 2017-06-09 | 22 | 180 |  |  |  |  | -6.49‰ | -6.49‰ | 2.00‰ | 2.00‰ | -6.49 ‰ | 2.00 ‰ | AMS | -8.40‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -601.03 ‰ | -601.03 ‰ | 2.9 | 3.03 | -596.11 | 2.93 ‰ | -596.11 | 2.93 ‰ |  |  | 2017-12-11 | Jocelyn Turnbull | 9:07 | 0.39897 | 0.39897 | 0.0029 | 0.00303 | 0.00285 | 0.00285 | 1.8 |  |  | 0.40492 | 0.00105 | 15082 | 2017-12-08 | 14:10 | rachel corran | 208498 | 0.00928 | 0.0093 | 0.00418 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:35 | Jocelyn Turnbull | 0 |  | 39.9 | 0.29% | 15 | 5 | 6 | ... | 0.38942 | 0.00087 |  |  |  |  | 0.99587 | 0.99587 | Yes | Yes |  | result deviates from consensus by ~4sigma. Can't see anything to suggest a swap, and d13C looks correct. wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 61.0475326605094584. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00928 | 0.00833 | 0.00306 | 0.3837 | 0.38‰ | 0.00274 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.39897 | 0 | 0.0029 | 0 | 0 | 0.00928 | 0 | 0.00418 | 0 | 0.3837 | 0 | 0.00274 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-05 | Julia Collins | 12:38 |  | Completed |  |  |  |  | 73522 | 2000000 | 73551 73570 73485 73420 73196 73029 73012 72897 72804 72525 | 73551~73570~73485~73420~73196~73029~73012~72897~72804~72525 | 73618 73621 73624 73639 73641 73648 | 73618~73621~73624~73639~73641~73648 | 3147 | 3147 3146 3143 3142 3139 3135 3132 3128 3126 3119 | 3147 3147 3147 3147 3147 3147 | 2017-12-08 | 15:08 | rachel corran |  | 1.13 mg | 0.13 | 1.13 | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 1.0019 | 0.0014 | 73618 | ST | XCAMS | EA | 0% |  | TW3147 measurement: * no problems with XCAMS * large samples * none processed via ST, except the C6 sucrose CalAMS error adjustments: * 0% system error * factor x1.03 across all cathodes, and 1.3 for cat#23 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: !! result deviates from consensus by ~4sigma !! cause unknown blanks: * one group has been identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2017-12-22 | Water | Complete | JOB-208498 |  | water line used for extraction, preparation method not documented but likely flushed with N2 gas, MN. | 64324 | Normal | 14047/12 |  |  | 2017-11-13 | 3141 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 1.0007 | 0.00145 | 73621 | EA | [25 empty] | 2017-11-13 | 1922 | Gas Splitting | Completed
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 1.0012 | 0.00139 | 73624 | ST | [25 empty] | 2017-11-13 | 3528 | Choose Gas Splitting | Completed
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 0.99773 | 0.00139 | 73639 | EA | [25 empty] | 2017-11-14 | 13532 | Graphite | Completed
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 0.99885 | 0.00141 | 73641 | ST | [25 empty] | 2017-11-17 | 5463 | C13 Measurement | Completed
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 | 0.99977 | 0.00141 | 73648 | EA | [25 empty] | 2017-12-11 | 15082 | AMS Submission Results | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 |  |  |  | ST | [25 empty] | 2017-12-11 | 18037 | Close Job | Completed
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 |  |  |  | ST
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 |  |  |  | EA
[163 empty] | 2017-06-28 | 0.00529 | 72525 | 3119 | 1.12 |  |  |  | EA
[7 empty] | 100.9 | 88.3472 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37799 | 37798 | 37799 | 111 | 110 | 111 |  |  | Julia Collins | -990.95 | 0.12 ‰ | 2017-12-06 | 2017-12-06 | 19 | 0 |  |  |  |  | 3.12‰ | 3.12‰ | 2.00‰ | 2.00‰ | 3.12 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -990.95 ‰ | -990.95 ‰ | 0.12 | 0.12 | -990.84 | 0.13 ‰ | -990.84 | 0.13 ‰ |  |  | 2017-12-11 | Jocelyn Turnbull | 9:07 | 0.0090462 | 0.0090462 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00905 | 0.00012 | 15097 | 2017-12-08 | 14:10 | rachel corran | 208552 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 30 | 5 | 6 | ... | 0.0087 | 0.00012 |  |  |  |  | 0.99587 | 0.99587 | Yes | Yes | [5 empty] | 0.0087 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00905 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0087 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-05 | Julia Collins | 12:39 |  | Completed |  |  |  |  | 73551 | 2000000 |  |  | 73618 73621 73624 73639 73641 73648 | 73618~73621~73624~73639~73641~73648 | 3147 |  | 3147 3147 3147 3147 3147 3147 | 2017-12-08 | 15:08 | rachel corran |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0019 | 0.0014 | 73618 |  | XCAMS | EA | 0% |  | TW3147 measurement: * no problems with XCAMS * large samples * none processed via ST, except the C6 sucrose CalAMS error adjustments: * 0% system error * factor x1.03 across all cathodes, and 1.3 for cat#23 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: !! result deviates from consensus by ~4sigma !! cause unknown blanks: * one group has been identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2018-01-10 | Water | Not Complete | JOB-208552 |  |  | 64323 | Normal | 14047/11 |  |  | 2017-11-15 | 3157 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00145 | 73621 | [26 empty] | 2017-11-15 | 1927 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00139 | 73624 | [26 empty] | 2017-11-15 | 3534 | Choose Gas Splitting | Completed
[168 empty] | 0.99773 | 0.00139 | 73639 | [26 empty] | 2017-11-17 | 13561 | Graphite | Completed
[168 empty] | 0.99885 | 0.00141 | 73641 | [26 empty] | 2017-11-21 | 5480 | C13 Measurement | Completed
[168 empty] | 0.99977 | 0.00141 | 73648 | [26 empty] | 2017-12-11 | 15097 | AMS Submission Results | Completed
[197 empty] | 2017-12-11 | 18034 | Close Job | Completed
[5 empty] | 4.778e-05 | 0.0093721 | 114.4 | 89.3163 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7106 | 7106 | 7106 | 61 | 60 | 63 | 7046 | 7167.2 | margaret norris | -587.14 | 3.11 ‰ | 2018-02-01 | 2017-08-05 | 20 | 180 |  |  |  |  | -7.57‰ | -7.57‰ | 2.00‰ | 2.00‰ | -7.57 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0032 | 0 | 0.8165 | -587.14 ‰ | -587.14 ‰ | 3.11 | 3.25 | -582.04 | 3.15 ‰ | -582.04 | 3.15 ‰ |  |  | 2018-02-07 | Cathy Ginnane | 14:31 | 0.41286 | 0.41286 | 0.00311 | 0.00325 | 0.00307 | 0.00307 | 2 |  |  | 0.41903 | 0.00133 | 15228 |  |  |  | 208688 | 0.00984 | 0.0098 | 0.00443 | 0.0044 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.29 | 0.31% | 8 | 5 | 6 | ... | 0.40299 | 0.00117 |  |  |  |  | 0.99607 | 0.99607 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 63.2310232321469861. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00984 | 0.00901 | 0.00275 | 0.39706 | 0.40‰ | 0.00295 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.41286 | 0 | 0.00311 | 0 | 0 | 0.00984 | 0 | 0.00443 | 0 | 0.39706 | 0 | 0.00295 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-01-30 | margaret norris | 10:43 |  | Completed |  |  |  |  | 73741 | 2000000 | 73742 73551 73570 73485 73420 73196 73029 73012 72897 72804 | 73742~73551~73570~73485~73420~73196~73029~73012~72897~72804 | 73774 73775 73789 73790 73791 73792 | 73774~73775~73789~73790~73791~73792 | 3151 | 3151 3147 3146 3143 3142 3139 3135 3132 3128 3126 | 3151 3151 3151 3151 3151 3151 | 2018-02-05 | 10:30 | Albert Zondervan |  | 1.00 mg | 0.13 | 1 | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 0.99939 | 0.00188 | 73774 | EA | XCAMS | ST | 0% |  | TW3151 measurement: * XCAMS: ion source a bit unstable * regular precision * large samples * no EA CalAMS error adjustments: * 0% system error * 1.04 global error multiplier * 1.4 error multiplier for cat#4 (TP73715) category line (method) unkn/blan/seco window [d] ----------¦--------------¦---------------¦---------- inorganic carbonate 5/1/1 180 water water 15/1/1 180 organic AAA 4/1/2 180 organic CE-Rafter 1/1/0 180 unknowns: * no failures blanks: * results in keeping with previous analyses secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x IAEA-C6: result consistent with consensus value * 1x FIRI-I: result consistent with consensus value primaries: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-01-31 | Water | Cancel | JOB-208688 |  | No results, only testers. TEST for SIL new C13 device; MS bottle number is wrong, shown as MS#20, changed it to actual bottle MS#26; Clarification: MS#26 was for C13 only and didn't receive its own job number. MS#20 was on graphite. Sort of fixed up job for measurement. Two samples with TP 73741 will come up in AMS- please delete cancelled job. | 64435 | Normal | 14047/12 |  |  | 2018-01-11 | 3203 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 1.0015 | 0.00187 | 73775 | ST | [25 empty] | 2018-01-11 | 1944 | Gas Splitting | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 1.0002 | 0.00188 | 73789 | EA | [25 empty] | 2018-01-11 | 3557 | Choose Gas Splitting | Completed
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 0.99915 | 0.00189 | 73790 | ST | [25 empty] | 2018-01-12 | 13751 | Graphite | Completed
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 0.99931 | 0.00179 | 73791 | EA | [25 empty] | 2018-01-24 | 5568 | C13 Measurement | Cancelled
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 | 1.0004 | 0.00187 | 73792 | ST | [25 empty] | 2018-02-07 | 15228 | AMS Submission Results | Completed
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 |  |  |  | ST | [25 empty] | 2018-03-15 | 18377 | Close Job | Completed
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 |  |  |  | EA
[163 empty] | 2017-09-06 | 0.01008 | 72897 | 3128 | 0.98 |  |  |  | EA
[163 empty] | 2017-08-24 | 0.01433 | 72804 | 3126 | 0.87 |  |  |  | ST
[7 empty] | 114.9 | 88.2886 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36002 | 36002 | 36002 | 140 | 140 | 140 |  |  | margaret norris | -988.69 | 0.20 ‰ | 2018-02-01 | 2018-02-01 | 20 | 0 |  |  |  |  | 3.98‰ | 3.98‰ | 2.00‰ | 2.00‰ | 3.98 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0032 | 0 | 0.8165 | -988.69 ‰ | -988.69 ‰ | 0.2 | 0.2 | -988.55 | 0.20 ‰ | -988.55 | 0.20 ‰ |  |  | 2018-02-07 | Cathy Ginnane | 14:31 | 0.011313 | 0.011313 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 1.5 |  |  | 0.01131 | 0.0002 | 15229 |  |  |  | 208689 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.13 | 0.02% | 18 | 5 | 5 | ... | 0.01088 | 0.00019 |  |  |  |  | 0.99607 | 0.99607 | Yes | Yes |  | Not a great blank but not unusual |  |  |  | 0.01088 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.01131 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01088 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-01-30 | margaret norris | 10:43 |  | Completed |  |  |  |  | 73742 | 2000000 |  |  | 73774 73775 73789 73790 73791 73792 | 73774~73775~73789~73790~73791~73792 | 3151 |  | 3151 3151 3151 3151 3151 3151 | 2018-02-05 | 10:30 | Albert Zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99939 | 0.00188 | 73774 |  | XCAMS | ST | 0% |  | TW3151 measurement: * XCAMS: ion source a bit unstable * regular precision * large samples * no EA CalAMS error adjustments: * 0% system error * 1.04 global error multiplier * 1.4 error multiplier for cat#4 (TP73715) category line (method) unkn/blan/seco window [d] ----------¦--------------¦---------------¦---------- inorganic carbonate 5/1/1 180 water water 15/1/1 180 organic AAA 4/1/2 180 organic CE-Rafter 1/1/0 180 unknowns: * no failures blanks: * results in keeping with previous analyses secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x IAEA-C6: result consistent with consensus value * 1x FIRI-I: result consistent with consensus value primaries: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-31 | Water |  | JOB-208689 |  |  | 64431 | Normal | 14047/11 |  |  | 2018-01-11 | 3200 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00187 | 73775 | [26 empty] | 2018-01-11 | 1945 | Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00188 | 73789 | [26 empty] | 2018-01-11 | 3558 | Choose Gas Splitting | Completed
[168 empty] | 0.99915 | 0.00189 | 73790 | [26 empty] | 2018-01-12 | 13752 | Graphite | Completed
[168 empty] | 0.99931 | 0.00179 | 73791 | [26 empty] | 2018-02-07 | 15229 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00187 | 73792 | [26 empty] | 2018-03-15 | 18378 | Close Job | Completed
[7 empty] | 122.3 | 88.36 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50278 | 50278 | 50278 | 393 | 392 | 393 |  |  | margaret norris | -998.09 | 0.09 ‰ | 2018-02-01 | 2018-02-01 | 22 | 0 |  |  |  |  | 3.17‰ | 3.17‰ | 2.00‰ | 2.00‰ | 3.17 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0032 | 0 | 0.8165 | -998.09 ‰ | -998.09 ‰ | 0.09 | 0.09 | -998.06 | 0.09 ‰ | -998.06 | 0.09 ‰ |  |  | 2018-02-07 | Cathy Ginnane | 14:31 | 0.0019132 | 0.0019132 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.4 |  |  | 0.00191 | 9e-05 | 15235 |  |  |  | 208842 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 37 | 5 | 5 | ... | 0.00184 | 9e-05 |  |  |  |  | 0.99607 | 0.99607 | Yes | Yes |  | Submitted -8 per mil C13 is the value for R14047/2, travertine. Actual C13 should have been about 2.5, so the AMS agrees well with expected value. Submitted value has now been updated in the R number file |  |  |  | 0.00184 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00191 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00184 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-01-30 | margaret norris | 10:43 |  | Completed |  |  |  |  | 73725 | 2000000 |  |  | 73774 73775 73789 73790 73791 73792 | 73774~73775~73789~73790~73791~73792 | 3151 |  | 3151 3151 3151 3151 3151 3151 | 2018-02-05 | 10:30 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99939 | 0.00188 | 73774 |  | XCAMS | ST | 0% |  | TW3151 measurement: * XCAMS: ion source a bit unstable * regular precision * large samples * no EA CalAMS error adjustments: * 0% system error * 1.04 global error multiplier * 1.4 error multiplier for cat#4 (TP73715) category line (method) unkn/blan/seco window [d] ----------¦--------------¦---------------¦---------- inorganic carbonate 5/1/1 180 water water 15/1/1 180 organic AAA 4/1/2 180 organic CE-Rafter 1/1/0 180 unknowns: * no failures blanks: * results in keeping with previous analyses secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x IAEA-C6: result consistent with consensus value * 1x FIRI-I: result consistent with consensus value primaries: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-29 | Carbonate | Complete | JOB-208842 |  |  | 64421 |  | 14047/1 |  |  | 2018-01-08 | 3197 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00187 | 73775 | [26 empty] | 2018-01-08 | 1942 | Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00188 | 73789 | [26 empty] | 2018-01-08 | 3554 | Choose Gas Splitting | Completed
[168 empty] | 0.99915 | 0.00189 | 73790 | [26 empty] | 2018-01-10 | 13735 | Graphite | Completed
[168 empty] | 0.99931 | 0.00179 | 73791 | [26 empty] | 2018-02-07 | 15235 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00187 | 73792 | [26 empty] | 2018-03-15 | 18374 | Close Job | Completed
[7 empty] | 77.6 | 90.2597 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42745 | 42744 | 42745 | 256 | 256 | 256 |  |  | julia collins | -995.11 | 0.16 ‰ | 2018-03-09 | 2018-03-09 | 38 | 0 |  |  |  |  | -1.39‰ | -1.39‰ | 2.00‰ | 2.00‰ | -1.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.11 ‰ | -995.11 ‰ | 0.16 | 0.16 | -995.05 | 0.16 ‰ | -995.05 | 0.16 ‰ |  |  | 2018-03-14 | margaret norris | 11:29 | 0.004887 | 0.004887 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00489 | 0.00016 | 15316 |  |  |  | 209040 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.49 | 0.02% | 13 | 5 | 5 | ... | 0.0047 | 0.00015 |  |  |  |  | 1.0129 | 1.0129 | Yes | Yes |  | d13C out by >3.5 per mil, not a great background |  |  |  | 0.0047 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00489 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.0047 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-02-14 | julia collins | 8:11 |  | Completed |  |  |  |  | 73840 | 2000000 |  |  | 73815 73817 73819 73846 73847 73893 | 73815~73817~73819~73846~73847~73893 | 3153 |  | 3153 3153 3153 3153 3153 3153 | 2018-03-12 | 13:18 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99955 | 0.00158 | 73815 |  | SSAMS | EA | 0% |  | TW3153 measurement: * Relative to original wheel make-up, two positions are new: TP74018,TP73880. Neither one is used in RLIMS calculations for the unknowns. * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Regular precision has been achieved. * File 'context' from ANU was used to prepare the CalAMS input file. CalAMS error adjustments: * no system error * global multiplication factor 1.05 due to excess variability * individual multiplication factor 1.3 for cat#3 (TP73904) due to excess variability primaries: * OxI cathodes: 6x no failures secondaries, relative to their respective expectation values for DEL14C: * 1x IAEA-C2: result is 2-sigma away from consensus value blanks: * Four groups have been identified to blank-correct: (ground)waters: 26x unknowns + 1x secondaries + 1x blanks window = 180 days, capturing 9x blanks textiles (via AAA): 2x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 6x blanks shells: 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 5x blanks tree bark (via CE-Rafter): 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 8x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-03-26 | Water |  | JOB-209040 |  |  | 64550 | Normal | 14047/11 |  |  | 2018-01-29 | 3225 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00161 | 73817 | [26 empty] | 2018-01-29 | 1972 | Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00159 | 73819 | [26 empty] | 2018-01-29 | 3609 | Choose Gas Splitting | Completed
[168 empty] | 0.99878 | 0.00159 | 73846 | [26 empty] | 2018-01-30 | 13850 | Graphite | Completed
[168 empty] | 0.99875 | 0.00158 | 73847 | [26 empty] | 2018-03-14 | 15316 | AMS Submission Results | Completed
[168 empty] | 1.0027 | 0.00165 | 73893 | [26 empty] | 2018-03-15 | 18395 | Close Job | Completed
[5 empty] | 4.896e-05 | 0.0084918 | 82.5 | 90.628 | 0.069 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7236 | 7235 | 7236 | 55 | 54 | 58 | 7181.2 | 7290.5 | julia collins | -593.73 | 2.76 ‰ | 2018-03-09 | 2017-09-10 | 38 | 180 |  |  |  |  | -5.45‰ | -5.45‰ | 2.00‰ | 2.00‰ | -5.45 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -593.73 ‰ | -593.73 ‰ | 2.76 | 2.91 | -588.72 | 2.80 ‰ | -588.72 | 2.80 ‰ |  |  | 2018-03-14 | margaret norris | 11:29 | 0.40627 | 0.40627 | 0.00276 | 0.00291 | 0.00271 | 0.00271 | 2.1 |  |  | 0.41181 | 0.00116 | 15317 |  |  |  | 209041 | 0.00874 | 0.0087 | 0.00393 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 40.63 | 0.28% | 29 | 5 | 5 | ... | 0.39605 | 0.00099 |  |  |  |  | 1.0129 | 1.0129 | Yes | Yes |  | result is 2-sigma away from consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.5815793517541047. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00874 | 0.00817 | 0.00281 | 0.39072 | 0.39‰ | 0.00261 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.40627 | 0 | 0.00276 | 0 | 0 | 0.00874 | 0 | 0.00393 | 0 | 0.39072 | 0 | 0.00261 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-02-14 | julia collins | 8:11 |  | Completed |  |  |  |  | 73841 | 2000000 | 73840 73742 73551 73570 73485 73420 73196 73029 73012 | 73840~73742~73551~73570~73485~73420~73196~73029~73012 | 73815 73817 73819 73846 73847 73893 | 73815~73817~73819~73846~73847~73893 | 3153 | 3153 3151 3147 3146 3143 3142 3139 3135 3132 | 3153 3153 3153 3153 3153 3153 | 2018-03-12 | 13:18 | Albert Zondervan |  | 1.06 mg | 0.13 | 1.06 | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 0.99955 | 0.00158 | 73815 | EA | SSAMS | EA | 0% |  | TW3153 measurement: * Relative to original wheel make-up, two positions are new: TP74018,TP73880. Neither one is used in RLIMS calculations for the unknowns. * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Regular precision has been achieved. * File 'context' from ANU was used to prepare the CalAMS input file. CalAMS error adjustments: * no system error * global multiplication factor 1.05 due to excess variability * individual multiplication factor 1.3 for cat#3 (TP73904) due to excess variability primaries: * OxI cathodes: 6x no failures secondaries, relative to their respective expectation values for DEL14C: * 1x IAEA-C2: result is 2-sigma away from consensus value blanks: * Four groups have been identified to blank-correct: (ground)waters: 26x unknowns + 1x secondaries + 1x blanks window = 180 days, capturing 9x blanks textiles (via AAA): 2x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 6x blanks shells: 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 5x blanks tree bark (via CE-Rafter): 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 8x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-03-26 | Water |  | JOB-209041 |  |  | 64552 | Normal | 14047/12 |  |  | 2018-01-29 | 3224 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 0.9993 | 0.00161 | 73817 | ST | [25 empty] | 2018-01-29 | 1973 | Gas Splitting | Completed
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 1.0011 | 0.00159 | 73819 | EA | [25 empty] | 2018-01-29 | 3611 | Choose Gas Splitting | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 0.99878 | 0.00159 | 73846 | ST | [25 empty] | 2018-01-30 | 13851 | Graphite | Completed
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 0.99875 | 0.00158 | 73847 | ST | [25 empty] | 2018-03-14 | 15317 | AMS Submission Results | Completed
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 | 1.0027 | 0.00165 | 73893 | EA | [25 empty] | 2018-03-15 | 18396 | Close Job | Completed
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 |  |  |  | EA
[163 empty] | 2017-09-29 | 0.01377 | 73029 | 3135 | 0.67 |  |  |  | ST
[163 empty] | 2017-09-15 | 0.00894 | 73012 | 3132 | 0.74 |  |  |  | EA
[7 empty] | 82.3 | 90.1846 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42542 | 42542 | 42542 | 216 | 216 | 216 |  |  | julia collins | -994.99 | 0.13 ‰ | 2018-03-09 | 2018-03-09 | 29 | 0 |  |  |  |  | 5.30‰ | 5.30‰ | 2.00‰ | 2.00‰ | 5.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.89443 | -994.99 ‰ | -994.99 ‰ | 0.13 | 0.13 | -994.93 | 0.14 ‰ | -994.93 | 0.14 ‰ |  |  | 2018-03-29 | Jocelyn Turnbull | 10:31 | 0.0050118 | 0.0050118 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2.3 |  |  | 0.00501 | 0.00013 | 15370 | 2018-03-26 | 11:08 | rachel corran | 208857 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.5 | 0.01% | 38 | 5 | 5 | ... | 0.00482 | 0.00013 |  |  |  |  | 1.0188 | 1.0188 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00482 | 0.00‰ | 0.00013 | 0.00‰ | No | 1 | 1.0007 | 0.00106 | 0.00106 | 0.00501 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00482 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-02-21 | julia collins | 11:30 |  | Completed |  |  |  |  | 73885 | 2000000 |  |  | 73946 73947 73948 73949 73950 | 73946~73947~73948~73949~73950 | 3155 |  | 3155 3155 3155 3155 3155 |  |  |  |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0034 | 0.00239 | 73946 |  | SSAMS | FLASK | 0.14% |  | TW3155 measurement: * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Previously (during measurement of TW3154) signal from the low-energy side Faraday cups was lost. While these signals are good for exercising quality control, they are not essential in deriving the final ratio-to-standard values. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.14% Note that many C13 values were measured at NIWA, and only cat#0,3,38,39 were measured here JCT - I believe that the C13s were measured on the 14CO samples before they were diluted at NIWA, and therefore should not be reflective of the AMS measured value - and the AMS value should be close to the kapuni d13C since that is the dilutant gas. primaries: * OxI cathodes: 6x no failures secondaries: 0 secondaries in wheel blanks: * 2 groups have been identified to blank-correct: waters 5x unknowns only window = 180 days, capturing 9 blanks, no water blanks were in this wheel CO2 from atmospheric CO: 27x unknowns only window = use of dilution blanks used for CO dilution (capturing only the 4x blanks measured in ~2016) * Instead of the usual Kapuni cat#38 for tuning, a IAEA-C1 marble blank (carbonate line processed) was put there. * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-30 | Carbonate |  | JOB-208857 |  | being used as a background target instead of kapuni | 64582 |  | 14047/1 |  |  | 2018-02-07 | 3242 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99829 | 0.00235 | 73947 | [26 empty] | 2018-02-07 | 1982 | Gas Splitting | Completed
[168 empty] | 0.99999 | 0.00236 | 73948 | [26 empty] | 2018-02-07 | 3620 | Choose Gas Splitting | Completed
[168 empty] | 1.0019 | 0.00237 | 73949 | [26 empty] | 2018-02-08 | 13895 | Graphite | Completed
[168 empty] | 0.99994 | 0.00237 | 73950 | [26 empty] | 2018-03-29 | 15370 | AMS Submission Results | Completed
[197 empty] | 2018-03-29 | 18449 | Close Job | Completed
[7 empty] | 71.9 | 90.5707 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34269 | 34269 | 34269 | 125 | 124 | 125 |  |  | julia collins | -985.96 | 0.22 ‰ | 2018-04-13 | 2018-04-13 | 52 | 0 |  |  |  |  | 3.39‰ | 3.39‰ | 2.00‰ | 2.00‰ | 3.39 ‰ | 2.00 ‰ | AMS | 2.22‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -985.96 ‰ | -985.96 ‰ | 0.22 | 0.22 | -985.79 | 0.22 ‰ | -985.79 | 0.22 ‰ |  |  | 2018-04-13 | Jocelyn Turnbull | 15:08 | 0.014037 | 0.014037 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.1 |  |  | 0.01404 | 0.00022 | 15427 |  |  |  | 209076 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.4 | 0.02% | 8 | 5 | 6 | ... | 0.0135 | 0.00021 |  |  |  |  | 1.0212 | 1.0212 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.0135 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01404 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.0135 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-03-14 | julia collins | 7:50 |  | Completed |  |  |  |  | 73983 | 2000000 |  |  | 73973 73985 73986 73987 73988 73989 | 73973~73985~73986~73987~73988~73989 | 3156 |  | 3156 3156 3156 3156 3156 3156 | 2018-04-13 | 14:50 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0009 | 0.00162 | 73973 |  | SSAMS | ST | 0% |  | TW3156 measurement: * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Previously (during measurement of TW3154) signal from the low-energy side Faraday cups was lost. While these signals are good for exercising quality control, they are not essential in deriving the final ratio-to-standard values. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, both agreeing with their consensus value blanks: * 2 groups have been identified to blank-correct: waters 24x unknowns + 1x secondaries + 1x blank window = 180 days organics: 4x unknowns + 2x secondaries + 1x blank window = 180 days unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2018-03-28 | Water |  | JOB-209076 |  |  | 64622 | Normal | 14047/11 |  |  | 2018-02-19 | 3270 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99806 | 0.00156 | 73985 | [26 empty] | 2018-02-19 | 2004 | Gas Splitting | Completed
[168 empty] | 0.99949 | 0.00156 | 73986 | [26 empty] | 2018-02-19 | 3648 | Choose Gas Splitting | Completed
[168 empty] | 0.99917 | 0.00154 | 73987 | [26 empty] | 2018-02-20 | 13993 | Graphite | Completed
[168 empty] | 1.0012 | 0.00155 | 73988 | [26 empty] | 2018-02-23 | 5642 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00153 | 73989 | [26 empty] | 2018-04-13 | 15427 | AMS Submission Results | Completed
[197 empty] | 2018-04-13 | 18544 | Close Job | Completed
[5 empty] | 5.1888e-05 | 0.008314 | 77.3 | 91.6075 | 0.0706 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7071 | 7070 | 7071 | 53 | 53 | 56 | 7018 | 7124 | julia collins | -585.31 | 2.74 ‰ | 2018-04-13 | 2017-10-15 | 52 | 180 |  |  |  |  | -7.97‰ | -7.97‰ | 2.00‰ | 2.00‰ | -7.97 ‰ | 2.00 ‰ | AMS | -8.46‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -585.31 ‰ | -585.31 ‰ | 2.74 | 2.89 | -580.19 | 2.77 ‰ | -580.19 | 2.77 ‰ |  |  | 2018-04-13 | Jocelyn Turnbull | 15:08 | 0.41469 | 0.41469 | 0.00274 | 0.00289 | 0.00268 | 0.00268 | 1.7 |  |  | 0.42012 | 0.00119 | 15428 |  |  |  | 209077 | 0.00868 | 0.0087 | 0.00391 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.47 | 0.27% | 27 | 5 | 6 | ... | 0.40404 | 0.00102 |  |  |  |  | 1.0212 | 1.0212 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 56.0062264602487563. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00868 | 0.00799 | 0.00294 | 0.39882 | 0.40‰ | 0.00258 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41469 | 0 | 0.00274 | 0 | 0 | 0.00868 | 0 | 0.00391 | 0 | 0.39882 | 0 | 0.00258 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-03-14 | julia collins | 7:50 |  | Completed |  |  |  |  | 73984 | 2000000 | 73983 73840 73742 73551 73570 73485 73420 73196 | 73983~73840~73742~73551~73570~73485~73420~73196 | 73973 73985 73986 73987 73988 73989 | 73973~73985~73986~73987~73988~73989 | 3156 | 3156 3153 3151 3147 3146 3143 3142 3139 | 3156 3156 3156 3156 3156 3156 | 2018-04-13 | 14:50 | Albert Zondervan |  | 1.12 mg | 0.13 | 1.12 | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 1.0009 | 0.00162 | 73973 | EA | SSAMS | ST | 0% |  | TW3156 measurement: * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Previously (during measurement of TW3154) signal from the low-energy side Faraday cups was lost. While these signals are good for exercising quality control, they are not essential in deriving the final ratio-to-standard values. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, both agreeing with their consensus value blanks: * 2 groups have been identified to blank-correct: waters 24x unknowns + 1x secondaries + 1x blank window = 180 days organics: 4x unknowns + 2x secondaries + 1x blank window = 180 days unknowns: * no failures | -8.5 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-03-28 | Water | Complete | JOB-209077 |  |  | 64623 | Normal | 14047/12 |  |  | 2018-02-19 | 3271 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 0.99806 | 0.00156 | 73985 | ST | [25 empty] | 2018-02-19 | 2003 | Gas Splitting | Completed
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 0.99949 | 0.00156 | 73986 | ST | [25 empty] | 2018-02-19 | 3646 | Choose Gas Splitting | Completed
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 0.99917 | 0.00154 | 73987 | EA | [25 empty] | 2018-02-20 | 13994 | Graphite | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 1.0012 | 0.00155 | 73988 | EA | [25 empty] | 2018-02-23 | 5650 | C13 Measurement | Completed
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 | 1.0011 | 0.00153 | 73989 | ST | [25 empty] | 2018-04-13 | 15428 | AMS Submission Results | Completed
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 |  |  |  | EA | [25 empty] | 2018-04-13 | 18545 | Close Job | Completed
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 |  |  |  | ST
[7 empty] | 78.1 | 91.5912 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36736 | 36736 | 36736 | 121 | 121 | 121 |  |  | julia collins | -989.67 | 0.16 ‰ | 2018-05-13 | 2018-05-13 | 51 | 0 |  |  |  |  | 3.87‰ | 3.87‰ | 2.00‰ | 2.00‰ | 3.87 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0 | 0.8165 | -989.67 ‰ | -989.67 ‰ | 0.16 | 0.16 | -989.55 | 0.16 ‰ | -989.55 | 0.16 ‰ |  |  | 2018-05-15 | Jocelyn Turnbull | 11:29 | 0.010325 | 0.010325 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.8 |  |  | 0.01033 | 0.00016 | 15535 |  |  |  | 209461 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.03 | 0.02% | 20 | 5 | 5 | ... | 0.00993 | 0.00015 |  |  |  |  | 1.0332 | 1.0332 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00993 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.01033 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00993 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-09 | julia collins | 11:04 |  | Completed |  |  |  |  | 74073 | 2000000 |  |  | 74116 74117 74120 74121 74122 74123 | 74116~74117~74120~74121~74122~74123 | 3159 |  | 3159 3159 3159 3159 3159 3159 | 2018-05-15 | 10:39 | Albert Zondervan |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0006 | 0.00152 | 74116 |  | SSAMS | ST | 0% |  | TW3159 measurement: * The wheel was measured on 12+13 May 2018 with SSAMS at ANU because of a continued problem with our ion source. NOTE: In RLIMS, the field 'Date Run' has been imported with incorrect values. JCT - corrected to 13/5/2018 data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 3 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 8x unknowns + 1x secondaries + 1x blank window = 180 days carbonates 4x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via AAA: 14x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period unknowns: * no failures JCT - blanks are not at all good, but to be expected for having sat for so long before measurement. Secondaries look fine with the corrections assigned and no really old samples, so all ok. |  | RRL | [5 empty] | Background Inorganic |  | 2018-04-18 | Water |  | JOB-209461 |  |  | 64783 | Normal | 14047/11 | KIP |  | 2018-03-12 | 3295 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99843 | 0.00153 | 74117 | [26 empty] | 2018-03-12 | 2025 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00151 | 74120 | [26 empty] | 2018-03-12 | 3670 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00151 | 74121 | [26 empty] | 2018-03-23 | 14083 | Graphite | Completed
[168 empty] | 0.99815 | 0.00151 | 74122 | [26 empty] | 2018-05-15 | 15535 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.0015 | 74123 | [26 empty] | 2018-05-15 | 18743 | Close Job | Completed
[5 empty] | 4.9235e-05 | 0.0085144 | 82.8 | 92.6336 | 0.0703 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7083 | 7083 | 7083 | 53 | 53 | 56 | 7030 | 7137 | julia collins | -585.95 | 2.76 ‰ | 2018-05-13 | 2017-11-14 | 51 | 180 |  |  |  |  | -7.42‰ | -7.42‰ | 2.00‰ | 2.00‰ | -7.42 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0 | 0.8165 | -585.95 ‰ | -585.95 ‰ | 2.76 | 2.91 | -580.84 | 2.79 ‰ | -580.84 | 2.79 ‰ |  |  | 2018-05-15 | Jocelyn Turnbull | 11:29 | 0.41405 | 0.41405 | 0.00276 | 0.00291 | 0.0027 | 0.0027 | 1.9 |  |  | 0.41957 | 0.00114 | 15536 |  |  |  | 209462 | 0.00882 | 0.0088 | 0.00397 | 0.004 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.4 | 0.28% | 1 | 5 | 5 | ... | 0.40351 | 0.00096 |  |  |  |  | 1.0332 | 1.0332 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. | result consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 56.4559515318325906. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00882 | 0.00819 | 0.00278 | 0.3982 | 0.40‰ | 0.0026 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.41405 | 0 | 0.00276 | 0 | 0 | 0.00882 | 0 | 0.00397 | 0 | 0.3982 | 0 | 0.0026 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-09 | julia collins | 11:04 |  | Completed |  |  |  |  | 74074 | 2000000 | 73983 74073 73840 73742 73551 73570 73485 73420 73196 | 73983~74073~73840~73742~73551~73570~73485~73420~73196 | 74116 74117 74120 74121 74122 74123 | 74116~74117~74120~74121~74122~74123 | 3159 | 3156 3159 3153 3151 3147 3146 3143 3142 3139 | 3159 3159 3159 3159 3159 3159 | 2018-05-15 | 10:39 | Albert Zondervan |  | 1.05 mg | 0.13 | 1.05 | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 | 1.0006 | 0.00152 | 74116 | EA | SSAMS | ST | 0% |  | TW3159 measurement: * The wheel was measured on 12+13 May 2018 with SSAMS at ANU because of a continued problem with our ion source. NOTE: In RLIMS, the field 'Date Run' has been imported with incorrect values. JCT - corrected to 13/5/2018 data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 3 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 8x unknowns + 1x secondaries + 1x blank window = 180 days carbonates 4x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via AAA: 14x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period unknowns: * no failures JCT - blanks are not at all good, but to be expected for having sat for so long before measurement. Secondaries look fine with the corrections assigned and no really old samples, so all ok. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-04-18 | Water | Complete | JOB-209462 |  |  | 64784 | Normal | 14047/12 | KIP |  | 2018-03-12 | 3296 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 0.99843 | 0.00153 | 74117 | ST | [25 empty] | 2018-03-12 | 2026 | Gas Splitting | Completed
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 1.0013 | 0.00151 | 74120 | ST | [25 empty] | 2018-03-12 | 3672 | Choose Gas Splitting | Completed
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 1.0007 | 0.00151 | 74121 | EA | [25 empty] | 2018-03-23 | 14084 | Graphite | Completed
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 0.99815 | 0.00151 | 74122 | EA | [25 empty] | 2018-05-15 | 15536 | AMS Submission Results | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 | 1.0008 | 0.0015 | 74123 | ST | [25 empty] | 2018-05-15 | 18744 | Close Job | Completed
[163 empty] | 2017-11-14 | 0.01099 | 73485 | 3143 | 1.1 |  |  |  | EA
[163 empty] | 2017-11-08 | 0.00585 | 73420 | 3142 | 1.12 |  |  |  | ST
[163 empty] | 2017-10-24 | 0.00721 | 73196 | 3139 | 1.23 |  |  |  | ST
[7 empty] | 80.5 | 91.6998 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40797 | 40796 | 40797 | 161 | 160 | 161 |  |  | julia collins | -993.77 | 0.12 ‰ | 2018-05-13 | 2018-05-13 | 37 | 0 |  |  |  |  | 2.69‰ | 2.69‰ | 2.00‰ | 2.00‰ | 2.69 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0 | 0.8165 | -993.77 ‰ | -993.77 ‰ | 0.12 | 0.12 | -993.69 | 0.13 ‰ | -993.69 | 0.13 ‰ |  |  | 2018-05-15 | Jocelyn Turnbull | 11:29 | 0.0062284 | 0.0062284 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.2 |  |  | 0.00623 | 0.00012 | 15564 |  |  |  | 209621 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 13 | 5 | 5 | ... | 0.00599 | 0.00012 |  |  |  |  | 1.0332 | 1.0332 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00599 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00623 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00599 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-09 | julia collins | 11:04 |  | Completed |  |  |  |  | 74137 | 2000000 |  |  | 74116 74117 74120 74121 74122 74123 | 74116~74117~74120~74121~74122~74123 | 3159 |  | 3159 3159 3159 3159 3159 3159 | 2018-05-15 | 10:39 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0006 | 0.00152 | 74116 |  | SSAMS | ST | 0% |  | TW3159 measurement: * The wheel was measured on 12+13 May 2018 with SSAMS at ANU because of a continued problem with our ion source. NOTE: In RLIMS, the field 'Date Run' has been imported with incorrect values. JCT - corrected to 13/5/2018 data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 3 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 8x unknowns + 1x secondaries + 1x blank window = 180 days carbonates 4x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via AAA: 14x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period unknowns: * no failures JCT - blanks are not at all good, but to be expected for having sat for so long before measurement. Secondaries look fine with the corrections assigned and no really old samples, so all ok. |  | RRL | [5 empty] | Background Inorganic |  | 2018-04-18 | Carbonate | Complete | JOB-209621 |  | repat job for 209537 | 64782 |  | 14047/1 | KIP |  | 2018-04-04 | 3317 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99843 | 0.00153 | 74117 | [26 empty] | 2018-04-04 | 2045 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00151 | 74120 | [26 empty] | 2018-04-04 | 3701 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00151 | 74121 | [26 empty] | 2018-04-06 | 14147 | Graphite | Completed
[168 empty] | 0.99815 | 0.00151 | 74122 | [26 empty] | 2018-05-15 | 15564 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.0015 | 74123 | [26 empty] | 2018-05-15 | 18745 | Close Job | Completed
[7 empty] | 72.4 | 91.0234 | 0.0017 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37008 | 37007 | 37008 | 134 | 133 | 134 |  |  | Julia Collins | -990.02 | 0.17 ‰ | 2018-05-24 | 2018-05-24 | 37 | 0 |  |  |  |  | 6.65‰ | 6.65‰ | 2.00‰ | 2.00‰ | 6.65 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0019 | 0 | 0.8165 | -990.02 ‰ | -990.02 ‰ | 0.17 | 0.17 | -989.89 | 0.17 ‰ | -989.89 | 0.17 ‰ |  |  | 2018-05-24 | Jocelyn Turnbull | 16:18 | 0.009982 | 0.009982 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.00998 | 0.00017 | 15606 | 2018-05-24 | 15:09 | Albert Zondervan | 209652 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1 | 0.02% | 22 | 5 | 5 | ... | 0.0096 | 0.00016 |  |  |  |  | 1.0296 | 1.0296 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.0096 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00998 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.0096 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-27 | Julia Collins | 10:42 |  | Completed |  |  |  |  | 74147 | 2000000 |  |  | 74152 74153 74162 74163 74171 74172 | 74152~74153~74162~74163~74171~74172 | 3160 |  | 3160 3160 3160 3160 3160 3160 | 2018-05-24 | 15:26 | Albert Zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0007 | 0.00171 | 74152 |  | SSAMS | EA | 0% |  | TW3160 measurement: * The wheel was measured on 22/23ish May 2018 with SSAMS at ANU because of a continued problem with our ion source. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 4x unknowns + 1x secondaries + 1x blank window = 180 days organics via AAA: 23x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via Rafter Cellulose: 1x unknowns + 0x secondaries + 0x blank window = covering beyond ANU-SSAMS period as no rafter cellulose blanks during this time unknowns: * TP74174 (cat#32) gives a low output, different trend in 13/12, and large deviation from IRMS del13C: REJECT |  | RRL | [5 empty] | Background Inorganic |  | 2018-06-06 | Water |  | JOB-209652 |  |  | 64855 | Normal | 14047/11 |  |  | 2018-04-11 | 3325 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99763 | 0.00173 | 74153 | [26 empty] | 2018-04-11 | 2055 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.0017 | 74162 | [26 empty] | 2018-04-11 | 3715 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00173 | 74163 | [26 empty] | 2018-04-20 | 14157 | Graphite | Completed
[168 empty] | 0.99872 | 0.00171 | 74171 | [26 empty] | 2018-05-24 | 15606 | AMS Submission Results | Completed
[168 empty] | 1.0016 | 0.00173 | 74172 | [26 empty] | 2018-05-28 | 18836 | Close Job | Completed
[5 empty] | 5.935e-05 | 0.0091091 | 70.8 | 91.9686 | 0.0694 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7208 | 7207 | 7208 | 58 | 58 | 61 | 7150 | 7266.3 | Julia Collins | -592.33 | 2.95 ‰ | 2018-05-24 | 2017-11-25 | 37 | 180 |  |  |  |  | -3.70‰ | -3.70‰ | 2.00‰ | 2.00‰ | -3.70 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0019 | 0 | 0.8165 | -592.33 ‰ | -592.33 ‰ | 2.95 | 3.09 | -587.29 | 2.99 ‰ | -587.29 | 2.99 ‰ |  |  | 2018-05-24 | Jocelyn Turnbull | 16:18 | 0.40767 | 0.40767 | 0.00295 | 0.00309 | 0.0029 | 0.0029 | 1.7 |  |  | 0.41354 | 0.00125 | 15607 | 2018-05-24 | 15:09 | Albert Zondervan | 209653 | 0.00928 | 0.0093 | 0.00418 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 40.77 | 0.29% | 8 | 5 | 5 | ... | 0.39771 | 0.00109 |  |  |  |  | 1.0296 | 1.0296 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 60.8594736862566621. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00928 | 0.00876 | 0.00274 | 0.39207 | 0.39‰ | 0.00279 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.40767 | 0 | 0.00295 | 0 | 0 | 0.00928 | 0 | 0.00418 | 0 | 0.39207 | 0 | 0.00279 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-27 | Julia Collins | 10:42 |  | Completed |  |  |  |  | 74148 | 2000000 | 74147 74073 73983 73840 73742 73551 73570 | 74147~74073~73983~73840~73742~73551~73570 | 74152 74153 74162 74163 74171 74172 | 74152~74153~74162~74163~74171~74172 | 3160 | 3160 3159 3156 3153 3151 3147 3146 | 3160 3160 3160 3160 3160 3160 | 2018-05-24 | 15:26 | Albert Zondervan |  | 1.20 mg | 0.13 | 1.2 | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 1.0007 | 0.00171 | 74152 | EA | SSAMS | EA | 0% |  | TW3160 measurement: * The wheel was measured on 22/23ish May 2018 with SSAMS at ANU because of a continued problem with our ion source. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 4x unknowns + 1x secondaries + 1x blank window = 180 days organics via AAA: 23x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via Rafter Cellulose: 1x unknowns + 0x secondaries + 0x blank window = covering beyond ANU-SSAMS period as no rafter cellulose blanks during this time unknowns: * TP74174 (cat#32) gives a low output, different trend in 13/12, and large deviation from IRMS del13C: REJECT |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-06-06 | Water |  | JOB-209653 |  |  | 64856 | Normal | 14047/12 |  |  | 2018-04-11 | 3326 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 | 0.99763 | 0.00173 | 74153 | EA | [25 empty] | 2018-04-11 | 2054 | Gas Splitting | Completed
[163 empty] | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 1.0003 | 0.0017 | 74162 | ST | [25 empty] | 2018-04-11 | 3713 | Choose Gas Splitting | Completed
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 1.001 | 0.00173 | 74163 | EA | [25 empty] | 2018-04-20 | 14158 | Graphite | Completed
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 0.99872 | 0.00171 | 74171 | ST | [25 empty] | 2018-05-24 | 15607 | AMS Submission Results | Completed
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 | 1.0016 | 0.00173 | 74172 | ST | [25 empty] | 2018-05-28 | 18837 | Close Job | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 |  |  |  | EA
[7 empty] | 117.9 | 88.2268 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45610 | 45610 | 45610 | 244 | 244 | 244 |  |  | margaret norris | -996.58 | 0.10 ‰ | 2018-05-23 | 2018-05-23 | 19 | 0 |  |  |  |  | 2.03‰ | 2.03‰ | 2.00‰ | 2.00‰ | 2.03 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -996.58 ‰ | -996.58 ‰ | 0.1 | 0.1 | -996.54 | 0.11 ‰ | -996.54 | 0.11 ‰ |  |  | 2018-05-25 | Jocelyn Turnbull | 10:42 | 0.0034209 | 0.0034209 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.1 |  |  | 0.00342 | 0.0001 | 15610 |  |  |  | 209742 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.34 | 0.01% | 8 | 5 | 6 | ... | 0.00329 | 0.0001 |  |  |  |  | 0.99344 | 0.99344 | Yes | Yes | [5 empty] | 0.00329 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00342 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00329 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-05-08 | margaret norris | 11:14 |  | Completed |  |  |  |  | 74233 | 2000000 |  |  | 74217 74224 74225 74226 74228 74230 | 74217~74224~74225~74226~74228~74230 | 3161 |  | 3161 3161 3161 3161 3161 3161 | 2018-05-24 | 18:38 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 1.0021 | 0.00168 | 74217 |  | XCAMS | ST | 0% |  | TW3161 measurement: * first wheel with unknowns since TW3152 * regular precision achieved with ease * no problems with XCAMS * The significant difference between AMS and IRMS del13C values for 7 out of 8 cathodes in this wheel correlates with beam intensity. A similarly significant difference for one other cathode does not show such a connection. There is no connection with sample type or size. The runs for each cathode separately shows the connection with beam intensity. This is fractionation in XCAMS due to a combination of factors to do with beam optical transmission and tuning. CalAMS error adjustments: * 0% system error * factor 1.3 for TP74242 (cat#29) due to excess run-to-run variance caused by one outlier run primaries: * OxI cathodes: 6x via ST no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates: 30x unknowns + 1x secondaries + 1x blanks because graphite was stored for a long time, only the blank in this wheel was used for blank correction unknowns: * no failures, but some have a bit large IRMS-AMS difference in del13C |  | RRL | [5 empty] | Background Inorganic |  | 2018-05-08 | Carbonate |  | JOB-209742 |  |  | 64895 |  | 14047/1 |  |  | 2018-04-26 | 3333 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99804 | 0.0017 | 74224 | [26 empty] | 2018-04-26 | 2066 | Gas Splitting | Completed
[168 empty] | 0.99917 | 0.00167 | 74225 | [26 empty] | 2018-04-26 | 3727 | Choose Gas Splitting | Completed
[168 empty] | 0.9995 | 0.00173 | 74226 | [26 empty] | 2018-05-04 | 14243 | Graphite | Completed
[168 empty] | 1.0001 | 0.00171 | 74228 | [26 empty] | 2018-05-04 | 5776 | C13 Measurement | Cancelled
[168 empty] | 1.001 | 0.0017 | 74230 | [26 empty] | 2018-05-25 | 15610 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18826 | Close Job | Completed
[5 empty] | 5.1544e-05 | 0.008557 | 84.9 | 89.4267 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7135 | 7135 | 7135 | 53 | 53 | 56 | 7082.5 | 7188.5 | margaret norris | -588.63 | 2.72 ‰ | 2018-05-25 | 2017-11-26 | 21 | 180 |  |  |  |  | -7.76‰ | -7.76‰ | 2.00‰ | 2.00‰ | -7.76 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -588.63 ‰ | -588.63 ‰ | 2.72 | 2.87 | -583.55 | 2.75 ‰ | -583.55 | 2.75 ‰ |  |  | 2018-05-28 | Jocelyn Turnbull | 14:39 | 0.41137 | 0.41137 | 0.00272 | 0.00287 | 0.00266 | 0.00266 | 2.5 |  |  | 0.41699 | 0.00093 | 15703 |  |  |  | 209759 | 0.00894 | 0.0089 | 0.00402 | 0.004 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.14 | 0.27% | 29 | 5 | 6 | ... | 0.40103 | 0.00073 |  |  |  |  | 0.99711 | 0.99711 | Yes | Yes |  | result is consistent with expectation value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 56.0182157920192598. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00894 | 0.00823 | 0.0027 | 0.39563 | 0.40‰ | 0.00256 | 0.00‰ | Yes | 1 | 1 | 0.00046 | 0.00046 | 0.41137 | 0 | 0.00272 | 0 | 0 | 0.00894 | 0 | 0.00402 | 0 | 0.39563 | 0 | 0.00256 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-05-24 | margaret norris | 10:37 |  | Completed |  |  |  |  | 74232 | 2000000 | 74147 74073 73983 73840 73742 73551 73570 74231 | 74147~74073~73983~73840~73742~73551~73570~74231 | 74282 74284 74285 74286 74287 74288 | 74282~74284~74285~74286~74287~74288 | 3165 | 3160 3159 3156 3153 3151 3147 3146 3165 | 3165 3165 3165 3165 3165 3165 | 2018-05-28 | 12:22 | Albert Zondervan |  | 0.81 mg | 0.13 | 0.81 | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 1.0009 | 0.00114 | 74282 | EA | XCAMS | ST | 0% |  | TW3165 measurement: * no problems with XCAMS * The complete tune prior measurement has resolved two issues relative to the previous wheel (TW3161): The graphite cathodes last longer (= higher efficiency) due to the a reduced sputter rate. The mode=5 ratio shows no beam current dependency: Runs with sputtering into the stubby do not have to be rejected. * Measurement was extended beyond the demand for regular precision, to prove that high precision can be achieved on 0.5 mgC graphites. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x via ST no failures secondaries: * 1x IAEA-C2 via water-line: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates: 6x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 2x blanks waters: 1x unknowns + 1x secondaries + 1x blanks 180 days, including SSAMS runs: capturing 7x blanks organics via AAA: 20x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-06-21 | Water |  | JOB-209759 |  |  | 64962 | Normal | 14047/12 |  |  | 2018-05-02 | 3353 | Water CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 0.99818 | 0.00115 | 74284 | EA | [25 empty] | 2018-05-03 | 3356 | Water CO2 Evolution | Completed
[163 empty] | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 | 1.0001 | 0.00113 | 74285 | ST | [25 empty] | 2018-05-03 | 2079 | Gas Splitting | Completed
[163 empty] | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 0.99929 | 0.00111 | 74286 | EA | [25 empty] | 2018-05-03 | 3748 | Choose Gas Splitting | Completed
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 1.0002 | 0.00114 | 74287 | ST | [25 empty] | 2018-05-04 | 14242 | Graphite | Completed
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 | 1.0013 | 0.00114 | 74288 | ST | [25 empty] | 2018-05-28 | 15703 | AMS Submission Results | Completed
[163 empty] | 2017-12-06 | 0.0087 | 73551 | 3147 | 0.83 |  |  |  | EA | [25 empty] | 2018-05-28 | 18942 | Close Job | Completed
[163 empty] | 2017-12-04 | 0.00764 | 73570 | 3146 | 0.96 |  |  |  | ST
[7 empty] | 71.7 | 88.9012 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40018 | 40017 | 40018 | 122 | 121 | 122 |  |  | margaret norris | -993.14 | 0.10 ‰ | 2018-05-25 | 2017-11-26 | 21 | 180 |  |  |  |  | -1.89‰ | -1.89‰ | 2.00‰ | 2.00‰ | -1.89 ‰ | 2.00 ‰ | AMS | 1.86‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -993.14 ‰ | -993.14 ‰ | 0.1 | 0.1 | -993.05 | 0.11 ‰ | -993.05 | 0.11 ‰ |  |  | 2018-05-28 | Jocelyn Turnbull | 14:39 | 0.0068626 | 0.0068626 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00686 | 0.0001 | 15708 |  |  |  | 209758 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.69 | 0.01% | 15 | 5 | 6 | ... | 0.0066 | 0.0001 |  |  |  |  | 0.99711 | 0.99711 | Yes | Yes | [5 empty] | 0.0066 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00046 | 0.00046 | 0.00686 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.0066 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-05-24 | margaret norris | 10:37 |  | Completed |  |  |  |  | 74231 | 2000000 |  |  | 74282 74284 74285 74286 74287 74288 | 74282~74284~74285~74286~74287~74288 | 3165 |  | 3165 3165 3165 3165 3165 3165 | 2018-05-28 | 12:22 | Albert Zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0009 | 0.00114 | 74282 |  | XCAMS | ST | 0% |  | TW3165 measurement: * no problems with XCAMS * The complete tune prior measurement has resolved two issues relative to the previous wheel (TW3161): The graphite cathodes last longer (= higher efficiency) due to the a reduced sputter rate. The mode=5 ratio shows no beam current dependency: Runs with sputtering into the stubby do not have to be rejected. * Measurement was extended beyond the demand for regular precision, to prove that high precision can be achieved on 0.5 mgC graphites. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x via ST no failures secondaries: * 1x IAEA-C2 via water-line: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates: 6x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 2x blanks waters: 1x unknowns + 1x secondaries + 1x blanks 180 days, including SSAMS runs: capturing 7x blanks organics via AAA: 20x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures | 1.9 | RRL | [5 empty] | Background Inorganic |  | 2018-06-21 | Water |  | JOB-209758 |  | d13C lower than it should be | 64961 | Normal | 14047/11 |  |  | 2018-05-02 | 3352 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99818 | 0.00115 | 74284 | [26 empty] | 2018-05-02 | 2078 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00113 | 74285 | [26 empty] | 2018-05-02 | 3746 | Choose Gas Splitting | Completed
[168 empty] | 0.99929 | 0.00111 | 74286 | [26 empty] | 2018-05-04 | 14241 | Graphite | Completed
[168 empty] | 1.0002 | 0.00114 | 74287 | [26 empty] | 2018-05-07 | 5778 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00114 | 74288 | [26 empty] | 2018-05-28 | 15708 | AMS Submission Results | Completed
[197 empty] | 2018-05-28 | 18941 | Close Job | Completed
[7 empty] | 75.2 | 89.0906 | 0.0007 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44347 | 44347 | 44347 | 209 | 208 | 209 |  |  | jenny dahl | -996 | 0.10 ‰ | 2018-06-09 | 2017-12-11 | 2 | 180 |  |  |  |  | 0.52‰ | 0.52‰ | 2.00‰ | 2.00‰ | 0.52 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.89443 | -996.00 ‰ | -996.00 ‰ | 0.1 | 0.1 | -995.95 | 0.11 ‰ | -995.95 | 0.11 ‰ |  |  | 2018-06-14 | Jocelyn Turnbull | 10:53 | 0.0040032 | 0.0040032 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 6.2 |  |  | 0.004 | 0.0001 | 15785 |  |  |  | 209463 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 33 | 5 | 3 | ... | 0.00385 | 0.0001 |  |  |  |  | 1.0016 | 1.0017 | Yes | Yes | [5 empty] | 0.00385 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.004 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00385 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-06-08 | jenny dahl | 13:16 |  | Completed |  |  |  |  | 74433 | 2000000 |  |  | 74387 74388 74389 74390 74424 | 74387~74388~74389~74390~74424 | 3167 |  | 3167 3167 3167 3167 3167 |  |  |  |  | 0.31 mg |  | 0.31 | [5 empty] | 0.99881 | 0.00149 | 74387 |  | XCAMS | ST | 0% |  | TW3167 measurement: * no problems with XCAMS * small and large samples blank-corrected separately: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: all 9x via ST 5x large: no failures, used for normalisation in CalAMS 4x small: no failures, d = 1.47 ugC secondaries: * 1x IAEA-C2 via water-line: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates, large: 2x unknowns + 0x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 3x blanks organics via AAA, large: 2x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 3x blanks carbonates, small: 1x unknowns + 2x secondaries + 1x blanks only blanks from this wheel taken m = 0.53 ugC organics via AAA, small: 14x unknowns + 4x secondaries + 4x blanks only blanks from this wheel taken m = 0.90 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-03-14 | Carbonate |  | JOB-209463 |  | SLUSHIE small std | 65041 |  | 14047/1 |  |  | 2018-03-16 | 3312 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00144 | 74388 | [26 empty] | 2018-06-07 | 14443 | Graphite | Completed
[168 empty] | 0.99907 | 0.00149 | 74389 | [26 empty] | 2018-06-14 | 15785 | AMS Submission Results | Completed
[168 empty] | 0.99915 | 0.0015 | 74390 | [26 empty] | 2018-06-14 | 19126 | Close Job | Completed
[168 empty] | 1.0022 | 0.00148 | 74424
[7 empty] | 66 | 89.1509 | 0.0004 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49371 | 49371 | 49371 | 312 | 311 | 312 |  |  | jenny dahl | -997.86 | 0.08 ‰ | 2018-06-09 | 2017-12-11 | 3 | 180 |  |  |  |  | -0.15‰ | -0.15‰ | 2.00‰ | 2.00‰ | -0.15 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.89443 | -997.86 ‰ | -997.86 ‰ | 0.08 | 0.08 | -997.83 | 0.08 ‰ | -997.83 | 0.08 ‰ |  |  | 2018-06-14 | Jocelyn Turnbull | 10:53 | 0.002142 | 0.002142 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 13.3 |  |  | 0.00214 | 8e-05 | 15796 |  |  |  | 209788 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.21 | 0.01% | 23 | 5 | 3 | ... | 0.00206 | 8e-05 |  |  |  |  | 1.0016 | 1.0017 | Yes | Yes | [5 empty] | 0.00206 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00214 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00206 | 0 | 8e-05 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2018-06-08 | jenny dahl | 13:16 |  | Completed |  |  |  |  | 74421 | 2000000 |  |  | 74387 74388 74389 74390 74424 | 74387~74388~74389~74390~74424 | 3167 |  | 3167 3167 3167 3167 3167 |  |  |  |  | 0.16 mg |  | 0.16 | [5 empty] | 0.99881 | 0.00149 | 74387 |  | XCAMS | ST | 0% |  | TW3167 measurement: * no problems with XCAMS * small and large samples blank-corrected separately: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: all 9x via ST 5x large: no failures, used for normalisation in CalAMS 4x small: no failures, d = 1.47 ugC secondaries: * 1x IAEA-C2 via water-line: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates, large: 2x unknowns + 0x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 3x blanks organics via AAA, large: 2x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 3x blanks carbonates, small: 1x unknowns + 2x secondaries + 1x blanks only blanks from this wheel taken m = 0.53 ugC organics via AAA, small: 14x unknowns + 4x secondaries + 4x blanks only blanks from this wheel taken m = 0.90 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-05-31 | Carbonate |  | JOB-209788 |  | SLUSHIE to match small size | 65042 |  | 14047/1 |  |  | 2018-05-10 | 3372 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00144 | 74388 | [26 empty] | 2018-06-06 | 14431 | Graphite | Completed
[168 empty] | 0.99907 | 0.00149 | 74389 | [26 empty] | 2018-06-14 | 15796 | AMS Submission Results | Completed
[168 empty] | 0.99915 | 0.0015 | 74390 | [26 empty] | 2018-06-14 | 19134 | Close Job | Completed
[168 empty] | 1.0022 | 0.00148 | 74424
[7 empty] | 98.5 | 88.3235 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50021 | 50020 | 50021 | 296 | 295 | 296 |  |  | Julia Collins | -998.02 | 0.07 ‰ | 2018-06-15 | 2017-12-17 | 4 | 180 |  |  |  |  | 4.90‰ | 4.90‰ | 2.00‰ | 2.00‰ | 4.90 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -998.02 ‰ | -998.02 ‰ | 0.07 | 0.07 | -998 | 0.07 ‰ | -998 | 0.07 ‰ |  |  | 2018-06-18 | Jocelyn Turnbull | 14:27 | 0.0019756 | 0.0019756 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.5 |  |  | 0.00198 | 7e-05 | 15852 | 2018-06-18 | 12:12 | Albert Zondervan | 209760 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 1 | 5 | 4 | ... | 0.0019 | 7e-05 |  |  |  |  | 0.99737 | 0.99737 | Yes | Yes | [5 empty] | 0.0019 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00198 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0019 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-06-13 | Julia Collins | 13:55 |  | Completed |  |  |  |  | 74470 | 2000000 |  |  | 74464 74465 74466 74493 74494 74499 | 74464~74465~74466~74493~74494~74499 | 3169 |  | 3169 3169 3169 3169 3169 3169 | 2018-06-18 | 12:30 | Albert Zondervan |  | 1.26 mg |  | 1.26 | [5 empty] | 0.99907 | 0.00181 | 74464 |  | XCAMS | ST | 0.11% |  | TW3169 measurement: * no problems with XCAMS CalAMS error adjustments: * 0.11% system error * small error factors for cathodes #11 (x1.3) and 19 (x1.2) primaries: * OxI cathodes: all 6x via ST no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * two group has been identified to blank-correct: carbonates: 17x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 4x blanks organics via AAA, large: 11x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 3x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-05-17 | Carbonate |  | JOB-209760 |  | started d13c process but bottle missing cancel c13 process | 65119 |  | 14047/1 |  |  | 2018-05-23 | 3376 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99945 | 0.0018 | 74465 | [26 empty] | 2018-06-11 | 14480 | Graphite | Completed
[168 empty] | 0.99717 | 0.00181 | 74466 | [26 empty] | 2018-06-19 | 5858 | C13 Measurement | Cancelled
[168 empty] | 1.0013 | 0.00183 | 74493 | [26 empty] | 2018-06-18 | 15852 | AMS Submission Results | Completed
[168 empty] | 1.0022 | 0.00182 | 74494 | [26 empty] | 2018-06-19 | 19220 | Close Job | Completed
[168 empty] | 1.0009 | 0.00182 | 74499
[7 empty] | 105.2 | 88.5179 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47594 | 47594 | 47594 | 281 | 280 | 281 |  |  | jenny dahl | -997.33 | 0.09 ‰ | 2018-06-21 | 2017-12-23 | 6 | 180 |  |  |  |  | 2.90‰ | 2.90‰ | 2.00‰ | 2.00‰ | 2.90 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19 | 0 | 0.8165 | -997.33 ‰ | -997.33 ‰ | 0.09 | 0.09 | -997.29 | 0.09 ‰ | -997.29 | 0.09 ‰ |  |  | 2018-06-25 | Jocelyn Turnbull | 9:40 | 0.0026723 | 0.0026723 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00267 | 9e-05 | 15951 | 2018-06-22 | 17:43 | Albert Zondervan | 209864 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.27 | 0.01% | 37 | 5 | 6 | ... | 0.00257 | 9e-05 |  |  |  |  | 0.99757 | 0.99757 | Yes | Yes | [5 empty] | 0.00257 | 0.00‰ | 9e-05 | 0.00‰ | No | 1 | 1.0001 | 0.00114 | 0.00114 | 0.00267 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00257 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-06-20 | jenny dahl | 15:11 |  | Completed |  |  |  |  | 74539 | 2000000 |  |  | 74294 74537 74569 74571 74579 74589 | 74294~74537~74569~74571~74579~74589 | 3171 |  | 3171 3171 3171 3171 3171 3171 | 2018-06-22 | 18:11 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0014 | 0.00275 | 74294 |  | XCAMS | EA | 0.21% |  | TW3171 measurement: * no problems with XCAMS CalAMS error adjustments: * 0.21% system error * small error factors for cathodes #35 (x1.3) primaries: * OxI cathodes: all 6x via EA no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * four groups have been identified to blank-correct: organics prepared by U Otago: 3x unknowns + 0x secondaries + 1x blanks organics via AAA: 8x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 4x blanks carbonates: 15x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 6x blanks waters: 1x unknowns + 0x secondaries + 0x blanks 180 days, including SSAMS runs: capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2018-06-06 | Carbonate |  | JOB-209864 |  |  | 65199 |  | 14047/1 |  |  | 2018-06-13 | 3411 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0024 | 0.00276 | 74537 | [26 empty] | 2018-06-13 | 2125 | Gas Splitting | Completed
[168 empty] | 1.001 | 0.00275 | 74569 | [26 empty] | 2018-06-13 | 3837 | Choose Gas Splitting | Completed
[168 empty] | 0.99665 | 0.00273 | 74571 | [26 empty] | 2018-06-15 | 14549 | Graphite | Completed
[168 empty] | 0.99688 | 0.00274 | 74579 | [26 empty] | 2018-06-18 | 5872 | C13 Measurement | Completed
[168 empty] | 1.003 | 0.00315 | 74589 | [26 empty] | 2018-06-25 | 15951 | AMS Submission Results | Completed
[197 empty] | 2018-06-25 | 19317 | Close Job | Completed
[5 empty] | 5.9681e-05 | 0.0089301 | 103.4 | 89.9847 | 0.0705 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7017 | 7016 | 7017 | 59 | 58 | 61 | 6958.3 | 7075.3 | jenny dahl | -582.5 | 3.04 ‰ | 2018-07-12 | 2018-01-13 | 9 | 180 |  |  |  |  | -10.66‰ | -10.66‰ | 2.00‰ | 2.00‰ | -10.66 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -582.50 ‰ | -582.50 ‰ | 3.04 | 3.18 | -577.35 | 3.08 ‰ | -577.35 | 3.08 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:01 | 0.4175 | 0.4175 | 0.00304 | 0.00318 | 0.00299 | 0.00299 | 2 |  |  | 0.42352 | 0.00132 | 16039 | 2018-07-13 | 14:25 | Albert Zondervan | 209938 | 0.00967 | 0.0097 | 0.00435 | 0.0044 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.75 | 0.3% | 36 | 5 | 5 | ... | 0.40731 | 0.00131 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes |  | result is 2-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 61.2430027209171037. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00967 | 0.00859 | 0.00311 | 0.40152 | 0.40‰ | 0.00294 | 0.00‰ | Yes | 1 | 0.99999 | 0.00097 | 0.00097 | 0.4175 | 0 | 0.00304 | 0 | 0 | 0.00967 | 0 | 0.00435 | 0 | 0.40152 | 0 | 0.00294 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-04 | jenny dahl | 8:32 |  | Completed |  |  |  |  | 74711 | 2000000 | 74710 74231 74147 74073 73983 73840 73742 | 74710~74231~74147~74073~73983~73840~73742 | 74587 74591 74597 74606 74607 74608 | 74587~74591~74597~74606~74607~74608 | 3174 | 3174 3165 3160 3159 3156 3153 3151 | 3174 3174 3174 3174 3174 3174 | 2018-07-13 | 14:32 | Albert Zondervan |  | 0.94 mg | 0.13 | 0.94 | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 0.99795 | 0.00241 | 74587 | ST | XCAMS | EA | 0.15% |  | TW3174 measurement: * A mistake was made during graphite loading: One funnel must have been used for loading two cathodes. One of the 40 funnels was left over after loading all 40 graphites. It is unknown which cathodes were possibly compromised. The resulting data do not show a cross-contamination problem. * There were a few issues with XCAMS: There is a small step in 13/12 for some cathodes halfway in the measurement. The tune for TW3173 has already shown a possible issue with the scan of ESX reg#1 for 13C. They are probably associated with the slight lack of repeatability amongst the standards. These need to be sorted before high-precision is attempted: A source clean is warranted. After his return, Albert to look into the ESX reg#1 profile of 13C+ transmission into MFC 04-2. CalAMS error adjustments: * small system error: 0.15% * small error factors x1.3 on cat#33 primaries: * OxI cathodes: 6x from flask no failures secondaries: * IAEA-C2: 1x result ~ 2 sigma younger than consensus value blanks: * One groups have been identified to blank-correct: waters: window = 180 days, capturing 7x blanks unknowns: * no failures in wheel other: * It appears the field 'Standard Prep Type' in RLIMS is no longer filled in. Is this no longer needed? |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-07-19 | Water | Complete | JOB-209938 |  |  | 65280 | Normal | 14047/12 |  |  | 2018-07-02 | 3451 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 1.0013 | 0.00238 | 74591 | EA | [25 empty] | 2018-07-02 | 2157 | Gas Splitting | Completed
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 0.99931 | 0.00237 | 74597 | ST | [25 empty] | 2018-07-02 | 3888 | Choose Gas Splitting | Completed
[163 empty] | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 | 0.99676 | 0.00236 | 74606 | ST | [25 empty] | 2018-07-03 | 14721 | Graphite | Completed
[163 empty] | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 1.0018 | 0.0024 | 74607 | EA | [25 empty] | 2018-07-17 | 16039 | AMS Submission Results | Completed
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 | 1.0028 | 0.00236 | 74608 | ST | [25 empty] | 2018-07-17 | 19404 | Close Job | Completed
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 |  |  |  | EA
[7 empty] | 107.1 | 89.1413 | 0.0022 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34913 | 34913 | 34913 | 128 | 128 | 128 |  |  | jenny dahl | -987.04 | 0.21 ‰ | 2018-07-12 | 2018-01-13 | 9 | 180 |  |  |  |  | -1.30‰ | -1.30‰ | 2.00‰ | 2.00‰ | -1.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -987.04 ‰ | -987.04 ‰ | 0.21 | 0.21 | -986.88 | 0.21 ‰ | -986.88 | 0.21 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:01 | 0.012956 | 0.012956 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2 |  |  | 0.01296 | 0.00021 | 16048 | 2018-07-13 | 14:25 | Albert Zondervan | 209984 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 1.3 | 0.02% | 37 | 5 | 5 | ... | 0.01246 | 0.0002 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes | [5 empty] | 0.01246 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 0.99999 | 0.00097 | 0.00097 | 0.01296 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01246 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-04 | jenny dahl | 8:32 |  | Completed |  |  |  |  | 74710 | 2000000 |  |  | 74587 74591 74597 74606 74607 74608 | 74587~74591~74597~74606~74607~74608 | 3174 |  | 3174 3174 3174 3174 3174 3174 | 2018-07-13 | 14:32 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 0.99795 | 0.00241 | 74587 |  | XCAMS | EA | 0.15% |  | TW3174 measurement: * A mistake was made during graphite loading: One funnel must have been used for loading two cathodes. One of the 40 funnels was left over after loading all 40 graphites. It is unknown which cathodes were possibly compromised. The resulting data do not show a cross-contamination problem. * There were a few issues with XCAMS: There is a small step in 13/12 for some cathodes halfway in the measurement. The tune for TW3173 has already shown a possible issue with the scan of ESX reg#1 for 13C. They are probably associated with the slight lack of repeatability amongst the standards. These need to be sorted before high-precision is attempted: A source clean is warranted. After his return, Albert to look into the ESX reg#1 profile of 13C+ transmission into MFC 04-2. CalAMS error adjustments: * small system error: 0.15% * small error factors x1.3 on cat#33 primaries: * OxI cathodes: 6x from flask no failures secondaries: * IAEA-C2: 1x result ~ 2 sigma younger than consensus value blanks: * One groups have been identified to blank-correct: waters: window = 180 days, capturing 7x blanks unknowns: * no failures in wheel other: * It appears the field 'Standard Prep Type' in RLIMS is no longer filled in. Is this no longer needed? |  | RRL | [5 empty] | Background Inorganic |  | 2018-07-24 | Water |  | JOB-209984 |  |  | 65279 | Normal | 14047/11 |  |  | 2018-07-02 | 3450 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00238 | 74591 | [26 empty] | 2018-07-02 | 2155 | Gas Splitting | Completed
[168 empty] | 0.99931 | 0.00237 | 74597 | [26 empty] | 2018-07-02 | 3885 | Choose Gas Splitting | Completed
[168 empty] | 0.99676 | 0.00236 | 74606 | [26 empty] | 2018-07-03 | 14720 | Graphite | Completed
[168 empty] | 1.0018 | 0.0024 | 74607 | [26 empty] | 2018-07-17 | 16048 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00236 | 74608 | [26 empty] | 2018-07-17 | 19414 | Close Job | Completed
[5 empty] | 5.522e-05 | 0.0089363 | 98 | 90.0941 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7123 | 7122 | 7123 | 59 | 58 | 61 | 7064.3 | 7181.7 | Cathy Ginnane | -587.98 | 3.01 ‰ | 2018-07-14 | 2018-01-15 | 8 | 180 |  |  |  |  | -11.66‰ | -11.66‰ | 2.00‰ | 2.00‰ | -11.66 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -587.98 ‰ | -587.98 ‰ | 3.01 | 3.15 | -582.89 | 3.05 ‰ | -582.89 | 3.05 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:21 | 0.41202 | 0.41202 | 0.00301 | 0.00315 | 0.00296 | 0.00296 | 1.9 |  |  | 0.41801 | 0.00129 | 16067 | 2018-07-17 | 11:04 | Jocelyn Turnbull | 209985 | 0.00954 | 0.0095 | 0.00429 | 0.0043 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.2 | 0.3% | 36 | 5 | 5 | ... | 0.40201 | 0.00113 |  |  |  |  | 1.0006 | 1.0006 | Yes | Yes |  | ok wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 61.4359277504204598. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00954 | 0.00859 | 0.0029 | 0.39625 | 0.40‰ | 0.00285 | 0.00‰ | No | 1 | 0.99998 | 0.0008 | 0.0008 | 0.41202 | 0 | 0.00301 | 0 | 0 | 0.00954 | 0 | 0.00429 | 0 | 0.39625 | 0 | 0.00285 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-11 | Cathy Ginnane | 15:24 |  | Completed |  |  |  |  | 74759 | 2000000 | 74758 74710 74231 74147 74073 73983 73840 73742 | 74758~74710~74231~74147~74073~73983~73840~73742 | 74703 74704 74712 74713 74760 74761 | 74703~74704~74712~74713~74760~74761 | 3175 | 3175 3174 3165 3160 3159 3156 3153 3151 | 3175 3175 3175 3175 3175 3175 | 2018-07-17 | 11:16 | Jocelyn Turnbull |  | 1.04 mg | 0.13 | 1.04 | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 0.99969 | 0.00194 | 74703 | ST | XCAMS | EA |  |  |  |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-07-24 | Water |  | JOB-209985 |  |  | 65320 | Normal | 14047/12 |  |  | 2018-07-04 | 3454 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 0.99996 | 0.00197 | 74704 | EA | [25 empty] | 2018-07-04 | 2163 | Gas Splitting | Completed
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 0.99776 | 0.00198 | 74712 | ST | [25 empty] | 2018-07-04 | 3899 | Choose Gas Splitting | Completed
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 0.99851 | 0.00197 | 74713 | ST | [25 empty] | 2018-07-06 | 14769 | Graphite | Completed
[163 empty] | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 | 1.0034 | 0.00195 | 74760 | EA | [25 empty] | 2018-07-17 | 16067 | AMS Submission Results | Completed
[163 empty] | 2018-04-13 | 0.0135 | 73983 | 3156 | 0.95 | 1.0005 | 0.00194 | 74761 | ST | [25 empty] | 2018-07-17 | 19415 | Close Job | Completed
[163 empty] | 2018-03-09 | 0.0047 | 73840 | 3153 | 0.85 |  |  |  | EA
[163 empty] | 2018-02-01 | 0.01088 | 73742 | 3151 | 1.23 |  |  |  | EA
[7 empty] | 101.6 | 88.8725 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37863 | 37863 | 37863 | 130 | 130 | 130 |  |  | Cathy Ginnane | -991.03 | 0.15 ‰ | 2018-07-14 | 2018-01-15 | 8 | 180 |  |  |  |  | 1.93‰ | 1.93‰ | 2.00‰ | 2.00‰ | 1.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -991.03 ‰ | -991.03 ‰ | 0.15 | 0.15 | -990.92 | 0.15 ‰ | -990.92 | 0.15 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:21 | 0.0089734 | 0.0089734 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.8 |  |  | 0.00897 | 0.00015 | 16091 | 2018-07-17 | 11:04 | Jocelyn Turnbull | 210223 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 37 | 5 | 5 | ... | 0.00863 | 0.00014 |  |  |  |  | 1.0006 | 1.0006 | Yes | Yes | [5 empty] | 0.00863 | 0.01‰ | 0.00014 | 0.00‰ | No | 1 | 0.99998 | 0.0008 | 0.0008 | 0.00897 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00863 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-11 | Cathy Ginnane | 15:25 |  | Completed |  |  |  |  | 74758 | 2000000 |  |  | 74703 74704 74712 74713 74760 74761 | 74703~74704~74712~74713~74760~74761 | 3175 |  | 3175 3175 3175 3175 3175 3175 | 2018-07-17 | 11:16 | Jocelyn Turnbull |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99969 | 0.00194 | 74703 |  | XCAMS | EA |  |  |  |  | RRL | [5 empty] | Background Inorganic |  | 2018-08-13 | Water |  | JOB-210223 |  |  | 65319 | Normal | 14047/11 |  |  | 2018-07-04 | 3453 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99996 | 0.00197 | 74704 | [26 empty] | 2018-07-04 | 2161 | Gas Splitting | Completed
[168 empty] | 0.99776 | 0.00198 | 74712 | [26 empty] | 2018-07-04 | 3896 | Choose Gas Splitting | Completed
[168 empty] | 0.99851 | 0.00197 | 74713 | [26 empty] | 2018-07-06 | 14768 | Graphite | Completed
[168 empty] | 1.0034 | 0.00195 | 74760 | [26 empty] | 2018-07-17 | 16091 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00194 | 74761 | [26 empty] | 2018-07-17 | 19418 | Close Job | Completed
[7 empty] | 125.8 | 88.4675 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52474 | 52473 | 52474 | 459 | 459 | 459 |  |  | Julia Collins | -998.54 | 0.08 ‰ | 2018-08-06 | 2018-02-07 | 17 | 180 |  |  |  |  | 4.33‰ | 4.33‰ | 2.00‰ | 2.00‰ | 4.33 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0008 | 0 | 0.8165 | -998.54 ‰ | -998.54 ‰ | 0.08 | 0.08 | -998.53 | 0.08 ‰ | -998.53 | 0.08 ‰ |  |  | 2018-08-13 | Jocelyn Turnbull | 12:26 | 0.0014557 | 0.0014557 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.8 |  |  | 0.00146 | 8e-05 | 16309 |  |  |  | 210358 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 15 | 5 | 5 | ... | 0.0014 | 8e-05 |  |  |  |  | 0.99844 | 0.99844 | Yes | Yes | [5 empty] | 0.0014 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00146 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0014 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-08-03 | Julia Collins | 9:17 |  | Completed |  |  |  |  | 74870 | 2000000 |  |  | 74820 74823 74846 74847 74878 74879 | 74820~74823~74846~74847~74878~74879 | 3181 |  | 3181 3181 3181 3181 3181 3181 | 2018-08-07 | 13:07 | Albert Zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0003 | 0.00152 | 74820 |  | XCAMS | EA | 0% |  | TW3181 measurement: * no problems with XCAMS CalAMS error adjustments: * 0% system error * small error factors for cathodes #5,#35 (x1.3) primaries: * OxI cathodes: all 6x via EA no failures secondaries: * 1x IAEA-C2: result is NOT consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * three groups have been identified to blank-correct: organics via AAA: 12x unknowns + 2x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 5x blanks organics via CE-Rafter: 4x unknowns + 0x secondaries + 1x blanks 360 days: capturing 9x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-07-30 | Carbonate |  | JOB-210358 |  |  | 65568 |  | 14047/1 |  |  | 2018-07-19 | 3474 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99967 | 0.00162 | 74823 | [26 empty] | 2018-07-19 | 2174 | Gas Splitting | Completed
[168 empty] | 0.99939 | 0.00166 | 74846 | [26 empty] | 2018-07-19 | 3910 | Choose Gas Splitting | Completed
[168 empty] | 0.99981 | 0.00162 | 74847 | [26 empty] | 2018-07-20 | 14880 | Graphite | Completed
[168 empty] | 0.99952 | 0.00208 | 74878 | [26 empty] | 2018-08-13 | 16309 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.00161 | 74879 | [26 empty] | 2018-08-13 | 19720 | Close Job | Completed
[5 empty] | 7.0863e-05 | 0.0083369 | 117.8 | 89.6734 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7149 | 7149 | 7149 | 55 | 55 | 58 | 7094.2 | 7204.8 | jenny dahl | -589.34 | 2.83 ‰ | 2018-08-13 | 2018-02-14 | 12 | 180 |  |  |  |  | -8.57‰ | -8.57‰ | 2.00‰ | 2.00‰ | -8.57 ‰ | 2.00 ‰ | AMS | -8.30‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -589.34 ‰ | -589.34 ‰ | 2.83 | 2.97 | -584.27 | 2.86 ‰ | -584.27 | 2.86 ‰ |  |  | 2018-08-16 | Jocelyn Turnbull | 15:21 | 0.41066 | 0.41066 | 0.00283 | 0.00297 | 0.00278 | 0.00278 | 2 |  |  | 0.41627 | 0.00123 | 16356 |  |  |  | 210224 | 0.00892 | 0.0089 | 0.00401 | 0.004 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.07 | 0.28% | 1 | 5 | 6 | ... | 0.40034 | 0.00106 |  |  |  |  | 0.99904 | 0.99904 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 58.1850095183356886. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00892 | 0.00801 | 0.00251 | 0.39494 | 0.39‰ | 0.00267 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.41066 | 0 | 0.00283 | 0 | 0 | 0.00892 | 0 | 0.00401 | 0 | 0.39494 | 0 | 0.00267 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-08-09 | jenny dahl | 9:26 |  | Completed |  |  |  |  | 74966 | 2000000 | 74967 74758 74710 74231 | 74967~74758~74710~74231 | 74881 74882 74907 74912 75033 75034 | 74881~74882~74907~74912~75033~75034 | 3183 | 3183 3175 3174 3165 | 3183 3183 3183 3183 3183 3183 | 2018-08-14 | 16:23 | Albert Zondervan |  | 1.07 mg | 0.13 | 1.07 | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 1.0004 | 0.00177 | 74881 | ST | XCAMS | EA | 0% |  | TW3183 measurement: * no problems with XCAMS CalAMS error adjustments: * 0% system error primaries: * OxI cathodes: all 6x via EA no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6: result is NOT consistent with expectation value * 1x FIRI-I: result is consistent with expectation value * 1x FIRI whalebone: result is consistent with expectation value blanks: * four groups have been identified to blank-correct: organics via AAA: 14x unknowns + 3x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks waters: 8x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 4x blanks carbonates: 2x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks unknowns: * TP74905: 8 permil difference in IRMS d13C to AMS checked carbonate and graphite processes for this sample and found nothing unusual; can't see a problem in XCAMS raw data either | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-08-13 | Water |  | JOB-210224 |  | PC was left open between the left and right sides of the line during sample transfer. Consultation with JT concurs that unlikely to be any cross-contamination but be aware if sample results look odd | 65612 | Normal | 14047/12 |  |  | 2018-07-31 | 3482 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 1.0004 | 0.00172 | 74882 | EA | [25 empty] | 2018-07-31 | 2184 | Gas Splitting | Completed
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 0.99798 | 0.00163 | 74907 | EA | [25 empty] | 2018-07-31 | 3922 | Choose Gas Splitting | Completed
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 1.0006 | 0.00165 | 74912 | EA | [25 empty] | 2018-08-01 | 14976 | Graphite | Completed
[168 empty] | 1.0002 | 0.00172 | 75033 | [26 empty] | 2018-08-02 | 5985 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00177 | 75034 | [26 empty] | 2018-08-16 | 16356 | AMS Submission Results | Completed
[197 empty] | 2018-08-17 | 19767 | Close Job | Completed
[7 empty] | 113.3 | 88.7041 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38492 | 38492 | 38492 | 161 | 161 | 161 |  |  | jenny dahl | -991.7 | 0.17 ‰ | 2018-08-13 | 2018-02-14 | 12 | 180 |  |  |  |  | 2.26‰ | 2.26‰ | 2.00‰ | 2.00‰ | 2.26 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -991.70 ‰ | -991.70 ‰ | 0.17 | 0.17 | -991.6 | 0.17 ‰ | -991.6 | 0.17 ‰ |  |  | 2018-08-16 | Jocelyn Turnbull | 15:21 | 0.0082976 | 0.0082976 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.0083 | 0.00017 | 16361 |  |  |  | 210353 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.83 | 0.02% | 20 | 5 | 6 | ... | 0.00798 | 0.00016 |  |  |  |  | 0.99904 | 0.99904 | Yes | Yes | [5 empty] | 0.00798 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0083 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00798 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-08-09 | jenny dahl | 9:26 |  | Completed |  |  |  |  | 74967 | 2000000 |  |  | 74881 74882 74907 74912 75033 75034 | 74881~74882~74907~74912~75033~75034 | 3183 |  | 3183 3183 3183 3183 3183 3183 | 2018-08-14 | 16:23 | Albert Zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0004 | 0.00177 | 74881 |  | XCAMS | EA | 0% |  | TW3183 measurement: * no problems with XCAMS CalAMS error adjustments: * 0% system error primaries: * OxI cathodes: all 6x via EA no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6: result is NOT consistent with expectation value * 1x FIRI-I: result is consistent with expectation value * 1x FIRI whalebone: result is consistent with expectation value blanks: * four groups have been identified to blank-correct: organics via AAA: 14x unknowns + 3x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks waters: 8x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 4x blanks carbonates: 2x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks unknowns: * TP74905: 8 permil difference in IRMS d13C to AMS checked carbonate and graphite processes for this sample and found nothing unusual; can't see a problem in XCAMS raw data either | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2018-08-30 | Water | Complete | JOB-210353 |  | PC was left open between the left and right sides of the line during sample transfer. Consultation with JT concurs that unlikely to be any cross-contamination but be aware if sample results look odd | 65611 | Normal | 14047/11 |  |  | 2018-07-31 | 3485 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00172 | 74882 | [26 empty] | 2018-07-31 | 2185 | Gas Splitting | Completed
[168 empty] | 0.99798 | 0.00163 | 74907 | [26 empty] | 2018-07-31 | 3924 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00165 | 74912 | [26 empty] | 2018-08-01 | 14977 | Graphite | Completed
[168 empty] | 1.0002 | 0.00172 | 75033 | [26 empty] | 2018-08-02 | 5981 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00177 | 75034 | [26 empty] | 2018-08-16 | 16361 | AMS Submission Results | Completed
[197 empty] | 2018-08-17 | 19769 | Close Job | Completed
[7 empty] | 62.3 | 89.0681 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45182 | 45182 | 45182 | 208 | 208 | 208 |  |  | Jeremy thompson | -996.39 | 0.09 ‰ | 2018-08-20 | 2018-02-21 | 13 | 180 |  |  |  |  | 2.44‰ | 2.44‰ | 2.00‰ | 2.00‰ | 2.44 ‰ | 2.00 ‰ | AMS | 2.66‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0 | 0.8165 | -996.39 ‰ | -996.39 ‰ | 0.09 | 0.09 | -996.35 | 0.09 ‰ | -996.35 | 0.09 ‰ |  |  | 2018-08-27 | Jocelyn Turnbull | 15:54 | 0.0036081 | 0.0036081 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00361 | 9e-05 | 16442 | 2018-08-23 | 11:45 | Albert Zondervan | 210452 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 32 | 5 | 6 | ... | 0.00347 | 9e-05 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes | [5 empty] | 0.00347 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00361 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00347 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-08-17 | Jeremy thompson | 9:46 |  | Completed |  |  |  |  | 75021 | 2000000 |  |  | 75066 75067 75073 75075 75076 75077 | 75066~75067~75073~75075~75076~75077 | 3185 |  | 3185 3185 3185 3185 3185 3185 | 2018-08-27 | 14:46 | margaret norris |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0015 | 0.00168 | 75066 |  | XCAMS | ST | 0% |  | TW3185 measurement: * SNICS indexer got stuck for a bit, but was resolved with NEC's help * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC * Some samples required additional customisation: Chris Moy small sample blank-corrected with his own blank only, both processed elsewhere. Jompra sample not blank-corrected because no blank was provided and it was processed elsewhere. CalAMS error adjustments: * none needed primaries: * OxI cathodes: all 9x via ST 6x large: no failures, used for normalisation in CalAMS 2x small: no failures, d = 1.45 ugC using 6x small OxIs TW3167,3185. secondaries: * 1x IAEA-C2 carbonate, large: result is 'older' than consensus value by 3-sigma * 1x IAEA-C6: result looks ok * 1x FIRI-I, small: result is consistent with consensus value * 1x FIRI-I, large: result is 'older' than consensus value by 2.5-sigma blanks: * Three groups have been identified to blank-correct: carbonates, large: 15x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 7x blanks organics via AAA, large: 5x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 5x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 2x blanks blanks used: from this wheel and TW3167 m = 0.83 ugC unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2018-08-20 | Carbonate |  | JOB-210452 |  |  | 65651 |  | 14047/1 |  |  | 2018-08-06 | 3493 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00164 | 75067 | [26 empty] | 2018-08-06 | 2200 | Gas Splitting | Completed
[168 empty] | 0.99953 | 0.00152 | 75073 | [26 empty] | 2018-08-06 | 3940 | Choose Gas Splitting | Completed
[168 empty] | 0.99795 | 0.00187 | 75075 | [26 empty] | 2018-08-07 | 15031 | Graphite | Completed
[168 empty] | 1.001 | 0.00155 | 75076 | [26 empty] | 2018-08-09 | 6016 | C13 Measurement | Completed
[168 empty] | 0.99836 | 0.00152 | 75077 | [26 empty] | 2018-08-27 | 16442 | AMS Submission Results | Completed
[197 empty] | 2018-08-29 | 19856 | Close Job | Completed
[7 empty] | 124.7 | 88.4656 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52589 | 52589 | 52589 | 349 | 349 | 349 |  |  | margaret norris | -998.57 | 0.06 ‰ | 2018-09-07 | 2018-03-11 | 4 | 180 |  |  |  |  | 0.91‰ | 0.91‰ | 2.00‰ | 2.00‰ | 0.91 ‰ | 2.00 ‰ | AMS | 1.73‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0 | 0.8165 | -998.57 ‰ | -998.57 ‰ | 0.06 | 0.06 | -998.55 | 0.06 ‰ | -998.55 | 0.06 ‰ |  |  | 2018-09-12 | Cathy Ginnane | 9:18 | 0.0014349 | 0.0014349 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.6 |  |  | 0.00143 | 6e-05 | 16592 | 2018-09-09 | 14:21 | Albert Zondervan | 210461 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.14 | 0.01% | 6 | 5 | 7 | ... | 0.00138 | 6e-05 |  |  |  |  | 0.99501 | 0.99501 | Yes | Yes | [5 empty] | 0.00138 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00143 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00138 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-09-06 | margaret norris | 11:52 |  | Completed |  |  |  |  | 75256 | 2000000 |  |  | 75223 75235 75236 75237 75249 75250 | 75223~75235~75236~75237~75249~75250 | 3189 |  | 3189 3189 3189 3189 3189 3189 | 2018-09-09 | 15:05 | Albert Zondervan |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99861 | 0.00161 | 75223 |  | XCAMS | ST | 0% |  | TW3189 measurement: * no problems with XCAMS CalAMS error adjustments: * 0% system error * uncertainty multiplier = 1.3 for cat#13 (TP75227), cat#20 (TP75239) primaries: * OxI cathodes: all 6x via ST no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6 (via EA): result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value * 1x FIRI whalebone: result is consistent with expectation value blanks: * four groups have been identified to blank-correct: carbonates: 3x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 7x blanks waters: 4x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 2x blanks organics via AAA: 17x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 6x blanks unknowns: * TP75238: The result of this repeat (job210592) is consistent with that of the original (job210354, TP74905). This time there is almost no difference between IRMS del13C to AMS del13C. * no failures | 1.7 | RRL | [5 empty] | Background Inorganic |  | 2018-08-21 | Carbonate | Complete | JOB-210461 |  | new treatment from bulk acid etch and ground C13 was out of normal range | 65813 |  | 14047/1 |  |  | 2018-09-03 | 5491 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00165 | 75235 | [26 empty] | 2018-09-03 | 3510 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99936 | 0.00165 | 75236 | [26 empty] | 2018-09-06 | 2224 | Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00218 | 75237 | [26 empty] | 2018-09-03 | 3975 | Choose Gas Splitting | Completed
[168 empty] | 1.0016 | 0.00164 | 75249 | [26 empty] | 2018-09-04 | 15266 | Graphite | Completed
[168 empty] | 0.99998 | 0.00164 | 75250 | [26 empty] | 2018-09-06 | 6058 | C13 Measurement | Completed
[197 empty] | 2018-09-12 | 16592 | AMS Submission Results | Completed
[197 empty] | 2018-09-12 | 19973 | Close Job | Completed
[10 empty] | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | [11 empty] | margaret norris |  |  |  | ? |  | 180 | [9 empty] | 0.20 ‰ | C13 |  | 0.2‰ | C13 | [5 empty] | -1000.00 ‰ |  |  |  | -1000 | [6 empty] | margaret norris | [7 empty] | 2.6 |  |  |  |  | 16593 |  |  |  | 210719 |  |  |  |  | 45% |  | 2/10/19 | 15:28:48 | Jocelyn Turnbull | [5 empty] | 5 | 7 | ... | [9 empty] | Yes | [21 empty] | 0 | 0 |  |  | 0 |  | Regular | Yes |  |  |  | 2018-09-06 | margaret norris | 11:52 |  | Started |  |  |  |  | 75256 | 2000000 | [11 empty] | 0.81 mg |  | 0.81 | [14 empty] | 1.7 | RRL | [5 empty] | Background Inorganic |  | 2018-09-27 | Carbonate |  | JOB-210719 |  | C13 was out of normal range extra job created in error when trying to create job for d13c split only |  |  | 14047/1 |  |  | 2018-09-03 | 5491 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[197 empty] | 2018-09-03 | 3510 | Carbonate CO2 Evolution | Completed
[197 empty] | 2018-09-06 | 2224 | Gas Splitting | Completed
[197 empty] | 2018-09-03 | 3975 | Choose Gas Splitting | Completed
[197 empty] | 2018-09-06 | 19968 | Close Job | Cancelled
[7 empty] | 107.2 | 88.8211 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51301 | 51301 | 51301 | 298 | 297 | 297 |  |  | Julia Collins | -998.32 | 0.06 ‰ | 2018-10-08 | 2018-04-11 | 14 | 180 |  |  |  |  | 0.48‰ | 0.48‰ | 2.00‰ | 2.00‰ | 0.48 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0014 | 0 | 0.8165 | -998.32 ‰ | -998.32 ‰ | 0.06 | 0.06 | -998.29 | 0.06 ‰ | -998.29 | 0.06 ‰ |  |  | 2018-10-12 | Jocelyn Turnbull | 14:08 | 0.0016845 | 0.0016845 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.5 |  |  | 0.00168 | 6e-05 | 16818 | 2018-10-11 | 10:36 | Rachel Corran | 210703 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 8 | 5 | 5 | ... | 0.00162 | 6e-05 |  |  |  |  | 0.99858 | 0.99858 | Yes | Yes | [5 empty] | 0.00162 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00168 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00162 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-09-27 | Julia Collins | 15:28 |  | Completed |  |  |  |  | 75423 | 2000000 |  |  | 75404 75405 75453 75456 75473 75475 | 75404~75405~75453~75456~75473~75475 | 3194 |  | 3194 3194 3194 3194 3194 3194 | 2018-10-11 | 11:36 | Rachel Corran |  | 0.78 mg |  | 0.78 | [5 empty] | 0.99902 | 0.00135 | 75404 |  | XCAMS | EA | 0% |  | TW3194 measurement: * unstable accelerator (shut down part way through measurement, but restarted next day and ran ok) CalAMS error adjustments: * 0.00% system error primaries: * OxI cathodes: 6x no cathode failures secondaries: * IAEA-C2: * FIRI-I: blanks: * three groups identified to blank-correct: organics: window = 180 days, capturing 7x blanks waters: window = 180 days, capturing 7x blanks carbonates: window = 180 days, capturing 8x blanks unknowns: * no failures in wheel |  | RRL | [5 empty] | Background Inorganic |  | 2018-10-26 | Carbonate |  | JOB-210703 |  | In breakseal | 66051 |  | 14047/1 |  |  | 2018-09-12 | 3516 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.001 | 0.00144 | 75405 | [26 empty] | 2018-09-12 | 2228 | Gas Splitting | Completed
[168 empty] | 0.9995 | 0.0014 | 75453 | [26 empty] | 2018-09-12 | 3979 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00144 | 75456 | [26 empty] | 2018-09-24 | 15433 | Graphite | Completed
[168 empty] | 1.0012 | 0.00148 | 75473 | [26 empty] | 2018-10-12 | 16818 | AMS Submission Results | Completed
[168 empty] | 0.99906 | 0.00147 | 75475 | [26 empty] | 2018-10-12 | 20252 | Close Job | Completed
[7 empty] | 97.4 | 88.622 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33531 | 33530 | 33530 | 114 | 113 | 113 |  |  | Julia Collins | -984.61 | 0.22 ‰ | 2018-10-29 | 2018-05-02 | 21 | 180 |  |  |  |  | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0017 | 0 | 0.8165 | -984.61 ‰ | -984.61 ‰ | 0.22 | 0.22 | -984.42 | 0.22 ‰ | -984.42 | 0.22 ‰ |  |  | 2018-11-05 | Jocelyn Turnbull | 13:58 | 0.015389 | 0.015389 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.3 |  |  | 0.01539 | 0.00022 | 16995 |  |  |  | 209937 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 1.54 | 0.02% | 1 | 5 | 5 | ... | 0.0148 | 0.00021 |  |  |  |  | 0.99885 | 0.99885 | Yes | Yes | [5 empty] | 0.0148 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01539 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.0148 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-10-24 | Julia Collins | 15:03 |  | Completed |  |  |  |  | 75583 | 2000000 |  |  | 75434 75435 75447 75448 75449 75605 | 75434~75435~75447~75448~75449~75605 | 3199 |  | 3199 3199 3199 3199 3199 3199 | 2018-11-02 | 9:02 | Albert Zondervan |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99908 | 0.00173 | 75434 |  | XCAMS | ST | 0% |  | TW3199 measurement: * The problem with cooling-water temperature stability, causing transmission variations, was recognised easily; corresponding outlier data were rejected. * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x, all via ST 6x large: no failures, used for normalisation in CalAMS 1x small: no failures used for calculation of DCC d = 2.41 ugC using this wheel's small OxI only because reduction in rts is significantly higher than in previous wheels secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value blanks: * Four groups have been identified to blank-correct: waters, large: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, large: 9x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 11x unknowns + 0x secondaries + 2x blanks blanks used: from this wheel only because blank level is high m = 1.55 ugC, double the previous value reported in TW3185 organics via CE, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures JCT - checked small sample corrections. Modern contamination is higher because this wheel had longer days as graphite than normal. So larger m than typical is expected. d is also higher than normal, but is based on a single small OxI target that is only marginally higher than in previous wheels. Seems reasonable to use d based on this wheel, with the associated larger error bar, to be consistent with using m from this wheel alone. Even though there is no obvious reason why d would be larger than normal. |  | RRL | [5 empty] | Background Inorganic |  | 2018-07-19 | Water |  | JOB-209937 |  |  | 66211 | Normal | 14047/11 |  |  | 2018-10-05 | 3512 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99841 | 0.00176 | 75435 | [26 empty] | 2018-10-05 | 2247 | Gas Splitting | Completed
[168 empty] | 1.0018 | 0.00181 | 75447 | [26 empty] | 2018-10-05 | 4005 | Choose Gas Splitting | Completed
[168 empty] | 0.99903 | 0.00179 | 75448 | [26 empty] | 2018-10-08 | 15593 | Graphite | Completed
[168 empty] | 1.0015 | 0.00173 | 75449 | [26 empty] | 2018-11-05 | 16995 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00175 | 75605 | [26 empty] | 2018-11-06 | 20462 | Close Job | Completed
[5 empty] | 5.7992e-05 | 0.0093091 | 93.8 | 89.6089 | 0.0699 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7108 | 7107 | 7107 | 61 | 60 | 63 | 7047 | 7169 | Julia Collins | -587.21 | 3.13 ‰ | 2018-10-29 | 2018-05-02 | 21 | 180 |  |  |  |  | -8.05‰ | -8.05‰ | 2.00‰ | 2.00‰ | -8.05 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0017 | 0 | 0.8165 | -587.21 ‰ | -587.21 ‰ | 3.13 | 3.27 | -582.12 | 3.17 ‰ | -582.12 | 3.17 ‰ |  |  | 2018-11-05 | Jocelyn Turnbull | 13:58 | 0.41279 | 0.41279 | 0.00313 | 0.00327 | 0.00309 | 0.00309 | 2.1 |  |  | 0.41906 | 0.00129 | 17007 |  |  |  | 210850 | 0.01 | 0.01 | 0.0045 | 0.0045 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.28 | 0.31% | 19 | 5 | 5 | ... | 0.40302 | 0.00112 |  |  |  |  | 0.99885 | 0.99885 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 63.6230848315219654. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01 | 0.00895 | 0.00279 | 0.39699 | 0.40‰ | 0.00297 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.41279 | 0 | 0.00313 | 0 | 0 | 0.01 | 0 | 0.0045 | 0 | 0.39699 | 0 | 0.00297 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-10-24 | Julia Collins | 15:03 |  | Completed |  |  |  |  | 75584 | 2000000 | 75583 74967 74758 74710 74231 74147 74073 | 75583~74967~74758~74710~74231~74147~74073 | 75434 75435 75447 75448 75449 75605 | 75434~75435~75447~75448~75449~75605 | 3199 | 3199 3183 3175 3174 3165 3160 3159 | 3199 3199 3199 3199 3199 3199 | 2018-11-02 | 9:02 | Albert Zondervan |  | 0.99 mg | 0.13 | 0.99 | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 0.99908 | 0.00173 | 75434 | ST | XCAMS | ST | 0% |  | TW3199 measurement: * The problem with cooling-water temperature stability, causing transmission variations, was recognised easily; corresponding outlier data were rejected. * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x, all via ST 6x large: no failures, used for normalisation in CalAMS 1x small: no failures used for calculation of DCC d = 2.41 ugC using this wheel's small OxI only because reduction in rts is significantly higher than in previous wheels secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value blanks: * Four groups have been identified to blank-correct: waters, large: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, large: 9x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 11x unknowns + 0x secondaries + 2x blanks blanks used: from this wheel only because blank level is high m = 1.55 ugC, double the previous value reported in TW3185 organics via CE, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures JCT - checked small sample corrections. Modern contamination is higher because this wheel had longer days as graphite than normal. So larger m than typical is expected. d is also higher than normal, but is based on a single small OxI target that is only marginally higher than in previous wheels. Seems reasonable to use d based on this wheel, with the associated larger error bar, to be consistent with using m from this wheel alone. Even though there is no obvious reason why d would be larger than normal. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-11-27 | Water |  | JOB-210850 |  |  | 66212 | Normal | 14047/12 |  |  | 2018-10-05 | 3559 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 0.99841 | 0.00176 | 75435 | EA | [25 empty] | 2018-10-05 | 2248 | Gas Splitting | Completed
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 1.0018 | 0.00181 | 75447 | ST | [25 empty] | 2018-10-05 | 4007 | Choose Gas Splitting | Completed
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 0.99903 | 0.00179 | 75448 | EA | [25 empty] | 2018-10-08 | 15594 | Graphite | Completed
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 1.0015 | 0.00173 | 75449 | EA | [25 empty] | 2018-11-05 | 17007 | AMS Submission Results | Completed
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 1.0002 | 0.00175 | 75605 | EA | [25 empty] | 2018-11-06 | 20476 | Close Job | Completed
[163 empty] | 2018-05-13 | 0.00993 | 74073 | 3159 | 1.13 |  |  |  | ST
[7 empty] | 98.9 | 88.5585 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51919 | 51919 | 51919 | 267 | 267 | 267 |  |  | Julia Collins | -998.44 | 0.05 ‰ | 2018-11-12 | 2018-05-16 | 41 | 180 |  |  |  |  | 3.74‰ | 3.74‰ | 2.00‰ | 2.00‰ | 3.74 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0017 | 0 | 0.8165 | -998.44 ‰ | -998.44 ‰ | 0.05 | 0.05 | -998.42 | 0.05 ‰ | -998.42 | 0.05 ‰ |  |  | 2018-11-20 | Jocelyn Turnbull | 9:38 | 0.0015597 | 0.0015597 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 1.9 |  |  | 0.00156 | 5e-05 | 17036 | 2018-11-19 | 13:11 | Albert Zondervan | 210761 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.16 | 0.01% | 8 | 5 | 5 | ... | 0.0015 | 5e-05 |  |  |  |  | 0.99888 | 0.99888 | Yes | Yes | [5 empty] | 0.0015 | 0.00‰ | 5e-05 | 0.00‰ | No | 1 | 1 | 0.00102 | 0.00102 | 0.00156 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.0015 | 0 | 5e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-10-31 | Julia Collins | 13:28 |  | Completed |  |  |  |  | 75518 | 2000000 |  |  | 75471 75555 75593 75600 75636 75640 | 75471~75555~75593~75600~75636~75640 | 3200 |  | 3200 3200 3200 3200 3200 3200 | 2018-11-19 | 13:25 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.9999 | 0.00246 | 75471 |  | XCAMS | EA | 0.21% |  | TW3200 measurement: * 2nd wheel measurement after AccelNET upgrade CalAMS error adjustments: * system error = 0.21% * global random error multiplication factor = 1.02 primaries: * OxI cathodes: 6x large: no failures Apparent outlier cat#35 was not rejected as calibration cathode, as there was no solid reason to do so. JCT -note that excluding this target would have only minimal effect on the absolute values of the results. secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value blanks: * Two groups have been identified to blank-correct: carbonates, large: 29x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks waters, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-11-12 | Carbonate |  | JOB-210761 |  |  | 66291 |  | 14047/1 |  |  | 2018-10-01 | 3548 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0008 | 0.00249 | 75555 | [26 empty] | 2018-10-01 | 2241 | Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00252 | 75593 | [26 empty] | 2018-10-01 | 3997 | Choose Gas Splitting | Completed
[168 empty] | 1.0008 | 0.00251 | 75600 | [26 empty] | 2018-10-02 | 15528 | Graphite | Completed
[168 empty] | 1.0013 | 0.00252 | 75636 | [26 empty] | 2018-11-20 | 17036 | AMS Submission Results | Completed
[168 empty] | 0.99511 | 0.0025 | 75640 | [26 empty] | 2018-11-21 | 20590 | Close Job | Completed
[7 empty] | 117.3 | 88.1779 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39088 | 39087 | 39087 | 152 | 151 | 151 |  |  | Julia Collins | -992.3 | 0.15 ‰ | 2018-11-15 | 2018-05-19 | 16 | 180 |  |  |  |  | 4.96‰ | 4.96‰ | 2.00‰ | 2.00‰ | 4.96 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -992.30 ‰ | -992.30 ‰ | 0.15 | 0.15 | -992.2 | 0.15 ‰ | -992.2 | 0.15 ‰ |  |  | 2018-11-20 | Jocelyn Turnbull | 9:45 | 0.0077049 | 0.0077049 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.6 |  |  | 0.0077 | 0.00015 | 17076 | 2018-11-19 | 16:44 | Albert Zondervan | 210849 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.77 | 0.01% | 1 | 5 | 5 | ... | 0.00741 | 0.00014 |  |  |  |  | 0.99579 | 0.99579 | Yes | Yes | [5 empty] | 0.00741 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0077 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00741 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-12 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 75742 | 2000000 |  |  | 75715 75716 75717 75794 75795 75796 | 75715~75716~75717~75794~75795~75796 | 3201 |  | 3201 3201 3201 3201 3201 3201 | 2018-11-19 | 17:11 | Albert Zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 1.0003 | 0.0017 | 75715 |  | XCAMS | ST |  |  | TW3201 measurement: * XCAMS ok * raw 14C/C ratios low because of known effect in detector setup * large samples only CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I celluse: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks woods via ABBA, for Chris Moy (R41230): 4x unknowns + 0x secondaries + 1x blanks 10 days: capturing 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-11-27 | Water |  | JOB-210849 |  |  | 66330 | Normal | 14047/11 |  |  | 2018-10-29 | 3584 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0025 | 0.00175 | 75716 | [26 empty] | 2018-10-29 | 2270 | Gas Splitting | Completed
[168 empty] | 0.99955 | 0.00172 | 75717 | [26 empty] | 2018-10-29 | 4044 | Choose Gas Splitting | Completed
[168 empty] | 0.99913 | 0.00175 | 75794 | [26 empty] | 2018-10-30 | 15752 | Graphite | Completed
[168 empty] | 1.0001 | 0.00172 | 75795 | [26 empty] | 2018-11-20 | 17076 | AMS Submission Results | Completed
[168 empty] | 0.99856 | 0.00171 | 75796 | [26 empty] | 2018-11-21 | 20597 | Close Job | Completed
[5 empty] | 5.7408e-05 | 0.0089219 | 117.4 | 89.3251 | 0.0661 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7138 | 7137 | 7137 | 59 | 59 | 61 | 7078.6 | 7197 | Julia Collins | -588.74 | 3.03 ‰ | 2018-11-15 | 2018-05-19 | 16 | 180 |  |  |  |  | -7.95‰ | -7.95‰ | 2.00‰ | 2.00‰ | -7.95 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -588.74 ‰ | -588.74 ‰ | 3.03 | 3.17 | -583.66 | 3.07 ‰ | -583.66 | 3.07 ‰ |  |  | 2018-11-20 | Jocelyn Turnbull | 9:45 | 0.41126 | 0.41126 | 0.00303 | 0.00317 | 0.00298 | 0.00298 | 2.1 |  |  | 0.41732 | 0.00126 | 17081 | 2018-11-19 | 16:44 | Albert Zondervan | 210915 | 0.00964 | 0.0096 | 0.00434 | 0.0043 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.13 | 0.3% | 22 | 5 | 5 | ... | 0.40135 | 0.00109 |  |  |  |  | 0.99579 | 0.99579 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 61.9184648816679916. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00964 | 0.00858 | 0.00296 | 0.39552 | 0.40‰ | 0.00287 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.41126 | 0 | 0.00303 | 0 | 0 | 0.00964 | 0 | 0.00434 | 0 | 0.39552 | 0 | 0.00287 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-12 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 75743 | 2000000 | 75742 75583 74967 74758 74710 74231 74147 | 75742~75583~74967~74758~74710~74231~74147 | 75715 75716 75717 75794 75795 75796 | 75715~75716~75717~75794~75795~75796 | 3201 | 3201 3199 3183 3175 3174 3165 3160 | 3201 3201 3201 3201 3201 3201 | 2018-11-19 | 17:11 | Albert Zondervan |  | 0.99 mg | 0.13 | 0.99 | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 1.0003 | 0.0017 | 75715 | EA | XCAMS | ST |  |  | TW3201 measurement: * XCAMS ok * raw 14C/C ratios low because of known effect in detector setup * large samples only CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I celluse: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks woods via ABBA, for Chris Moy (R41230): 4x unknowns + 0x secondaries + 1x blanks 10 days: capturing 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-12-10 | Water |  | JOB-210915 |  |  | 66331 | Normal | 14047/12 |  |  | 2018-10-29 | 3585 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 1.0025 | 0.00175 | 75716 | ST | [25 empty] | 2018-10-29 | 2271 | Gas Splitting | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 0.99955 | 0.00172 | 75717 | EA | [25 empty] | 2018-10-29 | 4046 | Choose Gas Splitting | Completed
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 0.99913 | 0.00175 | 75794 | EA | [25 empty] | 2018-10-30 | 15753 | Graphite | Completed
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 1.0001 | 0.00172 | 75795 | EA | [25 empty] | 2018-11-20 | 17081 | AMS Submission Results | Completed
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 0.99856 | 0.00171 | 75796 | ST | [25 empty] | 2018-11-21 | 20601 | Close Job | Completed
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 |  |  |  | ST
[7 empty] | 100.6 | 88.5094 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35936 | 35936 | 35936 | 154 | 153 | 153 |  |  | jenny dahl | -988.59 | 0.22 ‰ | 2018-11-19 | 2018-05-23 | 13 | 180 |  |  |  |  | 3.90‰ | 3.90‰ | 2.00‰ | 2.00‰ | 3.90 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.70711 | -988.59 ‰ | -988.59 ‰ | 0.22 | 0.22 | -988.45 | 0.22 ‰ | -988.45 | 0.22 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 11:07 | 0.011407 | 0.011407 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 1.7 |  |  | 0.01141 | 0.00022 | 17176 | 2018-11-20 | 12:57 | Albert Zondervan | 211055 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.14 | 0.02% | 1 | 5 | 6 | ... | 0.01097 | 0.00021 |  |  |  |  | 0.99848 | 0.99848 | Yes | Yes | [5 empty] | 0.01097 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.01141 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01097 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:40 |  | Completed |  |  |  |  | 75841 | 2000000 |  |  | 75819 75820 75822 75825 75840 75856 75875 75876 | 75819~75820~75822~75825~75840~75856~75875~75876 | 3202 |  | 3202 3202 3202 3202 3202 3202 3202 3202 | 2018-11-20 | 13:46 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99902 | 0.00194 | 75819 |  | XCAMS | ST | 0% |  | TW3202 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.32 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures AMS category changed from 'standard as unknown' to 'primary standard OxI' NOT used for calculation of DCC, because their rts equals that of large OxIs d = 2.51 ugC extracted from the one small OxI in TW3199 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters, large: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA, large: 2x unknowns + 1x secondaries + 0x blanks 180 days: capturing 8x blanks organics via AAA, small, via Excel: 4x unknowns + 1x secondaries + 2x blanks blanks used: 2x from this wheel + 2x from TW3199 m = 1.32 ugC unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2018-12-27 | Water |  | JOB-211055 |  |  | 66406 | Normal | 14047/11 |  |  | 2018-11-05 | 3613 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00196 | 75820 | [26 empty] | 2018-11-05 | 2293 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00195 | 75822 | [26 empty] | 2018-11-05 | 4076 | Choose Gas Splitting | Completed
[168 empty] | 1.0031 | 0.00201 | 75825 | [26 empty] | 2018-11-06 | 15851 | Graphite | Completed
[168 empty] | 0.99842 | 0.00198 | 75840 | [26 empty] | 2018-11-07 | 6222 | C13 Measurement | Completed
[168 empty] | 0.99898 | 0.00198 | 75856 | [26 empty] | 2018-11-21 | 17176 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.002 | 75875 | [26 empty] | 2018-11-22 | 20617 | Close Job | Completed
[168 empty] | 0.99971 | 0.002 | 75876
[5 empty] | 5.5521e-05 | 0.0090826 | 97.5 | 89.5386 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7100 | 7099 | 7099 | 61 | 60 | 63 | 7039.2 | 7161.1 | jenny dahl | -586.81 | 3.13 ‰ | 2018-11-19 | 2018-05-23 | 13 | 180 |  |  |  |  | -7.64‰ | -7.64‰ | 2.00‰ | 2.00‰ | -7.64 ‰ | 2.00 ‰ | AMS | -8.12‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.70711 | -586.81 ‰ | -586.81 ‰ | 3.13 | 3.27 | -581.71 | 3.17 ‰ | -581.71 | 3.17 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 11:07 | 0.41319 | 0.41319 | 0.00313 | 0.00327 | 0.00309 | 0.00309 | 2.3 |  |  | 0.41934 | 0.00142 | 17177 | 2018-11-20 | 12:57 | Albert Zondervan | 211056 | 0.00981 | 0.0098 | 0.00441 | 0.0044 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.32 | 0.31% | 22 | 5 | 6 | ... | 0.40329 | 0.00126 |  |  |  |  | 0.99848 | 0.99848 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 63.5673752122460625. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00981 | 0.00873 | 0.00278 | 0.39738 | 0.40‰ | 0.00297 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.41319 | 0 | 0.00313 | 0 | 0 | 0.00981 | 0 | 0.00441 | 0 | 0.39738 | 0 | 0.00297 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:40 |  | Completed |  |  |  |  | 75842 | 2000000 | 75841 75742 75583 74967 74758 74710 74231 74147 | 75841~75742~75583~74967~74758~74710~74231~74147 | 75819 75820 75822 75825 75840 75856 75875 75876 | 75819~75820~75822~75825~75840~75856~75875~75876 | 3202 | 3202 3201 3199 3183 3175 3174 3165 3160 | 3202 3202 3202 3202 3202 3202 3202 3202 | 2018-11-20 | 13:46 | Albert Zondervan |  | 0.83 mg | 0.13 | 0.83 | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 0.99902 | 0.00194 | 75819 | EA | XCAMS | ST | 0% |  | TW3202 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.32 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures AMS category changed from 'standard as unknown' to 'primary standard OxI' NOT used for calculation of DCC, because their rts equals that of large OxIs d = 2.51 ugC extracted from the one small OxI in TW3199 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters, large: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA, large: 2x unknowns + 1x secondaries + 0x blanks 180 days: capturing 8x blanks organics via AAA, small, via Excel: 4x unknowns + 1x secondaries + 2x blanks blanks used: 2x from this wheel + 2x from TW3199 m = 1.32 ugC unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-12-27 | Water |  | JOB-211056 |  |  | 66407 | Normal | 14047/12 |  |  | 2018-11-05 | 3614 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 0.99956 | 0.00196 | 75820 | ST | [25 empty] | 2018-11-05 | 2292 | Gas Splitting | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 1.0012 | 0.00195 | 75822 | EA | [25 empty] | 2018-11-05 | 4074 | Choose Gas Splitting | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 1.0031 | 0.00201 | 75825 | EA | [25 empty] | 2018-11-06 | 15852 | Graphite | Completed
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 0.99842 | 0.00198 | 75840 | EA | [25 empty] | 2018-11-07 | 6223 | C13 Measurement | Completed
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 | 0.99898 | 0.00198 | 75856 | ST | [25 empty] | 2018-11-21 | 17177 | AMS Submission Results | Completed
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 | 1.0001 | 0.002 | 75875 | ST | [25 empty] | 2018-11-22 | 20618 | Close Job | Completed
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 | 0.99971 | 0.002 | 75876 | ST
[7 empty] | 110.5 | 87.9367 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39969 | 39969 | 39969 | 181 | 181 | 181 |  |  | jenny dahl | -993.1 | 0.16 ‰ | 2018-11-20 | 2018-05-24 | 13 | 180 |  |  |  |  | 6.13‰ | 6.13‰ | 2.00‰ | 2.00‰ | 6.13 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -993.10 ‰ | -993.10 ‰ | 0.16 | 0.16 | -993.01 | 0.16 ‰ | -993.01 | 0.16 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 10:59 | 0.0069042 | 0.0069042 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.5 |  |  | 0.0069 | 0.00016 | 17222 | 2018-11-21 | 9:46 | Albert Zondervan | 211057 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.69 | 0.02% | 8 | 5 | 5 | ... | 0.00664 | 0.00015 |  |  |  |  | 0.99422 | 0.99422 | Yes | Yes | [5 empty] | 0.00664 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.0069 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00664 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:47 |  | Completed |  |  |  |  | 75854 | 2000000 |  |  | 75849 75851 75853 75904 75905 75906 | 75849~75851~75853~75904~75905~75906 | 3204 |  | 3204 3204 3204 3204 3204 3204 | 2018-11-21 | 10:25 | Albert Zondervan |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0007 | 0.00194 | 75849 |  | XCAMS | EA | 0% |  | TW3204 measurement: * large samples CalAMS error adjustments: * factor 1.3 for cat#17 (TP75850) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is 1.5 sigma older than consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I: result agrees with consensus value blanks: * Three groups have been identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days: capturing 9x blanks organics via CE-Rafter: 2x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-12-27 | Water |  | JOB-211057 |  |  | 66370 | Normal | 14047/11 |  |  | 2018-11-05 | 3615 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99999 | 0.002 | 75851 | [26 empty] | 2018-11-05 | 2294 | Gas Splitting | Completed
[168 empty] | 1.0033 | 0.00194 | 75853 | [26 empty] | 2018-11-05 | 4078 | Choose Gas Splitting | Completed
[168 empty] | 0.99807 | 0.00194 | 75904 | [26 empty] | 2018-11-07 | 15864 | Graphite | Completed
[168 empty] | 0.99968 | 0.00192 | 75905 | [26 empty] | 2018-11-21 | 17222 | AMS Submission Results | Completed
[168 empty] | 0.99838 | 0.00187 | 75906 | [26 empty] | 2018-11-22 | 20619 | Close Job | Completed
[5 empty] | 5.2306e-05 | 0.0088376 | 115 | 89.0944 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7242 | 7241 | 7241 | 60 | 60 | 62 | 7181.7 | 7302 | jenny dahl | -594.03 | 3.04 ‰ | 2018-11-20 | 2018-05-24 | 13 | 180 |  |  |  |  | -6.94‰ | -6.94‰ | 2.00‰ | 2.00‰ | -6.94 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -594.03 ‰ | -594.03 ‰ | 3.04 | 3.18 | -589.02 | 3.08 ‰ | -589.02 | 3.08 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 10:59 | 0.40597 | 0.40597 | 0.00304 | 0.00318 | 0.00299 | 0.00299 | 2 |  |  | 0.41196 | 0.00136 | 17223 | 2018-11-21 | 9:46 | Albert Zondervan | 211058 | 0.00945 | 0.0095 | 0.00425 | 0.0043 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 40.6 | 0.3% | 29 | 5 | 5 | ... | 0.39619 | 0.0012 |  |  |  |  | 0.99422 | 0.99422 | Yes | Yes |  | result is 1.5 sigma older than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 62.8283786361044606. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00945 | 0.0085 | 0.00281 | 0.39043 | 0.39‰ | 0.00288 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.40597 | 0 | 0.00304 | 0 | 0 | 0.00945 | 0 | 0.00425 | 0 | 0.39043 | 0 | 0.00288 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:47 |  | Completed |  |  |  |  | 75855 | 2000000 | 75854 75841 75742 75583 74967 74758 74710 74231 74147 | 75854~75841~75742~75583~74967~74758~74710~74231~74147 | 75849 75851 75853 75904 75905 75906 | 75849~75851~75853~75904~75905~75906 | 3204 | 3204 3202 3201 3199 3183 3175 3174 3165 3160 | 3204 3204 3204 3204 3204 3204 | 2018-11-21 | 10:25 | Albert Zondervan |  | 0.93 mg | 0.13 | 0.93 | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 1.0007 | 0.00194 | 75849 | EA | XCAMS | EA | 0% |  | TW3204 measurement: * large samples CalAMS error adjustments: * factor 1.3 for cat#17 (TP75850) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is 1.5 sigma older than consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I: result agrees with consensus value blanks: * Three groups have been identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days: capturing 9x blanks organics via CE-Rafter: 2x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2018-12-27 | Water |  | JOB-211058 |  |  | 66371 | Normal | 14047/12 |  |  | 2018-11-05 | 3616 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 0.99999 | 0.002 | 75851 | ST | [25 empty] | 2018-11-05 | 2291 | Gas Splitting | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 1.0033 | 0.00194 | 75853 | EA | [25 empty] | 2018-11-05 | 4072 | Choose Gas Splitting | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 0.99807 | 0.00194 | 75904 | EA | [25 empty] | 2018-11-07 | 15865 | Graphite | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 0.99968 | 0.00192 | 75905 | EA | [25 empty] | 2018-11-21 | 17223 | AMS Submission Results | Completed
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 | 0.99838 | 0.00187 | 75906 | ST | [25 empty] | 2018-11-22 | 20620 | Close Job | Completed
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 |  |  |  | ST
[163 empty] | 2018-05-25 | 0.0066 | 74231 | 3165 | 1.04 |  |  |  | ST
[163 empty] | 2018-05-24 | 0.0096 | 74147 | 3160 | 1.09 |  |  |  | EA
[5 empty] | 5.6484e-05 | 0.0079453 | 72.3 | 89.7364 | 0.0693 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7158 | 7158 | 7158 | 56 | 55 | 58 | 7102.3 | 7214.3 | jenny dahl | -589.79 | 2.86 ‰ | 2018-12-02 | 2018-06-05 | 9 | 180 |  |  |  |  | -8.51‰ | -8.51‰ | 2.00‰ | 2.00‰ | -8.51 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19 | 0 | 0.8165 | -589.79 ‰ | -589.79 ‰ | 2.86 | 3 | -584.73 | 2.89 ‰ | -584.73 | 2.89 ‰ |  |  | 2018-12-06 | Jocelyn Turnbull | 15:28 | 0.41021 | 0.41021 | 0.00286 | 0.003 | 0.00281 | 0.00281 | 1.9 |  |  | 0.41589 | 0.00125 | 17347 |  |  |  | 211403 | 0.00901 | 0.009 | 0.00405 | 0.0041 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.02 | 0.29% | 8 | 5 | 5 | ... | 0.39997 | 0.0012 |  |  |  |  | 0.99981 | 0.99981 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 58.8442745698521976. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00901 | 0.00764 | 0.00363 | 0.39451 | 0.39‰ | 0.00275 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.41021 | 0 | 0.00286 | 0 | 0 | 0.00901 | 0 | 0.00405 | 0 | 0.39451 | 0 | 0.00275 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-28 | jenny dahl | 15:01 |  | Completed |  |  |  |  | 75969 | 2000000 | 75971 75854 75841 75742 75583 74967 74758 74710 | 75971~75854~75841~75742~75583~74967~74758~74710 | 75993 75994 75995 75996 75997 75998 | 75993~75994~75995~75996~75997~75998 | 3207 | 3207 3204 3202 3201 3199 3183 3175 3174 | 3207 3207 3207 3207 3207 3207 | 2018-12-04 | 15:43 | Albert Zondervan |  | 1.10 mg | 0.13 | 1.1 | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 0.99754 | 0.00216 | 75993 | EA | XCAMS | EA | 0.13% |  | TW3207 measurement: * large water samples * SNICS indexer stopped measurement unexpectedly, but dataset is large enough CalAMS error adjustments: * factor 1.3 for cat#16,26 (TP75931,75949) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result looks ok blanks: * Three groups have been identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks IAEA-C1 blank in this wheel (TP75971) looks remarkably good unknowns: * no failures * TP75930: AMS-IRMS difference in del13C = 4 permil |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-01-02 | Water |  | JOB-211403 |  |  | 66572 | Normal | 14047/12 |  |  | 2018-11-21 | 3672 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 1.0034 | 0.00214 | 75994 | EA | [25 empty] | 2018-11-21 | 2322 | Gas Splitting | Completed
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 1.0001 | 0.00217 | 75995 | EA | [25 empty] | 2018-11-21 | 4114 | Choose Gas Splitting | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 0.99806 | 0.00215 | 75996 | ST | [25 empty] | 2018-11-23 | 15979 | Graphite | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 1.0013 | 0.00217 | 75997 | ST | [25 empty] | 2018-12-06 | 17347 | AMS Submission Results | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 | 0.99965 | 0.00215 | 75998 | ST | [25 empty] | 2018-12-07 | 20838 | Close Job | Completed
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 |  |  |  | EA
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 |  |  |  | EA
[7 empty] | 74.9 | 88.7507 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45808 | 45807 | 45807 | 275 | 275 | 275 |  |  | jenny dahl | -996.66 | 0.11 ‰ | 2018-12-02 | 2018-06-05 | 9 | 180 |  |  |  |  | 2.50‰ | 2.50‰ | 2.00‰ | 2.00‰ | 2.50 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19 | 0 | 0.8165 | -996.66 ‰ | -996.66 ‰ | 0.11 | 0.11 | -996.62 | 0.12 ‰ | -996.62 | 0.12 ‰ |  |  | 2018-12-06 | Jocelyn Turnbull | 15:28 | 0.0033377 | 0.0033377 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00334 | 0.00011 | 17354 |  |  |  | 211438 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 27 | 5 | 5 | ... | 0.00321 | 0.00011 |  |  |  |  | 0.99981 | 0.99981 | Yes | Yes |  | rather a good water blank! |  |  |  | 0.00321 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00334 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00321 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-28 | jenny dahl | 15:01 |  | Completed |  |  |  |  | 75971 | 2000000 |  |  | 75993 75994 75995 75996 75997 75998 | 75993~75994~75995~75996~75997~75998 | 3207 |  | 3207 3207 3207 3207 3207 3207 | 2018-12-04 | 15:43 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99754 | 0.00216 | 75993 |  | XCAMS | EA | 0.13% |  | TW3207 measurement: * large water samples * SNICS indexer stopped measurement unexpectedly, but dataset is large enough CalAMS error adjustments: * factor 1.3 for cat#16,26 (TP75931,75949) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result looks ok blanks: * Three groups have been identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks IAEA-C1 blank in this wheel (TP75971) looks remarkably good unknowns: * no failures * TP75930: AMS-IRMS difference in del13C = 4 permil |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-07 | Water |  | JOB-211438 |  |  | 66571 | Normal | 14047/11 |  |  | 2018-11-21 | 3673 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0034 | 0.00214 | 75994 | [26 empty] | 2018-11-21 | 2321 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00217 | 75995 | [26 empty] | 2018-11-21 | 4112 | Choose Gas Splitting | Completed
[168 empty] | 0.99806 | 0.00215 | 75996 | [26 empty] | 2018-11-23 | 15981 | Graphite | Completed
[168 empty] | 1.0013 | 0.00217 | 75997 | [26 empty] | 2018-12-06 | 17354 | AMS Submission Results | Completed
[168 empty] | 0.99965 | 0.00215 | 75998 | [26 empty] | 2018-12-07 | 20845 | Close Job | Completed
[7 empty] | 99.1 | 88.7236 | 0.002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35648 | 35648 | 35648 | 127 | 127 | 127 |  |  | Margaret Norris | -988.18 | 0.19 ‰ | 2018-12-07 | 2018-06-10 | 14 | 180 |  |  |  |  | 1.59‰ | 1.59‰ | 2.00‰ | 2.00‰ | 1.59 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -988.18 ‰ | -988.18 ‰ | 0.19 | 0.19 | -988.03 | 0.19 ‰ | -988.03 | 0.19 ‰ |  |  | 2018-12-10 | Jocelyn Turnbull | 10:15 | 0.011822 | 0.011822 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.5 |  |  | 0.01182 | 0.00019 | 17404 |  |  |  | 211402 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.18 | 0.02% | 10 | 5 | 5 | ... | 0.01137 | 0.00018 |  |  |  |  | 0.99859 | 0.99859 | Yes | Yes | [5 empty] | 0.01137 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.01182 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.01137 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-12-07 | Margaret Norris | 14:06 |  | Completed |  |  |  |  | 75968 | 2000000 |  |  | 76029 76035 76036 76037 76042 76043 | 76029~76035~76036~76037~76042~76043 | 3209 |  | 3209 3209 3209 3209 3209 3209 | 2018-12-09 | 12:43 | Albert Zondervan |  | 1.28 mg |  | 1.28 | [5 empty] | 1.0005 | 0.00149 | 76029 |  | XCAMS | ST | 0% |  | TW3209 measurement: * some EXPRESS samples * all large samples CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.4 for cat#17 (TP75966), 1.3 for cat#34 (TP76040) * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result agrees with consensus value * 1x FIRI-I: result is 1.5 sigma older than consensus value * 1x FIRI-L whalebone: result is 1.5 sigma younger than consensus value blanks: * Three groups have been identified to blank-correct: (ground)waters: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks bone via AAA: 6x unknowns + 3x secondaries + 0x blanks 180 days: capturing 8x blanks wood via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-02 | Water | Complete | JOB-211402 |  |  | 66651 | Normal | 14047/11 |  |  | 2018-11-21 | 3671 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99794 | 0.00149 | 76035 | [26 empty] | 2018-11-21 | 2323 | Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00151 | 76036 | [26 empty] | 2018-11-21 | 4116 | Choose Gas Splitting | Completed
[168 empty] | 1.0014 | 0.0015 | 76037 | [26 empty] | 2018-11-23 | 15978 | Graphite | Completed
[168 empty] | 0.99942 | 0.00148 | 76042 | [26 empty] | 2018-12-10 | 17404 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00149 | 76043 | [26 empty] | 2018-12-10 | 20895 | Close Job | Completed
[5 empty] | 5.4075e-05 | 0.0082689 | 97.8 | 89.6079 | 0.0697 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7118 | 7118 | 7118 | 56 | 56 | 58 | 7062.6 | 7174.8 | Margaret Norris | -587.76 | 2.88 ‰ | 2018-12-07 | 2018-06-10 | 14 | 180 |  |  |  |  | -8.29‰ | -8.29‰ | 2.00‰ | 2.00‰ | -8.29 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -587.76 ‰ | -587.76 ‰ | 2.88 | 3.02 | -582.67 | 2.91 ‰ | -582.67 | 2.91 ‰ |  |  | 2018-12-10 | Jocelyn Turnbull | 10:15 | 0.41224 | 0.41224 | 0.00288 | 0.00302 | 0.00283 | 0.00283 | 2.1 |  |  | 0.41805 | 0.00114 | 17417 |  |  |  | 211439 | 0.00927 | 0.0093 | 0.00417 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.22 | 0.29% | 23 | 5 | 5 | ... | 0.40205 | 0.00096 |  |  |  |  | 0.99859 | 0.99859 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 58.9369568968153746. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00927 | 0.00795 | 0.00349 | 0.39646 | 0.40‰ | 0.00272 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.41224 | 0 | 0.00288 | 0 | 0 | 0.00927 | 0 | 0.00417 | 0 | 0.39646 | 0 | 0.00272 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-12-07 | Margaret Norris | 14:06 |  | Completed |  |  |  |  | 75972 | 2000000 | 75968 75971 75854 75841 75742 75583 74967 74758 74710 | 75968~75971~75854~75841~75742~75583~74967~74758~74710 | 76029 76035 76036 76037 76042 76043 | 76029~76035~76036~76037~76042~76043 | 3209 | 3209 3207 3204 3202 3201 3199 3183 3175 3174 | 3209 3209 3209 3209 3209 3209 | 2018-12-09 | 12:43 | Albert Zondervan |  | 0.96 mg | 0.13 | 0.96 | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 1.0005 | 0.00149 | 76029 | EA | XCAMS | ST | 0% |  | TW3209 measurement: * some EXPRESS samples * all large samples CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.4 for cat#17 (TP75966), 1.3 for cat#34 (TP76040) * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result agrees with consensus value * 1x FIRI-I: result is 1.5 sigma older than consensus value * 1x FIRI-L whalebone: result is 1.5 sigma younger than consensus value blanks: * Three groups have been identified to blank-correct: (ground)waters: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks bone via AAA: 6x unknowns + 3x secondaries + 0x blanks 180 days: capturing 8x blanks wood via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-01-07 | Water |  | JOB-211439 |  |  | 66652 | Normal | 14047/12 |  |  | 2018-11-21 | 3674 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 0.99794 | 0.00149 | 76035 | EA | [25 empty] | 2018-11-21 | 2320 | Gas Splitting | Completed
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 1.0002 | 0.00151 | 76036 | EA | [25 empty] | 2018-11-21 | 4110 | Choose Gas Splitting | Completed
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 1.0014 | 0.0015 | 76037 | ST | [25 empty] | 2018-11-23 | 15982 | Graphite | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 0.99942 | 0.00148 | 76042 | ST | [25 empty] | 2018-12-10 | 17417 | AMS Submission Results | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 | 1.0006 | 0.00149 | 76043 | ST | [25 empty] | 2018-12-10 | 20897 | Close Job | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 |  |  |  | EA
[163 empty] | 2018-07-14 | 0.00863 | 74758 | 3175 | 1.09 |  |  |  | EA
[163 empty] | 2018-07-12 | 0.01246 | 74710 | 3174 | 0.94 |  |  |  | ST
[7 empty] | 77.8 | 88.7397 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 175 | 175 | 175 |  |  | Julia Collins | -994.28 | 0.12 ‰ | 2019-01-16 | 2018-07-20 | 6 | 180 |  |  |  |  | 2.84‰ | 2.84‰ | 2.00‰ | 2.00‰ | 2.84 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -994.28 ‰ | -994.28 ‰ | 0.12 | 0.12 | -994.21 | 0.13 ‰ | -994.21 | 0.13 ‰ |  |  | 2019-01-23 | Cathy Ginnane | 14:06 | 0.0057189 | 0.0057189 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00572 | 0.00012 | 17466 | 2019-01-18 | 9:05 | Albert Zondervan | 211585 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 8 | 5 | 5 | ... | 0.0055 | 0.00012 |  |  |  |  | 1 | 1 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00572 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-15 | Julia Collins | 13:47 |  | Completed |  |  |  |  | 76104 | 2000000 |  |  | 76072 76073 76074 76075 76115 76118 | 76072~76073~76074~76075~76115~76118 | 3210 |  | 3210 3210 3210 3210 3210 3210 | 2019-01-18 | 9:21 | Albert Zondervan |  | 1.11 mg |  | 1.11 | [5 empty] | 1.001 | 0.00158 | 76072 |  | XCAMS | ST | 0% |  | TW3210 measurement: * some EXPRESS samples * all large samples * 1st wheel of 2019 CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.3 for cat#23, 1.2 for cat#33 * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I: result is consistent with consensus value * 1x FIRI-L whalebone: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 23x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks bone&textile via AAA: 4x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-29 | Water | Not Complete | JOB-211585 |  |  | 66721 | Normal | 14047/11 |  |  | 2018-12-11 | 3708 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0009 | 0.00154 | 76073 | [26 empty] | 2018-12-11 | 2352 | Gas Splitting | Completed
[168 empty] | 0.99832 | 0.0015 | 76074 | [26 empty] | 2018-12-11 | 4153 | Choose Gas Splitting | Completed
[168 empty] | 0.99943 | 0.00152 | 76075 | [26 empty] | 2019-01-10 | 16114 | Graphite | Completed
[168 empty] | 0.99969 | 0.00156 | 76115 | [26 empty] | 2019-01-23 | 17466 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.00156 | 76118 | [26 empty] | 2019-01-30 | 21006 | Close Job | Completed
[5 empty] | 5.4579e-05 | 0.0073563 | 80.1 | 89.6747 | 0.07 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7085 | 7085 | 7085 | 52 | 51 | 54 | 7033.6 | 7137.4 | Julia Collins | -586.06 | 2.67 ‰ | 2019-01-16 | 2018-07-20 | 6 | 180 |  |  |  |  | -7.62‰ | -7.62‰ | 2.00‰ | 2.00‰ | -7.62 ‰ | 2.00 ‰ | AMS | -8.46‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -586.06 ‰ | -586.06 ‰ | 2.67 | 2.83 | -580.95 | 2.71 ‰ | -580.95 | 2.71 ‰ |  |  | 2019-01-23 | Cathy Ginnane | 14:06 | 0.41394 | 0.41394 | 0.00267 | 0.00283 | 0.00262 | 0.00262 | 2.1 |  |  | 0.41926 | 0.00116 | 17467 | 2019-01-18 | 9:05 | Albert Zondervan | 211587 | 0.00849 | 0.0085 | 0.00382 | 0.0038 | 45% | 5 | 5/07/22 | 14:06:36 | Jocelyn Turnbull | 0 |  | 41.39 | 0.27% | 22 | 5 | 7 | ... | 0.40321 | 0.00099 |  |  |  |  | 1 | 1 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 54.9716441693612471. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00849 | 0.00708 | 0.0037 | 0.3981 | 0.40‰ | 0.00252 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.41394 | 0 | 0.00267 | 0 | 0 | 0.00849 | 0 | 0.00382 | 0 | 0.3981 | 0 | 0.00252 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-15 | Julia Collins | 13:47 |  | Completed |  |  |  |  | 76105 | 2000000 | 76104 75968 75971 75854 75841 75742 75583 74967 | 76104~75968~75971~75854~75841~75742~75583~74967 | 76072 76073 76074 76075 76115 76118 | 76072~76073~76074~76075~76115~76118 | 3210 | 3210 3209 3207 3204 3202 3201 3199 3183 | 3210 3210 3210 3210 3210 3210 | 2019-01-18 | 9:21 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 1.001 | 0.00158 | 76072 | EA | XCAMS | ST | 0% |  | TW3210 measurement: * some EXPRESS samples * all large samples * 1st wheel of 2019 CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.3 for cat#23, 1.2 for cat#33 * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I: result is consistent with consensus value * 1x FIRI-L whalebone: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 23x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks bone&textile via AAA: 4x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | -8.5 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-01-29 | Water | Complete | JOB-211587 |  |  | 66722 | Normal | 14047/12 |  |  | 2018-12-11 | 3709 | Water CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 1.0009 | 0.00154 | 76073 | ST | [25 empty] | 2018-12-12 | 3710 | Water CO2 Evolution | Completed
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 0.99832 | 0.0015 | 76074 | ST | [25 empty] | 2018-12-12 | 2354 | Gas Splitting | Completed
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 0.99943 | 0.00152 | 76075 | ST | [25 empty] | 2018-12-12 | 4155 | Choose Gas Splitting | Completed
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 0.99969 | 0.00156 | 76115 | EA | [25 empty] | 2019-01-15 | 6341 | C13 Measurement | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 | 1.0008 | 0.00156 | 76118 | EA | [25 empty] | 2019-01-10 | 16115 | Graphite | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | ST | [25 empty] | 2019-01-23 | 17467 | AMS Submission Results | Completed
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 |  |  |  | ST | [25 empty] | 2019-01-30 | 21007 | Close Job | Completed
[7 empty] | 112.3 | 89.0279 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38373 | 38372 | 38372 | 148 | 148 | 148 |  |  | jenny dahl | -991.58 | 0.16 ‰ | 2019-01-29 | 2018-08-02 | 6 | 180 |  |  |  |  | 0.69‰ | 0.69‰ | 2.00‰ | 2.00‰ | 0.69 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -991.58 ‰ | -991.58 ‰ | 0.16 | 0.16 | -991.47 | 0.16 ‰ | -991.47 | 0.16 ‰ |  |  | 2019-02-04 | Jocelyn Turnbull | 15:37 | 0.0084223 | 0.0084223 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00842 | 0.00016 | 17541 |  |  |  | 211700 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 8 | 5 | 5 | ... | 0.0081 | 0.00015 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.0081 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1 | 0.00066 | 0.00066 | 0.00842 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.0081 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-25 | jenny dahl | 12:26 |  | Completed |  |  |  |  | 76214 | 2000000 |  |  | 76153 76154 76155 76183 76184 76185 | 76153~76154~76155~76183~76184~76185 | 3212 |  | 3212 3212 3212 3212 3212 3212 |  |  |  |  | 0.89 mg |  | 0.89 | [5 empty] | 1.0008 | 0.0016 | 76153 |  | XCAMS | EA | 0.11% |  | TW3212 measurement: * all large samples CalAMS error adjustments: * single-cathode factors: 1.3 for cat#39 * system error: 0.11% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-02-08 | Water |  | JOB-211700 |  |  | 66771 | Normal | 14047/11 |  |  | 2018-12-18 | 3722 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9996 | 0.00161 | 76154 | [26 empty] | 2018-12-18 | 2365 | Gas Splitting | Completed
[168 empty] | 1.0019 | 0.0016 | 76155 | [26 empty] | 2018-12-18 | 4166 | Choose Gas Splitting | Completed
[168 empty] | 1.0015 | 0.00166 | 76183 | [26 empty] | 2019-01-23 | 16224 | Graphite | Completed
[168 empty] | 0.9987 | 0.00158 | 76184 | [26 empty] | 2019-02-04 | 17541 | AMS Submission Results | Completed
[168 empty] | 0.99783 | 0.00159 | 76185 | [26 empty] | 2019-02-07 | 21069 | Close Job | Completed
[5 empty] | 5.1506e-05 | 0.0074729 | 109.4 | 89.8035 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7125 | 7125 | 7125 | 50 | 50 | 53 | 7074.9 | 7175.7 | jenny dahl | -588.11 | 2.58 ‰ | 2019-01-29 | 2018-08-02 | 6 | 180 |  |  |  |  | -7.96‰ | -7.96‰ | 2.00‰ | 2.00‰ | -7.96 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -588.11 ‰ | -588.11 ‰ | 2.58 | 2.75 | -583.02 | 2.62 ‰ | -583.02 | 2.62 ‰ |  |  | 2019-02-04 | Jocelyn Turnbull | 15:37 | 0.41189 | 0.41189 | 0.00258 | 0.00275 | 0.00253 | 0.00253 | 2 |  |  | 0.4172 | 0.00093 | 17542 |  |  |  | 211701 | 0.00844 | 0.0084 | 0.0038 | 0.0038 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.19 | 0.26% | 27 | 5 | 5 | ... | 0.40123 | 0.00085 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 53.5442084889933423. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00844 | 0.00719 | 0.00347 | 0.39613 | 0.40‰ | 0.00247 | 0.00‰ | No | 1 | 1 | 0.00066 | 0.00066 | 0.41189 | 0 | 0.00258 | 0 | 0 | 0.00844 | 0 | 0.0038 | 0 | 0.39613 | 0 | 0.00247 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-25 | jenny dahl | 12:26 |  | Completed |  |  |  |  | 76215 | 2000000 | 76214 76104 75968 75971 75854 75841 75742 75583 74967 | 76214~76104~75968~75971~75854~75841~75742~75583~74967 | 76153 76154 76155 76183 76184 76185 | 76153~76154~76155~76183~76184~76185 | 3212 | 3212 3210 3209 3207 3204 3202 3201 3199 3183 | 3212 3212 3212 3212 3212 3212 | 2019-02-04 | 11:35 | Albert Zondervan |  | 1.08 mg | 0.13 | 1.08 | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 1.0008 | 0.0016 | 76153 | EA | XCAMS | EA | 0.11% |  | TW3212 measurement: * all large samples CalAMS error adjustments: * single-cathode factors: 1.3 for cat#39 * system error: 0.11% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-02-08 | Water |  | JOB-211701 |  |  | 66772 | Normal | 14047/12 |  |  | 2018-12-18 | 3723 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 0.9996 | 0.00161 | 76154 | ST | [25 empty] | 2018-12-18 | 2366 | Gas Splitting | Completed
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 1.0019 | 0.0016 | 76155 | ST | [25 empty] | 2018-12-18 | 4167 | Choose Gas Splitting | Completed
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 1.0015 | 0.00166 | 76183 | ST | [25 empty] | 2019-01-23 | 16225 | Graphite | Completed
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 0.9987 | 0.00158 | 76184 | EA | [25 empty] | 2019-02-04 | 17542 | AMS Submission Results | Completed
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 | 0.99783 | 0.00159 | 76185 | EA | [25 empty] | 2019-02-07 | 21070 | Close Job | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | ST
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | ST
[163 empty] | 2018-08-13 | 0.00798 | 74967 | 3183 | 1.05 |  |  |  | EA
[7 empty] | 137.9 | 88.4346 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48990 | 48990 | 48990 | 260 | 260 | 260 |  |  | jenny dahl | -997.75 | 0.07 ‰ | 2019-02-15 | 2018-08-19 | 15 | 180 |  |  |  |  | 3.73‰ | 3.73‰ | 2.00‰ | 2.00‰ | 3.73 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -997.75 ‰ | -997.75 ‰ | 0.07 | 0.07 | -997.73 | 0.07 ‰ | -997.73 | 0.07 ‰ |  |  | 2019-02-20 | Jocelyn Turnbull | 16:45 | 0.002246 | 0.002246 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.00225 | 7e-05 | 17666 | 2019-02-16 | 16:18 | Albert Zondervan | 211853 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.22 | 0.01% | 15 | 5 | 4 | ... | 0.00216 | 7e-05 |  |  |  |  | 0.99747 | 0.99747 | Yes | Yes | [5 empty] | 0.00216 | 0.00‰ | 7e-05 | 0.00‰ | No | 1 | 1 | 0.00081 | 0.00081 | 0.00225 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00216 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-02-13 | jenny dahl | 11:39 |  | Completed |  |  |  |  | 76285 | 2000000 |  |  | 76298 76299 76311 76312 76313 76314 | 76298~76299~76311~76312~76313~76314 | 3215 |  | 3215 3215 3215 3215 3215 3215 | 2019-02-18 | 10:04 | Albert Zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0013 | 0.00231 | 76298 |  | XCAMS | ST | 0.12% |  | TW3215 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0.12% * random error multiplication factor for cat#11 (TP76330) = 1.3 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is 2-sigma older than consensus value blanks: * Two groups have been identified to blank-correct: carbonates: 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks organics via AAA: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures * cat#37 (TP76271): 5 permil difference between AMS and IRMS del13C | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2019-03-14 | Carbonate | Complete | JOB-211853 |  |  | 66891 |  | 14047/1 |  |  | 2019-01-31 | 3790 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0036 | 0.00196 | 76299 | [26 empty] | 2019-02-01 | 16295 | Graphite | Completed
[168 empty] | 0.99928 | 0.0019 | 76311 | [26 empty] | 2019-02-04 | 6401 | C13 Measurement | Completed
[168 empty] | 0.99878 | 0.00192 | 76312 | [26 empty] | 2019-02-20 | 17666 | AMS Submission Results | Completed
[168 empty] | 0.99889 | 0.00194 | 76313 | [26 empty] | 2019-02-22 | 21248 | Close Job | Completed
[168 empty] | 0.99892 | 0.0019 | 76314
[7 empty] | 79.5 | 88.9028 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42953 | 42952 | 42952 | 192 | 191 | 191 |  |  | jenny dahl | -995.24 | 0.11 ‰ | 2019-02-22 | 2018-08-26 | 11 | 180 |  |  |  |  | 2.09‰ | 2.09‰ | 2.00‰ | 2.00‰ | 2.09 ‰ | 2.00 ‰ | AMS | 2.69‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -995.24 ‰ | -995.24 ‰ | 0.11 | 0.11 | -995.18 | 0.12 ‰ | -995.18 | 0.12 ‰ |  |  | 2019-02-25 | Jocelyn Turnbull | 11:54 | 0.0047623 | 0.0047623 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.7 |  |  | 0.00476 | 0.00011 | 17715 |  |  |  | 210848 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.48 | 0.01% | 6 | 5 | 4 | ... | 0.00458 | 0.00011 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.00458 | 0.00‰ | 0.00011 | 0.00‰ | No | 1 | 0.99998 | 0.00127 | 0.00127 | 0.00476 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00458 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-02-15 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 76348 | 2000000 |  |  | 76186 76187 76188 76194 76372 76394 | 76186~76187~76188~76194~76372~76394 | 3217 |  | 3217 3217 3217 3217 3217 3217 | 2019-02-25 | 10:27 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99924 | 0.0031 | 76186 |  | XCAMS | EA | 0.27% |  | TW3217 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * global factor: 1.03 * single-cathode factors: 1.3 for cat#8 (TP76363) * system error: 0.27% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result is 1.5-sigma younger than consensus value * 1x IAEA-C2 water: result is 2.5-sigma younger than consensus value blanks: * groups identified to blank-correct: carbonates: 26x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2018-11-27 | Carbonate | Complete | JOB-210848 |  | No C13 split taken on carbonate line | 66971 |  | 14047/1 |  |  | 2019-02-08 | 3789 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00313 | 76187 | [26 empty] | 2019-02-11 | 16358 | Graphite | Completed
[168 empty] | 1.0057 | 0.00311 | 76188 | [26 empty] | 2019-02-14 | 6416 | C13 Measurement | Completed
[168 empty] | 0.99656 | 0.00309 | 76194 | [26 empty] | 2019-02-25 | 17715 | AMS Submission Results | Completed
[168 empty] | 0.999 | 0.0031 | 76372 | [26 empty] | 2019-02-26 | 21280 | Close Job | Completed
[168 empty] | 0.99911 | 0.00317 | 76394
[7 empty] | 86.9 | 88.775 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36319 | 36318 | 36318 | 121 | 120 | 120 |  |  | jenny dahl | -989.12 | 0.16 ‰ | 2019-02-22 | 2018-08-26 | 10 | 180 |  |  |  |  | 3.53‰ | 3.53‰ | 2.00‰ | 2.00‰ | 3.53 ‰ | 2.00 ‰ | AMS | 2.64‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -989.12 ‰ | -989.12 ‰ | 0.16 | 0.16 | -988.99 | 0.17 ‰ | -988.99 | 0.17 ‰ |  |  | 2019-02-25 | Jocelyn Turnbull | 11:54 | 0.010876 | 0.010876 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.01088 | 0.00016 | 17719 |  |  |  | 211586 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.09 | 0.02% | 29 | 5 | 8 | ... | 0.01046 | 0.00016 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.01046 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 0.99998 | 0.00127 | 0.00127 | 0.01088 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.01046 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-02-15 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 76353 | 2000000 |  |  | 76186 76187 76188 76194 76372 76394 | 76186~76187~76188~76194~76372~76394 | 3217 |  | 3217 3217 3217 3217 3217 3217 | 2019-02-25 | 10:27 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99924 | 0.0031 | 76186 |  | XCAMS | EA | 0.27% |  | TW3217 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * global factor: 1.03 * single-cathode factors: 1.3 for cat#8 (TP76363) * system error: 0.27% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result is 1.5-sigma younger than consensus value * 1x IAEA-C2 water: result is 2.5-sigma younger than consensus value blanks: * groups identified to blank-correct: carbonates: 26x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2019-01-29 | Water | Complete | JOB-211586 |  |  | 66972 | Normal | 14047/11 |  |  | 2018-12-14 | 3718 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00313 | 76187 | [26 empty] | 2019-01-23 | 3773 | Water CO2 Evolution | Cancelled
[168 empty] | 1.0057 | 0.00311 | 76188 | [26 empty] | 2019-01-23 | 3775 | Water CO2 Evolution | Completed
[168 empty] | 0.99656 | 0.00309 | 76194 | [26 empty] | 2019-01-23 | 2418 | Gas Splitting | Completed
[168 empty] | 0.999 | 0.0031 | 76372 | [26 empty] | 2019-01-23 | 4228 | Choose Gas Splitting | Completed
[168 empty] | 0.99911 | 0.00317 | 76394 | [26 empty] | 2019-01-28 | 6383 | C13 Measurement | Completed
[197 empty] | 2019-02-12 | 16363 | Graphite | Completed
[197 empty] | 2019-02-25 | 17719 | AMS Submission Results | Completed
[197 empty] | 2019-02-26 | 21253 | Close Job | Completed
[5 empty] | 5.1427e-05 | 0.0077298 | 80.9 | 89.8244 | 0.0703 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 6990 | 6989 | 6989 | 52 | 52 | 55 | 6937.3 | 7042.3 | jenny dahl | -581.1 | 2.74 ‰ | 2019-02-22 | 2018-08-26 | 10 | 180 |  |  |  |  | -8.19‰ | -8.19‰ | 2.00‰ | 2.00‰ | -8.19 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -581.10 ‰ | -581.10 ‰ | 2.74 | 2.9 | -575.93 | 2.77 ‰ | -575.93 | 2.77 ‰ |  |  | 2019-02-25 | Jocelyn Turnbull | 11:54 | 0.4189 | 0.4189 | 0.00274 | 0.0029 | 0.00268 | 0.00268 | 4 |  |  | 0.42432 | 0.00119 | 17720 |  |  |  | 211588 | 0.00872 | 0.0087 | 0.00392 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.89 | 0.27% | 15 | 5 | 7 | ... | 0.40808 | 0.0015 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes |  | result is 2.5-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 55.5232835700234809. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00872 | 0.00743 | 0.00352 | 0.40287 | 0.40‰ | 0.0028 | 0.00‰ | No | 1 | 0.99998 | 0.00127 | 0.00127 | 0.4189 | 0 | 0.00274 | 0 | 0 | 0.00872 | 0 | 0.00392 | 0 | 0.40287 | 0 | 0.0028 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-02-15 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 76354 | 2000000 | 76353 76214 76104 75968 75971 75854 75841 75742 75583 | 76353~76214~76104~75968~75971~75854~75841~75742~75583 | 76186 76187 76188 76194 76372 76394 | 76186~76187~76188~76194~76372~76394 | 3217 | 3217 3212 3210 3209 3207 3204 3202 3201 3199 | 3217 3217 3217 3217 3217 3217 | 2019-02-25 | 10:27 | Albert Zondervan |  | 0.52 mg | 0.13 | 0.52 | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 0.99924 | 0.0031 | 76186 | ST | XCAMS | EA | 0.27% |  | TW3217 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * global factor: 1.03 * single-cathode factors: 1.3 for cat#8 (TP76363) * system error: 0.27% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result is 1.5-sigma younger than consensus value * 1x IAEA-C2 water: result is 2.5-sigma younger than consensus value blanks: * groups identified to blank-correct: carbonates: 26x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-01-29 | Water | Complete | JOB-211588 |  |  | 66973 | Normal | 14047/12 |  |  | 2018-12-14 | 3719 | Water CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 1.0003 | 0.00313 | 76187 | ST | [25 empty] | 2019-01-23 | 3774 | Water CO2 Evolution | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 1.0057 | 0.00311 | 76188 | ST | [25 empty] | 2019-01-23 | 2417 | Gas Splitting | Completed
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 0.99656 | 0.00309 | 76194 | EA | [25 empty] | 2019-01-23 | 4226 | Choose Gas Splitting | Completed
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 0.999 | 0.0031 | 76372 | EA | [25 empty] | 2019-02-12 | 16364 | Graphite | Completed
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 | 0.99911 | 0.00317 | 76394 | ST | [25 empty] | 2019-02-18 | 6429 | C13 Measurement | Completed
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | ST | [25 empty] | 2019-02-25 | 17720 | AMS Submission Results | Completed
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA | [25 empty] | 2019-02-26 | 21254 | Close Job | Completed
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | EA
[7 empty] | 118 | 88.6531 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38353 | 38352 | 38352 | 148 | 148 | 148 |  |  | Julia Collins | -991.56 | 0.16 ‰ | 2019-03-07 | 2018-09-08 | 15 | 180 |  |  |  |  | 3.38‰ | 3.38‰ | 2.00‰ | 2.00‰ | 3.38 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0046 | 0 | 0.8165 | -991.56 ‰ | -991.56 ‰ | 0.16 | 0.16 | -991.45 | 0.16 ‰ | -991.45 | 0.16 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:33 | 0.0084431 | 0.0084431 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00844 | 0.00016 | 17793 | 2019-03-11 | 9:42 | Albert Zondervan | 212062 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 8 | 5 | 4 | ... | 0.00812 | 0.00015 |  |  |  |  | 0.99958 | 0.99958 | Yes | Yes | [5 empty] | 0.00812 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00844 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00812 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-04 | Julia Collins | 14:12 |  | Completed |  |  |  |  | 76446 | 2000000 |  |  | 76399 76400 76401 76480 76483 76486 | 76399~76400~76401~76480~76483~76486 | 3218 |  | 3218 3218 3218 3218 3218 3218 | 2019-03-18 | 15:08 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99966 | 0.00167 | 76399 |  | XCAMS | EA | 0% |  | TW3218 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% * uncertainty multiplier for cat#36 (TP76442) = 1.4 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks atmos-CH4 as CO2 in break-seals: 13x unknowns + 0x secondaries + 0x blanks Since blank-correction must involve NIWA's process blank as well and since it has to be done separately, only the first step has been done here, by manual entering the graphitisation blank using 117x Kapuni (R40699/1) measurements. unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-04-04 | Water | Complete | JOB-212062 |  |  | 67011 | Normal | 14047/11 | CH4 |  | 2019-02-19 | 3846 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00195 | 76400 | [26 empty] | 2019-02-20 | 16456 | Graphite | Completed
[168 empty] | 0.99954 | 0.00192 | 76401 | [26 empty] | 2019-02-21 | 6464 | C13 Measurement | Completed
[168 empty] | 0.99953 | 0.00169 | 76480 | [26 empty] | 2019-03-21 | 17793 | AMS Submission Results | Completed
[168 empty] | 0.99774 | 0.00169 | 76483 | [26 empty] | 2019-03-21 | 21458 | Close Job | Completed
[168 empty] | 1.0035 | 0.0017 | 76486
[5 empty] | 4.884e-05 | 0.0077998 | 117.3 | 89.4564 | 0.0668 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7144 | 7144 | 7144 | 54 | 54 | 57 | 7090.2 | 7198.7 | Julia Collins | -589.08 | 2.78 ‰ | 2019-03-07 | 2018-09-08 | 14 | 180 |  |  |  |  | -5.63‰ | -5.63‰ | 2.00‰ | 2.00‰ | -5.63 ‰ | 2.00 ‰ | AMS | -8.06‰ | 0.2‰ | C13 | No | -19 | -19.0046 | 0 | 0.8165 | -589.08 ‰ | -589.08 ‰ | 2.78 | 2.93 | -584.01 | 2.81 ‰ | -584.01 | 2.81 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:33 | 0.41092 | 0.41092 | 0.00278 | 0.00293 | 0.00272 | 0.00272 | 1.8 |  |  | 0.41637 | 0.00126 | 17794 | 2019-03-11 | 9:42 | Albert Zondervan | 212063 | 0.00866 | 0.0087 | 0.0039 | 0.0039 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.09 | 0.28% | 27 | 5 | 4 | ... | 0.40043 | 0.00109 |  |  |  |  | 0.99958 | 0.99958 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 57.2055752450762073. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00866 | 0.00749 | 0.00333 | 0.39519 | 0.40‰ | 0.00262 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.41092 | 0 | 0.00278 | 0 | 0 | 0.00866 | 0 | 0.0039 | 0 | 0.39519 | 0 | 0.00262 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-04 | Julia Collins | 14:12 |  | Completed |  |  |  |  | 76450 | 2000000 | 76446 76353 76214 76104 75968 75971 75854 75841 75742 75583 | 76446~76353~76214~76104~75968~75971~75854~75841~75742~75583 | 76399 76400 76401 76480 76483 76486 | 76399~76400~76401~76480~76483~76486 | 3218 | 3218 3217 3212 3210 3209 3207 3204 3202 3201 3199 | 3218 3218 3218 3218 3218 3218 | 2019-03-18 | 15:08 | Albert Zondervan |  | 1.11 mg | 0.13 | 1.11 | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 | 0.99966 | 0.00167 | 76399 | ST | XCAMS | EA | 0% |  | TW3218 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% * uncertainty multiplier for cat#36 (TP76442) = 1.4 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks atmos-CH4 as CO2 in break-seals: 13x unknowns + 0x secondaries + 0x blanks Since blank-correction must involve NIWA's process blank as well and since it has to be done separately, only the first step has been done here, by manual entering the graphitisation blank using 117x Kapuni (R40699/1) measurements. unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-04-04 | Water | Complete | JOB-212063 |  |  | 67012 | Normal | 14047/12 | CH4 |  | 2019-02-19 | 3847 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 1 | 0.00195 | 76400 | ST | [25 empty] | 2019-02-21 | 16460 | Graphite | Completed
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 0.99954 | 0.00192 | 76401 | ST | [25 empty] | 2019-02-26 | 6465 | C13 Measurement | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 0.99953 | 0.00169 | 76480 | EA | [25 empty] | 2019-03-21 | 17794 | AMS Submission Results | Completed
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 0.99774 | 0.00169 | 76483 | EA | [25 empty] | 2019-03-21 | 21459 | Close Job | Completed
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 | 1.0035 | 0.0017 | 76486 | ST
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 |  |  |  | ST
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | EA
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | EA
[7 empty] | 94.5 | 88.7104 | 0.0025 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33678 | 33678 | 33678 | 127 | 127 | 127 |  |  | Julia Collins | -984.89 | 0.24 ‰ | 2019-03-14 | 2018-09-15 | 13 | 180 |  |  |  |  | 2.61‰ | 2.61‰ | 2.00‰ | 2.00‰ | 2.61 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -984.89 ‰ | -984.89 ‰ | 0.24 | 0.24 | -984.71 | 0.24 ‰ | -984.71 | 0.24 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:41 | 0.015108 | 0.015108 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 3.5 |  |  | 0.01511 | 0.00024 | 17870 | 2019-03-15 | 16:48 | Albert Zondervan | 212111 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.51 | 0.02% | 15 | 5 | 7 | ... | 0.01453 | 0.00023 |  |  |  |  | 0.99946 | 0.99946 | Yes | Yes | [5 empty] | 0.01453 | 0.01‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.01511 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01453 | 0 | 0.00023 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-13 | Julia Collins | 14:28 |  | Completed |  |  |  |  | 76506 | 2000000 |  |  | 76491 76524 76525 76589 76599 76600 | 76491~76524~76525~76589~76599~76600 | 3220 |  | 3220 3220 3220 3220 3220 3220 | 2019-03-18 | 14:12 | Albert Zondervan |  | 0.57 mg |  | 0.57 | [5 empty] | 0.99959 | 0.00195 | 76491 |  | XCAMS | EA | 0% |  | TW3220 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is 1.5-sigma younger than consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 16x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2019-04-08 | Water | Not Complete | JOB-212111 |  | C13 split taken on RG20, not waterline. Higher non-cons when opening to add hydrogen on graphite than others. | 67091 | Normal | 14047/11 |  |  | 2019-02-26 | 3858 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00197 | 76524 | [26 empty] | 2019-02-26 | 3860 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99786 | 0.00196 | 76525 | [26 empty] | 2019-02-28 | 3861 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.00193 | 76589 | [26 empty] | 2019-03-01 | 2442 | Gas Splitting | Cancelled
[168 empty] | 0.99918 | 0.00199 | 76599 | [26 empty] | 2019-03-01 | 16516 | Graphite | Completed
[168 empty] | 1.0005 | 0.00205 | 76600 | [26 empty] | 2019-03-01 | 6480 | C13 Measurement | Completed
[197 empty] | 2019-03-21 | 17870 | AMS Submission Results | Completed
[197 empty] | 2019-03-21 | 21465 | Close Job | Completed
[5 empty] | 4.7852e-05 | 0.0080924 | 95.9 | 89.5888 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7051 | 7051 | 7051 | 58 | 57 | 60 | 6993.9 | 7109.1 | Julia Collins | -584.3 | 2.98 ‰ | 2019-03-14 | 2018-09-15 | 15 | 180 |  |  |  |  | -7.22‰ | -7.22‰ | 2.00‰ | 2.00‰ | -7.22 ‰ | 2.00 ‰ | AMS | -8.18‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -584.30 ‰ | -584.30 ‰ | 2.98 | 3.13 | -579.17 | 3.02 ‰ | -579.17 | 3.02 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:41 | 0.4157 | 0.4157 | 0.00298 | 0.00313 | 0.00293 | 0.00293 | 3.7 |  |  | 0.42144 | 0.00144 | 17871 | 2019-03-15 | 16:48 | Albert Zondervan | 212112 | 0.00919 | 0.0092 | 0.00414 | 0.0041 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.57 | 0.3% | 29 | 5 | 4 | ... | 0.40531 | 0.00128 |  |  |  |  | 0.99946 | 0.99946 | Yes | Yes |  | result is 1.5-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 60.3891046582006705. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00919 | 0.00778 | 0.00362 | 0.39979 | 0.40‰ | 0.00282 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.4157 | 0 | 0.00298 | 0 | 0 | 0.00919 | 0 | 0.00414 | 0 | 0.39979 | 0 | 0.00282 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-13 | Julia Collins | 14:28 |  | Completed |  |  |  |  | 76472 | 2000000 | 76506 76446 76353 76214 76104 75968 75971 75854 75841 75742 75583 | 76506~76446~76353~76214~76104~75968~75971~75854~75841~75742~75583 | 76491 76524 76525 76589 76599 76600 | 76491~76524~76525~76589~76599~76600 | 3220 | 3220 3218 3217 3212 3210 3209 3207 3204 3202 3201 3199 | 3220 3220 3220 3220 3220 3220 | 2019-03-18 | 14:12 | Albert Zondervan |  | 0.56 mg | 0.13 | 0.56 | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 0.99959 | 0.00195 | 76491 | ST | XCAMS | EA | 0% |  | TW3220 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is 1.5-sigma younger than consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 16x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-04-08 | Water | Complete | JOB-212112 |  |  | 67092 | Normal | 14047/12 |  |  | 2019-02-26 | 3859 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 | 1.0013 | 0.00197 | 76524 | ST | [25 empty] | 2019-02-27 | 16482 | Graphite | Completed
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 0.99786 | 0.00196 | 76525 | ST | [25 empty] | 2019-03-01 | 6471 | C13 Measurement | Completed
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 1.0016 | 0.00193 | 76589 | EA | [25 empty] | 2019-03-21 | 17871 | AMS Submission Results | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 0.99918 | 0.00199 | 76599 | EA | [25 empty] | 2019-03-21 | 21466 | Close Job | Completed
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 | 1.0005 | 0.00205 | 76600 | ST
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 |  |  |  | ST
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 |  |  |  | EA
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | EA
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | EA
[7 empty] | 76.1 | 88.8777 | 0.0005 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47138 | 47138 | 47138 | 236 | 236 | 236 |  |  | Julia Collins | -997.17 | 0.08 ‰ | 2019-04-08 | 2018-10-10 | 12 | 180 |  |  |  |  | -1.00‰ | -1.00‰ | 2.00‰ | 2.00‰ | -1.00 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0008 | 0 | 0.8165 | -997.17 ‰ | -997.17 ‰ | 0.08 | 0.08 | -997.14 | 0.08 ‰ | -997.14 | 0.08 ‰ |  |  | 2019-04-11 | Margaret Norris | 12:10 | 0.0028282 | 0.0028282 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 6.3 |  |  | 0.00283 | 8e-05 | 18049 |  |  |  | 212145 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 11 | 5 | 4 | ... | 0.00272 | 8e-05 |  |  |  |  | 0.99773 | 0.99773 | Yes | Yes | [5 empty] | 0.00272 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00283 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00272 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-05 | Julia Collins | 14:51 |  | Completed |  |  |  |  | 76700 | 2000000 |  |  | 76648 76649 76650 76653 76654 76655 | 76648~76649~76650~76653~76654~76655 | 3225 |  | 3225 3225 3225 3225 3225 3225 | 2019-04-11 | 11:21 | Albert Zondervan |  | 0.33 mg |  | 0.33 | [5 empty] | 0.99899 | 0.00151 | 76648 |  | XCAMS | ST | 0% |  | TW3225 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 9x, all via ST, all used for normalisation in CalAMS 6x large: no failures 3x small: no failures d = 3.77 ugC extracted from the 4x small OxI in TW3199,3202,3225 secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is slightly more than 2-sigma older than consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: carbonates, large: 11x unknowns + 1x secondaries + 2x blanks 180 days: capturing 5x blanks organics via AAA, large: 6x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 4x unknowns + 1x secondaries + 3x blanks blanks used: 3x from this wheel + 2x from TW3202 + 2x from TW3199 m = 1.38 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-04-25 | Carbonate | Complete | JOB-212145 |  | SLUSHIE aim for 0.25mg | 67304 |  | 14047/1 |  |  | 2019-02-28 | 5805 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.001 | 0.00151 | 76649 | [26 empty] | 2019-03-25 | 3901 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99825 | 0.00151 | 76650 | [26 empty] | 2019-03-27 | 16710 | Graphite | Completed
[168 empty] | 1.0009 | 0.0015 | 76653 | [26 empty] | 2019-04-11 | 18049 | AMS Submission Results | Completed
[168 empty] | 0.99888 | 0.0015 | 76654 | [26 empty] | 2019-04-11 | 21642 | Close Job | Completed
[168 empty] | 1.002 | 0.00153 | 76655
[7 empty] | 82.3 | 88.4095 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47287 | 47287 | 47287 | 241 | 240 | 240 |  |  | Julia Collins | -997.22 | 0.08 ‰ | 2019-04-08 | 2018-10-10 | 31 | 180 |  |  |  |  | 4.29‰ | 4.29‰ | 2.00‰ | 2.00‰ | 4.29 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0008 | 0 | 0.8165 | -997.22 ‰ | -997.22 ‰ | 0.08 | 0.08 | -997.19 | 0.08 ‰ | -997.19 | 0.08 ‰ |  |  | 2019-04-11 | Margaret Norris | 12:10 | 0.0027763 | 0.0027763 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00278 | 8e-05 | 18053 |  |  |  | 212171 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 31 | 5 | 5 | ... | 0.00267 | 8e-05 |  |  |  |  | 0.99773 | 0.99773 | Yes | Yes | [5 empty] | 0.00267 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00278 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00267 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-05 | Julia Collins | 14:51 |  | Completed |  |  |  |  | 76555 | 2000000 |  |  | 76648 76649 76650 76653 76654 76655 | 76648~76649~76650~76653~76654~76655 | 3225 |  | 3225 3225 3225 3225 3225 3225 | 2019-04-11 | 11:21 | Albert Zondervan |  | 1.03 mg |  | 1.03 | [5 empty] | 0.99899 | 0.00151 | 76648 |  | XCAMS | ST | 0% |  | TW3225 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 9x, all via ST, all used for normalisation in CalAMS 6x large: no failures 3x small: no failures d = 3.77 ugC extracted from the 4x small OxI in TW3199,3202,3225 secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is slightly more than 2-sigma older than consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: carbonates, large: 11x unknowns + 1x secondaries + 2x blanks 180 days: capturing 5x blanks organics via AAA, large: 6x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 4x unknowns + 1x secondaries + 3x blanks blanks used: 3x from this wheel + 2x from TW3202 + 2x from TW3199 m = 1.38 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-05-01 | Carbonate |  | JOB-212171 |  |  | 67308 |  | 14047/1 |  |  | 2019-03-07 | 3867 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.001 | 0.00151 | 76649 | [26 empty] | 2019-03-07 | 2446 | Gas Splitting | Completed
[168 empty] | 0.99825 | 0.00151 | 76650 | [26 empty] | 2019-03-07 | 4267 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.0015 | 76653 | [26 empty] | 2019-03-08 | 16565 | Graphite | Completed
[168 empty] | 0.99888 | 0.0015 | 76654 | [26 empty] | 2019-04-11 | 18053 | AMS Submission Results | Completed
[168 empty] | 1.002 | 0.00153 | 76655 | [26 empty] | 2019-04-11 | 21643 | Close Job | Completed
[7 empty] | 87.9 | 88.6826 | 0.0028 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 32834 | 32834 | 32834 | 134 | 134 | 134 |  |  | Julia Collins | -983.22 | 0.28 ‰ | 2019-04-11 | 2018-10-13 | 28 | 180 |  |  |  |  | 2.21‰ | 2.21‰ | 2.00‰ | 2.00‰ | 2.21 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -983.22 ‰ | -983.22 ‰ | 0.28 | 0.28 | -983.01 | 0.28 ‰ | -983.01 | 0.28 ‰ |  |  | 2019-04-15 | Cathy Ginnane | 10:01 | 0.016782 | 0.016782 | 0.00028 | 0.00028 | 0.00028 | 0.00028 | 2.8 |  |  | 0.01678 | 0.00028 | 18077 |  |  |  | 212206 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.68 | 0.03% | 6 | 5 | 4 | ... | 0.01614 | 0.00027 |  |  |  |  | 0.99875 | 0.99875 | Yes | Yes |  | Youngest blank in a long time. Not a good blank. Sitting in graphite for three weeks--possibly contributing? |  |  |  | 0.01614 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00082 | 0.00082 | 0.01678 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01614 | 0 | 0.00027 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-10 | Julia Collins | 9:55 |  | Completed |  |  |  |  | 76607 | 2000000 |  |  | 76712 76715 76717 76739 76742 76745 | 76712~76715~76717~76739~76742~76745 | 3226 |  | 3226 3226 3226 3226 3226 3226 | 2019-04-12 | 12:53 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 0.99692 | 0.002 | 76712 |  | XCAMS | EA | 0% |  | TW3226 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x FIRI-I cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 12x blanks organics via AAA: 7x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks organics via CE-Rafter (cellulose): 6x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2019-05-03 | Water | Complete | JOB-212206 |  | C13 split sealed as it was not needed at the time | 67339 | Normal | 14047/11 |  |  | 2019-03-11 | 3878 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00201 | 76715 | [26 empty] | 2019-03-19 | 16617 | Graphite | Completed
[168 empty] | 0.99958 | 0.002 | 76717 | [26 empty] | 2019-03-28 | 6505 | C13 Measurement | Completed
[168 empty] | 1.002 | 0.00205 | 76739 | [26 empty] | 2019-04-15 | 18077 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00201 | 76742 | [26 empty] | 2019-04-17 | 21704 | Close Job | Completed
[168 empty] | 1.0001 | 0.002 | 76745
[5 empty] | 4.7172e-05 | 0.0083377 | 85.6 | 89.5785 | 0.0698 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7069 | 7069 | 7069 | 61 | 60 | 63 | 7008.7 | 7130.6 | Julia Collins | -585.24 | 3.15 ‰ | 2019-04-11 | 2018-10-13 | 28 | 180 |  |  |  |  | -7.81‰ | -7.81‰ | 2.00‰ | 2.00‰ | -7.81 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -585.24 ‰ | -585.24 ‰ | 3.15 | 3.28 | -580.12 | 3.18 ‰ | -580.12 | 3.18 ‰ |  |  | 2019-04-15 | Cathy Ginnane | 10:01 | 0.41476 | 0.41476 | 0.00315 | 0.00328 | 0.0031 | 0.0031 | 3.1 |  |  | 0.42087 | 0.00146 | 18086 |  |  |  | 212235 | 0.00977 | 0.0098 | 0.0044 | 0.0044 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.48 | 0.31% | 22 | 5 | 5 | ... | 0.40476 | 0.0013 |  |  |  |  | 0.99875 | 0.99875 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 63.5429218639386871. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00977 | 0.00801 | 0.00399 | 0.39889 | 0.40‰ | 0.00298 | 0.00‰ | Yes | 1 | 1 | 0.00082 | 0.00082 | 0.41476 | 0 | 0.00315 | 0 | 0 | 0.00977 | 0 | 0.0044 | 0 | 0.39889 | 0 | 0.00298 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-10 | Julia Collins | 9:55 |  | Completed |  |  |  |  | 76608 | 2000000 | 76607 76506 76446 76353 76214 76104 75968 75971 75854 75841 75742 75583 | 76607~76506~76446~76353~76214~76104~75968~75971~75854~75841~75742~75583 | 76712 76715 76717 76739 76742 76745 | 76712~76715~76717~76739~76742~76745 | 3226 | 3226 3220 3218 3217 3212 3210 3209 3207 3204 3202 3201 3199 | 3226 3226 3226 3226 3226 3226 | 2019-04-12 | 12:53 | Albert Zondervan |  | 0.69 mg | 0.13 | 0.69 | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 0.99692 | 0.002 | 76712 | ST | XCAMS | EA | 0% |  | TW3226 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x FIRI-I cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 12x blanks organics via AAA: 7x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks organics via CE-Rafter (cellulose): 6x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-05-07 | Water |  | JOB-212235 |  | Take C13 split on water line and create new job for it | 67342 | Normal | 14047/12 |  |  | 2019-03-13 | 3887 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 1.0001 | 0.00201 | 76715 | ST | [25 empty] | 2019-03-13 | 2452 | Gas Splitting | Completed
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 | 0.99958 | 0.002 | 76717 | ST | [25 empty] | 2019-03-13 | 4275 | Choose Gas Splitting | Completed
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 1.002 | 0.00205 | 76739 | EA | [25 empty] | 2019-03-19 | 16618 | Graphite | Completed
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 1.0014 | 0.00201 | 76742 | EA | [25 empty] | 2019-04-15 | 18086 | AMS Submission Results | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 | 1.0001 | 0.002 | 76745 | ST | [25 empty] | 2019-04-17 | 21705 | Close Job | Completed
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 |  |  |  | ST
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 |  |  |  | EA
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 |  |  |  | EA
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | EA
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | EA
[7 empty] | 75 | 88.8684 | 0.0031 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 32012 | 32011 | 32011 | 90 | 89 | 89 |  |  | Julia Collins | -981.41 | 0.21 ‰ | 2019-04-24 | 2018-10-26 | 9 | 180 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -981.41 ‰ | -981.41 ‰ | 0.21 | 0.21 | -981.18 | 0.21 ‰ | -981.18 | 0.21 ‰ |  |  | 2019-05-07 | Jocelyn Turnbull | 12:49 | 0.018592 | 0.018592 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2.3 |  |  | 0.01859 | 0.00021 | 18150 | 2019-04-26 | 13:00 | Albert Zondervan | 212255 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.86 | 0.02% | 8 | 5 | 5 | ... | 0.01788 | 0.0002 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes | [5 empty] | 0.01788 | 0.02‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01859 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01788 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-18 | Julia Collins | 13:23 |  | Completed |  |  |  |  | 76841 | 2000000 |  |  | 76773 76774 76775 76869 76871 76894 | 76773~76774~76775~76869~76871~76894 | 3228 |  | 3228 3228 3228 3228 3228 3228 | 2019-05-07 | 12:32 | Albert Zondervan |  | 0.80 mg |  | 0.8 | [5 empty] | 0.99885 | 0.00159 | 76773 |  | XCAMS | ST | 0% |  | TW3228 measurement: * no issues with XCAMS * Cathodes had to be transferred to another wheel because of persistent ticking on cat#38,39. This resolved the issue completely. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1.04 * cathode run-to-run variability factor = 1.3 (cat#25), 1.2 (cat#36) primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 25x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 3x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks: 1x Queets 'A' from this wheel 7x Kauri Renton Rd from previous wheels unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-05-13 | Water |  | JOB-212255 |  |  | 67451 | Normal | 14047/11 |  |  | 2019-03-18 | 3892 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0006 | 0.00156 | 76774 | [26 empty] | 2019-03-18 | 2457 | Gas Splitting | Completed
[168 empty] | 0.99864 | 0.00161 | 76775 | [26 empty] | 2019-03-18 | 4285 | Choose Gas Splitting | Completed
[168 empty] | 0.9999 | 0.00155 | 76869 | [26 empty] | 2019-04-15 | 16851 | Graphite | Completed
[168 empty] | 0.99943 | 0.00156 | 76871 | [26 empty] | 2019-05-07 | 18150 | AMS Submission Results | Completed
[168 empty] | 1.0025 | 0.00155 | 76894 | [26 empty] | 2019-05-13 | 21849 | Close Job | Completed
[5 empty] | 4.6003e-05 | 0.0088395 | 74.4 | 89.7806 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7141 | 7141 | 7141 | 63 | 62 | 65 | 7079 | 7204.3 | Julia Collins | -588.94 | 3.21 ‰ | 2019-04-24 | 2018-10-26 | 9 | 180 |  |  |  |  | -7.97‰ | -7.97‰ | 2.00‰ | 2.00‰ | -7.97 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -588.94 ‰ | -588.94 ‰ | 3.21 | 3.34 | -583.86 | 3.25 ‰ | -583.86 | 3.25 ‰ |  |  | 2019-05-07 | Jocelyn Turnbull | 12:49 | 0.41106 | 0.41106 | 0.00321 | 0.00334 | 0.00316 | 0.00316 | 2 |  |  | 0.41759 | 0.00119 | 18151 | 2019-04-26 | 13:00 | Albert Zondervan | 212256 | 0.01039 | 0.0104 | 0.00468 | 0.0047 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.11 | 0.32% | 29 | 5 | 5 | ... | 0.40161 | 0.00102 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 65.2073242672899839. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01039 | 0.0085 | 0.00443 | 0.39533 | 0.40‰ | 0.00304 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41106 | 0 | 0.00321 | 0 | 0 | 0.01039 | 0 | 0.00468 | 0 | 0.39533 | 0 | 0.00304 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-18 | Julia Collins | 13:23 |  | Completed |  |  |  |  | 76842 | 2000000 | 76841 76607 76506 76446 76353 76214 76104 75968 75971 75854 75841 75742 75583 | 76841~76607~76506~76446~76353~76214~76104~75968~75971~75854~75841~75742~75583 | 76773 76774 76775 76869 76871 76894 | 76773~76774~76775~76869~76871~76894 | 3228 | 3228 3226 3220 3218 3217 3212 3210 3209 3207 3204 3202 3201 3199 | 3228 3228 3228 3228 3228 3228 | 2019-05-07 | 12:32 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 | 0.99885 | 0.00159 | 76773 | ST | XCAMS | ST | 0% |  | TW3228 measurement: * no issues with XCAMS * Cathodes had to be transferred to another wheel because of persistent ticking on cat#38,39. This resolved the issue completely. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1.04 * cathode run-to-run variability factor = 1.3 (cat#25), 1.2 (cat#36) primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 25x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 3x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks: 1x Queets 'A' from this wheel 7x Kauri Renton Rd from previous wheels unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-05-13 | Water |  | JOB-212256 |  |  | 67452 | Normal | 14047/12 |  |  | 2019-03-18 | 3893 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 1.0006 | 0.00156 | 76774 | ST | [25 empty] | 2019-03-18 | 2456 | Gas Splitting | Completed
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 0.99864 | 0.00161 | 76775 | ST | [25 empty] | 2019-03-18 | 4283 | Choose Gas Splitting | Completed
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 | 0.9999 | 0.00155 | 76869 | EA | [25 empty] | 2019-04-15 | 16852 | Graphite | Completed
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 0.99943 | 0.00156 | 76871 | EA | [25 empty] | 2019-05-07 | 18151 | AMS Submission Results | Completed
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 | 1.0025 | 0.00155 | 76894 | ST | [25 empty] | 2019-05-13 | 21850 | Close Job | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 |  |  |  | ST
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 |  |  |  | EA
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 |  |  |  | EA
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 |  |  |  | EA
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | EA
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA
[163 empty] | 2018-10-29 | 0.0148 | 75583 | 3199 | 0.86 |  |  |  | ST
[7 empty] | 62.5 | 88.6224 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38895 | 38894 | 38894 | 159 | 158 | 158 |  |  | Margaret Norris | -992.11 | 0.16 ‰ | 2019-05-14 | 2018-11-15 | 7 | 180 |  |  |  |  | 3.50‰ | 3.50‰ | 2.00‰ | 2.00‰ | 3.50 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -18.9413 | 0 | 0.8165 | -992.11 ‰ | -992.11 ‰ | 0.16 | 0.16 | -992.01 | 0.16 ‰ | -992.01 | 0.16 ‰ |  |  | 2019-05-22 | Jocelyn Turnbull | 9:35 | 0.007892 | 0.007892 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00789 | 0.00016 | 18309 | 2019-05-17 | 11:00 | Albert Zondervan | 212412 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.79 | 0.02% | 8 | 5 | 6 | ... | 0.00759 | 0.00015 |  |  |  |  | 0.99936 | 0.99929 | No | Yes | [5 empty] | 0.00759 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.00072 | 0.00789 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00759 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-13 | Margaret Norris | 10:44 |  | Completed |  |  |  |  | 76956 | 2000000 |  |  | 76887 76968 76969 76970 76986 76989 | 76887~76968~76969~76970~76986~76989 | 3232 |  | 3232 3232 3232 3232 3232 3232 | 2019-05-17 | 11:02 | Albert Zondervan |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99914 | 0.00175 | 76887 |  | XCAMS | EA | 0% |  | TW3232 measurement: * SNICS ion source ticky during first rotation; run data rejected. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value * 1x IAEA-C6 ANU sucrose result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-06-03 | Water | Not Complete | JOB-212412 |  |  | 67541 | Normal | 14047/11 |  |  | 2019-05-06 | 5868 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00176 | 76968 | [26 empty] | 2019-05-06 | 3943 | Water CO2 Evolution | Completed
[168 empty] | 1.0017 | 0.00176 | 76969 | [26 empty] | 2019-05-06 | 2517 | Gas Splitting | Completed
[168 empty] | 1.0008 | 0.00173 | 76970 | [26 empty] | 2019-05-06 | 4362 | Choose Gas Splitting | Completed
[168 empty] | 0.99874 | 0.00179 | 76986 | [26 empty] | 2019-05-07 | 16966 | Graphite | Completed
[168 empty] | 0.99928 | 0.00173 | 76989 | [26 empty] | 2019-05-22 | 18309 | AMS Submission Results | Completed
[197 empty] | 2019-05-22 | 21954 | Close Job | Completed
[5 empty] | 4.5054e-05 | 0.0084816 | 60.2 | 89.5983 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7114 | 7114 | 7114 | 60 | 60 | 62 | 7054.4 | 7174.9 | Margaret Norris | -587.56 | 3.09 ‰ | 2019-05-14 | 2018-11-15 | 7 | 180 |  |  |  |  | -7.43‰ | -7.43‰ | 2.00‰ | 2.00‰ | -7.43 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -18.9413 | 0 | 0.8165 | -587.56 ‰ | -587.56 ‰ | 3.09 | 3.23 | -582.46 | 3.13 ‰ | -582.46 | 3.13 ‰ |  |  | 2019-05-22 | Jocelyn Turnbull | 9:35 | 0.41244 | 0.41244 | 0.00309 | 0.00323 | 0.00305 | 0.00305 | 2 |  |  | 0.41862 | 0.0013 | 18318 | 2019-05-17 | 11:00 | Albert Zondervan | 212236 | 0.00984 | 0.0098 | 0.00443 | 0.0044 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.24 | 0.31% | 29 | 5 | 6 | ... | 0.4026 | 0.00114 |  |  |  |  | 0.99936 | 0.99929 | No | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 62.9056342702385709. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00984 | 0.00815 | 0.00428 | 0.39666 | 0.40‰ | 0.00293 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.00072 | 0.41244 | 0 | 0.00309 | 0 | 0 | 0.00984 | 0 | 0.00443 | 0 | 0.39666 | 0 | 0.00293 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-13 | Margaret Norris | 10:44 |  | Completed |  |  |  |  | 76964 | 2000000 | 76956 76841 76607 76506 76446 76353 76214 76104 75968 75971 75854 75841 75742 | 76956~76841~76607~76506~76446~76353~76214~76104~75968~75971~75854~75841~75742 | 76887 76968 76969 76970 76986 76989 | 76887~76968~76969~76970~76986~76989 | 3232 | 3232 3228 3226 3220 3218 3217 3212 3210 3209 3207 3204 3202 3201 | 3232 3232 3232 3232 3232 3232 | 2019-05-17 | 11:02 | Albert Zondervan |  | 1.03 mg | 0.13 | 1.03 | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 | 0.99914 | 0.00175 | 76887 | ST | XCAMS | EA | 0% |  | TW3232 measurement: * SNICS ion source ticky during first rotation; run data rejected. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value * 1x IAEA-C6 ANU sucrose result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-05-07 | Water |  | JOB-212236 |  | Take C13 split on water line and create new job for it. Seal both in breakseals | 67542 | Normal | 14047/12 |  |  | 2019-03-13 | 3888 | Water CO2 Evolution | Cancelled | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 | 1.0003 | 0.00176 | 76968 | ST | [25 empty] | 2019-03-15 | 3891 | Water CO2 Evolution | Completed
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 1.0017 | 0.00176 | 76969 | EA | [25 empty] | 2019-03-15 | 2455 | Gas Splitting | Completed
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 1.0008 | 0.00173 | 76970 | EA | [25 empty] | 2019-03-15 | 4281 | Choose Gas Splitting | Completed
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 | 0.99874 | 0.00179 | 76986 | ST | [25 empty] | 2019-05-07 | 16974 | Graphite | Completed
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 | 0.99928 | 0.00173 | 76989 | ST | [25 empty] | 2019-05-22 | 18318 | AMS Submission Results | Completed
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 |  |  |  | EA | [25 empty] | 2019-05-22 | 21946 | Close Job | Completed
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 |  |  |  | EA
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 |  |  |  | EA
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 |  |  |  | EA
[163 empty] | 2018-11-20 | 0.00664 | 75854 | 3204 | 1.2 |  |  |  | EA
[163 empty] | 2018-11-19 | 0.01097 | 75841 | 3202 | 1.12 |  |  |  | ST
[163 empty] | 2018-11-15 | 0.00741 | 75742 | 3201 | 1.16 |  |  |  | EA
[7 empty] | 82 | 88.6192 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33920 | 33919 | 33919 | 102 | 101 | 101 |  |  | jenny dahl | -985.34 | 0.19 ‰ | 2019-05-28 | 2018-11-29 | 8 | 180 |  |  |  |  | 3.62‰ | 3.62‰ | 2.00‰ | 2.00‰ | 3.62 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.89443 | -985.34 ‰ | -985.34 ‰ | 0.19 | 0.19 | -985.16 | 0.19 ‰ | -985.16 | 0.19 ‰ |  |  | 2019-06-05 | Jocelyn Turnbull | 14:03 | 0.014661 | 0.014661 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 3.1 |  |  | 0.01466 | 0.00019 | 18350 | 2019-05-30 | 10:30 | Albert Zondervan | 212238 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.47 | 0.02% | 17 | 5 | 5 | ... | 0.0141 | 0.00018 |  |  |  |  | 0.99944 | 0.99944 | Yes | Yes | [5 empty] | 0.0141 | 0.01‰ | 0.00018 | 0.00‰ | No | 1 | 0.99998 | 0.00098 | 0.00098 | 0.01466 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0141 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-27 | jenny dahl | 7:55 |  | Completed |  |  |  |  | 77013 | 2000000 |  |  | 77024 77025 77026 77027 77028 | 77024~77025~77026~77027~77028 | 3233 |  | 3233 3233 3233 3233 3233 | 2019-05-30 | 12:10 | Albert Zondervan |  | 0.67 mg |  | 0.67 | [5 empty] | 0.99938 | 0.00222 | 77024 |  | XCAMS | ST | 0.16% |  | TW3233 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0.16% system error * for excess run-to-run variability: x1.2 (cat#7), x1.3 (cat#32) primaries: * OxI cathodes: 7x, all via ST, all used for normalisation in CalAMS 5x large: no failures 2x small (subm = Rees): no failures d = 1.49 ugC extracted from the 4x small OxI in TW3185,3233 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result differs from consensus value by 4-sigma: ~450 years too old blanks: * Four groups have been identified to blank-correct: carbonates, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 4x blanks organics via AAA, large: 1x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks organics via AAA, small, via Excel: 2x unknowns + 1x secondaries + 3x blanks blanks used: Queets'A' only because it could not be resolved why the most recent lot of small KauriRRd appear worse than the few small Queets 'A': 2x Queets from this wheel + 1x from TW3185 m = 0.61 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-05-07 | Water | Complete | JOB-212238 |  | Take C13 split on water line and create new job for it. Seal both into breakseals | 67623 | Normal | 14047/11 |  |  | 2019-03-13 | 3890 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.0022 | 77025 | [26 empty] | 2019-03-13 | 2454 | Gas Splitting | Completed
[168 empty] | 0.99634 | 0.00221 | 77026 | [26 empty] | 2019-03-13 | 4279 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00218 | 77027 | [26 empty] | 2019-05-20 | 17023 | Graphite | Completed
[168 empty] | 1.0017 | 0.0022 | 77028 | [26 empty] | 2019-06-05 | 18350 | AMS Submission Results | Completed
[197 empty] | 2019-06-17 | 22048 | Close Job | Completed
[5 empty] | 4.9243e-05 | 0.0090135 | 81.7 | 89.633 | 0.0663 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7160 | 7159 | 7159 | 64 | 64 | 66 | 7095.7 | 7223.9 | jenny dahl | -589.86 | 3.27 ‰ | 2019-05-28 | 2018-11-29 | 6 | 180 |  |  |  |  | -7.73‰ | -7.73‰ | 2.00‰ | 2.00‰ | -7.73 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.89443 | -589.86 ‰ | -589.86 ‰ | 3.27 | 3.4 | -584.8 | 3.31 ‰ | -584.8 | 3.31 ‰ |  |  | 2019-06-05 | Jocelyn Turnbull | 14:04 | 0.41014 | 0.41014 | 0.00327 | 0.0034 | 0.00323 | 0.00323 | 2 |  |  | 0.41683 | 0.00117 | 18354 | 2019-05-30 | 10:30 | Albert Zondervan | 212413 | 0.01064 | 0.0106 | 0.00479 | 0.0048 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.01 | 0.33% | 35 | 5 | 5 | ... | 0.40088 | 0.00119 |  |  |  |  | 0.99944 | 0.99944 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 66.62486712490355. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01064 | 0.00866 | 0.00463 | 0.39444 | 0.39‰ | 0.00317 | 0.00‰ | No | 1 | 0.99998 | 0.00098 | 0.00098 | 0.41014 | 0 | 0.00327 | 0 | 0 | 0.01064 | 0 | 0.00479 | 0 | 0.39444 | 0 | 0.00317 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-27 | jenny dahl | 7:55 |  | Completed |  |  |  |  | 77036 | 2000000 | 77013 76956 76841 76607 76506 76446 76353 76214 76104 75968 75971 | 77013~76956~76841~76607~76506~76446~76353~76214~76104~75968~75971 | 77024 77025 77026 77027 77028 | 77024~77025~77026~77027~77028 | 3233 | 3233 3232 3228 3226 3220 3218 3217 3212 3210 3209 3207 | 3233 3233 3233 3233 3233 | 2019-05-30 | 12:10 | Albert Zondervan |  | 1.03 mg | 0.13 | 1.03 | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 | 0.99938 | 0.00222 | 77024 | EA | XCAMS | ST | 0.16% |  | TW3233 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0.16% system error * for excess run-to-run variability: x1.2 (cat#7), x1.3 (cat#32) primaries: * OxI cathodes: 7x, all via ST, all used for normalisation in CalAMS 5x large: no failures 2x small (subm = Rees): no failures d = 1.49 ugC extracted from the 4x small OxI in TW3185,3233 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result differs from consensus value by 4-sigma: ~450 years too old blanks: * Four groups have been identified to blank-correct: carbonates, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 4x blanks organics via AAA, large: 1x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks organics via AAA, small, via Excel: 2x unknowns + 1x secondaries + 3x blanks blanks used: Queets'A' only because it could not be resolved why the most recent lot of small KauriRRd appear worse than the few small Queets 'A': 2x Queets from this wheel + 1x from TW3185 m = 0.61 ugC unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-06-03 | Water |  | JOB-212413 |  |  | 67624 | Normal | 14047/12 |  |  | 2019-05-21 | 3968 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 | 1.0011 | 0.0022 | 77025 | ST | [25 empty] | 2019-05-21 | 2527 | Gas Splitting | Completed
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 | 0.99634 | 0.00221 | 77026 | ST | [25 empty] | 2019-05-21 | 4377 | Choose Gas Splitting | Completed
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 1.0013 | 0.00218 | 77027 | EA | [25 empty] | 2019-05-22 | 17046 | Graphite | Completed
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 1.0017 | 0.0022 | 77028 | EA | [25 empty] | 2019-06-05 | 18354 | AMS Submission Results | Completed
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 |  |  |  | EA | [25 empty] | 2019-06-17 | 22054 | Close Job | Completed
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 |  |  |  | EA
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 |  |  |  | EA
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 |  |  |  | ST
[163 empty] | 2018-12-07 | 0.01137 | 75968 | 3209 | 1.28 |  |  |  | EA
[163 empty] | 2018-12-02 | 0.00321 | 75971 | 3207 | 1.01 |  |  |  | ST
[7 empty] | 109.2 | 88.4737 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50455 | 50454 | 50454 | 312 | 312 | 312 |  |  | Julia Collins | -998.13 | 0.07 ‰ | 2019-06-10 | 2018-12-12 | 11 | 180 |  |  |  |  | 3.49‰ | 3.49‰ | 2.00‰ | 2.00‰ | 3.49 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -998.13 ‰ | -998.13 ‰ | 0.07 | 0.07 | -998.11 | 0.07 ‰ | -998.11 | 0.07 ‰ |  |  | 2019-06-14 | Jocelyn Turnbull | 9:20 | 0.0018716 | 0.0018716 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.1 |  |  | 0.00187 | 7e-05 | 18460 | 2019-06-11 | 16:34 | Albert Zondervan | 212568 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 6 | 5 | 5 | ... | 0.0018 | 7e-05 |  |  |  |  | 0.99767 | 0.99767 | Yes | Yes | [5 empty] | 0.0018 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00187 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0018 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-07 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 77098 | 2000000 |  |  | 76991 76994 77057 77058 77060 77061 | 76991~76994~77057~77058~77060~77061 | 3235 |  | 3235 3235 3235 3235 3235 3235 | 2019-06-11 | 16:47 | Albert Zondervan |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99769 | 0.00158 | 76991 |  | XCAMS | EA | 0% |  | TW3235 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result agrees with consensus value * 1x IAEA-C2 water: result agrees with consensus value blanks: * groups identified to blank-correct: carbonates: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-06-26 | Carbonate | Not Complete | JOB-212568 |  |  | 67703 |  | 14047/1 |  |  | 2019-05-21 | 3976 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99965 | 0.00158 | 76994 | [26 empty] | 2019-05-21 | 2526 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00159 | 77057 | [26 empty] | 2019-05-21 | 4375 | Choose Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00161 | 77058 | [26 empty] | 2019-05-30 | 17108 | Graphite | Completed
[168 empty] | 1.0017 | 0.0016 | 77060 | [26 empty] | 2019-06-14 | 18460 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00159 | 77061 | [26 empty] | 2019-06-17 | 22063 | Close Job | Completed
[5 empty] | 5.1881e-05 | 0.0090273 | 109.8 | 89.4633 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7145 | 7145 | 7145 | 64 | 63 | 66 | 7081.9 | 7209.6 | Julia Collins | -589.15 | 3.27 ‰ | 2019-06-10 | 2018-12-12 | 11 | 180 |  |  |  |  | -7.61‰ | -7.61‰ | 2.00‰ | 2.00‰ | -7.61 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -589.15 ‰ | -589.15 ‰ | 3.27 | 3.4 | -584.07 | 3.31 ‰ | -584.07 | 3.31 ‰ |  |  | 2019-06-14 | Jocelyn Turnbull | 9:20 | 0.41085 | 0.41085 | 0.00327 | 0.0034 | 0.00322 | 0.00322 | 2.2 |  |  | 0.41752 | 0.0012 | 18465 | 2019-06-11 | 16:34 | Albert Zondervan | 212575 | 0.01059 | 0.0106 | 0.00477 | 0.0048 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.09 | 0.33% | 37 | 5 | 5 | ... | 0.40154 | 0.00103 |  |  |  |  | 0.99767 | 0.99767 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 66.3936909036312009. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01059 | 0.00867 | 0.00483 | 0.39513 | 0.40‰ | 0.0031 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41085 | 0 | 0.00327 | 0 | 0 | 0.01059 | 0 | 0.00477 | 0 | 0.39513 | 0 | 0.0031 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-07 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 77114 | 2000000 | 77113 77013 76956 76841 76607 76506 76446 76353 76214 76104 | 77113~77013~76956~76841~76607~76506~76446~76353~76214~76104 | 76991 76994 77057 77058 77060 77061 | 76991~76994~77057~77058~77060~77061 | 3235 | 3235 3233 3232 3228 3226 3220 3218 3217 3212 3210 | 3235 3235 3235 3235 3235 3235 | 2019-06-11 | 16:47 | Albert Zondervan |  | 0.93 mg | 0.13 | 0.93 | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 | 0.99769 | 0.00158 | 76991 | ST | XCAMS | EA | 0% |  | TW3235 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result agrees with consensus value * 1x IAEA-C2 water: result agrees with consensus value blanks: * groups identified to blank-correct: carbonates: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-06-26 | Water | Complete | JOB-212575 |  |  | 67705 | Normal | 14047/12 |  |  | 2019-05-29 | 3995 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 | 0.99965 | 0.00158 | 76994 | EA | [25 empty] | 2019-05-29 | 2532 | Gas Splitting | Completed
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 | 1.0003 | 0.00159 | 77057 | EA | [25 empty] | 2019-05-29 | 4385 | Choose Gas Splitting | Completed
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 | 0.99922 | 0.00161 | 77058 | EA | [25 empty] | 2019-05-30 | 17124 | Graphite | Completed
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 1.0017 | 0.0016 | 77060 | EA | [25 empty] | 2019-06-14 | 18465 | AMS Submission Results | Completed
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 | 1.0015 | 0.00159 | 77061 | EA | [25 empty] | 2019-06-17 | 22068 | Close Job | Completed
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 |  |  |  | ST
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 |  |  |  | EA
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 |  |  |  | ST
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 |  |  |  | EA
[7 empty] | 111.7 | 88.3948 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45252 | 45251 | 45251 | 280 | 280 | 280 |  |  | Julia Collins | -996.42 | 0.12 ‰ | 2019-06-10 | 2018-12-12 | 11 | 180 |  |  |  |  | 4.39‰ | 4.39‰ | 2.00‰ | 2.00‰ | 4.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.42 ‰ | -996.42 ‰ | 0.12 | 0.12 | -996.38 | 0.13 ‰ | -996.38 | 0.13 ‰ |  |  | 2019-06-14 | Jocelyn Turnbull | 9:20 | 0.0035769 | 0.0035769 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00358 | 0.00012 | 18466 | 2019-06-11 | 16:34 | Albert Zondervan | 212582 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 27 | 5 | 5 | ... | 0.00344 | 0.00012 |  |  |  |  | 0.99767 | 0.99767 | Yes | Yes | [5 empty] | 0.00344 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00358 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00344 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-07 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 77113 | 2000000 |  |  | 76991 76994 77057 77058 77060 77061 | 76991~76994~77057~77058~77060~77061 | 3235 |  | 3235 3235 3235 3235 3235 3235 | 2019-06-11 | 16:47 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99769 | 0.00158 | 76991 |  | XCAMS | EA | 0% |  | TW3235 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result agrees with consensus value * 1x IAEA-C2 water: result agrees with consensus value blanks: * groups identified to blank-correct: carbonates: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-06-26 | Water |  | JOB-212582 |  |  | 67704 | Normal | 14047/11 |  |  | 2019-05-29 | 3994 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00158 | 76994 | [26 empty] | 2019-05-29 | 2531 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00159 | 77057 | [26 empty] | 2019-05-29 | 4383 | Choose Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00161 | 77058 | [26 empty] | 2019-05-30 | 17123 | Graphite | Completed
[168 empty] | 1.0017 | 0.0016 | 77060 | [26 empty] | 2019-06-14 | 18466 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00159 | 77061 | [26 empty] | 2019-06-17 | 22074 | Close Job | Completed
[7 empty] | 105.3 | 88.3287 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47563 | 47562 | 47562 | 280 | 280 | 280 |  |  | Julia Collins | -997.32 | 0.09 ‰ | 2019-06-13 | 2018-12-15 | 8 | 180 |  |  |  |  | 1.22‰ | 1.22‰ | 2.00‰ | 2.00‰ | 1.22 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -997.32 ‰ | -997.32 ‰ | 0.09 | 0.09 | -997.28 | 0.09 ‰ | -997.28 | 0.09 ‰ |  |  | 2019-06-14 | Jocelyn Turnbull | 14:56 | 0.0026827 | 0.0026827 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.4 |  |  | 0.00268 | 9e-05 | 18491 | 2019-06-14 | 12:32 | Albert Zondervan | 212569 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.27 | 0.01% | 15 | 5 | 6 | ... | 0.00258 | 9e-05 |  |  |  |  | 0.99378 | 0.99378 | Yes | Yes | [5 empty] | 0.00258 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00268 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00258 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-11 | Julia Collins | 13:50 |  | Completed |  |  |  |  | 77136 | 2000000 |  |  | 77062 77064 77115 77118 77121 77126 | 77062~77064~77115~77118~77121~77126 | 3236 |  | 3236 3236 3236 3236 3236 3236 | 2019-06-14 | 14:40 | Albert Zondervan |  | 0.86 mg |  | 0.86 | [5 empty] | 1.0001 | 0.00161 | 77062 |  | XCAMS | EA | 0% |  | TW3236 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * excess run-to-run variability: factor 1.3 for cat#12,18 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result agrees with consensus value * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result is within 2-sigma equal to consensus value blanks: * groups identified to blank-correct: carbonates: 22x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks organics (charcoal) via AAA: 3x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 3x unknowns + 0x secondaries + 1x blanks 250 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2019-06-26 | Carbonate | Complete | JOB-212569 |  |  | 67743 |  | 14047/1 |  |  | 2019-06-04 | 4007 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99863 | 0.00164 | 77064 | [26 empty] | 2019-06-04 | 2535 | Gas Splitting | Completed
[168 empty] | 0.99932 | 0.00165 | 77115 | [26 empty] | 2019-06-04 | 4389 | Choose Gas Splitting | Completed
[168 empty] | 1.0024 | 0.00164 | 77118 | [26 empty] | 2019-06-05 | 17146 | Graphite | Completed
[168 empty] | 1.0005 | 0.00163 | 77121 | [26 empty] | 2019-06-07 | 6620 | C13 Measurement | Completed
[168 empty] | 0.99906 | 0.00164 | 77126 | [26 empty] | 2019-06-14 | 18491 | AMS Submission Results | Completed
[197 empty] | 2019-06-17 | 22064 | Close Job | Completed
[7 empty] | 95.2 | 88.5289 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43275 | 43274 | 43274 | 219 | 219 | 219 |  |  | Julia Collins | -995.42 | 0.12 ‰ | 2019-06-20 | 2018-12-22 | 9 | 180 |  |  |  |  | 2.79‰ | 2.79‰ | 2.00‰ | 2.00‰ | 2.79 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -995.42 ‰ | -995.42 ‰ | 0.12 | 0.12 | -995.37 | 0.13 ‰ | -995.37 | 0.13 ‰ |  |  | 2019-06-21 | Jocelyn Turnbull | 16:34 | 0.0045751 | 0.0045751 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00458 | 0.00012 | 18544 | 2019-06-21 | 14:49 | Albert Zondervan | 212710 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 8 | 5 | 5 | ... | 0.0044 | 0.00012 |  |  |  |  | 0.99759 | 0.99759 | Yes | Yes | [5 empty] | 0.0044 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00078 | 0.00078 | 0.00458 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0044 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-18 | Julia Collins | 13:15 |  | Completed |  |  |  |  | 77166 | 2000000 |  |  | 77168 77173 77174 77175 77218 77219 | 77168~77173~77174~77175~77218~77219 | 3237 |  | 3237 3237 3237 3237 3237 3237 | 2019-06-21 | 15:59 | Albert Zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 0.99835 | 0.00186 | 77168 |  | XCAMS | EA | 0% |  | TW3237 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * excess run-to-run variability: none * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result is within 2-sigma equal to consensus value * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: carbonates: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA: 6x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 4x unknowns + 0x secondaries + 1x blanks 300 days: capturing 5x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-07-19 | Carbonate |  | JOB-212710 |  |  | 67783 |  | 14047/1 |  |  | 2019-06-10 | 4031 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00186 | 77173 | [26 empty] | 2019-06-10 | 2539 | Gas Splitting | Completed
[168 empty] | 1.0026 | 0.00194 | 77174 | [26 empty] | 2019-06-10 | 4396 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00191 | 77175 | [26 empty] | 2019-06-11 | 17176 | Graphite | Completed
[168 empty] | 1.0003 | 0.00196 | 77218 | [26 empty] | 2019-06-21 | 18544 | AMS Submission Results | Completed
[168 empty] | 0.99778 | 0.00195 | 77219 | [26 empty] | 2019-06-24 | 22180 | Close Job | Completed
[7 empty] | 94.3 | 87.9677 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37358 | 37358 | 37358 | 157 | 157 | 157 |  |  | jenny dahl | -990.44 | 0.19 ‰ | 2019-07-09 | 2019-01-10 | 11 | 180 |  |  |  |  | 5.11‰ | 5.11‰ | 2.00‰ | 2.00‰ | 5.11 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -990.44 ‰ | -990.44 ‰ | 0.19 | 0.19 | -990.33 | 0.19 ‰ | -990.33 | 0.19 ‰ |  |  | 2019-07-15 | Cathy Ginnane | 14:20 | 0.0095557 | 0.0095557 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.6 |  |  | 0.00956 | 0.00019 | 18639 | 2019-07-10 | 16:45 | albert zondervan | 212714 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.96 | 0.02% | 10 | 5 | 5 | ... | 0.00919 | 0.00018 |  |  |  |  | 0.99356 | 0.99356 | Yes | Yes |  | AMS d13C 2.5 per mil heavy |  |  |  | 0.00919 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00956 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00919 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-05 | jenny dahl | 13:51 |  | Completed |  |  |  |  | 77273 | 2000000 |  |  | 77346 77347 77358 77359 77360 77361 | 77346~77347~77358~77359~77360~77361 | 3240 |  | 3240 3240 3240 3240 3240 3240 | 2019-07-10 | 17:02 | albert zondervan |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0017 | 0.00169 | 77346 |  | XCAMS | ST | 0% |  | TW3240 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * excess run-to-run variability: zero * system error: zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 carbonate, waters: result agrees with consensus value * 1x IAEA-C6 sucrose (changed to tuning not std as unknown; CG), lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 6x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-07-19 | Water |  | JOB-212714 |  |  | 67911 | Normal | 14047/11 |  |  | 2019-06-27 | 4069 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99823 | 0.00177 | 77347 | [26 empty] | 2019-06-27 | 2552 | Gas Splitting | Completed
[168 empty] | 1.0008 | 0.00176 | 77358 | [26 empty] | 2019-06-27 | 4414 | Choose Gas Splitting | Completed
[168 empty] | 0.9997 | 0.0018 | 77359 | [26 empty] | 2019-06-28 | 17283 | Graphite | Completed
[168 empty] | 0.99837 | 0.00176 | 77360 | [26 empty] | 2019-07-15 | 18639 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00174 | 77361 | [26 empty] | 2019-07-16 | 22325 | Close Job | Completed
[5 empty] | 4.9996e-05 | 0.009065 | 91.3 | 89.0986 | 0.0699 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7128 | 7127 | 7127 | 64 | 63 | 66 | 7064.3 | 7191.8 | jenny dahl | -588.24 | 3.27 ‰ | 2019-07-09 | 2019-01-10 | 11 | 180 |  |  |  |  | -7.65‰ | -7.65‰ | 2.00‰ | 2.00‰ | -7.65 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -588.24 ‰ | -588.24 ‰ | 3.27 | 3.4 | -583.16 | 3.31 ‰ | -583.16 | 3.31 ‰ |  |  | 2019-07-15 | Cathy Ginnane | 14:20 | 0.41176 | 0.41176 | 0.00327 | 0.0034 | 0.00322 | 0.00322 | 1.9 |  |  | 0.41833 | 0.00132 | 18640 | 2019-07-10 | 16:45 | albert zondervan | 212715 | 0.01046 | 0.0105 | 0.00471 | 0.0047 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.18 | 0.33% | 11 | 5 | 5 | ... | 0.40232 | 0.00116 |  |  |  |  | 0.99356 | 0.99356 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 66.2622747167770445. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01046 | 0.00871 | 0.0046 | 0.396 | 0.40‰ | 0.0031 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.41176 | 0 | 0.00327 | 0 | 0 | 0.01046 | 0 | 0.00471 | 0 | 0.396 | 0 | 0.0031 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-05 | jenny dahl | 13:51 |  | Completed |  |  |  |  | 77274 | 2000000 | 77273 77113 77013 76956 76841 76607 76506 76446 76353 76214 76104 | 77273~77113~77013~76956~76841~76607~76506~76446~76353~76214~76104 | 77346 77347 77358 77359 77360 77361 | 77346~77347~77358~77359~77360~77361 | 3240 | 3240 3235 3233 3232 3228 3226 3220 3218 3217 3212 3210 | 3240 3240 3240 3240 3240 3240 | 2019-07-10 | 17:02 | albert zondervan |  | 1.08 mg | 0.13 | 1.08 | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 | 1.0017 | 0.00169 | 77346 | ST | XCAMS | ST | 0% |  | TW3240 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * excess run-to-run variability: zero * system error: zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 carbonate, waters: result agrees with consensus value * 1x IAEA-C6 sucrose (changed to tuning not std as unknown; CG), lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 6x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-07-19 | Water |  | JOB-212715 |  |  | 67914 | Normal | 14047/12 |  |  | 2019-06-27 | 4070 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 | 0.99823 | 0.00177 | 77347 | EA | [25 empty] | 2019-06-27 | 2553 | Gas Splitting | Completed
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 | 1.0008 | 0.00176 | 77358 | EA | [25 empty] | 2019-06-27 | 4416 | Choose Gas Splitting | Completed
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 | 0.9997 | 0.0018 | 77359 | EA | [25 empty] | 2019-06-28 | 17284 | Graphite | Completed
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 | 0.99837 | 0.00176 | 77360 | EA | [25 empty] | 2019-07-15 | 18640 | AMS Submission Results | Completed
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 | 1.001 | 0.00174 | 77361 | EA | [25 empty] | 2019-07-16 | 22326 | Close Job | Completed
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 |  |  |  | ST
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 |  |  |  | EA
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01 |  |  |  | ST
[163 empty] | 2019-01-29 | 0.0081 | 76214 | 3212 | 0.89 |  |  |  | EA
[163 empty] | 2019-01-16 | 0.0055 | 76104 | 3210 | 1.11 |  |  |  | ST
[7 empty] | 92.9 | 88.7338 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44452 | 44452 | 44452 | 275 | 274 | 274 | 44182.2 | 44732.1 | Julia Collins | -996.05 | 0.14 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.47‰ | 2.47‰ | 2.00‰ | 2.00‰ | 2.47 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.05 ‰ | -996.05 ‰ | 0.14 | 0.14 | -996 | 0.14 ‰ | -996 | 0.14 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0039512 | 0.0039512 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.6 |  |  | 0.00395 | 0.00014 | 18770 | 2019-07-23 | 15:56 | albert zondervan | 212849 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 30 | 5 | 5 | ... | 0.0038 | 0.00013 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.0038 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00395 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0038 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77319 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water | Complete | JOB-212849 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 1 new material prepared in job 212871 | 68095 | Normal | 14047/11 |  |  | 2019-07-01 | 4075 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-01 | 2554 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-01 | 4418 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-02 | 17329 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18770 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22490 | Close Job | Completed
[197 empty] | 2019-09-19 | 9374 | Invoice | Completed
[7 empty] | 96.6 | 88.7946 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 45159 | 45159 | 45159 | 277 | 276 | 276 | 44886.8 | 45441 | Julia Collins | -996.38 | 0.12 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 1.79‰ | 1.79‰ | 2.00‰ | 2.00‰ | 1.79 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.38 ‰ | -996.38 ‰ | 0.12 | 0.12 | -996.34 | 0.13 ‰ | -996.34 | 0.13 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0036185 | 0.0036185 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00362 | 0.00012 | 18771 | 2019-07-23 | 15:56 | albert zondervan | 212851 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 29 | 5 | 5 | ... | 0.00348 | 0.00012 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00348 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00362 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00348 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77354 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212851 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 3 new material prepared in job 212871 | 68096 | Normal | 14047/11 |  |  | 2019-07-02 | 4086 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 2567 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-02 | 4432 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-04 | 17364 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18771 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22491 | Close Job | Completed
[197 empty] | 2019-09-19 | 9376 | Invoice | Completed
[7 empty] | 95.9 | 88.7357 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 45045 | 45044 | 45044 | 296 | 295 | 295 | 44754 | 45345.9 | Julia Collins | -996.33 | 0.14 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.33 ‰ | -996.33 ‰ | 0.14 | 0.14 | -996.28 | 0.14 ‰ | -996.28 | 0.14 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0036705 | 0.0036705 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00367 | 0.00014 | 18772 | 2019-07-23 | 15:56 | albert zondervan | 212852 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 31 | 5 | 3 | ... | 0.00353 | 0.00013 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00353 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00367 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00353 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77320 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.24 mg |  | 1.24 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212852 | Water blank test JCT 9/19 | 2/7/19 This sample was too large on graphite. Gas was transferred back to bottle and split as a C13 to drop the size (CG). Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 1 new material prepared in job 212871 | 68097 | Normal | 14047/11 |  |  | 2019-07-01 | 4076 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 17330 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-08-07 | 18772 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-08-08 | 22492 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-09-19 | 9377 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 98.2 | 88.5721 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41028 | 41027 | 41027 | 193 | 193 | 193 | 40837.1 | 41223.6 | Julia Collins | -993.95 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 4.30‰ | 4.30‰ | 2.00‰ | 2.00‰ | 4.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -993.95 ‰ | -993.95 ‰ | 0.15 | 0.15 | -993.87 | 0.15 ‰ | -993.87 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0060516 | 0.0060516 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00605 | 0.00015 | 18773 | 2019-07-23 | 15:56 | albert zondervan | 212853 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 36 | 5 | 5 | ... | 0.00582 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00582 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00605 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00582 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77355 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212853 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 2 new material prepared in job 212871 | 68098 | Normal | 14047/11 |  |  | 2019-07-02 | 4081 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 2559 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-02 | 4424 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-04 | 17365 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18773 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22493 | Close Job | Completed
[197 empty] | 2019-09-19 | 9378 | Invoice | Completed
[7 empty] | 94.7 | 88.8006 | 0.0013 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 39252 | 39252 | 39252 | 177 | 177 | 177 | 39077 | 39431.1 | Julia Collins | -992.45 | 0.17 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 1.72‰ | 1.72‰ | 2.00‰ | 2.00‰ | 1.72 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -992.45 ‰ | -992.45 ‰ | 0.17 | 0.17 | -992.36 | 0.17 ‰ | -992.36 | 0.17 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0075489 | 0.0075489 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.6 |  |  | 0.00755 | 0.00017 | 18774 | 2019-07-23 | 15:56 | albert zondervan | 212854 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.75 | 0.02% | 32 | 5 | 3 | ... | 0.00726 | 0.00016 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00726 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00755 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00726 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77362 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.24 mg |  | 1.24 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water | Not Complete | JOB-212854 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 3 new material prepared in job 212871 | 68099 | Normal | 14047/11 |  |  | 2019-07-03 | 4088 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-05 | 17372 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-08-07 | 18774 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-08-08 | 22494 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-09-19 | 9379 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 96.7 | 88.6529 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48110 | 48110 | 48110 | 300 | 299 | 299 |  |  | Julia Collins | -997.49 | 0.09 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 3.39‰ | 3.39‰ | 2.00‰ | 2.00‰ | 3.39 ‰ | 2.00 ‰ | AMS | 3.07‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -997.49 ‰ | -997.49 ‰ | 0.09 | 0.09 | -997.46 | 0.09 ‰ | -997.46 | 0.09 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0025059 | 0.0025059 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00251 | 9e-05 | 18775 | 2019-07-23 | 15:56 | albert zondervan | 212855 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.25 | 0.01% | 33 | 5 | 6 | ... | 0.00241 | 9e-05 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00241 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00251 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00241 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77321 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 3.1 | RRL | [5 empty] | Background Inorganic |  | 2019-08-14 | Water |  | JOB-212855 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel as regular standard run in set 1 new material prepared in job 212871 | 68100 | Normal | 14047/11 |  |  | 2019-07-01 | 4077 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-01 | 2555 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-01 | 4419 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-02 | 17331 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-07-09 | 6694 | C13 Measurement | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-07 | 18775 | AMS Submission Results | Completed
[197 empty] | 2019-08-08 | 22495 | Close Job | Completed
[197 empty] | 2019-09-19 | 9380 | Invoice | Completed
[7 empty] | 91.8 | 88.6935 | 0.0005 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 46877 | 46876 | 46876 | 286 | 285 | 285 | 46596.1 | 47168 | Julia Collins | -997.08 | 0.10 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -997.08 ‰ | -997.08 ‰ | 0.1 | 0.1 | -997.04 | 0.11 ‰ | -997.04 | 0.11 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0029218 | 0.0029218 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00292 | 0.0001 | 18776 | 2019-07-23 | 15:56 | albert zondervan | 212856 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.29 | 0.01% | 34 | 5 | 5 | ... | 0.00281 | 0.0001 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00281 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00292 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00281 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77322 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.21 mg |  | 1.21 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212856 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel as regular standard run in set 1 new material prepared in job 212871 | 68101 | Normal | 14047/11 |  |  | 2019-07-01 | 4078 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-01 | 2556 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-01 | 4420 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-02 | 17332 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18776 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22496 | Close Job | Completed
[197 empty] | 2019-09-19 | 9381 | Invoice | Completed
[7 empty] | 99.6 | 88.5219 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41734 | 41734 | 41734 | 211 | 210 | 210 | 41526.2 | 41948.3 | Julia Collins | -994.46 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 4.87‰ | 4.87‰ | 2.00‰ | 2.00‰ | 4.87 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -994.46 ‰ | -994.46 ‰ | 0.15 | 0.15 | -994.39 | 0.15 ‰ | -994.39 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0055421 | 0.0055421 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00554 | 0.00015 | 18777 | 2019-07-23 | 15:56 | albert zondervan | 212857 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 37 | 5 | 6 | ... | 0.00533 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00533 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00554 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00533 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77356 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212857 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel as regular standard run in set 2 new material prepared in job 212871 | 68102 | Normal | 14047/11 |  |  | 2019-07-02 | 4082 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 2563 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-02 | 4428 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-04 | 17366 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-07-09 | 6704 | C13 Measurement | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-07 | 18777 | AMS Submission Results | Completed
[197 empty] | 2019-08-08 | 22497 | Close Job | Completed
[197 empty] | 2019-09-19 | 9382 | Invoice | Completed
[7 empty] | 100.2 | 88.7416 | 0.0026 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33332 | 33332 | 33332 | 132 | 132 | 132 | 33200.9 | 33465.7 | Julia Collins | -984.23 | 0.26 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 2.39‰ | 2.39‰ | 2.00‰ | 2.00‰ | 2.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -984.23 ‰ | -984.23 ‰ | 0.26 | 0.26 | -984.03 | 0.26 ‰ | -984.03 | 0.26 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.015774 | 0.015774 | 0.00026 | 0.00026 | 0.00026 | 0.00026 | 1.8 |  |  | 0.01577 | 0.00026 | 18778 | 2019-07-23 | 15:56 | albert zondervan | 212859 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.58 | 0.03% | 19 | 5 | 5 | ... | 0.01517 | 0.00025 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.01517 | 0.02‰ | 0.00025 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.01577 | 0 | 0.00026 | 0 | 0 |  |  | 0 | 0 | 0.01517 | 0 | 0.00025 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77357 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212859 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run in set 1 new material prepared in job 212871 | 68103 | Normal | 14047/11 |  |  | 2019-07-02 | 4083 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 2561 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-02 | 4426 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-04 | 17367 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18778 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22498 | Close Job | Completed
[197 empty] | 2019-09-19 | 9384 | Invoice | Completed
[7 empty] | 100.5 | 88.5317 | 0.0032 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31847 | 31847 | 31847 | 119 | 118 | 118 | 31729.4 | 31967.1 | Julia Collins | -981.02 | 0.28 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -981.02 ‰ | -981.02 ‰ | 0.28 | 0.28 | -980.79 | 0.28 ‰ | -980.79 | 0.28 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.018976 | 0.018976 | 0.00028 | 0.00028 | 0.00028 | 0.00028 | 1.5 |  |  | 0.01898 | 0.00028 | 18779 | 2019-07-23 | 15:56 | albert zondervan | 212860 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.9 | 0.03% | 20 | 5 | 3 | ... | 0.01825 | 0.00027 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.01825 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.01898 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01825 | 0 | 0.00027 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77363 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212860 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run in set 2 new material prepared in job 212871 | 68104 | Normal | 14047/11 |  |  | 2019-07-03 | 4084 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-05 | 17373 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-08-07 | 18779 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-08-08 | 22499 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-09-19 | 9385 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 100.8 | 88.6505 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42329 | 42328 | 42328 | 227 | 227 | 227 | 42104.6 | 42559.1 | Julia Collins | -994.85 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 3.42‰ | 3.42‰ | 2.00‰ | 2.00‰ | 3.42 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.85 ‰ | -994.85 ‰ | 0.15 | 0.15 | -994.79 | 0.15 ‰ | -994.79 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.005147 | 0.005147 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2 |  |  | 0.00515 | 0.00015 | 18780 | 2019-07-23 | 15:56 | albert zondervan | 212861 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 25 | 5 | 5 | ... | 0.00495 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00495 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00515 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00495 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77364 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212861 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run set 3 new material prepared in job 212871 | 68105 | Normal | 14047/11 |  |  | 2019-07-03 | 4091 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-03 | 2570 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-03 | 4443 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-05 | 17374 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18780 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22500 | Close Job | Completed
[197 empty] | 2019-09-19 | 9386 | Invoice | Completed
[7 empty] | 98.5 | 88.6483 | 0.0093 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23267 | 23266 | 23266 | 67 | 66 | 66 | 23200.2 | 23333.3 | Julia Collins | -944.78 | 0.46 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 3.44‰ | 3.44‰ | 2.00‰ | 2.00‰ | 3.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -944.78 ‰ | -944.78 ‰ | 0.46 | 0.46 | -944.09 | 0.46 ‰ | -944.09 | 0.46 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.055223 | 0.055223 | 0.00046 | 0.00046 | 0.00046 | 0.00046 | 1.7 |  |  | 0.05522 | 0.00046 | 18781 | 2019-07-23 | 15:56 | albert zondervan | 212862 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.52 | 0.05% | 22 | 5 | 3 | ... | 0.05311 | 0.00044 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.05311 | 0.05‰ | 0.00044 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05522 | 0 | 0.00046 | 0 | 0 |  |  | 0 | 0 | 0.05311 | 0 | 0.00044 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77323 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212862 | Water blank test JCT 9/19 | 2/7/19 This sample was too large on graphite. Gas was transferred back to bottle and split as a C13 to drop the size (CG). Made to test water blanks Jeremy to run large vessel no flushing run in set 1 new material prepared in job 212871 | 68106 | Normal | 14047/11 |  |  | 2019-07-01 | 4079 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 17333 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-08-07 | 18781 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-08-08 | 22501 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-09-19 | 9387 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 102 | 88.7564 | 0.0087 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23762 | 23762 | 23762 | 66 | 65 | 65 | 23696.8 | 23828.7 | Julia Collins | -948.08 | 0.43 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -948.08 ‰ | -948.08 ‰ | 0.43 | 0.43 | -947.44 | 0.43 ‰ | -947.44 | 0.43 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.051917 | 0.051917 | 0.00043 | 0.00043 | 0.00043 | 0.00043 | 1.6 |  |  | 0.05192 | 0.00043 | 18782 | 2019-07-23 | 15:56 | albert zondervan | 212863 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.19 | 0.04% | 23 | 5 | 5 | ... | 0.04993 | 0.00041 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.04993 | 0.05‰ | 0.00041 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05192 | 0 | 0.00043 | 0 | 0 |  |  | 0 | 0 | 0.04993 | 0 | 0.00041 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77365 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212863 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel no flushing run in set 2 new material prepared in job 212871 | 68107 | Normal | 14047/11 |  |  | 2019-07-02 | 4085 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-02 | 2564 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-02 | 4429 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-05 | 17375 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18782 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22502 | Close Job | Completed
[197 empty] | 2019-09-19 | 9388 | Invoice | Completed
[7 empty] | 100 | 88.7555 | 0.0089 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23613 | 23612 | 23612 | 65 | 64 | 64 | 23548.2 | 23677.7 | Julia Collins | -947.11 | 0.43 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.23‰ | 2.23‰ | 2.00‰ | 2.00‰ | 2.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -947.11 ‰ | -947.11 ‰ | 0.43 | 0.43 | -946.45 | 0.43 ‰ | -946.45 | 0.43 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.052894 | 0.052894 | 0.00043 | 0.00043 | 0.00043 | 0.00043 | 1.7 |  |  | 0.05289 | 0.00043 | 18783 | 2019-07-23 | 15:56 | albert zondervan | 212864 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.29 | 0.04% | 24 | 5 | 5 | ... | 0.05087 | 0.00041 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.05087 | 0.05‰ | 0.00041 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05289 | 0 | 0.00043 | 0 | 0 |  |  | 0 | 0 | 0.05087 | 0 | 0.00041 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77366 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | JOB-212864 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel no flushing run in set 3 new material prepared in job 212871 | 68108 | Normal | 14047/11 |  |  | 2019-07-03 | 4090 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-03 | 2571 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-03 | 4444 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-05 | 17376 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-08-07 | 18783 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-08 | 22503 | Close Job | Completed
[197 empty] | 2019-09-19 | 9389 | Invoice | Completed
[7 empty] | 95.7 | 88.7557 | 0.0012 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 39649 | 39648 | 39648 | 186 | 186 | 186 | 39465.1 | 39837.2 | Julia Collins | -992.82 | 0.17 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.23‰ | 2.23‰ | 2.00‰ | 2.00‰ | 2.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -992.82 ‰ | -992.82 ‰ | 0.17 | 0.17 | -992.73 | 0.17 ‰ | -992.73 | 0.17 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.007185 | 0.007185 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.00718 | 0.00017 | 18794 | 2019-07-23 | 15:56 | albert zondervan | 212895 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.72 | 0.02% | 26 | 5 | 6 | ... | 0.00691 | 0.00016 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00691 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00718 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00691 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77373 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.17 mg |  | 1.17 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-28 | Water | Complete | JOB-212895 | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 2 new material prepared in job 212871 Don't trust result due to vessel # being wrong looks to be large vessel from extract time and vessel number | 68109 | Normal | 14047/11 |  |  | 2019-07-02 | 4080 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [26 empty] | 2019-07-03 | 4087 | Water CO2 Evolution | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [26 empty] | 2019-07-03 | 2573 | Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [26 empty] | 2019-07-03 | 4445 | Choose Gas Splitting | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [26 empty] | 2019-07-05 | 17383 | Graphite | Completed
[168 empty] | 1 | 0.00191 | 77444 | [26 empty] | 2019-08-07 | 18794 | AMS Submission Results | Completed
[197 empty] | 2019-08-08 | 22504 | Close Job | Completed
[197 empty] | 2019-09-19 | 9392 | Invoice | Completed
[7 empty] | 91.4 | 88.3784 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41352 | 41351 | 41351 | 172 | 172 | 172 |  |  | Julia Collins | -994.19 | 0.12 ‰ | 2019-08-03 | 2019-02-04 | 10 | 180 |  |  |  |  | 2.67‰ | 2.67‰ | 2.00‰ | 2.00‰ | 2.67 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -994.19 ‰ | -994.19 ‰ | 0.12 | 0.12 | -994.12 | 0.13 ‰ | -994.12 | 0.13 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.0058125 | 0.0058125 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00581 | 0.00012 | 18804 | 2019-08-05 | 11:38 | Jeremy thompson | 212768 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 1 | 5 | 4 | ... | 0.00559 | 0.00012 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes | [5 empty] | 0.00559 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.00581 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00559 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:30 |  | Completed |  |  |  |  | 77484 | 2000000 |  |  | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 |  | 3245 3245 3245 3245 3245 3245 | 2019-08-05 | 13:33 | Jeremy thompson |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99964 | 0.00213 | 77527 |  | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2019-08-02 | Carbonate |  | JOB-212768 |  |  | 68045 |  | 14047/1 |  |  | 2019-07-22 | 4100 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9995 | 0.0021 | 77528 | [26 empty] | 2019-07-24 | 17494 | Graphite | Completed
[168 empty] | 0.99635 | 0.0021 | 77529 | [26 empty] | 2019-07-26 | 6712 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.0021 | 77530 | [26 empty] | 2019-08-07 | 18804 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00208 | 77557 | [26 empty] | 2019-08-08 | 22481 | Close Job | Completed
[168 empty] | 1.0021 | 0.00211 | 77558
[7 empty] | 93.2 | 88.3795 | 0.0023 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34329 | 34328 | 34328 | 107 | 107 | 107 |  |  | Julia Collins | -986.07 | 0.19 ‰ | 2019-08-03 | 2019-02-04 | 2 | 180 |  |  |  |  | 2.65‰ | 2.65‰ | 2.00‰ | 2.00‰ | 2.65 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -986.07 ‰ | -986.07 ‰ | 0.19 | 0.19 | -985.89 | 0.19 ‰ | -985.89 | 0.19 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.013933 | 0.013933 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2 |  |  | 0.01393 | 0.00019 | 18809 | 2019-08-05 | 11:38 | Jeremy thompson | 212936 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.39 | 0.02% | 36 | 5 | 4 | ... | 0.0134 | 0.00018 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes |  | bad result. Jeremy to look into |  |  |  | 0.0134 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.01393 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0134 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 77552 | 2000000 |  |  | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 |  | 3245 3245 3245 3245 3245 3245 | 2019-08-05 | 13:33 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 0.99964 | 0.00213 | 77527 |  | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2019-09-09 | Water | Complete | JOB-212936 | Water blank test JCT 9/19 | run as normal with new ss carbosorb trap use for blank corr of TW3244 waters | 68055 | Normal | 14047/11 |  |  | 2019-07-31 | 4116 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9995 | 0.0021 | 77528 | [26 empty] | 2019-08-01 | 17562 | Graphite | Completed
[168 empty] | 0.99635 | 0.0021 | 77529 | [26 empty] | 2019-08-05 | 6723 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.0021 | 77530 | [26 empty] | 2019-08-07 | 18809 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00208 | 77557 | [26 empty] | 2019-08-08 | 22505 | Close Job | Completed
[168 empty] | 1.0021 | 0.00211 | 77558 | [26 empty] | 2019-09-19 | 9393 | Invoice | Completed
[5 empty] | 5.3499e-05 | 0.010652 | 88.3 | 89.4784 | 0.0687 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7221 | 7220 | 7220 | 71 | 71 | 73 | 7149.7 | 7292.3 | Julia Collins | -592.97 | 3.61 ‰ | 2019-08-03 | 2019-02-04 | 2 | 180 |  |  |  |  | -9.66‰ | -9.66‰ | 2.00‰ | 2.00‰ | -9.66 ‰ | 2.00 ‰ | AMS | -7.86‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -592.97 ‰ | -592.97 ‰ | 3.61 | 3.73 | -587.95 | 3.66 ‰ | -587.95 | 3.66 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.40703 | 0.40703 | 0.00361 | 0.00373 | 0.00357 | 0.00357 | 2.1 |  |  | 0.4145 | 0.0012 | 18810 | 2019-08-05 | 11:38 | Jeremy thompson | 212937 | 0.01182 | 0.0118 | 0.00532 | 0.0053 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 40.7 | 0.36% | 8 | 5 | 4 | ... | 0.39864 | 0.00117 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 73.5502337602371891. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01182 | 0.01023 | 0.00438 | 0.39145 | 0.39‰ | 0.00348 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.40703 | 0 | 0.00361 | 0 | 0 | 0.01182 | 0 | 0.00532 | 0 | 0.39145 | 0 | 0.00348 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 77551 | 2000000 | 77552 77553 77273 77113 77013 76956 76841 76607 76506 76446 76353 | 77552~77553~77273~77113~77013~76956~76841~76607~76506~76446~76353 | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 | 3245 3245 3240 3235 3233 3232 3228 3226 3220 3218 3217 | 3245 3245 3245 3245 3245 3245 | 2019-08-07 | 11:18 | Cathy Ginnane |  | 1.05 mg | 0.13 | 1.05 | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 0.99964 | 0.00213 | 77527 | EA | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | -7.9 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-09-09 | Water |  | JOB-212937 |  |  | 68059 | Normal | 14047/12 |  |  | 2019-07-30 | 4113 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 | 0.9995 | 0.0021 | 77528 | EA | [25 empty] | 2019-08-01 | 17561 | Graphite | Completed
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 | 0.99635 | 0.0021 | 77529 | EA | [25 empty] | 2019-08-05 | 6727 | C13 Measurement | Completed
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 | 1.0019 | 0.0021 | 77530 | EA | [25 empty] | 2019-08-07 | 18810 | AMS Submission Results | Completed
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 | 1.0004 | 0.00208 | 77557 | ST | [25 empty] | 2019-08-08 | 22506 | Close Job | Completed
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 | 1.0021 | 0.00211 | 77558 | EA
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | ST
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 |  |  |  | EA
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 |  |  |  | ST
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 |  |  |  | EA
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01
[7 empty] | 93.3 | 88.4422 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33359 | 33358 | 33358 | 100 | 100 | 100 |  |  | Julia Collins | -984.28 | 0.20 ‰ | 2019-08-03 | 2019-02-04 | 2 | 180 |  |  |  |  | 1.94‰ | 1.94‰ | 2.00‰ | 2.00‰ | 1.94 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -984.28 ‰ | -984.28 ‰ | 0.2 | 0.2 | -984.08 | 0.20 ‰ | -984.08 | 0.20 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.015722 | 0.015722 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2 |  |  | 0.01572 | 0.0002 | 18827 | 2019-08-05 | 11:38 | Jeremy thompson | 212991 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.57 | 0.02% | 22 | 5 | 4 | ... | 0.01512 | 0.00019 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes |  | bad result. jeremy to look into |  |  |  | 0.01512 | 0.02‰ | 0.00019 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.01572 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01512 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 77553 | 2000000 |  |  | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 |  | 3245 3245 3245 3245 3245 3245 | 2019-08-05 | 13:33 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 0.99964 | 0.00213 | 77527 |  | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2019-09-17 | Water | Complete | JOB-212991 | Water blank test JCT 9/19 | run as normal with new ss carbosorb trap use for blank corr of TW3244 waters | 68056 | Normal | 14047/11 |  |  | 2019-07-31 | 4117 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9995 | 0.0021 | 77528 | [26 empty] | 2019-08-01 | 17563 | Graphite | Completed
[168 empty] | 0.99635 | 0.0021 | 77529 | [26 empty] | 2019-08-05 | 6726 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.0021 | 77530 | [26 empty] | 2019-08-07 | 18827 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00208 | 77557 | [26 empty] | 2019-08-08 | 22513 | Close Job | Completed
[168 empty] | 1.0021 | 0.00211 | 77558 | [26 empty] | 2019-09-19 | 9394 | Invoice | Completed
[7 empty] | 79.9 | 88.1653 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38503 | 38502 | 38502 | 161 | 161 | 161 |  |  | jenny dahl | -991.71 | 0.17 ‰ | 2019-08-16 | 2019-02-17 | 1 | 180 |  |  |  |  | 6.14‰ | 6.14‰ | 2.00‰ | 2.00‰ | 6.14 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.89443 | -991.71 ‰ | -991.71 ‰ | 0.17 | 0.17 | -991.61 | 0.17 ‰ | -991.61 | 0.17 ‰ |  |  | 2019-08-20 | Cathy Ginnane | 11:05 | 0.0082872 | 0.0082872 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.6 |  |  | 0.00829 | 0.00017 | 18954 | 2019-08-19 | 14:01 | albert zondervan | 213051 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.83 | 0.02% | 37 | 5 | 5 | ... | 0.00797 | 0.00016 |  |  |  |  | 0.99682 | 0.99682 | Yes | Yes |  | d13C is more than 3.5 per mil heavy. odd c13 result |  |  |  | 0.00797 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00829 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00797 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-16 | jenny dahl | 12:31 |  | Completed |  |  |  |  | 77657 | 2000000 |  |  | 77589 77590 77633 77634 77635 | 77589~77590~77633~77634~77635 | 3247 |  | 3247 3247 3247 3247 3247 | 2019-08-19 | 14:23 | albert zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 0.99992 | 0.00165 | 77589 |  | XCAMS | EA | 0% |  | TW3247 measurement: * one problem with XCAMS: indexer error * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA 1x (TP77591) has failed because its rts = 0.990 and has been rejected as an outlier A complete retune mid-way did not improve its rts value. secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value * 1x FIRI-L whalebone: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 3x secondaries + 1x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 5x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-09-23 | Water |  | JOB-213051 |  |  | 68123 | Normal | 14047/11 |  |  | 2019-08-13 | 4133 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00166 | 77590 | [26 empty] | 2019-08-13 | 2589 | Gas Splitting | Completed
[168 empty] | 0.99998 | 0.0017 | 77633 | [26 empty] | 2019-08-13 | 4465 | Choose Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00174 | 77634 | [26 empty] | 2019-08-15 | 17667 | Graphite | Completed
[168 empty] | 0.99766 | 0.00176 | 77635 | [26 empty] | 2019-08-20 | 18954 | AMS Submission Results | Completed
[197 empty] | 2019-08-22 | 22576 | Close Job | Completed
[5 empty] | 4.4735e-05 | 0.0086198 | 73 | 89.3805 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7138 | 7137 | 7137 | 65 | 65 | 67 | 7072.6 | 7203.1 | jenny dahl | -588.74 | 3.34 ‰ | 2019-08-16 | 2019-02-17 | 1 | 180 |  |  |  |  | -7.53‰ | -7.53‰ | 2.00‰ | 2.00‰ | -7.53 ‰ | 2.00 ‰ | AMS | -8.08‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.89443 | -588.74 ‰ | -588.74 ‰ | 3.34 | 3.47 | -583.66 | 3.38 ‰ | -583.66 | 3.38 ‰ |  |  | 2019-08-20 | Cathy Ginnane | 11:05 | 0.41126 | 0.41126 | 0.00334 | 0.00347 | 0.0033 | 0.0033 | 2.8 |  |  | 0.41805 | 0.00126 | 18955 | 2019-08-19 | 14:01 | albert zondervan | 213052 | 0.0108 | 0.0108 | 0.00486 | 0.0049 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.13 | 0.33% | 18 | 5 | 4 | ... | 0.40205 | 0.00109 |  |  |  |  | 0.99682 | 0.99682 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 67.685575584935395. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0108 | 0.00827 | 0.00484 | 0.39552 | 0.40‰ | 0.00317 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.41126 | 0 | 0.00334 | 0 | 0 | 0.0108 | 0 | 0.00486 | 0 | 0.39552 | 0 | 0.00317 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-16 | jenny dahl | 12:31 |  | Completed |  |  |  |  | 77658 | 2000000 | 77657 77552 77553 77321 77273 77113 77013 76956 76841 76607 76506 76446 76353 | 77657~77552~77553~77321~77273~77113~77013~76956~76841~76607~76506~76446~76353 | 77589 77590 77633 77634 77635 | 77589~77590~77633~77634~77635 | 3247 | 3247 3245 3245 3244 3240 3235 3233 3232 3228 3226 3220 3218 3217 | 3247 3247 3247 3247 3247 | 2019-08-19 | 14:23 | albert zondervan |  | 0.68 mg | 0.13 | 0.68 | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 0.99992 | 0.00165 | 77589 | EA | XCAMS | EA | 0% |  | TW3247 measurement: * one problem with XCAMS: indexer error * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA 1x (TP77591) has failed because its rts = 0.990 and has been rejected as an outlier A complete retune mid-way did not improve its rts value. secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value * 1x FIRI-L whalebone: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 3x secondaries + 1x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 5x blanks unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-09-23 | Water |  | JOB-213052 |  |  | 68126 | Normal | 14047/12 |  |  | 2019-08-13 | 4134 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 1.0017 | 0.00166 | 77590 | EA | [25 empty] | 2019-08-15 | 17668 | Graphite | Completed
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 | 0.99998 | 0.0017 | 77633 | EA | [25 empty] | 2019-08-19 | 6751 | C13 Measurement | Completed
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 | 1.0005 | 0.00174 | 77634 | EA | [25 empty] | 2019-08-20 | 18955 | AMS Submission Results | Completed
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 | 0.99766 | 0.00176 | 77635 | ST | [25 empty] | 2019-08-22 | 22590 | Close Job | Completed
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 |  |  |  | EA
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 |  |  |  | ST
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 |  |  |  | EA
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | ST
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 |  |  |  | EA
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 |  |  |  | EA
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02 |  |  |  | EA
[163 empty] | 2019-02-22 | 0.01046 | 76353 | 3217 | 1.01
[5 empty] | 3.9188e-05 | 0.0069112 | 105.7 | 89.0852 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7207 | 7206 | 7206 | 62 | 61 | 64 | 7145.4 | 7268.5 | jenny dahl | -592.27 | 3.12 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | -6.42‰ | -6.42‰ | 2.00‰ | 2.00‰ | -6.42 ‰ | 2.00 ‰ | AMS | -8.04‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -592.27 ‰ | -592.27 ‰ | 3.12 | 3.26 | -587.23 | 3.16 ‰ | -587.23 | 3.16 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.40773 | 0.40773 | 0.00312 | 0.00326 | 0.00308 | 0.00308 | 1.8 |  |  | 0.4142 | 0.00105 | 18978 | 2019-08-26 | 15:01 | Jeremy thompson | 212992 | 0.01023 | 0.0102 | 0.0046 | 0.0046 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 40.77 | 0.31% | 19 | 5 | 4 | ... | 0.39835 | 0.00087 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 64.1333862473911242. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01023 | 0.00663 | 0.0053 | 0.39213 | 0.39‰ | 0.00296 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.40773 | 0 | 0.00312 | 0 | 0 | 0.01023 | 0 | 0.0046 | 0 | 0.39213 | 0 | 0.00296 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77652 | 2000000 | 77653 77657 77552 77553 77321 77273 77113 77013 76956 76841 76607 76506 76446 | 77653~77657~77552~77553~77321~77273~77113~77013~76956~76841~76607~76506~76446 | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 | 3249 3247 3245 3245 3244 3240 3235 3233 3232 3228 3226 3220 3218 | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.05 mg | 0.13 | 1.05 | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 0.99907 | 0.00136 | 77615 | EA | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-09-17 | Water |  | JOB-212992 |  |  | 68246 | Normal | 14047/12 |  |  | 2019-08-13 | 4131 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 0.99948 | 0.00137 | 77616 | EA | [25 empty] | 2019-08-15 | 17662 | Graphite | Completed
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 1.002 | 0.00136 | 77617 | EA | [25 empty] | 2019-08-19 | 6747 | C13 Measurement | Completed
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 | 0.99799 | 0.00137 | 77624 | ST | [25 empty] | 2019-08-29 | 18978 | AMS Submission Results | Completed
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 | 0.99994 | 0.00135 | 77626 | EA | [25 empty] | 2019-08-30 | 22659 | Close Job | Completed
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 | 1.0015 | 0.00137 | 77629 | ST
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 |  |  |  | EA
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 |  |  |  | ST
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 |  |  |  | EA
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | EA
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 |  |  |  | EA
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57 |  |  |  | ST
[163 empty] | 2019-03-07 | 0.00812 | 76446 | 3218 | 1.02
[7 empty] | 105.9 | 88.4117 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46219 | 46218 | 46218 | 184 | 184 | 184 |  |  | jenny dahl | -996.83 | 0.07 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 1.14‰ | 1.14‰ | 2.00‰ | 2.00‰ | 1.14 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -996.83 ‰ | -996.83 ‰ | 0.07 | 0.07 | -996.79 | 0.07 ‰ | -996.79 | 0.07 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0031714 | 0.0031714 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.6 |  |  | 0.00317 | 7e-05 | 18979 | 2019-08-26 | 15:01 | Jeremy thompson | 213006 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 24 | 5 | 3 | ... | 0.00305 | 7e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00305 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00317 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00305 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77653 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.27 mg |  | 1.27 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-09-17 | Water | Complete | JOB-213006 |  | use for blank correction of ZIP wheel | 68233 | Normal | 14047/11 |  |  | 2019-08-14 | 4132 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-15 | 17663 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18979 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22660 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 107.4 | 88.1079 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42443 | 42443 | 42443 | 148 | 148 | 148 | 42296.3 | 42592.6 | jenny dahl | -994.93 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.60‰ | 4.60‰ | 2.00‰ | 2.00‰ | 4.60 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.93 ‰ | -994.93 ‰ | 0.09 | 0.09 | -994.86 | 0.09 ‰ | -994.86 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0050742 | 0.0050742 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.7 |  |  | 0.00507 | 9e-05 | 18982 | 2019-08-26 | 15:01 | Jeremy thompson | 213053 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 23 | 5 | 3 | ... | 0.00488 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00488 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00507 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00488 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77662 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water | Not Complete | JOB-213053 | Water blank test JCT 9/19 | sample too large at graphite. had to transfer back to bottle and split. LARGE vessel 7 as normal blank TEST BLANK set 1 | 68234 | Normal | 14047/11 |  |  | 2019-08-14 | 4138 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17672 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18982 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22661 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9395 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 105.4 | 88.1144 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42643 | 42643 | 42643 | 152 | 151 | 151 | 42492.6 | 42796.4 | jenny dahl | -995.05 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.05 ‰ | -995.05 ‰ | 0.09 | 0.09 | -994.99 | 0.09 ‰ | -994.99 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0049494 | 0.0049494 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00495 | 9e-05 | 18983 | 2019-08-26 | 15:01 | Jeremy thompson | 213054 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 26 | 5 | 5 | ... | 0.00476 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00476 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00495 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00476 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77659 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213054 | Water blank test JCT 9/19 | Large vessel 5 as normal blank TEST BLANK set 1 | 68235 | Normal | 14047/11 |  |  | 2019-08-14 | 4139 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-14 | 2590 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-14 | 4467 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-15 | 17669 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-08-29 | 18983 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [26 empty] | 2019-08-30 | 22662 | Close Job | Completed
[197 empty] | 2019-09-19 | 9396 | Invoice | Completed
[7 empty] | 104 | 88.3837 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43441 | 43440 | 43440 | 168 | 167 | 167 | 43274.8 | 43610.3 | jenny dahl | -995.52 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 1.46‰ | 1.46‰ | 2.00‰ | 2.00‰ | 1.46 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.52 ‰ | -995.52 ‰ | 0.09 | 0.09 | -995.46 | 0.09 ‰ | -995.46 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0044815 | 0.0044815 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.5 |  |  | 0.00448 | 9e-05 | 18984 | 2019-08-26 | 15:01 | Jeremy thompson | 213055 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.45 | 0.01% | 27 | 5 | 3 | ... | 0.00431 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00431 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00448 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00431 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77660 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.29 mg |  | 1.29 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213055 | Water blank test JCT 9/19 | Large vessel 27 as normal blank TEST BLANK set 1 | 68236 | Normal | 14047/11 |  |  | 2019-08-14 | 4140 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-15 | 17670 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18984 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22663 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9397 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 104.7 | 88.4929 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41439 | 41438 | 41438 | 145 | 145 | 145 | 41294.6 | 41585.2 | jenny dahl | -994.25 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 0.23‰ | 0.23‰ | 2.00‰ | 2.00‰ | 0.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.25 ‰ | -994.25 ‰ | 0.1 | 0.1 | -994.18 | 0.11 ‰ | -994.18 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0057501 | 0.0057501 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00575 | 0.0001 | 18985 | 2019-08-26 | 15:01 | Jeremy thompson | 213056 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 29 | 5 | 3 | ... | 0.00553 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00553 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00575 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00553 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77661 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.28 mg |  | 1.28 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213056 | Water blank test JCT 9/19 | Large vessel 4 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68237 | Normal | 14047/11 |  |  | 2019-08-14 | 4141 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-15 | 17671 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18985 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22664 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9398 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 105 | 88.3036 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41630 | 41629 | 41629 | 149 | 148 | 148 | 41482.3 | 41779.8 | jenny dahl | -994.39 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 2.37‰ | 2.37‰ | 2.00‰ | 2.00‰ | 2.37 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.39 ‰ | -994.39 ‰ | 0.1 | 0.1 | -994.32 | 0.11 ‰ | -994.32 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0056149 | 0.0056149 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.5 |  |  | 0.00561 | 0.0001 | 18986 | 2019-08-26 | 15:01 | Jeremy thompson | 213057 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 25 | 5 | 5 | ... | 0.0054 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.0054 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00561 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.0054 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77663 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.25 mg |  | 1.25 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213057 | Water blank test JCT 9/19 | Large vessel 32 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68238 | Normal | 14047/11 |  |  | 2019-08-14 | 4142 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-14 | 2591 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-14 | 4468 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-16 | 17673 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-08-29 | 18986 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [26 empty] | 2019-08-30 | 22665 | Close Job | Completed
[197 empty] | 2019-09-19 | 9399 | Invoice | Completed
[7 empty] | 104.7 | 88.0268 | 0.0004 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 49771 | 49770 | 49770 | 246 | 245 | 245 | 49528.5 | 50020.5 | jenny dahl | -997.96 | 0.06 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 5.52‰ | 5.52‰ | 2.00‰ | 2.00‰ | 5.52 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -997.96 ‰ | -997.96 ‰ | 0.06 | 0.06 | -997.94 | 0.06 ‰ | -997.94 | 0.06 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.002038 | 0.002038 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.6 |  |  | 0.00204 | 6e-05 | 18987 | 2019-08-26 | 15:01 | Jeremy thompson | 213058 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 30 | 5 | 3 | ... | 0.00196 | 6e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00196 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00204 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00196 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77664 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.20 mg |  | 1.2 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213058 | Water blank test JCT 9/19 | sample was oversized on graphite- had to return to bottle and split to reduce size. Large vessel 34 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68239 | Normal | 14047/11 |  |  | 2019-08-14 | 4143 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17674 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18987 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22666 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9400 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 103.4 | 88.271 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40837 | 40837 | 40837 | 135 | 134 | 134 | 40703.4 | 40973 | jenny dahl | -993.8 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 2.74‰ | 2.74‰ | 2.00‰ | 2.00‰ | 2.74 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.80 ‰ | -993.80 ‰ | 0.1 | 0.1 | -993.73 | 0.11 ‰ | -993.73 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0061972 | 0.0061972 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.8 |  |  | 0.0062 | 0.0001 | 18988 | 2019-08-26 | 15:01 | Jeremy thompson | 213059 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 31 | 5 | 5 | ... | 0.00596 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00596 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.0062 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00596 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77665 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213059 | Water blank test JCT 9/19 | LARGE vessel 7 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68240 | Normal | 14047/11 |  |  | 2019-08-15 | 4144 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-15 | 2593 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-15 | 4470 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-16 | 17675 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-08-29 | 18988 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [26 empty] | 2019-08-30 | 22667 | Close Job | Completed
[197 empty] | 2019-09-19 | 9401 | Invoice | Completed
[7 empty] | 103 | 88.3448 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37589 | 37588 | 37588 | 117 | 116 | 116 | 37472.8 | 37706.7 | jenny dahl | -990.71 | 0.14 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 1.90‰ | 1.90‰ | 2.00‰ | 2.00‰ | 1.90 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -990.71 ‰ | -990.71 ‰ | 0.14 | 0.14 | -990.6 | 0.14 ‰ | -990.6 | 0.14 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0092854 | 0.0092854 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00929 | 0.00014 | 18989 | 2019-08-26 | 15:01 | Jeremy thompson | 213060 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 32 | 5 | 5 | ... | 0.00893 | 0.00013 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00893 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00929 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00893 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77666 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213060 | Water blank test JCT 9/19 | LARGE vessel 27 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68241 | Normal | 14047/11 |  |  | 2019-08-15 | 4145 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-15 | 2594 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-15 | 4471 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-16 | 17676 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-08-29 | 18989 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [26 empty] | 2019-08-30 | 22668 | Close Job | Completed
[197 empty] | 2019-09-19 | 9402 | Invoice | Completed
[7 empty] | 104.3 | 87.9373 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44537 | 44537 | 44537 | 171 | 170 | 170 | 44368.3 | 44710.2 | jenny dahl | -996.09 | 0.08 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 6.54‰ | 6.54‰ | 2.00‰ | 2.00‰ | 6.54 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -996.09 ‰ | -996.09 ‰ | 0.08 | 0.08 | -996.04 | 0.08 ‰ | -996.04 | 0.08 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0039096 | 0.0039096 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.6 |  |  | 0.00391 | 8e-05 | 18990 | 2019-08-26 | 15:01 | Jeremy thompson | 213061 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.39 | 0.01% | 33 | 5 | 3 | ... | 0.00376 | 8e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes |  | ~4 per mil difference in C13 |  |  |  | 0.00376 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00391 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00376 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77673 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.22 mg |  | 1.22 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213061 | Water blank test JCT 9/19 | sample was oversized at graphite. Had to return to bottle and split to reduce size. LARGE vessel 5 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68242 | Normal | 14047/11 |  |  | 2019-08-15 | 4146 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17683 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18990 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22669 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9403 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 103.3 | 88.1528 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37607 | 37606 | 37606 | 117 | 117 | 117 | 37490.6 | 37725 | jenny dahl | -990.74 | 0.14 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.08‰ | 4.08‰ | 2.00‰ | 2.00‰ | 4.08 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -990.74 ‰ | -990.74 ‰ | 0.14 | 0.14 | -990.62 | 0.14 ‰ | -990.62 | 0.14 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0092646 | 0.0092646 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.6 |  |  | 0.00926 | 0.00014 | 18991 | 2019-08-26 | 15:01 | Jeremy thompson | 213062 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 34 | 5 | 3 | ... | 0.00891 | 0.00013 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00891 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00926 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00891 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77674 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213062 | Water blank test JCT 9/19 | LARGE vessel 32 as normal blank TEST BLANK set 3 | 68243 | Normal | 14047/11 |  |  | 2019-08-15 | 4147 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17684 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18991 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22670 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9404 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 104.9 | 88.0792 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40783 | 40783 | 40783 | 134 | 133 | 133 | 40650.5 | 40918.3 | jenny dahl | -993.76 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.92‰ | 4.92‰ | 2.00‰ | 2.00‰ | 4.92 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.76 ‰ | -993.76 ‰ | 0.1 | 0.1 | -993.68 | 0.11 ‰ | -993.68 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0062388 | 0.0062388 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00624 | 0.0001 | 18992 | 2019-08-26 | 15:01 | Jeremy thompson | 213063 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 36 | 5 | 3 | ... | 0.006 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.006 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00624 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.006 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77675 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | JOB-213063 | Water blank test JCT 9/19 | split on graphite line prior to transfer. don't trust measured pressure. LARGE vessel 4 as normal blank TEST BLANK set 3 | 68244 | Normal | 14047/11 |  |  | 2019-08-15 | 4148 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17685 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18992 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22671 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9405 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 102.7 | 88.341 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40468 | 40468 | 40468 | 142 | 141 | 141 | 40327.9 | 40611.1 | jenny dahl | -993.51 | 0.11 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 1.95‰ | 1.95‰ | 2.00‰ | 2.00‰ | 1.95 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.51 ‰ | -993.51 ‰ | 0.11 | 0.11 | -993.43 | 0.12 ‰ | -993.43 | 0.12 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0064883 | 0.0064883 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.6 |  |  | 0.00649 | 0.00011 | 18993 | 2019-08-26 | 15:01 | Jeremy thompson | 213064 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.65 | 0.01% | 37 | 5 | 3 | ... | 0.00624 | 0.00011 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00624 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00649 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00624 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77676 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.27 mg |  | 1.27 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water | Complete | JOB-213064 | Water blank test JCT 9/19 | LARGE vessel 34 as normal blank TEST BLANK set 3 | 68245 | Normal | 14047/11 |  |  | 2019-08-15 | 4149 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [26 empty] | 2019-08-16 | 17686 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [26 empty] | 2019-08-29 | 18993 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [26 empty] | 2019-08-30 | 22672 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [26 empty] | 2019-09-19 | 9406 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 89.4 | 88.3683 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42152 | 42152 | 42152 | 206 | 206 | 206 |  |  | Julia Collins | -994.74 | 0.14 ‰ | 2019-09-09 | 2019-03-13 | 13 | 180 |  |  |  |  | 2.92‰ | 2.92‰ | 2.00‰ | 2.00‰ | 2.92 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -994.74 ‰ | -994.74 ‰ | 0.14 | 0.14 | -994.67 | 0.14 ‰ | -994.67 | 0.14 ‰ |  |  | 2019-09-12 | Jocelyn Turnbull | 14:54 | 0.0052614 | 0.0052614 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00526 | 0.00014 | 19075 | 2019-09-12 | 13:18 | Jeremy thompson | 213281 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.53 | 0.01% | 27 | 5 | 4 | ... | 0.00506 | 0.00013 |  |  |  |  | 0.99591 | 0.99591 | Yes | Yes | [5 empty] | 0.00506 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00526 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00506 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-06 | Julia Collins | 12:51 |  | Completed |  |  |  |  | 77728 | 2000000 |  |  | 77746 77747 77748 77763 77765 77768 | 77746~77747~77748~77763~77765~77768 | 3251 |  | 3251 3251 3251 3251 3251 3251 | 2019-09-12 | 13:41 | Jeremy thompson |  | 1.09 mg |  | 1.09 | [5 empty] | 1 | 0.00173 | 77746 |  | XCAMS | EA | 0% |  | TW3251 measurement: * no problem with XCAMS * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA no failures secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2019-10-17 | Water |  | JOB-213281 |  |  | 68313 | Normal | 14047/11 |  |  | 2019-08-26 | 4172 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00175 | 77747 | [26 empty] | 2019-08-27 | 17738 | Graphite | Completed
[168 empty] | 1.0002 | 0.00176 | 77748 | [26 empty] | 2019-08-30 | 6775 | C13 Measurement | Completed
[168 empty] | 0.99975 | 0.00176 | 77763 | [26 empty] | 2019-09-12 | 19075 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00173 | 77765 | [26 empty] | 2019-09-13 | 22746 | Close Job | Completed
[168 empty] | 1.0001 | 0.00171 | 77768
[5 empty] | 3.8784e-05 | 0.0066807 | 91.1 | 89.3698 | 0.0686 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7164 | 7163 | 7163 | 62 | 61 | 64 | 7102.2 | 7225.5 | Julia Collins | -590.07 | 3.14 ‰ | 2019-09-09 | 2019-03-13 | 13 | 180 |  |  |  |  | -8.32‰ | -8.32‰ | 2.00‰ | 2.00‰ | -8.32 ‰ | 2.00 ‰ | AMS | -7.74‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -590.07 ‰ | -590.07 ‰ | 3.14 | 3.28 | -585.01 | 3.18 ‰ | -585.01 | 3.18 ‰ |  |  | 2019-09-12 | Jocelyn Turnbull | 14:54 | 0.40993 | 0.40993 | 0.00314 | 0.00328 | 0.0031 | 0.0031 | 2.3 |  |  | 0.41622 | 0.00129 | 19076 | 2019-09-12 | 13:18 | Jeremy thompson | 213283 | 0.00999 | 0.01 | 0.0045 | 0.0045 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 40.99 | 0.31% | 5 | 5 | 4 | ... | 0.40029 | 0.00113 |  |  |  |  | 0.99591 | 0.99591 | Yes | Yes |  | Result agrees with consensus wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 64.2118517285026435. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00999 | 0.00641 | 0.00547 | 0.39424 | 0.39‰ | 0.00298 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.40993 | 0 | 0.00314 | 0 | 0 | 0.00999 | 0 | 0.0045 | 0 | 0.39424 | 0 | 0.00298 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-06 | Julia Collins | 12:51 |  | Completed |  |  |  |  | 77729 | 2000000 | 77728 77653 77657 77552 77553 77321 77273 77113 77013 76956 76841 76607 76506 | 77728~77653~77657~77552~77553~77321~77273~77113~77013~76956~76841~76607~76506 | 77746 77747 77748 77763 77765 77768 | 77746~77747~77748~77763~77765~77768 | 3251 | 3251 3249 3247 3245 3245 3244 3240 3235 3233 3232 3228 3226 3220 | 3251 3251 3251 3251 3251 3251 | 2019-09-12 | 13:41 | Jeremy thompson |  | 0.92 mg | 0.13 | 0.92 | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 1 | 0.00173 | 77746 | EA | XCAMS | EA | 0% |  | TW3251 measurement: * no problem with XCAMS * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA no failures secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | -7.7 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-10-17 | Water | Complete | JOB-213283 |  |  | 68314 | Normal | 14047/12 |  |  | 2019-08-26 | 4173 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 0.99965 | 0.00175 | 77747 | EA | [25 empty] | 2019-08-27 | 17739 | Graphite | Completed
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 1.0002 | 0.00176 | 77748 | ST | [25 empty] | 2019-08-30 | 6776 | C13 Measurement | Completed
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 0.99975 | 0.00176 | 77763 | EA | [25 empty] | 2019-09-12 | 19076 | AMS Submission Results | Completed
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 | 1.0003 | 0.00173 | 77765 | ST | [25 empty] | 2019-09-13 | 22747 | Close Job | Completed
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 | 1.0001 | 0.00171 | 77768 | EA
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 |  |  |  | ST
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 |  |  |  | EA
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 |  |  |  | EA
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 |  |  |  | EA
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | ST
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78 |  |  |  | EA
[163 empty] | 2019-03-14 | 0.01453 | 76506 | 3220 | 0.57
[7 empty] | 93.9 | 88.9972 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40959 | 40959 | 40959 | 178 | 177 | 177 |  |  | jenny dahl | -993.9 | 0.14 ‰ | 2019-09-20 | 2019-03-24 | 7 | 180 |  |  |  |  | 0.92‰ | 0.92‰ | 2.00‰ | 2.00‰ | 0.92 ‰ | 2.00 ‰ | AMS | 2.68‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -993.90 ‰ | -993.90 ‰ | 0.14 | 0.14 | -993.82 | 0.14 ‰ | -993.82 | 0.14 ‰ |  |  | 2019-09-27 | Jocelyn Turnbull | 11:26 | 0.0061036 | 0.0061036 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.0061 | 0.00014 | 19147 | 2019-09-24 | 11:42 | albert zondervan | 213325 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 16 | 5 | 6 | ... | 0.00587 | 0.00013 |  |  |  |  | 1.001 | 1.001 | Yes | Yes | [5 empty] | 0.00587 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.0061 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00587 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-20 | jenny dahl | 12:34 |  | Completed |  |  |  |  | 77838 | 2000000 |  |  | 77778 77782 77784 77785 77786 77787 | 77778~77782~77784~77785~77786~77787 | 3253 |  | 3253 3253 3253 3253 3253 3253 | 2019-09-24 | 12:49 | albert zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99966 | 0.00171 | 77778 |  | XCAMS | ST | 0% |  | TW3253 measurement: * small & large samples * XCAMS: IML electrode 1 mm closer to ionizer, as part of UCI-cathode tests As observed in TW3234, a bias in delta13C is apparent in comparing the IRMS and AMS values. This effect appears sample size dependent and is probably generated by the modified different geometry in the ion source. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, small: result is NOT consistent with consensus value: 2.5-sigma older than consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks large organics via AAA: 6x unknowns + 0x secondaries + 2x blanks 180 days: capturing 10x blanks: 1x Queets A from this wheel 1x Queets A from TW3228 1x Kauri Renton Rd from this wheel 7x Kauri Renton Rd from previous wheels small organics via AAA: 6x unknowns + 1x secondaries + 4x blanks Excel file 'to blank-correct small samples.xlsm': 2x other wheels selected also, to assist with determining d, m parameters: TW3249, 3233 d = 1.5 ug C m = 0.81 ug C, derived from 5x Queets A blanks + 3x Kauri Renton Road blanks (Queets A on its own would result in m = 0.72 ug C; KRR on its own: m = 1.02 ug C.) unknowns: * no failures * Gary Wilson submitted unknowns and blanks processed at Otago U. They fall into both our small and large sample ranges. Consistency of their Queets A blanks with our Kauri Renton Rd blanks was tested in both sample size ranges. | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-10-21 | Water |  | JOB-213325 |  |  | 68393 | Normal | 14047/11 |  |  | 2019-09-12 | 4191 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.00172 | 77782 | [26 empty] | 2019-09-12 | 2622 | Gas Splitting | Completed
[168 empty] | 0.99982 | 0.0017 | 77784 | [26 empty] | 2019-09-12 | 4508 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00163 | 77785 | [26 empty] | 2019-09-13 | 17848 | Graphite | Completed
[168 empty] | 1.0003 | 0.00171 | 77786 | [26 empty] | 2019-09-18 | 6792 | C13 Measurement | Completed
[168 empty] | 0.99819 | 0.00168 | 77787 | [26 empty] | 2019-09-27 | 19147 | AMS Submission Results | Completed
[197 empty] | 2019-10-01 | 22877 | Close Job | Completed
[5 empty] | 3.7742e-05 | 0.0064258 | 99.5 | 89.8695 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7082 | 7081 | 7081 | 57 | 56 | 59 | 7025.3 | 7138.9 | jenny dahl | -585.88 | 2.93 ‰ | 2019-09-20 | 2019-03-24 | 7 | 180 |  |  |  |  | -8.80‰ | -8.80‰ | 2.00‰ | 2.00‰ | -8.80 ‰ | 2.00 ‰ | AMS | -8.12‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -585.88 ‰ | -585.88 ‰ | 2.93 | 3.07 | -580.77 | 2.97 ‰ | -580.77 | 2.97 ‰ |  |  | 2019-09-27 | Jocelyn Turnbull | 11:26 | 0.41412 | 0.41412 | 0.00293 | 0.00307 | 0.00288 | 0.00288 | 2.2 |  |  | 0.41995 | 0.00126 | 19148 | 2019-09-24 | 11:42 | albert zondervan | 213326 | 0.00932 | 0.0093 | 0.00419 | 0.0042 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.41 | 0.29% | 36 | 5 | 4 | ... | 0.40388 | 0.00109 |  |  |  |  | 1.001 | 1.001 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 59.6463897176221192. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00932 | 0.00617 | 0.0054 | 0.39827 | 0.40‰ | 0.00277 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.41412 | 0 | 0.00293 | 0 | 0 | 0.00932 | 0 | 0.00419 | 0 | 0.39827 | 0 | 0.00277 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-20 | jenny dahl | 12:34 |  | Completed |  |  |  |  | 77839 | 2000000 | 77838 77728 77653 77657 77552 77553 77321 77273 77113 77013 76956 76841 76607 | 77838~77728~77653~77657~77552~77553~77321~77273~77113~77013~76956~76841~76607 | 77778 77782 77784 77785 77786 77787 | 77778~77782~77784~77785~77786~77787 | 3253 | 3253 3251 3249 3247 3245 3245 3244 3240 3235 3233 3232 3228 3226 | 3253 3253 3253 3253 3253 3253 | 2019-09-24 | 12:49 | albert zondervan |  | 0.96 mg | 0.13 | 0.96 | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 0.99966 | 0.00171 | 77778 | EA | XCAMS | ST | 0% |  | TW3253 measurement: * small & large samples * XCAMS: IML electrode 1 mm closer to ionizer, as part of UCI-cathode tests As observed in TW3234, a bias in delta13C is apparent in comparing the IRMS and AMS values. This effect appears sample size dependent and is probably generated by the modified different geometry in the ion source. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, small: result is NOT consistent with consensus value: 2.5-sigma older than consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks large organics via AAA: 6x unknowns + 0x secondaries + 2x blanks 180 days: capturing 10x blanks: 1x Queets A from this wheel 1x Queets A from TW3228 1x Kauri Renton Rd from this wheel 7x Kauri Renton Rd from previous wheels small organics via AAA: 6x unknowns + 1x secondaries + 4x blanks Excel file 'to blank-correct small samples.xlsm': 2x other wheels selected also, to assist with determining d, m parameters: TW3249, 3233 d = 1.5 ug C m = 0.81 ug C, derived from 5x Queets A blanks + 3x Kauri Renton Road blanks (Queets A on its own would result in m = 0.72 ug C; KRR on its own: m = 1.02 ug C.) unknowns: * no failures * Gary Wilson submitted unknowns and blanks processed at Otago U. They fall into both our small and large sample ranges. Consistency of their Queets A blanks with our Kauri Renton Rd blanks was tested in both sample size ranges. | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-10-21 | Water | Not Complete | JOB-213326 |  |  | 68394 | Normal | 14047/12 |  |  | 2019-09-12 | 4192 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 1.0011 | 0.00172 | 77782 | ST | [25 empty] | 2019-09-13 | 17849 | Graphite | Completed
[163 empty] | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 0.99982 | 0.0017 | 77784 | EA | [25 empty] | 2019-09-23 | 6823 | C13 Measurement | Completed
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 1.0009 | 0.00163 | 77785 | ST | [25 empty] | 2019-09-27 | 19148 | AMS Submission Results | Completed
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 1.0003 | 0.00171 | 77786 | EA | [25 empty] | 2019-10-01 | 22878 | Close Job | Completed
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 | 0.99819 | 0.00168 | 77787 | ST
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 |  |  |  | EA
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 |  |  |  | EA
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 |  |  |  | EA
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 |  |  |  | ST
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 |  |  |  | EA
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | ST
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78
[7 empty] | 104.6 | 88.7704 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43862 | 43861 | 43861 | 216 | 216 | 216 |  |  | Julia Collins | -995.75 | 0.11 ‰ | 2019-09-27 | 2019-03-31 | 0 | 180 |  |  |  |  | 1.02‰ | 1.02‰ | 2.00‰ | 2.00‰ | 1.02 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -995.75 ‰ | -995.75 ‰ | 0.11 | 0.11 | -995.69 | 0.12 ‰ | -995.69 | 0.12 ‰ |  |  | 2019-09-30 | Jocelyn Turnbull | 11:16 | 0.0042528 | 0.0042528 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.1 |  |  | 0.00425 | 0.00011 | 19174 | 2019-09-30 | 9:14 | jeremy thompson | 213282 |  |  |  |  | 45% | 5 | 12/11/20 | 09:31:58 | albert zondervan | 0 |  | 0.43 | 0.01% | 8 | 5 | 4 | ... | 0.00409 | 0.00011 |  |  |  |  | 0.99855 | 0.99855 | Yes | Yes | [5 empty] | 0.00409 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00425 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00409 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-27 | Julia Collins | 13:15 |  | Completed |  |  |  |  | 77973 | 2000000 |  |  | 77853 77854 77855 77856 77857 77861 | 77853~77854~77855~77856~77857~77861 | 3254 |  | 3254 3254 3254 3254 3254 3254 | 2019-09-30 | 9:30 | Jeremy thompson |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99998 | 0.00164 | 77853 |  | XCAMS | ST | 0% |  | TW3254 measurement: * all large samples in NEC cathode holders * no problem with XCAMS * SNICS in modified geometry for UCI-cathode holder tests * regular precision reached CalAMS error adjustments: * system error: 0% possible problem: * sample-size dependent difference in del13C measured by IRMS and AMS * cause is unknown primaries: * OxI cathodes: 6x, all via ST no failures secondaries: * 1x IAEA-C6 sucrose, not used beyond tuning * 1x IAEA-C2 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 31x unknowns + 1x secondaries + 1x blanks 180 days: capturing 14x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-10-17 | Water |  | JOB-213282 |  |  | 68419 | Normal | 14047/11 |  |  | 2019-09-26 | 4222 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00163 | 77854 | [26 empty] | 2019-09-27 | 17983 | Graphite | Completed
[168 empty] | 1 | 0.00166 | 77855 | [26 empty] | 2019-09-30 | 6837 | C13 Measurement | Completed
[168 empty] | 1 | 0.00167 | 77856 | [26 empty] | 2019-09-30 | 19174 | AMS Submission Results | Completed
[168 empty] | 0.99823 | 0.00167 | 77857 | [26 empty] | 2019-10-01 | 22873 | Close Job | Completed
[168 empty] | 1.0011 | 0.00171 | 77861
[5 empty] | 3.5841e-05 | 0.0062124 | 108.1 | 89.6564 | 0.0688 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7163 | 7163 | 7163 | 56 | 55 | 58 | 7107.6 | 7219.2 | Julia Collins | -590.05 | 2.85 ‰ | 2019-09-27 | 2019-03-31 | 0 | 180 |  |  |  |  | -8.87‰ | -8.87‰ | 2.00‰ | 2.00‰ | -8.87 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -590.05 ‰ | -590.05 ‰ | 2.85 | 2.99 | -584.99 | 2.88 ‰ | -584.99 | 2.88 ‰ |  |  | 2019-09-30 | Jocelyn Turnbull | 11:16 | 0.40995 | 0.40995 | 0.00285 | 0.00299 | 0.0028 | 0.0028 | 1.9 |  |  | 0.41559 | 0.00124 | 19175 | 2019-09-30 | 9:14 | jeremy thompson | 213284 | 0.00895 | 0.009 | 0.00403 | 0.004 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 40.99 | 0.28% | 29 | 5 | 4 | ... | 0.39968 | 0.00107 |  |  |  |  | 0.99855 | 0.99855 | Yes | Yes |  | Result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 58.6530493032807663. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00895 | 0.00596 | 0.00537 | 0.39426 | 0.39‰ | 0.00269 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.40995 | 0 | 0.00285 | 0 | 0 | 0.00895 | 0 | 0.00403 | 0 | 0.39426 | 0 | 0.00269 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-27 | Julia Collins | 13:15 |  | Completed |  |  |  |  | 77974 | 2000000 | 77973 77838 77728 77653 77657 77552 77553 77321 77273 77113 77013 76956 76841 76607 | 77973~77838~77728~77653~77657~77552~77553~77321~77273~77113~77013~76956~76841~76607 | 77853 77854 77855 77856 77857 77861 | 77853~77854~77855~77856~77857~77861 | 3254 | 3254 3253 3251 3249 3247 3245 3245 3244 3240 3235 3233 3232 3228 3226 | 3254 3254 3254 3254 3254 3254 | 2019-09-30 | 9:30 | Jeremy thompson |  | 1.02 mg | 0.13 | 1.02 | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 0.99998 | 0.00164 | 77853 | EA | XCAMS | ST | 0% |  | TW3254 measurement: * all large samples in NEC cathode holders * no problem with XCAMS * SNICS in modified geometry for UCI-cathode holder tests * regular precision reached CalAMS error adjustments: * system error: 0% possible problem: * sample-size dependent difference in del13C measured by IRMS and AMS * cause is unknown primaries: * OxI cathodes: 6x, all via ST no failures secondaries: * 1x IAEA-C6 sucrose, not used beyond tuning * 1x IAEA-C2 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 31x unknowns + 1x secondaries + 1x blanks 180 days: capturing 14x blanks unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2019-10-17 | Water | Not Complete | JOB-213284 |  |  | 68420 | Normal | 14047/12 |  |  | 2019-09-26 | 4223 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 1.0007 | 0.00163 | 77854 | ST | [25 empty] | 2019-09-27 | 17984 | Graphite | Completed
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 1 | 0.00166 | 77855 | EA | [25 empty] | 2019-09-30 | 6833 | C13 Measurement | Completed
[163 empty] | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 1 | 0.00167 | 77856 | ST | [25 empty] | 2019-09-30 | 19175 | AMS Submission Results | Completed
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 0.99823 | 0.00167 | 77857 | EA | [25 empty] | 2019-10-01 | 22874 | Close Job | Completed
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 | 1.0011 | 0.00171 | 77861 | ST
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 |  |  |  | EA
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 |  |  |  | EA
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 |  |  |  | EA
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12 |  |  |  | ST
[163 empty] | 2019-05-28 | 0.0141 | 77013 | 3233 | 0.67 |  |  |  | EA
[163 empty] | 2019-05-14 | 0.00759 | 76956 | 3232 | 1.06 |  |  |  | ST
[163 empty] | 2019-04-24 | 0.01788 | 76841 | 3228 | 0.8 |  |  |  | ST
[163 empty] | 2019-04-11 | 0.01614 | 76607 | 3226 | 0.78
[7 empty] | 104.8 | 88.4917 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46432 | 46432 | 46432 | 297 | 297 | 297 |  |  | Julia Collins | -996.91 | 0.11 ‰ | 2019-10-08 | 2019-04-11 | 4 | 180 |  |  |  |  | 2.05‰ | 2.05‰ | 2.00‰ | 2.00‰ | 2.05 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0.82 | 0.8165 | -996.91 ‰ | -996.91 ‰ | 0.11 | 0.11 | -996.87 | 0.12 ‰ | -996.87 | 0.12 ‰ |  |  | 2019-10-11 | jenny dahl | 8:29 | 0.0030882 | 0.0030882 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.3 |  |  | 0.00309 | 0.00011 | 19218 | 2019-10-09 | 13:59 | Jeremy thompson | 213445 |  |  |  |  | 45% | 5 | 11/10/19 | 08:28:36 | jenny dahl | 0 |  | 0.31 | 0.01% | 11 | 5 | 5 | ... | 0.00297 | 0.00011 |  |  |  |  | 0.99643 | 0.99643 | Yes | Yes | [5 empty] | 0.00297 | 0.00‰ | 0.00011 | 0.00‰ | No | 1 | 0.99999 | 0.00095 | 0.00095 | 0.00309 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00297 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-04 | Julia Collins | 13:39 |  | Completed |  |  |  |  | 78005 | 2000000 |  |  | 77940 77941 77942 77943 77960 77966 | 77940~77941~77942~77943~77960~77966 | 3255 |  | 3255 3255 3255 3255 3255 3255 | 2019-10-09 | 14:13 | Jeremy thompson |  | 0.80 mg |  | 0.8 | [5 empty] | 1.0025 | 0.00231 | 77940 |  | XCAMS | EA | 0.14% |  | TW3255 measurement: * all large samples in NEC cathode holders * no problem with XCAMS * regular precision reached CalAMS error adjustments: * system error: 1.4% primaries: * OxI cathodes: 6x, all via ST no failures secondaries: * 1x IAEA-C2 Carbonate: result is consistent with consensus value * 1x FIRI-I cellulose: result is 1 sigma high * 1x IAEA-C6 sucrose, not used beyond tuning blanks: * Two groups identified to blank-correct: carbonate: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks organics via AAA: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks: unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-11-06 | Carbonate |  | JOB-213445 |  |  | 68493 |  | 14047/1 |  |  | 2019-10-01 | 4238 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99917 | 0.00236 | 77941 | [26 empty] | 2019-10-01 | 2631 | Gas Splitting | Completed
[168 empty] | 0.99598 | 0.00234 | 77942 | [26 empty] | 2019-10-01 | 4524 | Choose Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00232 | 77943 | [26 empty] | 2019-10-04 | 18015 | Graphite | Completed
[168 empty] | 1.0019 | 0.00233 | 77960 | [26 empty] | 2019-10-11 | 19218 | AMS Submission Results | Completed
[168 empty] | 0.99976 | 0.00235 | 77966 | [26 empty] | 2019-10-11 | 22961 | Close Job | Completed
[7 empty] | 81.6 | 88.6593 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46298 | 46297 | 46297 | 239 | 239 | 239 |  |  | Julia Collins | -996.86 | 0.09 ‰ | 2019-10-16 | 2019-04-19 | 5 | 180 |  |  |  |  | 4.17‰ | 4.17‰ | 2.00‰ | 2.00‰ | 4.17 ‰ | 2.00 ‰ | AMS | 2.76‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -996.86 ‰ | -996.86 ‰ | 0.09 | 0.09 | -996.82 | 0.09 ‰ | -996.82 | 0.09 ‰ |  |  | 2019-10-18 | Jocelyn Turnbull | 10:58 | 0.0031402 | 0.0031402 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00314 | 9e-05 | 19301 | 2019-10-18 | 9:14 | albert zondervan | 213577 |  |  |  |  | 45% | 5 | 18/10/19 | 10:57:53 | Jocelyn Turnbull | 0 |  | 0.31 | 0.01% | 1 | 5 | 6 | ... | 0.00302 | 9e-05 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes | [5 empty] | 0.00302 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00314 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00302 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-15 | Julia Collins | 13:36 |  | Completed |  |  |  |  | 78037 | 2000000 |  |  | 77968 77970 78022 78023 78074 78076 | 77968~77970~78022~78023~78074~78076 | 3257 |  | 3257 3257 3257 3257 3257 3257 | 2019-10-18 | 9:35 | albert zondervan |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0016 | 0.00176 | 77968 |  | XCAMS | EA | 0% |  | TW3257 measurement: * all large samples * no problem with XCAMS * regular precision reached CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA no failures secondaries: * 1x IAEA-C2 Carbonate: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value * 1x IAEA-C6 sucrose, not used beyond tuning blanks: * four groups identified to blank-correct: carbonate: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks water: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 13x blanks organics via AAA: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks: organics via CE: 2x unknowns + 0x secondaries + 0x blanks 180 days: capturing 5x blanks: unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2019-11-20 | Carbonate | Complete | JOB-213577 |  |  | 68553 |  | 14047/1 |  |  | 2019-10-07 | 4252 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99812 | 0.00175 | 77970 | [26 empty] | 2019-10-07 | 2638 | Gas Splitting | Completed
[168 empty] | 0.99811 | 0.00174 | 78022 | [26 empty] | 2019-10-07 | 4549 | Choose Gas Splitting | Completed
[168 empty] | 0.99847 | 0.00175 | 78023 | [26 empty] | 2019-10-11 | 18047 | Graphite | Completed
[168 empty] | 1.0024 | 0.0018 | 78074 | [26 empty] | 2019-10-14 | 6860 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.00182 | 78076 | [26 empty] | 2019-10-18 | 19301 | AMS Submission Results | Completed
[197 empty] | 2019-10-22 | 23029 | Close Job | Completed
[7 empty] | 78.2 | 88.4498 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40455 | 40455 | 40455 | 231 | 231 | 231 | 40227.3 | 40690.1 | Julia Collins | -993.5 | 0.19 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 3.46‰ | 3.46‰ | 2.00‰ | 2.00‰ | 3.46 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -993.50 ‰ | -993.50 ‰ | 0.19 | 0.19 | -993.42 | 0.19 ‰ | -993.42 | 0.19 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0064987 | 0.0064987 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.9 |  |  | 0.0065 | 0.00019 | 19371 | 2019-10-25 | 8:47 | Jeremy thompson | 213679 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:31 | Jocelyn Turnbull | 0 |  | 0.65 | 0.02% | 9 | 5 | 3 | ... | 0.00625 | 0.00018 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00625 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.0065 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00625 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78153 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.08 mg |  | 1.08 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213679 | Charge to GCT | 3 minute extraction test water blanks to be run by Jeremy | 68593 | Normal | 14047/11 |  |  | 2019-10-21 | 4296 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-22 | 18163 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19371 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23119 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10102 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79.7 | 88.3921 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43385 | 43385 | 43385 | 278 | 277 | 277 | 43112.1 | 43667.6 | Julia Collins | -995.49 | 0.16 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 4.11‰ | 4.11‰ | 2.00‰ | 2.00‰ | 4.11 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -995.49 ‰ | -995.49 ‰ | 0.16 | 0.16 | -995.43 | 0.16 ‰ | -995.43 | 0.16 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0045127 | 0.0045127 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.8 |  |  | 0.00451 | 0.00016 | 19372 | 2019-10-25 | 8:47 | Jeremy thompson | 213680 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:32 | Jocelyn Turnbull | 0 |  | 0.45 | 0.02% | 8 | 5 | 3 | ... | 0.00434 | 0.00015 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00434 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00451 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00434 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78154 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.13 mg |  | 1.13 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213680 | Charge to GCT | 5 minute extraction test water blanks to be run by Jeremy | 68594 | Normal | 14047/11 |  |  | 2019-10-21 | 4297 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-22 | 18164 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19372 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23120 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10103 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 78.5 | 88.3622 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44954 | 44954 | 44954 | 315 | 315 | 315 | 44645 | 45275.4 | Julia Collins | -996.29 | 0.15 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 4.45‰ | 4.45‰ | 2.00‰ | 2.00‰ | 4.45 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.29 ‰ | -996.29 ‰ | 0.15 | 0.15 | -996.24 | 0.15 ‰ | -996.24 | 0.15 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0037121 | 0.0037121 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00371 | 0.00015 | 19373 | 2019-10-25 | 8:47 | Jeremy thompson | 213681 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 10 | 5 | 3 | ... | 0.00357 | 0.00014 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00357 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00371 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00357 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78155 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213681 | Charge to GCT | 7 minute extraction test water blanks to be run by Jeremy | 68595 | Normal | 14047/11 |  |  | 2019-10-21 | 4298 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-22 | 18165 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19373 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23121 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10104 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 77 | 88.4959 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37367 | 37367 | 37367 | 184 | 183 | 183 | 37185.5 | 37553 | Julia Collins | -990.45 | 0.22 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -990.45 ‰ | -990.45 ‰ | 0.22 | 0.22 | -990.34 | 0.22 ‰ | -990.34 | 0.22 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0095453 | 0.0095453 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 1.7 |  |  | 0.00955 | 0.00022 | 19374 | 2019-10-25 | 8:47 | Jeremy thompson | 213682 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 11 | 5 | 3 | ... | 0.00918 | 0.00021 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00918 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00955 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.00918 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78156 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.20 mg |  | 1.2 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213682 | Charge to GCT | 10 minute extraction test water blanks to be run by Jeremy | 68596 | Normal | 14047/11 |  |  | 2019-10-21 | 4299 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-22 | 18166 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19374 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23122 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10105 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79 | 88.451 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38373 | 38372 | 38372 | 198 | 198 | 198 | 38176.6 | 38573.4 | Julia Collins | -991.58 | 0.21 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 3.44‰ | 3.44‰ | 2.00‰ | 2.00‰ | 3.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -991.58 ‰ | -991.58 ‰ | 0.21 | 0.21 | -991.47 | 0.21 ‰ | -991.47 | 0.21 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0084223 | 0.0084223 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 1.9 |  |  | 0.00842 | 0.00021 | 19375 | 2019-10-25 | 8:47 | Jeremy thompson | 213683 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 12 | 5 | 3 | ... | 0.0081 | 0.0002 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.0081 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00842 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.0081 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78157 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213683 | Charge to GCT | 20 minute extraction test water blanks to be run by Jeremy | 68597 | Normal | 14047/11 |  |  | 2019-10-21 | 4300 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-22 | 18167 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19375 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23123 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10106 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 80.7 | 88.3629 | 0.001 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 234 | 233 | 233 | 41251.9 | 41719.4 | Julia Collins | -994.28 | 0.17 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.44‰ | 4.44‰ | 2.00‰ | 2.00‰ | 4.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.28 ‰ | -994.28 ‰ | 0.17 | 0.17 | -994.21 | 0.17 ‰ | -994.21 | 0.17 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0057189 | 0.0057189 | 0.00017 | 0.00017 | 0.00017 | 0.00017 |  |  |  | 0.00572 | 0.00017 | 19376 | 2019-10-25 | 8:47 | Jeremy thompson | 213684 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:34 | Jocelyn Turnbull | 0 |  | 0.57 | 0.02% | 13 | 5 | 3 | ... | 0.0055 | 0.00016 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00572 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00016 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78171 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213684 | Charge to GCT | 3 minute extraction test water blanks to be run by Jeremy | 68598 | Normal | 14047/11 |  |  | 2019-10-22 | 4301 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-23 | 18181 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19376 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23124 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10107 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 77.2 | 88.3651 | 0.0009 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41964 | 41963 | 41963 | 248 | 248 | 248 | 41719.4 | 42215.8 | Julia Collins | -994.61 | 0.17 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.42‰ | 4.42‰ | 2.00‰ | 2.00‰ | 4.42 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.61 ‰ | -994.61 ‰ | 0.17 | 0.17 | -994.55 | 0.17 ‰ | -994.55 | 0.17 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0053861 | 0.0053861 | 0.00017 | 0.00017 | 0.00017 | 0.00017 |  |  |  | 0.00539 | 0.00017 | 19377 | 2019-10-25 | 8:47 | Jeremy thompson | 213685 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:34 | Jocelyn Turnbull | 0 |  | 0.54 | 0.02% | 15 | 5 | 3 | ... | 0.00518 | 0.00016 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00518 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00539 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00518 | 0 | 0.00016 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78172 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213685 | Charge to GCT | 5 minute extraction test water blanks to be run by Jeremy | 68599 | Normal | 14047/11 |  |  | 2019-10-22 | 4302 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-23 | 18182 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19377 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23125 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10108 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79.6 | 88.3355 | 0.0018 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 36219 | 36219 | 36219 | 182 | 182 | 182 | 36039.3 | 36403.5 | Julia Collins | -988.99 | 0.25 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -988.99 ‰ | -988.99 ‰ | 0.25 | 0.25 | -988.85 | 0.25 ‰ | -988.85 | 0.25 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.011011 | 0.011011 | 0.00025 | 0.00025 | 0.00025 | 0.00025 |  |  |  | 0.01101 | 0.00025 | 19378 | 2019-10-25 | 8:47 | Jeremy thompson | 213686 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:35 | Jocelyn Turnbull | 0 |  | 1.1 | 0.02% | 16 | 5 | 3 | ... | 0.01059 | 0.00024 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.01059 | 0.01‰ | 0.00024 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.01101 | 0 | 0.00025 | 0 | 0 |  |  | 0 | 0 | 0.01059 | 0 | 0.00024 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78173 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213686 | Charge to GCT | 7 minute extraction test water blanks to be run by Jeremy | 68600 | Normal | 14047/11 |  |  | 2019-10-22 | 4303 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-23 | 18183 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19378 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23126 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10109 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 75.7 | 88.5405 | 0.0035 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31104 | 31103 | 31103 | 124 | 124 | 124 | 30980.3 | 31229.1 | Julia Collins | -979.18 | 0.32 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 2.43‰ | 2.43‰ | 2.00‰ | 2.00‰ | 2.43 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -979.18 ‰ | -979.18 ‰ | 0.32 | 0.32 | -978.93 | 0.33 ‰ | -978.93 | 0.33 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.020817 | 0.020817 | 0.00032 | 0.00032 | 0.00032 | 0.00032 |  |  |  | 0.02082 | 0.00032 | 19379 | 2019-10-25 | 8:47 | Jeremy thompson | 213687 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:36 | Jocelyn Turnbull | 0 |  | 2.08 | 0.03% | 6 | 5 | 3 | ... | 0.02002 | 0.00031 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.02002 | 0.02‰ | 0.00031 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.02082 | 0 | 0.00032 | 0 | 0 |  |  | 0 | 0 | 0.02002 | 0 | 0.00031 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78174 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213687 | Charge to GCT | 10 minute extraction test water blanks to be run by Jeremy | 68601 | Normal | 14047/11 |  |  | 2019-10-22 | 4304 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-23 | 18184 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19379 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23127 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10110 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79 | 88.5167 | 0.0026 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33407 | 33406 | 33406 | 144 | 144 | 144 | 33263.7 | 33552.3 | Julia Collins | -984.37 | 0.28 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 2.70‰ | 2.70‰ | 2.00‰ | 2.00‰ | 2.70 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -984.37 ‰ | -984.37 ‰ | 0.28 | 0.28 | -984.18 | 0.28 ‰ | -984.18 | 0.28 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.015628 | 0.015628 | 0.00028 | 0.00028 | 0.00028 | 0.00028 |  |  |  | 0.01563 | 0.00028 | 19380 | 2019-10-25 | 8:47 | Jeremy thompson | 213688 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:36 | Jocelyn Turnbull | 0 |  | 1.56 | 0.03% | 17 | 5 | 3 | ... | 0.01503 | 0.00027 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.01503 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.01563 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01503 | 0 | 0.00027 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78175 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | JOB-213688 | Charge to GCT | 20 minute extraction test water blanks to be run by Jeremy | 68602 | Normal | 14047/11 |  |  | 2019-10-22 | 4305 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [26 empty] | 2019-10-23 | 18185 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [26 empty] | 2019-10-25 | 19380 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [26 empty] | 2019-10-30 | 23128 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [26 empty] | 2020-02-24 | 10111 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 73 | 88.4822 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46036 | 46036 | 46036 | 257 | 257 | 257 |  |  | Julia Collins | -996.76 | 0.10 ‰ | 2019-10-25 | 2019-04-28 | 7 | 180 |  |  |  |  | 2.43‰ | 2.43‰ | 2.00‰ | 2.00‰ | 2.43 ‰ | 2.00 ‰ | AMS | 2.73‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -996.76 ‰ | -996.76 ‰ | 0.1 | 0.1 | -996.72 | 0.11 ‰ | -996.72 | 0.11 ‰ |  |  | 2019-10-29 | Jocelyn Turnbull | 10:19 | 0.0032442 | 0.0032442 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.2 |  |  | 0.00324 | 0.0001 | 19430 | 2019-10-29 | 9:58 | albert zondervan | 213578 |  |  |  |  | 45% | 5 | 29/10/19 | 10:19:27 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 8 | 5 | 4 | ... | 0.00312 | 0.0001 |  |  |  |  | 0.99671 | 0.99671 | Yes | Yes | [5 empty] | 0.00312 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00324 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00312 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-24 | Julia Collins | 13:25 |  | Completed |  |  |  |  | 78134 | 2000000 |  |  | 78151 78166 78167 78168 78169 78170 | 78151~78166~78167~78168~78169~78170 | 3260 |  | 3260 3260 3260 3260 3260 3260 | 2019-10-29 | 10:03 | albert zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99773 | 0.00172 | 78151 |  | XCAMS | ST | 0% |  | TW3260 measurement: * no problem with XCAMS CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x no failures secondaries: * 1x IAEA-C2 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: shells: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-11-20 | Carbonate |  | JOB-213578 |  |  | 68604 |  | 14047/1 |  |  | 2019-10-17 | 4295 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0017 | 0.00174 | 78166 | [26 empty] | 2019-10-18 | 18144 | Graphite | Completed
[168 empty] | 0.99996 | 0.00175 | 78167 | [26 empty] | 2019-10-22 | 6903 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00176 | 78168 | [26 empty] | 2019-10-29 | 19430 | AMS Submission Results | Completed
[168 empty] | 0.99997 | 0.00176 | 78169 | [26 empty] | 2019-10-30 | 23117 | Close Job | Completed
[168 empty] | 0.99981 | 0.00171 | 78170
[7 empty] | 69.5 | 87.5143 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45205 | 45205 | 45205 | 232 | 232 | 232 |  |  | Julia Collins | -996.4 | 0.10 ‰ | 2019-11-22 | 2019-05-26 | 7 | 180 |  |  |  |  | 3.53‰ | 3.53‰ | 2.00‰ | 2.00‰ | 3.53 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.40 ‰ | -996.40 ‰ | 0.1 | 0.1 | -996.36 | 0.11 ‰ | -996.36 | 0.11 ‰ |  |  | 2019-11-26 | Jocelyn Turnbull | 10:07 | 0.0035977 | 0.0035977 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.0036 | 0.0001 | 19519 | 2019-11-25 | 9:17 | Jeremy thompson | 213843 |  |  |  |  | 45% | 5 | 26/11/19 | 10:06:39 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 11 | 5 | 5 | ... | 0.00346 | 0.0001 |  |  |  |  | 0.98689 | 0.98689 | Yes | Yes | [5 empty] | 0.00346 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.0036 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00346 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-11-21 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 78232 | 2000000 |  |  | 78221 78222 78262 78263 78278 78279 | 78221~78222~78262~78263~78278~78279 | 3263 |  | 3263 3263 3263 3263 3263 3263 | 2019-11-25 | 9:36 | Jeremy thompson |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99923 | 0.00154 | 78221 |  | XCAMS | EA | 0% |  | TW3263 measurement: * large samples * XCAMS: - no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large Carbonates: 3x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks large organics via AAA: 25x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks: unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-12-18 | Carbonate |  | JOB-213843 |  |  | 68753 |  | 14047/1 |  |  | 2019-11-14 | 4312 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0009 | 0.00158 | 78222 | [26 empty] | 2019-11-14 | 2656 | Gas Splitting | Completed
[168 empty] | 1.0019 | 0.00158 | 78262 | [26 empty] | 2019-11-14 | 4584 | Choose Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00157 | 78263 | [26 empty] | 2019-11-15 | 18242 | Graphite | Completed
[168 empty] | 0.99811 | 0.00159 | 78278 | [26 empty] | 2019-11-26 | 19519 | AMS Submission Results | Completed
[168 empty] | 0.99954 | 0.0016 | 78279 | [26 empty] | 2019-11-27 | 23301 | Close Job | Completed
[7 empty] | 83.3 | 87.5887 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43553 | 43553 | 43553 | 189 | 189 | 189 |  |  | Julia Collins | -995.58 | 0.10 ‰ | 2019-11-27 | 2019-05-31 | 5 | 180 |  |  |  |  | 2.75‰ | 2.75‰ | 2.00‰ | 2.00‰ | 2.75 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -995.58 ‰ | -995.58 ‰ | 0.1 | 0.1 | -995.53 | 0.11 ‰ | -995.53 | 0.11 ‰ |  |  | 2019-11-29 | Jocelyn Turnbull | 14:47 | 0.0044191 | 0.0044191 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.1 |  |  | 0.00442 | 0.0001 | 19560 | 2019-11-29 | 13:44 | Jeremy thompson | 213903 |  |  |  |  | 45% | 5 | 29/11/19 | 14:46:36 | Jocelyn Turnbull | 0 |  | 0.44 | 0.01% | 22 | 5 | 5 | ... | 0.00425 | 0.0001 |  |  |  |  | 0.98696 | 0.98696 | Yes | Yes | [5 empty] | 0.00425 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00442 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00425 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-11-26 | Julia Collins | 13:11 |  | Completed |  |  |  |  | 78296 | 2000000 |  |  | 78212 78214 78216 78217 78218 78219 | 78212~78214~78216~78217~78218~78219 | 3264 |  | 3264 3264 3264 3264 3264 3264 | 2019-11-29 | 14:09 | Jeremy thompson |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0003 | 0.0015 | 78212 |  | XCAMS | ST | 0% |  | TW3264 measurement: * large samples * XCAMS: - no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x Whalebone, large: result is consistent with consensus value blanks: * three groups identified to blank-correct: large Carbonates: 13x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large Waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks large Organics via AAA: 10x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-01-06 | Carbonate |  | JOB-213903 |  |  | 68793 |  | 14047/1 |  |  | 2019-11-20 | 4324 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99964 | 0.0015 | 78214 | [26 empty] | 2019-11-20 | 2661 | Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00152 | 78216 | [26 empty] | 2019-11-20 | 4591 | Choose Gas Splitting | Completed
[168 empty] | 0.99871 | 0.00152 | 78217 | [26 empty] | 2019-11-22 | 18306 | Graphite | Completed
[168 empty] | 1.0029 | 0.00151 | 78218 | [26 empty] | 2019-11-29 | 19560 | AMS Submission Results | Completed
[168 empty] | 0.99824 | 0.00161 | 78219 | [26 empty] | 2019-12-02 | 23333 | Close Job | Completed
[7 empty] | 82.1 | 87.5748 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42443 | 42443 | 42443 | 165 | 164 | 164 |  |  | Julia Collins | -994.93 | 0.10 ‰ | 2019-11-27 | 2019-05-31 | 5 | 180 |  |  |  |  | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -994.93 ‰ | -994.93 ‰ | 0.1 | 0.1 | -994.86 | 0.11 ‰ | -994.86 | 0.11 ‰ |  |  | 2019-11-29 | Jocelyn Turnbull | 14:47 | 0.0050742 | 0.0050742 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00507 | 0.0001 | 19562 | 2019-11-29 | 13:44 | Jeremy thompson | 213929 |  |  |  |  | 45% | 5 | 29/11/19 | 14:46:38 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 1 | 5 | 5 | ... | 0.00488 | 0.0001 |  |  |  |  | 0.98696 | 0.98696 | Yes | Yes | [5 empty] | 0.00488 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00507 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00488 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-11-26 | Julia Collins | 13:11 |  | Completed |  |  |  |  | 78305 | 2000000 |  |  | 78212 78214 78216 78217 78218 78219 | 78212~78214~78216~78217~78218~78219 | 3264 |  | 3264 3264 3264 3264 3264 3264 | 2019-11-29 | 14:09 | Jeremy thompson |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0003 | 0.0015 | 78212 |  | XCAMS | ST | 0% |  | TW3264 measurement: * large samples * XCAMS: - no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x Whalebone, large: result is consistent with consensus value blanks: * three groups identified to blank-correct: large Carbonates: 13x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large Waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks large Organics via AAA: 10x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-01-14 | Water |  | JOB-213929 |  |  | 68794 | Normal | 14047/11 |  |  | 2019-11-21 | 4328 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99964 | 0.0015 | 78214 | [26 empty] | 2019-11-21 | 2663 | Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00152 | 78216 | [26 empty] | 2019-11-21 | 4595 | Choose Gas Splitting | Completed
[168 empty] | 0.99871 | 0.00152 | 78217 | [26 empty] | 2019-11-22 | 18315 | Graphite | Completed
[168 empty] | 1.0029 | 0.00151 | 78218 | [26 empty] | 2019-11-29 | 19562 | AMS Submission Results | Completed
[168 empty] | 0.99824 | 0.00161 | 78219 | [26 empty] | 2019-12-02 | 23334 | Close Job | Completed
[5 empty] | 3.6154e-05 | 0.0051148 | 81.3 | 88.5279 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7142 | 7142 | 7142 | 44 | 43 | 47 | 7098.2 | 7186.1 | Julia Collins | -588.97 | 2.25 ‰ | 2019-11-27 | 2019-05-31 | 5 | 180 |  |  |  |  | -7.89‰ | -7.89‰ | 2.00‰ | 2.00‰ | -7.89 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -588.97 ‰ | -588.97 ‰ | 2.25 | 2.43 | -583.89 | 2.28 ‰ | -583.89 | 2.28 ‰ |  |  | 2019-11-29 | Jocelyn Turnbull | 14:47 | 0.41103 | 0.41103 | 0.00225 | 0.00243 | 0.00218 | 0.00218 | 2.1 |  |  | 0.41528 | 0.00114 | 19563 | 2019-11-29 | 13:44 | Jeremy thompson | 213930 | 0.00677 | 0.0068 | 0.00305 | 0.0031 | 45% | 5 | 5/07/22 | 14:06:51 | Jocelyn Turnbull | 0 |  | 41.1 | 0.22% | 29 | 5 | 5 | ... | 0.39939 | 0.00097 |  |  |  |  | 0.98696 | 0.98696 | Yes | Yes |  | Result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 47.511522160086777. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00677 | 0.00491 | 0.00424 | 0.3953 | 0.40‰ | 0.0021 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.41103 | 0 | 0.00225 | 0 | 0 | 0.00677 | 0 | 0.00305 | 0 | 0.3953 | 0 | 0.0021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-11-26 | Julia Collins | 13:11 |  | Completed |  |  |  |  | 78306 | 2000000 | 78305 77973 77838 77728 77653 77657 77552 77553 77321 77273 77113 | 78305~77973~77838~77728~77653~77657~77552~77553~77321~77273~77113 | 78212 78214 78216 78217 78218 78219 | 78212~78214~78216~78217~78218~78219 | 3264 | 3264 3254 3253 3251 3249 3247 3245 3245 3244 3240 3235 | 3264 3264 3264 3264 3264 3264 | 2019-11-29 | 14:09 | Jeremy thompson |  | 0.99 mg | 0.13 | 0.99 | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 | 1.0003 | 0.0015 | 78212 | EA | XCAMS | ST | 0% |  | TW3264 measurement: * large samples * XCAMS: - no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x Whalebone, large: result is consistent with consensus value blanks: * three groups identified to blank-correct: large Carbonates: 13x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large Waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks large Organics via AAA: 10x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: unknowns: * no failures |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-01-14 | Water | Complete | JOB-213930 |  |  | 68795 | Normal | 14047/12 |  |  | 2019-11-21 | 4329 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 0.99964 | 0.0015 | 78214 | ST | [25 empty] | 2019-11-21 | 2664 | Gas Splitting | Completed
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 0.99997 | 0.00152 | 78216 | EA | [25 empty] | 2019-11-21 | 4597 | Choose Gas Splitting | Completed
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 0.99871 | 0.00152 | 78217 | EA | [25 empty] | 2019-11-22 | 18316 | Graphite | Completed
[163 empty] | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 1.0029 | 0.00151 | 78218 | EA | [25 empty] | 2019-11-29 | 19563 | AMS Submission Results | Completed
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 | 0.99824 | 0.00161 | 78219 | ST | [25 empty] | 2019-12-02 | 23335 | Close Job | Completed
[163 empty] | 2019-08-03 | 0.0134 | 77552 | 3245 | 0.99 |  |  |  | EA
[163 empty] | 2019-08-03 | 0.01512 | 77553 | 3245 | 0.99 |  |  |  | ST
[163 empty] | 2019-07-22 | 0.00241 | 77321 | 3244 | 1.09 |  |  |  | ST
[163 empty] | 2019-07-09 | 0.00919 | 77273 | 3240 | 1.2 |  |  |  | ST
[163 empty] | 2019-06-10 | 0.00344 | 77113 | 3235 | 1.12
[7 empty] | 106.9 | 87.2294 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40190 | 40189 | 40189 | 187 | 186 | 186 | 40005.5 | 40378.6 | Julia Collins | -993.28 | 0.16 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.94‰ | 2.94‰ | 2.00‰ | 2.00‰ | 2.94 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -993.28 ‰ | -993.28 ‰ | 0.16 | 0.16 | -993.2 | 0.16 ‰ | -993.2 | 0.16 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0067171 | 0.0067171 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00672 | 0.00016 | 19606 | 2019-12-09 | 11:49 | Jeremy thompson | 213869 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.67 | 0.02% | 1 | 5 | 4 | ... | 0.00646 | 0.00015 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00646 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00672 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00646 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78361 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water | Complete | JOB-213869 | Charge to GCT | test water blanks to be run by Jeremy | 68874 | Normal | 14047/11 |  |  | 2019-11-26 | 4332 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-28 | 18371 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6945 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19606 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23525 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10112 | Invoice | Completed
[7 empty] | 109.5 | 87.2901 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43706 | 43706 | 43706 | 231 | 231 | 231 | 43478.2 | 43940.6 | Julia Collins | -995.66 | 0.12 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.24‰ | 2.24‰ | 2.00‰ | 2.00‰ | 2.24 ‰ | 2.00 ‰ | AMS | 2.92‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -995.66 ‰ | -995.66 ‰ | 0.12 | 0.12 | -995.61 | 0.13 ‰ | -995.61 | 0.13 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0043359 | 0.0043359 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00434 | 0.00012 | 19607 | 2019-12-09 | 11:49 | Jeremy thompson | 213870 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.43 | 0.01% | 2 | 5 | 4 | ... | 0.00417 | 0.00012 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00417 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00434 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00417 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78362 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.03 mg |  | 1.03 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213870 | Charge to GCT | test water blanks to be run by Jeremy | 68875 | Normal | 14047/11 |  |  | 2019-11-26 | 4333 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-28 | 18372 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6946 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19607 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23526 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10113 | Invoice | Completed
[7 empty] | 107.7 | 87.2256 | 0.0029 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 32517 | 32517 | 32517 | 115 | 114 | 114 | 32403.2 | 32632.8 | Julia Collins | -982.54 | 0.25 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.98‰ | 2.98‰ | 2.00‰ | 2.00‰ | 2.98 ‰ | 2.00 ‰ | AMS | 2.69‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -982.54 ‰ | -982.54 ‰ | 0.25 | 0.25 | -982.33 | 0.25 ‰ | -982.33 | 0.25 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.017458 | 0.017458 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 2.1 |  |  | 0.01746 | 0.00025 | 19608 | 2019-12-09 | 11:49 | Jeremy thompson | 213871 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 1.75 | 0.02% | 3 | 5 | 4 | ... | 0.01679 | 0.00024 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.01679 | 0.02‰ | 0.00024 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01746 | 0 | 0.00025 | 0 | 0 |  |  | 0 | 0 | 0.01679 | 0 | 0.00024 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78363 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213871 | Charge to GCT | test water blanks to be run by Jeremy | 68876 | Normal | 14047/11 |  |  | 2019-11-26 | 4334 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-28 | 18373 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6947 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19608 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23527 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10114 | Invoice | Completed
[7 empty] | 108.2 | 87.2638 | 0.0024 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34186 | 34186 | 34186 | 130 | 129 | 129 | 34057.7 | 34316.9 | Julia Collins | -985.82 | 0.23 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.54‰ | 2.54‰ | 2.00‰ | 2.00‰ | 2.54 ‰ | 2.00 ‰ | AMS | 2.79‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -985.82 ‰ | -985.82 ‰ | 0.23 | 0.23 | -985.64 | 0.23 ‰ | -985.64 | 0.23 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.014183 | 0.014183 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 1.7 |  |  | 0.01418 | 0.00023 | 19609 | 2019-12-09 | 11:49 | Jeremy thompson | 213872 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 1.42 | 0.02% | 4 | 5 | 4 | ... | 0.01364 | 0.00022 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.01364 | 0.01‰ | 0.00022 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01418 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01364 | 0 | 0.00022 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78364 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213872 | Charge to GCT | test water blanks to be run by Jeremy | 68877 | Normal | 14047/11 |  |  | 2019-11-26 | 4335 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-28 | 18374 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6948 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19609 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23528 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10115 | Invoice | Completed
[7 empty] | 109 | 87.2403 | 0.0013 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38853 | 38852 | 38852 | 168 | 168 | 168 | 38686.1 | 39023 | Julia Collins | -992.07 | 0.17 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.81‰ | 2.81‰ | 2.00‰ | 2.00‰ | 2.81 ‰ | 2.00 ‰ | AMS | 2.81‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -992.07 ‰ | -992.07 ‰ | 0.17 | 0.17 | -991.97 | 0.17 ‰ | -991.97 | 0.17 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0079336 | 0.0079336 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2 |  |  | 0.00793 | 0.00017 | 19610 | 2019-12-09 | 11:49 | Jeremy thompson | 213873 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.79 | 0.02% | 11 | 5 | 4 | ... | 0.00763 | 0.00016 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00763 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00793 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00763 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78365 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water | Complete | JOB-213873 | Charge to GCT | test water blanks to be run by Jeremy | 68878 | Normal | 14047/11 |  |  | 2019-11-26 | 4336 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-28 | 18375 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6949 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19610 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23529 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10116 | Invoice | Completed
[7 empty] | 107 | 87.1711 | 0.0012 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40079 | 40078 | 40078 | 184 | 183 | 183 | 39896.9 | 40264.9 | Julia Collins | -993.19 | 0.16 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 3.61‰ | 3.61‰ | 2.00‰ | 2.00‰ | 3.61 ‰ | 2.00 ‰ | AMS | 2.88‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -993.19 ‰ | -993.19 ‰ | 0.16 | 0.16 | -993.11 | 0.16 ‰ | -993.11 | 0.16 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0068107 | 0.0068107 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00681 | 0.00016 | 19611 | 2019-12-09 | 11:49 | Jeremy thompson | 213874 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.68 | 0.02% | 6 | 5 | 4 | ... | 0.00655 | 0.00015 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00655 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00681 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00655 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78366 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.02 mg |  | 1.02 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213874 | Charge to GCT | test water blanks to be run by Jeremy | 68879 | Normal | 14047/11 |  |  | 2019-11-27 | 4337 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-29 | 18376 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6950 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19611 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23530 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10117 | Invoice | Completed
[7 empty] | 108 | 87.2661 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 236 | 235 | 235 | 42393.8 | 42865.5 | Julia Collins | -995.04 | 0.15 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.51‰ | 2.51‰ | 2.00‰ | 2.00‰ | 2.51 ‰ | 2.00 ‰ | AMS | 2.83‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.15 | 0.15 | -994.98 | 0.15 ‰ | -994.98 | 0.15 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0049598 | 0.0049598 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00496 | 0.00015 | 19612 | 2019-12-09 | 11:49 | Jeremy thompson | 213875 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.5 | 0.01% | 7 | 5 | 4 | ... | 0.00477 | 0.00014 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00496 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78367 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213875 | Charge to GCT | test water blanks to be run by Jeremy | 68880 | Normal | 14047/11 |  |  | 2019-11-27 | 4338 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-29 | 18377 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6951 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19612 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23531 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10118 | Invoice | Completed
[7 empty] | 107 | 87.24 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41871 | 41871 | 41871 | 215 | 214 | 214 | 41659.5 | 42088.8 | Julia Collins | -994.55 | 0.15 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.81‰ | 2.81‰ | 2.00‰ | 2.00‰ | 2.81 ‰ | 2.00 ‰ | AMS | 2.87‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -994.55 ‰ | -994.55 ‰ | 0.15 | 0.15 | -994.48 | 0.15 ‰ | -994.48 | 0.15 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0054485 | 0.0054485 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00545 | 0.00015 | 19613 | 2019-12-09 | 11:49 | Jeremy thompson | 213876 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.54 | 0.01% | 8 | 5 | 4 | ... | 0.00524 | 0.00014 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00524 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00545 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00524 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78368 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.03 mg |  | 1.03 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213876 | Charge to GCT | test water blanks to be run by Jeremy | 68881 | Normal | 14047/11 |  |  | 2019-11-27 | 4339 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-29 | 18378 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6952 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19613 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23532 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10119 | Invoice | Completed
[7 empty] | 107.4 | 87.2268 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38645 | 38644 | 38644 | 174 | 174 | 174 | 38472.4 | 38821.2 | Julia Collins | -991.86 | 0.18 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.97‰ | 2.97‰ | 2.00‰ | 2.00‰ | 2.97 ‰ | 2.00 ‰ | AMS | 2.82‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -991.86 ‰ | -991.86 ‰ | 0.18 | 0.18 | -991.76 | 0.18 ‰ | -991.76 | 0.18 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0081416 | 0.0081416 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 1.9 |  |  | 0.00814 | 0.00018 | 19614 | 2019-12-09 | 11:49 | Jeremy thompson | 213877 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.81 | 0.02% | 9 | 5 | 4 | ... | 0.00783 | 0.00017 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00783 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00814 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00783 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78369 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.00 mg |  | 1 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213877 | Charge to GCT | test water blanks to be run by Jeremy | 68882 | Normal | 14047/11 |  |  | 2019-11-27 | 4340 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-29 | 18379 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6953 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19614 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23533 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10120 | Invoice | Completed
[7 empty] | 107.6 | 87.2266 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38759 | 38758 | 38758 | 166 | 166 | 166 | 38593.8 | 38926.8 | Julia Collins | -991.97 | 0.17 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.97‰ | 2.97‰ | 2.00‰ | 2.00‰ | 2.97 ‰ | 2.00 ‰ | AMS | 2.86‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -991.97 ‰ | -991.97 ‰ | 0.17 | 0.17 | -991.87 | 0.17 ‰ | -991.87 | 0.17 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0080272 | 0.0080272 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2 |  |  | 0.00803 | 0.00017 | 19615 | 2019-12-09 | 11:49 | Jeremy thompson | 213878 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.8 | 0.02% | 12 | 5 | 4 | ... | 0.00772 | 0.00016 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00772 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00803 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00772 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78370 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | JOB-213878 | Charge to GCT | test water blanks to be run by Jeremy | 68883 | Normal | 14047/11 |  |  | 2019-11-27 | 4341 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-11-29 | 18380 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-10 | 6954 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19615 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-20 | 23534 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [26 empty] | 2020-02-24 | 10121 | Invoice | Completed
[7 empty] | 108.2 | 87.2828 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44977 | 44976 | 44976 | 248 | 248 | 248 |  |  | Julia Collins | -996.3 | 0.11 ‰ | 2019-12-05 | 2019-06-08 | 3 | 180 |  |  |  |  | 2.32‰ | 2.32‰ | 2.00‰ | 2.00‰ | 2.32 ‰ | 2.00 ‰ | AMS | 2.81‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -996.30 ‰ | -996.30 ‰ | 0.11 | 0.11 | -996.25 | 0.12 ‰ | -996.25 | 0.12 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0037017 | 0.0037017 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.1 |  |  | 0.0037 | 0.00011 | 19630 | 2019-12-09 | 11:49 | Jeremy thompson | 214070 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.37 | 0.01% | 19 | 5 | 4 | ... | 0.00356 | 0.00011 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00356 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.0037 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00356 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78399 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2020-01-21 | Carbonate | Complete | JOB-214070 |  |  | 68873 |  | 14047/1 |  |  | 2019-11-29 | 4354 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99799 | 0.00172 | 78388 | [26 empty] | 2019-12-02 | 18409 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [26 empty] | 2019-12-03 | 6960 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [26 empty] | 2019-12-09 | 19630 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [26 empty] | 2019-12-10 | 23427 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401
[7 empty] | 92 | 87.6373 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46140 | 46139 | 46139 | 261 | 260 | 260 |  |  | Julia Collins | -996.8 | 0.10 ‰ | 2020-01-28 | 2019-08-01 | 11 | 180 |  |  |  |  | 1.66‰ | 1.66‰ | 2.00‰ | 2.00‰ | 1.66 ‰ | 2.00 ‰ | AMS | 3.47‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -996.80 ‰ | -996.80 ‰ | 0.1 | 0.1 | -996.76 | 0.11 ‰ | -996.76 | 0.11 ‰ |  |  | 2020-01-29 | Jocelyn Turnbull | 18:39 | 0.0032026 | 0.0032026 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.0032 | 0.0001 | 19741 | 2020-01-30 | 8:45 | albert zondervan | 214068 |  |  |  |  | 45% | 5 | 29/01/20 | 18:39:37 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 8 | 5 | 4 | ... | 0.00308 | 0.0001 |  |  |  |  | 0.98643 | 0.98643 | Yes | Yes | [5 empty] | 0.00308 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.0032 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00308 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-01-24 | Julia Collins | 13:58 |  | Completed |  |  |  |  | 78522 | 2000000 |  |  | 78474 78476 78502 78503 78504 78505 | 78474~78476~78502~78503~78504~78505 | 3269 |  | 3269 3269 3269 3269 3269 3269 | 2020-01-30 | 8:52 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 0.99981 | 0.00169 | 78474 |  | XCAMS | ST | 0% |  | TW3269: * first measurement of 2020 * first measurement after major job on Pelletron * large calcite samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: large result is consistent with consensus value blanks: * One groups identified to blank-correct: large Carbonates: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | 3.5 | RRL | [5 empty] | Background Inorganic |  | 2020-01-21 | Carbonate | Not Complete | JOB-214068 |  | From Andy P - this value is out of range so the oxalic was used to pull data into LIMS. | 69009 |  | 14047/1 |  |  | 2020-01-16 | 4391 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99727 | 0.00167 | 78476 | [26 empty] | 2020-01-17 | 18532 | Graphite | Completed
[168 empty] | 1.0021 | 0.00171 | 78502 | [26 empty] | 2020-01-23 | 6998 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00165 | 78503 | [26 empty] | 2020-01-29 | 19741 | AMS Submission Results | Completed
[168 empty] | 0.99972 | 0.00173 | 78504 | [26 empty] | 2020-01-31 | 23558 | Close Job | Completed
[168 empty] | 1.0009 | 0.00173 | 78505
[7 empty] | 110.9 | 87.564 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48806 | 48806 | 48806 | 254 | 254 | 254 |  |  | Julia Collins | -997.7 | 0.07 ‰ | 2020-01-31 | 2019-08-04 | 8 | 180 |  |  |  |  | 4.46‰ | 4.46‰ | 2.00‰ | 2.00‰ | 4.46 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.8165 | -997.70 ‰ | -997.70 ‰ | 0.07 | 0.07 | -997.67 | 0.07 ‰ | -997.67 | 0.07 ‰ |  |  | 2020-02-03 | Jocelyn Turnbull | 11:05 | 0.0022979 | 0.0022979 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.0023 | 7e-05 | 19757 | 2020-02-03 | 8:48 | albert zondervan | 214069 |  |  |  |  | 45% | 5 | 3/02/20 | 11:06:38 | Jocelyn Turnbull | 0 |  | 0.23 | 0.01% | 1 | 5 | 4 | ... | 0.00221 | 7e-05 |  |  |  |  | 0.98836 | 0.98836 | Yes | Yes | [5 empty] | 0.00221 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.0023 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00221 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-01-29 | Julia Collins | 11:08 |  | Completed |  |  |  |  | 78534 | 2000000 |  |  | 78515 78516 78517 78547 78548 78549 | 78515~78516~78517~78547~78548~78549 | 3270 |  | 3270 3270 3270 3270 3270 3270 | 2020-02-03 | 9:39 | albert zondervan |  | 1.00 mg |  | 1 | [5 empty] | 0.99991 | 0.00156 | 78515 |  | XCAMS | EA | 0% |  | TW3270: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x FIRI-I: cellulose: large result is consistent with consensus value * 1x IAEA-C2 carbonate: large result is consistent with consensus value blanks: * Two groups identified to blank-correct: large carbonates: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks large organics: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-01-21 | Carbonate | Complete | JOB-214069 |  |  | 69025 |  | 14047/1 |  |  | 2020-01-21 | 4398 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0009 | 0.00153 | 78516 | [26 empty] | 2020-01-23 | 18544 | Graphite | Completed
[168 empty] | 1.0006 | 0.00157 | 78517 | [26 empty] | 2020-01-27 | 7014 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00161 | 78547 | [26 empty] | 2020-02-03 | 19757 | AMS Submission Results | Completed
[168 empty] | 0.99902 | 0.00158 | 78548 | [26 empty] | 2020-02-04 | 23606 | Close Job | Completed
[168 empty] | 0.99743 | 0.00158 | 78549
[7 empty] | 94.5 | 87.7367 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48556 | 48555 | 48555 | 281 | 281 | 281 |  |  | Julia Collins | -997.63 | 0.08 ‰ | 2020-02-04 | 2019-08-08 | 6 | 180 |  |  |  |  | 2.73‰ | 2.73‰ | 2.00‰ | 2.00‰ | 2.73 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.0045 | 0 | 0.8165 | -997.63 ‰ | -997.63 ‰ | 0.08 | 0.08 | -997.6 | 0.08 ‰ | -997.6 | 0.08 ‰ |  |  | 2020-02-07 | Jocelyn Turnbull | 9:45 | 0.0023707 | 0.0023707 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.8 |  |  | 0.00237 | 8e-05 | 19797 | 2020-02-07 | 15:49 | albert zondervan | 214226 |  |  |  |  | 45% | 5 | 7/02/20 | 09:45:08 | Jocelyn Turnbull | 0 |  | 0.24 | 0.01% | 1 | 5 | 5 | ... | 0.00228 | 8e-05 |  |  |  |  | 0.98861 | 0.98861 | Yes | Yes | [5 empty] | 0.00228 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 1 | 0.00101 | 0.00101 | 0.00237 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00228 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-04 | Julia Collins | 12:29 |  | Completed |  |  |  |  | 78570 | 2000000 |  |  | 78563 78572 78573 78574 78600 78601 | 78563~78572~78573~78574~78600~78601 | 3271 |  | 3271 3271 3271 3271 3271 3271 | 2020-02-07 | 16:15 | albert zondervan |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0026 | 0.00246 | 78563 |  | XCAMS | ST | 0.18% |  | TW3271: * large samples CalAMS error adjustments: * global factor: 1.06 * cat#34 (TP78625) factor: 1.3 * system error: 0.18% primaries: * OxI cathodes: 6x large no failures secondaries: * 1x FIRI-I: cellulose: result is consistent with consensus value * 1x IAEA-C2 carbonate: result is consistent with consensus value * 1x FIRI-L whalebone: result is consistent with consensus value blanks: * Three groups identified to blank-correct: carbonates: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks organics, excl R41390 series: 8x unknowns + 2x secondaries + 1x blanks 180 days, capturing 8x blanks organics R41390 series only: 9x unknowns + 0x secondaries + 1x blanks 180 days, capturing 2x blanks (R41390/1, R41330/5) unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-03-05 | Carbonate | Complete | JOB-214226 |  |  | 69065 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00249 | 78572 | [26 empty] | 2020-01-27 | 4415 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99813 | 0.00244 | 78573 | [26 empty] | 2020-01-29 | 18580 | Graphite | Completed
[168 empty] | 1.0004 | 0.00247 | 78574 | [26 empty] | 2020-01-31 | 7029 | C13 Measurement | Completed
[168 empty] | 1.0024 | 0.00249 | 78600 | [26 empty] | 2020-02-07 | 19797 | AMS Submission Results | Completed
[168 empty] | 0.99626 | 0.00245 | 78601 | [26 empty] | 2020-02-10 | 23648 | Close Job | Completed
[7 empty] | 73.6 | 88.1492 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43257 | 43256 | 43256 | 182 | 182 | 182 |  |  | jenny dahl | -995.41 | 0.10 ‰ | 2020-02-07 | 2019-08-11 | 4 | 180 |  |  |  |  | 5.43‰ | 5.43‰ | 2.00‰ | 2.00‰ | 5.43 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -995.41 ‰ | -995.41 ‰ | 0.1 | 0.1 | -995.36 | 0.11 ‰ | -995.36 | 0.11 ‰ |  |  | 2020-02-10 | Jocelyn Turnbull | 9:58 | 0.0045855 | 0.0045855 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00459 | 0.0001 | 19858 | 2020-02-10 | 11:37 | albert zondervan | 214270 |  |  |  |  | 45% | 5 | 10/02/20 | 09:57:40 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 22 | 5 | 5 | ... | 0.00441 | 0.0001 |  |  |  |  | 0.99593 | 0.99593 | Yes | Yes | [5 empty] | 0.00441 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00459 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00441 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-05 | jenny dahl | 15:00 |  | Completed |  |  |  |  | 78621 | 2000000 |  |  | 78551 78552 78553 78583 78584 78586 | 78551~78552~78553~78583~78584~78586 | 3272 |  | 3272 3272 3272 3272 3272 3272 | 2020-02-10 | 11:55 | albert zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99926 | 0.00153 | 78551 |  | XCAMS | EA | 0% |  | TW3272: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x FIRI-I: cellulose: result is consistent with consensus value * 1x IAEA-C2 carbonate: result is consistent with consensus value * 1x IAEA-C6 sucrose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-03-17 | Water |  | JOB-214270 |  |  | 69127 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99877 | 0.00154 | 78552 | [26 empty] | 2020-01-31 | 4450 | Water CO2 Evolution | Completed
[168 empty] | 1.0032 | 0.00157 | 78553 | [26 empty] | 2020-02-03 | 18631 | Graphite | Completed
[168 empty] | 1.0007 | 0.00153 | 78583 | [26 empty] | 2020-02-05 | 7065 | C13 Measurement | Completed
[168 empty] | 0.99885 | 0.00157 | 78584 | [26 empty] | 2020-02-10 | 19858 | AMS Submission Results | Completed
[168 empty] | 0.99931 | 0.00159 | 78586 | [26 empty] | 2020-02-12 | 23675 | Close Job | Completed
[5 empty] | 4.0969e-05 | 0.0045967 | 75.4 | 89.1881 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7092 | 7091 | 7091 | 36 | 35 | 40 | 7056.2 | 7128 | jenny dahl | -586.4 | 1.85 ‰ | 2020-02-07 | 2019-08-11 | 4 | 180 |  |  |  |  | -6.28‰ | -6.28‰ | 2.00‰ | 2.00‰ | -6.28 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -586.40 ‰ | -586.40 ‰ | 1.85 | 2.07 | -581.29 | 1.87 ‰ | -581.29 | 1.87 ‰ |  |  | 2020-02-10 | Jocelyn Turnbull | 9:58 | 0.4136 | 0.4136 | 0.00185 | 0.00207 | 0.00177 | 0.00177 | 1.7 |  |  | 0.41676 | 0.00116 | 19859 | 2020-02-10 | 11:37 | albert zondervan | 214271 | 0.00505 | 0.0051 | 0.00227 | 0.0023 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.36 | 0.18% | 23 | 5 | 4 | ... | 0.40081 | 0.00099 |  |  |  |  | 0.99593 | 0.99593 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.1855541305023258. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00505 | 0.00442 | 0.00156 | 0.39777 | 0.40‰ | 0.0017 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.4136 | 0 | 0.00185 | 0 | 0 | 0.00505 | 0 | 0.00227 | 0 | 0.39777 | 0 | 0.0017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-05 | jenny dahl | 15:00 |  | Completed |  |  |  |  | 78622 | 2000000 | 78621 78305 77973 77838 77728 77653 77657 | 78621~78305~77973~77838~77728~77653~77657 | 78551 78552 78553 78583 78584 78586 | 78551~78552~78553~78583~78584~78586 | 3272 | 3272 3264 3254 3253 3251 3249 3247 | 3272 3272 3272 3272 3272 3272 | 2020-02-10 | 11:55 | albert zondervan |  | 1.00 mg | 0.13 | 1 | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 0.99926 | 0.00153 | 78551 | EA | XCAMS | EA | 0% |  | TW3272: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x FIRI-I: cellulose: result is consistent with consensus value * 1x IAEA-C2 carbonate: result is consistent with consensus value * 1x IAEA-C6 sucrose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-03-17 | Water | Complete | JOB-214271 |  |  | 69128 | Normal | 14047/12 |  |  | 2020-01-31 | 4451 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 | 0.99877 | 0.00154 | 78552 | ST | [25 empty] | 2020-02-03 | 18632 | Graphite | Completed
[163 empty] | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 1.0032 | 0.00157 | 78553 | EA | [25 empty] | 2020-02-05 | 7066 | C13 Measurement | Completed
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 1.0007 | 0.00153 | 78583 | ST | [25 empty] | 2020-02-10 | 19859 | AMS Submission Results | Completed
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 0.99885 | 0.00157 | 78584 | ST | [25 empty] | 2020-02-12 | 23676 | Close Job | Completed
[163 empty] | 2019-08-23 | 0.00305 | 77653 | 3249 | 1.27 | 0.99931 | 0.00159 | 78586 | ST
[163 empty] | 2019-08-16 | 0.00797 | 77657 | 3247 | 1.16 |  |  |  | EA
[7 empty] | 64.4 | 88.2168 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42216 | 42215 | 42215 | 224 | 224 | 224 |  |  | Julia Collins | -994.78 | 0.15 ‰ | 2020-02-26 | 2019-08-30 | 1 | 180 |  |  |  |  | 4.43‰ | 4.43‰ | 2.00‰ | 2.00‰ | 4.43 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19 | 0 | 0.8165 | -994.78 ‰ | -994.78 ‰ | 0.15 | 0.15 | -994.72 | 0.15 ‰ | -994.72 | 0.15 ‰ |  |  | 2020-02-28 | Jocelyn Turnbull | 14:14 | 0.0052198 | 0.0052198 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2 |  |  | 0.00522 | 0.00015 | 19916 | 2020-02-28 | 10:57 | albert zondervan | 214225 |  |  |  |  | 45% | 5 | 2/03/20 | 11:14:45 | Margaret Norris | 0 |  | 0.52 | 0.01% | 1 | 5 | 5 | ... | 0.00502 | 0.00014 |  |  |  |  | 0.9957 | 0.9957 | Yes | Yes | [5 empty] | 0.00502 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.00522 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00502 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-25 | Julia Collins | 13:06 |  | Completed |  |  |  |  | 78739 | 2000000 |  |  | 78560 78561 78562 78590 78593 78697 | 78560~78561~78562~78590~78593~78697 | 3274 |  | 3274 3274 3274 3274 3274 3274 | 2020-02-28 | 11:01 | albert zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0002 | 0.00197 | 78560 |  | XCAMS | EA | 0% |  | TW3274: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: result is consistent with consensus value blanks: * One groups identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-03-05 | Water | Complete | JOB-214225 |  |  | 69145 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99809 | 0.00197 | 78561 | [26 empty] | 2020-02-21 | 4496 | Water CO2 Evolution | Completed
[168 empty] | 0.99983 | 0.00199 | 78562 | [26 empty] | 2020-02-25 | 18749 | Graphite | Completed
[168 empty] | 1.0004 | 0.00196 | 78590 | [26 empty] | 2020-02-26 | 7119 | C13 Measurement | Completed
[168 empty] | 1.0024 | 0.00198 | 78593 | [26 empty] | 2020-02-28 | 19916 | AMS Submission Results | Completed
[168 empty] | 0.99902 | 0.00202 | 78697 | [26 empty] | 2020-03-03 | 23779 | Close Job | Completed
[5 empty] | 4.8907e-05 | 0.0049916 | 65.1 | 89.2783 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7115 | 7115 | 7115 | 39 | 38 | 42 | 7076.8 | 7154.3 | Julia Collins | -587.61 | 1.99 ‰ | 2020-02-26 | 2019-08-30 | 5 | 180 |  |  |  |  | -7.51‰ | -7.51‰ | 2.00‰ | 2.00‰ | -7.51 ‰ | 2.00 ‰ | AMS | -8.39‰ | 0.2‰ | C13 | No | -19 | -19 | 0 | 0.8165 | -587.61 ‰ | -587.61 ‰ | 1.99 | 2.19 | -582.52 | 2.01 ‰ | -582.52 | 2.01 ‰ |  |  | 2020-02-28 | Jocelyn Turnbull | 14:16 | 0.41239 | 0.41239 | 0.00199 | 0.00219 | 0.00191 | 0.00191 | 2.3 |  |  | 0.41546 | 0.00142 | 19923 | 2020-02-28 | 10:57 | albert zondervan | 213969 | 0.00489 | 0.0049 | 0.0022 | 0.0022 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.24 | 0.2% | 22 | 5 | 4 | ... | 0.39956 | 0.00126 |  |  |  |  | 0.9957 | 0.9957 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 42.7211174601581166. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00489 | 0.0048 | 0.00061 | 0.39661 | 0.40‰ | 0.00184 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.41239 | 0 | 0.00199 | 0 | 0 | 0.00489 | 0 | 0.0022 | 0 | 0.39661 | 0 | 0.00184 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-25 | Julia Collins | 13:06 |  | Completed |  |  |  |  | 78734 | 2000000 | 78739 78621 78305 77973 77838 77728 | 78739~78621~78305~77973~77838~77728 | 78560 78561 78562 78590 78593 78697 | 78560~78561~78562~78590~78593~78697 | 3274 | 3274 3272 3264 3254 3253 3251 | 3274 3274 3274 3274 3274 3274 | 2020-02-28 | 11:01 | albert zondervan |  | 0.97 mg | 0.13 | 0.97 | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 1.0002 | 0.00197 | 78560 | EA | XCAMS | EA | 0% |  | TW3274: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: result is consistent with consensus value blanks: * One groups identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-01-16 | Water | Complete | JOB-213969 |  | Swapped out vessel connection as the previous one was dirty. | 69146 | Normal | 14047/12 |  |  | 2020-02-20 | 4484 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 0.99809 | 0.00197 | 78561 | ST | [25 empty] | 2020-02-21 | 18744 | Graphite | Completed
[163 empty] | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 | 0.99983 | 0.00199 | 78562 | ST | [25 empty] | 2020-02-24 | 7111 | C13 Measurement | Completed
[163 empty] | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 1.0004 | 0.00196 | 78590 | ST | [25 empty] | 2020-02-28 | 19923 | AMS Submission Results | Completed
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 1.0024 | 0.00198 | 78593 | EA | [25 empty] | 2020-03-03 | 23756 | Close Job | Completed
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 | 0.99902 | 0.00202 | 78697 | EA
[7 empty] | 98.3 | 88.3622 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40677 | 40676 | 40676 | 185 | 184 | 184 |  |  | Julia Collins | -993.68 | 0.15 ‰ | 2020-03-03 | 2019-09-05 | 5 | 180 |  |  |  |  | 4.38‰ | 4.38‰ | 2.00‰ | 2.00‰ | 4.38 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -993.68 ‰ | -993.68 ‰ | 0.15 | 0.15 | -993.6 | 0.15 ‰ | -993.6 | 0.15 ‰ |  |  | 2020-03-09 | Jocelyn Turnbull | 9:16 | 0.006322 | 0.006322 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00632 | 0.00015 | 19993 | 2020-03-06 | 10:24 | albert zondervan | 214396 |  |  |  |  | 45% | 5 | 9/03/20 | 09:16:06 | Jocelyn Turnbull | 0 |  | 0.63 | 0.01% | 20 | 5 | 4 | ... | 0.00608 | 0.00014 |  |  |  |  | 0.99729 | 0.99729 | Yes | Yes | [5 empty] | 0.00608 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00632 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00608 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-02 | Julia Collins | 10:48 |  | Completed |  |  |  |  | 78782 | 2000000 |  |  | 78705 78706 78707 78724 78725 78726 | 78705~78706~78707~78724~78725~78726 | 3275 |  | 3275 3275 3275 3275 3275 3275 | 2020-03-06 | 10:35 | albert zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 0.99865 | 0.00214 | 78705 |  | XCAMS | EA | 0.14% |  | TW3275: * large samples * no issues with XCAMS CalAMS error adjustments: * 0.14% system error primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: resulting CRA is higher than consensus value by 2.5 sigma * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-03-30 | Water | Complete | JOB-214396 |  |  | 69224 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00214 | 78706 | [26 empty] | 2020-02-26 | 4510 | Water CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00215 | 78707 | [26 empty] | 2020-02-27 | 18792 | Graphite | Completed
[168 empty] | 0.99743 | 0.00215 | 78724 | [26 empty] | 2020-03-09 | 19993 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00219 | 78725 | [26 empty] | 2020-03-10 | 23890 | Close Job | Completed
[168 empty] | 1.0037 | 0.00217 | 78726
[5 empty] | 4.6352e-05 | 0.005127 | 94.8 | 89.4283 | 0.0681 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7234 | 7234 | 7234 | 38 | 37 | 41 | 7196.7 | 7271.9 | Julia Collins | -593.66 | 1.90 ‰ | 2020-03-03 | 2019-09-05 | 5 | 180 |  |  |  |  | -7.59‰ | -7.59‰ | 2.00‰ | 2.00‰ | -7.59 ‰ | 2.00 ‰ | AMS | -8.42‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -593.66 ‰ | -593.66 ‰ | 1.9 | 2.11 | -588.64 | 1.93 ‰ | -588.64 | 1.93 ‰ |  |  | 2020-03-09 | Jocelyn Turnbull | 9:17 | 0.40634 | 0.40634 | 0.0019 | 0.00211 | 0.00183 | 0.00183 | 1.9 |  |  | 0.40954 | 0.00122 | 19994 | 2020-03-06 | 10:24 | albert zondervan | 214397 | 0.00506 | 0.0051 | 0.00228 | 0.0023 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 40.63 | 0.19% | 1 | 5 | 4 | ... | 0.39387 | 0.00119 |  |  |  |  | 0.99729 | 0.99729 | Yes | Yes |  | resulting CRA is higher than consensus value by 2.5 sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 41.7168282476548805. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00506 | 0.00493 | 0.00072 | 0.39079 | 0.39‰ | 0.00184 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.40634 | 0 | 0.0019 | 0 | 0 | 0.00506 | 0 | 0.00228 | 0 | 0.39079 | 0 | 0.00184 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-02 | Julia Collins | 10:48 |  | Completed |  |  |  |  | 78777 | 2000000 | 78782 78739 78621 78305 77973 77838 77728 | 78782~78739~78621~78305~77973~77838~77728 | 78705 78706 78707 78724 78725 78726 | 78705~78706~78707~78724~78725~78726 | 3275 | 3275 3274 3272 3264 3254 3253 3251 | 3275 3275 3275 3275 3275 3275 | 2020-03-06 | 10:35 | albert zondervan |  | 0.98 mg | 0.13 | 0.98 | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 0.99865 | 0.00214 | 78705 | EA | XCAMS | EA | 0.14% |  | TW3275: * large samples * no issues with XCAMS CalAMS error adjustments: * 0.14% system error primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: resulting CRA is higher than consensus value by 2.5 sigma * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-03-30 | Water |  | JOB-214397 |  |  | 69225 | Normal | 14047/12 |  |  | 2020-02-26 | 4511 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 0.9993 | 0.00214 | 78706 | ST | [25 empty] | 2020-02-27 | 18787 | Graphite | Completed
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 1.0011 | 0.00215 | 78707 | ST | [25 empty] | 2020-03-03 | 7139 | C13 Measurement | Completed
[163 empty] | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 | 0.99743 | 0.00215 | 78724 | ST | [25 empty] | 2020-03-09 | 19994 | AMS Submission Results | Completed
[163 empty] | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 1 | 0.00219 | 78725 | EA | [25 empty] | 2020-03-10 | 23891 | Close Job | Completed
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 | 1.0037 | 0.00217 | 78726 | EA
[163 empty] | 2019-09-09 | 0.00506 | 77728 | 3251 | 1.09 |  |  |  | EA
[7 empty] | 93.7 | 87.4727 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37771 | 37770 | 37770 | 147 | 147 | 147 |  |  | Julia Collins | -990.92 | 0.17 ‰ | 2020-03-12 | 2019-09-14 | 1 | 180 |  |  |  |  | 4.87‰ | 4.87‰ | 2.00‰ | 2.00‰ | 4.87 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19.0042 | 0 | 0.8165 | -990.92 ‰ | -990.92 ‰ | 0.17 | 0.17 | -990.81 | 0.17 ‰ | -990.81 | 0.17 ‰ |  |  | 2020-03-13 | Jocelyn Turnbull | 8:56 | 0.0090774 | 0.0090774 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.00908 | 0.00017 | 20076 | 2020-03-13 | 15:58 | albert zondervan | 214586 |  |  |  |  | 45% | 5 | 16/03/20 | 09:52:51 | Margaret Norris | 0 |  | 0.91 | 0.02% | 1 | 5 | 4 | ... | 0.00873 | 0.00016 |  |  |  |  | 0.98774 | 0.98774 | Yes | Yes | [5 empty] | 0.00873 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00908 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00873 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-11 | Julia Collins | 14:51 |  | Completed |  |  |  |  | 78893 | 2000000 |  |  | 78833 78834 78835 78836 78854 78872 | 78833~78834~78835~78836~78854~78872 | 3278 |  | 3278 3278 3278 3278 3278 3278 | 2020-03-13 | 17:35 | albert zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0005 | 0.00174 | 78833 |  | XCAMS | ST | 0% |  | TW3278 measurement: * small & large samples * XCAMS: no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 1x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is 2-sigma younger than consensus * 1x FIRI-I, small: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks large organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks, all Kauri Renton Rd small organics via AAA: 5x unknowns + 1x secondaries + 3x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheels selected also, to assist with determining d, m parameters: TW3253 d = 1.8 ug C m = 0.77 ug C, derived from 4x Queets A blanks + 3x Kauri Renton Road blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-05-05 | Water |  | JOB-214586 |  |  | 69344 | Normal | 14047/11 |  |  | 2020-03-10 | 4535 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99952 | 0.00178 | 78834 | [26 empty] | 2020-03-11 | 18903 | Graphite | Completed
[168 empty] | 0.99963 | 0.00176 | 78835 | [26 empty] | 2020-03-12 | 7162 | C13 Measurement | Completed
[168 empty] | 1.0008 | 0.00175 | 78836 | [26 empty] | 2020-03-13 | 20076 | AMS Submission Results | Completed
[168 empty] | 0.99825 | 0.00184 | 78854 | [26 empty] | 2020-03-16 | 23963 | Close Job | Completed
[168 empty] | 1.0013 | 0.00183 | 78872
[5 empty] | 4.7307e-05 | 0.0054299 | 92.5 | 88.4163 | 0.0698 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7052 | 7051 | 7051 | 39 | 39 | 43 | 7012.2 | 7091 | Julia Collins | -584.31 | 2.04 ‰ | 2020-03-12 | 2019-09-14 | 1 | 180 |  |  |  |  | -5.85‰ | -5.85‰ | 2.00‰ | 2.00‰ | -5.85 ‰ | 2.00 ‰ | AMS | -8.37‰ | 0.2‰ | C13 | No | -19 | -19.0042 | 0 | 0.8165 | -584.31 ‰ | -584.31 ‰ | 2.04 | 2.24 | -579.18 | 2.06 ‰ | -579.18 | 2.06 ‰ |  |  | 2020-03-13 | Jocelyn Turnbull | 8:57 | 0.41569 | 0.41569 | 0.00204 | 0.00224 | 0.00197 | 0.00197 | 1.8 |  |  | 0.41917 | 0.00129 | 20090 | 2020-03-13 | 15:58 | albert zondervan | 214584 | 0.00558 | 0.0056 | 0.00251 | 0.0025 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.57 | 0.2% | 19 | 5 | 6 | ... | 0.40313 | 0.00112 |  |  |  |  | 0.98774 | 0.98774 | Yes | Yes |  | result is 2-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 43.3411456939455567. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00558 | 0.00522 | 0.00156 | 0.39978 | 0.40‰ | 0.00189 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.41569 | 0 | 0.00204 | 0 | 0 | 0.00558 | 0 | 0.00251 | 0 | 0.39978 | 0 | 0.00189 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-11 | Julia Collins | 14:51 |  | Completed |  |  |  |  | 78892 | 2000000 | 78893 78782 78739 78621 78305 77973 77838 | 78893~78782~78739~78621~78305~77973~77838 | 78833 78834 78835 78836 78854 78872 | 78833~78834~78835~78836~78854~78872 | 3278 | 3278 3275 3274 3272 3264 3254 3253 | 3278 3278 3278 3278 3278 3278 | 2020-03-13 | 17:35 | albert zondervan |  | 1.01 mg | 0.13 | 1.01 | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 1.0005 | 0.00174 | 78833 | ST | XCAMS | ST | 0% |  | TW3278 measurement: * small & large samples * XCAMS: no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 1x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is 2-sigma younger than consensus * 1x FIRI-I, small: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks large organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks, all Kauri Renton Rd small organics via AAA: 5x unknowns + 1x secondaries + 3x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheels selected also, to assist with determining d, m parameters: TW3253 d = 1.8 ug C m = 0.77 ug C, derived from 4x Queets A blanks + 3x Kauri Renton Road blanks unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-05-05 | Water |  | JOB-214584 |  |  | 69345 | Normal | 14047/12 |  |  | 2020-03-10 | 4534 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 0.99952 | 0.00178 | 78834 | ST | [25 empty] | 2020-03-10 | 2751 | Gas Splitting | Completed
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 0.99963 | 0.00176 | 78835 | ST | [25 empty] | 2020-03-10 | 4711 | Choose Gas Splitting | Completed
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 1.0008 | 0.00175 | 78836 | EA | [25 empty] | 2020-03-11 | 18902 | Graphite | Completed
[163 empty] | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 | 0.99825 | 0.00184 | 78854 | EA | [25 empty] | 2020-03-12 | 7161 | C13 Measurement | Completed
[163 empty] | 2019-09-27 | 0.00409 | 77973 | 3254 | 0.96 | 1.0013 | 0.00183 | 78872 | EA | [25 empty] | 2020-03-13 | 20090 | AMS Submission Results | Completed
[163 empty] | 2019-09-20 | 0.00587 | 77838 | 3253 | 1.02 |  |  |  | ST | [25 empty] | 2020-03-16 | 23962 | Close Job | Completed
[7 empty] | 91.7 | 87.7883 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46088 | 46087 | 46087 | 207 | 207 | 207 |  |  | Margaret Norris | -996.78 | 0.08 ‰ | 2020-05-22 | 2019-11-24 | 64 | 180 |  |  |  |  | 2.94‰ | 2.94‰ | 2.00‰ | 2.00‰ | 2.94 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0012 | 0 | 0.89443 | -996.78 ‰ | -996.78 ‰ | 0.08 | 0.08 | -996.74 | 0.08 ‰ | -996.74 | 0.08 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 12:02 | 0.0032234 | 0.0032234 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.7 |  |  | 0.00322 | 8e-05 | 20156 | 2020-05-26 | 15:28 | albert zondervan | 214580 |  |  |  |  | 45% | 5 | 31/05/20 | 12:01:38 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 1 | 5 | 7 | ... | 0.0031 | 8e-05 |  |  |  |  | 0.98939 | 0.98939 | Yes | Yes | [5 empty] | 0.0031 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00322 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0031 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-24 | Margaret Norris | 13:07 |  | Completed |  |  |  |  | 78922 | 2000000 |  |  | 78934 78935 78943 78946 78949 | 78934~78935~78943~78946~78949 | 3280 |  | 3280 3280 3280 3280 3280 | 2020-05-26 | 18:11 | albert zondervan |  | 0.66 mg |  | 0.66 | [5 empty] | 0.99954 | 0.00144 | 78934 |  | XCAMS | EA | 0% |  | TW3280 measurement: * wheel stored under vacuum of source since before lockdown * first measurement after lockdown and subsequent Pelletron service * large samples * 3x ramped pyrolysis samples: results (to be) worked up separately by Cathy * quality of results was not affected by noteworthy deviations from regular behaviour: small, temporary drift in terminal voltage, which was resolved through more conditioning some variability amongst the OxIs in terms of 13C/12C ratio CalAMS error adjustments: * zero system error after rejection of one outlier OxI * global excess variability: factor 1.03 primaries: * OxI cathodes: 6x large: one outlier rejected (TP78954, 0.5 mgC) secondaries: * 1x FIRI-I: result is consistent with consensus value * 1x IAEA-C2: result is consistent with consensus value blanks: * two groups identified to blank-correct: organics, via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks waters: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-05-04 | Water |  | JOB-214580 |  | no CO2 in the bottle, no C13 result reported | 69425 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99757 | 0.00149 | 78935 |  | XCAMS | [24 empty] | 2020-03-18 | 4540 | Water CO2 Evolution | Completed
[168 empty] | 1.0002 | 0.00148 | 78943 | [26 empty] | 2020-03-18 | 2756 | Gas Splitting | Completed
[168 empty] | 1.0016 | 0.00143 | 78946 | [26 empty] | 2020-03-18 | 4717 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00143 | 78949 | [26 empty] | 2020-03-19 | 18932 | Graphite | Completed
[197 empty] | 2020-03-23 | 7165 | C13 Measurement | Cancelled
[197 empty] | 2020-05-31 | 20156 | AMS Submission Results | Completed
[197 empty] | 2020-06-02 | 24117 | Close Job | Completed
[5 empty] | 4.6616e-05 | 0.0048524 | 93.8 | 88.6847 | 0.0687 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7175 | 7175 | 7175 | 37 | 37 | 41 | 7138.5 | 7212.5 | Margaret Norris | -590.67 | 1.89 ‰ | 2020-05-22 | 2019-11-24 | 64 | 180 |  |  |  |  | -7.20‰ | -7.20‰ | 2.00‰ | 2.00‰ | -7.20 ‰ | 2.00 ‰ | AMS | -8.26‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.89443 | -590.67 ‰ | -590.67 ‰ | 1.89 | 2.1 | -585.62 | 1.91 ‰ | -585.62 | 1.91 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 12:08 | 0.40933 | 0.40933 | 0.00189 | 0.0021 | 0.00181 | 0.00181 | 3.1 |  |  | 0.41272 | 0.0011 | 20166 | 2020-05-26 | 15:28 | albert zondervan | 214581 | 0.00537 | 0.0054 | 0.00242 | 0.0024 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 40.93 | 0.19% | 16 | 5 | 4 | ... | 0.39692 | 0.00092 |  |  |  |  | 0.98939 | 0.98939 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 41.1936375521135605. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00537 | 0.00467 | 0.00191 | 0.39366 | 0.39‰ | 0.00174 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.40933 | 0 | 0.00189 | 0 | 0 | 0.00537 | 0 | 0.00242 | 0 | 0.39366 | 0 | 0.00174 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-24 | Margaret Norris | 13:07 |  | Completed |  |  |  |  | 78923 | 2000000 | 78922 78893 78782 78739 78621 78305 | 78922~78893~78782~78739~78621~78305 | 78934 78935 78943 78946 78949 | 78934~78935~78943~78946~78949 | 3280 | 3280 3278 3275 3274 3272 3264 | 3280 3280 3280 3280 3280 | 2020-05-26 | 18:11 | albert zondervan |  | 0.65 mg | 0.13 | 0.65 | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 0.99954 | 0.00144 | 78934 | ST | XCAMS | EA | 0% |  | TW3280 measurement: * wheel stored under vacuum of source since before lockdown * first measurement after lockdown and subsequent Pelletron service * large samples * 3x ramped pyrolysis samples: results (to be) worked up separately by Cathy * quality of results was not affected by noteworthy deviations from regular behaviour: small, temporary drift in terminal voltage, which was resolved through more conditioning some variability amongst the OxIs in terms of 13C/12C ratio CalAMS error adjustments: * zero system error after rejection of one outlier OxI * global excess variability: factor 1.03 primaries: * OxI cathodes: 6x large: one outlier rejected (TP78954, 0.5 mgC) secondaries: * 1x FIRI-I: result is consistent with consensus value * 1x IAEA-C2: result is consistent with consensus value blanks: * two groups identified to blank-correct: organics, via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks waters: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-05-04 | Water | Not Complete | JOB-214581 |  |  | 69435 | Normal | 14047/12 |  |  | 2020-03-18 | 4541 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 0.99757 | 0.00149 | 78935 | EA | XCAMS | [24 empty] | 2020-03-19 | 18933 | Graphite | Completed
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 1.0002 | 0.00148 | 78943 | EA | [25 empty] | 2020-03-23 | 7166 | C13 Measurement | Completed
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 1.0016 | 0.00143 | 78946 | EA | [25 empty] | 2020-05-31 | 20166 | AMS Submission Results | Completed
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 1.0009 | 0.00143 | 78949 | ST | [25 empty] | 2020-06-02 | 24118 | Close Job | Completed
[163 empty] | 2019-11-27 | 0.00488 | 78305 | 3264 | 1.25 |  |  |  | EA
[7 empty] | 99.8 | 87.7612 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42918 | 42917 | 42917 | 175 | 174 | 174 |  |  | jenny dahl | -995.22 | 0.10 ‰ | 2020-05-24 | 2019-11-26 | 12 | 180 |  |  |  |  | 1.35‰ | 1.35‰ | 2.00‰ | 2.00‰ | 1.35 ‰ | 2.00 ‰ | AMS | 1.83‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.89443 | -995.22 ‰ | -995.22 ‰ | 0.1 | 0.1 | -995.16 | 0.11 ‰ | -995.16 | 0.11 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 12:31 | 0.0047831 | 0.0047831 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.2 |  |  | 0.00478 | 0.0001 | 20192 | 2020-05-27 | 11:17 | albert zondervan | 214570 |  |  |  |  | 45% | 5 | 31/05/20 | 12:30:35 | Jocelyn Turnbull | 0 |  | 0.48 | 0.01% | 37 | 5 | 7 | ... | 0.0046 | 0.0001 |  |  |  |  | 0.98752 | 0.98752 | Yes | Yes | [5 empty] | 0.0046 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00478 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.0046 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-22 | jenny dahl | 11:53 |  | Completed |  |  |  |  | 78979 | 2000000 |  |  | 78959 78960 78961 79068 79070 | 78959~78960~78961~79068~79070 | 3281 |  | 3281 3281 3281 3281 3281 | 2020-05-27 | 11:24 | albert zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99882 | 0.0014 | 78959 |  | XCAMS | ST | 0% |  | TW3281 measurement: * large samples CalAMS error adjustments: * zero system error after rejection of one outlier OxI primaries: * OxI cathodes: 5x accepted 1x rejected (TP78962) because of 1.4% low in rts there is circumstantial evidence that cause is leak in manifold secondaries: * 1x IAEA-C2: result is 3-sigma older than consensus value blanks: * one groups identified to blank-correct: carbonate: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 1.8 | RRL | [5 empty] | Background Inorganic |  | 2020-04-30 | Carbonate |  | JOB-214570 |  | C13 is lighter than range - note of high non con during evolution | 69461 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00143 | 78960 | [26 empty] | 2020-05-01 | 4561 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0003 | 0.00141 | 78961 | [26 empty] | 2020-05-01 | 2777 | Gas Splitting | Completed
[168 empty] | 0.99817 | 0.00143 | 79068 | [26 empty] | 2020-05-01 | 4738 | Choose Gas Splitting | Completed
[168 empty] | 1.0021 | 0.00143 | 79070 | [26 empty] | 2020-05-06 | 7188 | C13 Measurement | Completed
[197 empty] | 2020-05-12 | 18989 | Graphite | Completed
[197 empty] | 2020-05-31 | 20192 | AMS Submission Results | Completed
[197 empty] | 2020-06-02 | 24116 | Close Job | Completed
[7 empty] | 90.2 | 87.977 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42814 | 42813 | 42813 | 155 | 155 | 155 |  |  | Margaret Norris | -995.15 | 0.09 ‰ | 2020-05-27 | 2019-11-29 | 14 | 180 |  |  |  |  | 0.85‰ | 0.85‰ | 2.00‰ | 2.00‰ | 0.85 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -995.15 ‰ | -995.15 ‰ | 0.09 | 0.09 | -995.09 | 0.09 ‰ | -995.09 | 0.09 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:06 | 0.0048454 | 0.0048454 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.1 |  |  | 0.00485 | 9e-05 | 20239 | 2020-05-29 | 9:36 | albert zondervan | 214961 |  |  |  |  | 45% | 5 | 31/05/20 | 13:05:36 | Jocelyn Turnbull | 0 |  | 0.48 | 0.01% | 37 | 5 | 5 | ... | 0.00466 | 9e-05 |  |  |  |  | 0.98946 | 0.98946 | Yes | Yes | [5 empty] | 0.00466 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00485 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00466 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-22 | Margaret Norris | 14:31 |  | Completed |  |  |  |  | 78998 | 2000000 |  |  | 79069 79071 79072 79091 79092 79098 | 79069~79071~79072~79091~79092~79098 | 3282 |  | 3282 3282 3282 3282 3282 3282 | 2020-05-29 | 9:46 | albert zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 1.0028 | 0.00143 | 79069 |  | XCAMS | ST | 0% |  | TW3282 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * two groups identified to blank-correct: carbonate: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks water: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-06-24 | Carbonate | Not Complete | JOB-214961 |  |  | 69508 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99972 | 0.00142 | 79071 | [26 empty] | 2020-05-12 | 4604 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99841 | 0.00143 | 79072 | [26 empty] | 2020-05-13 | 19008 | Graphite | Completed
[168 empty] | 0.99967 | 0.00144 | 79091 | [26 empty] | 2020-05-18 | 7227 | C13 Measurement | Completed
[168 empty] | 0.99857 | 0.00142 | 79092 | [26 empty] | 2020-05-31 | 20239 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.00143 | 79098 | [26 empty] | 2020-06-02 | 24134 | Close Job | Completed
[7 empty] | 89.2 | 87.6999 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37616 | 37615 | 37615 | 117 | 117 | 117 |  |  | Margaret Norris | -990.75 | 0.14 ‰ | 2020-05-27 | 2019-11-29 | 8 | 180 |  |  |  |  | 4.01‰ | 4.01‰ | 2.00‰ | 2.00‰ | 4.01 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -990.75 ‰ | -990.75 ‰ | 0.14 | 0.14 | -990.63 | 0.14 ‰ | -990.63 | 0.14 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:20 | 0.0092542 | 0.0092542 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2 |  |  | 0.00925 | 0.00014 | 20260 | 2020-05-29 | 9:36 | albert zondervan | 214869 |  |  |  |  | 45% | 5 | 31/05/20 | 13:20:27 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 36 | 5 | 4 | ... | 0.0089 | 0.00013 |  |  |  |  | 0.98946 | 0.98946 | Yes | Yes | [5 empty] | 0.0089 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00925 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0089 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-22 | Margaret Norris | 14:31 |  | Completed |  |  |  |  | 79054 | 2000000 |  |  | 79069 79071 79072 79091 79092 79098 | 79069~79071~79072~79091~79092~79098 | 3282 |  | 3282 3282 3282 3282 3282 3282 | 2020-05-29 | 9:46 | albert zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0028 | 0.00143 | 79069 |  | XCAMS | ST | 0% |  | TW3282 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * two groups identified to blank-correct: carbonate: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks water: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-05-14 | Water | Complete | JOB-214869 |  |  | 69529 | Normal | 14047/11 |  |  | 2020-05-18 | 4620 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99972 | 0.00142 | 79071 | [26 empty] | 2020-05-19 | 19064 | Graphite | Completed
[168 empty] | 0.99841 | 0.00143 | 79072 | [26 empty] | 2020-05-21 | 7245 | C13 Measurement | Completed
[168 empty] | 0.99967 | 0.00144 | 79091 | [26 empty] | 2020-05-31 | 20260 | AMS Submission Results | Completed
[168 empty] | 0.99857 | 0.00142 | 79092 | [26 empty] | 2020-06-02 | 24122 | Close Job | Completed
[168 empty] | 1.0008 | 0.00143 | 79098
[5 empty] | 4.8655e-05 | 0.0053741 | 87.7 | 88.6812 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7107 | 7107 | 7107 | 40 | 39 | 43 | 7067.6 | 7146.9 | Margaret Norris | -587.18 | 2.04 ‰ | 2020-05-27 | 2019-11-29 | 8 | 180 |  |  |  |  | -7.10‰ | -7.10‰ | 2.00‰ | 2.00‰ | -7.10 ‰ | 2.00 ‰ | AMS | -7.95‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -587.18 ‰ | -587.18 ‰ | 2.04 | 2.24 | -582.08 | 2.06 ‰ | -582.08 | 2.06 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:21 | 0.41282 | 0.41282 | 0.00204 | 0.00224 | 0.00197 | 0.00197 | 2.6 |  |  | 0.4166 | 0.00109 | 20261 | 2020-05-29 | 9:36 | albert zondervan | 214870 | 0.00604 | 0.006 | 0.00272 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.28 | 0.2% | 33 | 5 | 4 | ... | 0.40066 | 0.00091 |  |  |  |  | 0.98946 | 0.98946 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 43.5726621547934289. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00604 | 0.00517 | 0.00236 | 0.39702 | 0.40‰ | 0.00189 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.41282 | 0 | 0.00204 | 0 | 0 | 0.00604 | 0 | 0.00272 | 0 | 0.39702 | 0 | 0.00189 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-22 | Margaret Norris | 14:31 |  | Completed |  |  |  |  | 79055 | 2000000 | 79054 78922 78893 78782 78739 78621 | 79054~78922~78893~78782~78739~78621 | 79069 79071 79072 79091 79092 79098 | 79069~79071~79072~79091~79092~79098 | 3282 | 3282 3280 3278 3275 3274 3272 | 3282 3282 3282 3282 3282 3282 | 2020-05-29 | 9:46 | albert zondervan |  | 0.74 mg | 0.13 | 0.74 | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 1.0028 | 0.00143 | 79069 | EA | XCAMS | ST | 0% |  | TW3282 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * two groups identified to blank-correct: carbonate: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks water: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-05-14 | Water | Complete | JOB-214870 |  |  | 69530 | Normal | 14047/12 |  |  | 2020-05-18 | 4621 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 0.99972 | 0.00142 | 79071 | EA | [25 empty] | 2020-05-19 | 19065 | Graphite | Completed
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 0.99841 | 0.00143 | 79072 | EA | [25 empty] | 2020-05-21 | 7246 | C13 Measurement | Completed
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 0.99967 | 0.00144 | 79091 | ST | [25 empty] | 2020-05-31 | 20261 | AMS Submission Results | Completed
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 0.99857 | 0.00142 | 79092 | EA | [25 empty] | 2020-06-02 | 24123 | Close Job | Completed
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 | 1.0008 | 0.00143 | 79098
ST
[7 empty] | 89.9 | 87.7857 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40664 | 40663 | 40663 | 198 | 197 | 197 |  |  | Julia Collins | -993.67 | 0.16 ‰ | 2020-05-29 | 2019-12-01 | 7 | 180 |  |  |  |  | 3.74‰ | 3.74‰ | 2.00‰ | 2.00‰ | 3.74 ‰ | 2.00 ‰ | AMS | 2.27‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -993.67 ‰ | -993.67 ‰ | 0.16 | 0.16 | -993.59 | 0.16 ‰ | -993.59 | 0.16 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:46 | 0.0063323 | 0.0063323 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.7 |  |  | 0.00633 | 0.00016 | 20279 | 2020-05-31 | 12:25 | albert zondervan | 213968 |  |  |  |  | 45% | 5 | 31/05/20 | 13:46:32 | Jocelyn Turnbull | 0 |  | 0.63 | 0.02% | 1 | 5 | 4 | ... | 0.00609 | 0.00015 |  |  |  |  | 0.99015 | 0.99015 | Yes | Yes | [5 empty] | 0.00609 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00633 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00609 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-28 | Julia Collins | 11:15 |  | Completed |  |  |  |  | 79099 | 2000000 |  |  | 79083 79085 79089 79139 79140 79141 | 79083~79085~79089~79139~79140~79141 | 3284 |  | 3284 3284 3284 3284 3284 3284 | 2020-05-31 | 12:57 | albert zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.9998 | 0.00162 | 79083 |  | XCAMS | EA | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks wood via CE, Rafter method: 3x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-01-16 | Water | Complete | JOB-213968 |  |  | 69548 | Normal | 14047/11 |  |  | 2020-05-20 | 4633 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99872 | 0.00162 | 79085 | [26 empty] | 2020-05-22 | 19109 | Graphite | Completed
[168 empty] | 1.0019 | 0.00161 | 79089 | [26 empty] | 2020-05-25 | 7260 | C13 Measurement | Completed
[168 empty] | 1.0005 | 0.00173 | 79139 | [26 empty] | 2020-05-31 | 20279 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00164 | 79140 | [26 empty] | 2020-06-02 | 24098 | Close Job | Completed
[168 empty] | 0.99777 | 0.00165 | 79141
[5 empty] | 4.6447e-05 | 0.0054591 | 93.8 | 88.5922 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7186 | 7186 | 7186 | 41 | 41 | 45 | 7144.8 | 7227.5 | Julia Collins | -591.22 | 2.11 ‰ | 2020-05-29 | 2019-12-01 | 7 | 180 |  |  |  |  | -5.40‰ | -5.40‰ | 2.00‰ | 2.00‰ | -5.40 ‰ | 2.00 ‰ | AMS | -8.13‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -591.22 ‰ | -591.22 ‰ | 2.11 | 2.3 | -586.17 | 2.13 ‰ | -586.17 | 2.13 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 14:12 | 0.40878 | 0.40878 | 0.00211 | 0.0023 | 0.00204 | 0.00204 | 1.9 |  |  | 0.4126 | 0.00121 | 20317 | 2020-05-31 | 12:25 | albert zondervan | 215046 | 0.00605 | 0.0061 | 0.00272 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 40.88 | 0.21% | 22 | 5 | 4 | ... | 0.39681 | 0.00104 |  |  |  |  | 0.99015 | 0.99015 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 45.167483695497191. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00605 | 0.00525 | 0.00215 | 0.39314 | 0.39‰ | 0.00196 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.40878 | 0 | 0.00211 | 0 | 0 | 0.00605 | 0 | 0.00272 | 0 | 0.39314 | 0 | 0.00196 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-28 | Julia Collins | 11:15 |  | Completed |  |  |  |  | 79104 | 2000000 | 79099 79054 78922 78893 78782 78739 78621 | 79099~79054~78922~78893~78782~78739~78621 | 79083 79085 79089 79139 79140 79141 | 79083~79085~79089~79139~79140~79141 | 3284 | 3284 3282 3280 3278 3275 3274 3272 | 3284 3284 3284 3284 3284 3284 | 2020-05-31 | 12:57 | albert zondervan |  | 0.95 mg | 0.13 | 0.95 | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 0.9998 | 0.00162 | 79083 | EA | XCAMS | EA | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks wood via CE, Rafter method: 3x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-07-01 | Water | Complete | JOB-215046 |  |  | 69586 | Normal | 14047/12 |  |  | 2020-05-21 | 4641 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 0.99872 | 0.00162 | 79085 | EA | [25 empty] | 2020-05-22 | 19114 | Graphite | Completed
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 1.0019 | 0.00161 | 79089 | EA | [25 empty] | 2020-05-25 | 7269 | C13 Measurement | Completed
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 1.0005 | 0.00173 | 79139 | ST | [25 empty] | 2020-05-31 | 20317 | AMS Submission Results | Completed
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 1.0013 | 0.00164 | 79140 | EA | [25 empty] | 2020-06-02 | 24135 | Close Job | Completed
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 | 0.99777 | 0.00165 | 79141
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 |  |  |  | ST
EA
[5 empty] | 4.3034e-05 | 0.0053413 | 99.2 | 88.759 | 0.069 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7152 | 7151 | 7151 | 41 | 40 | 44 | 7111.3 | 7192.5 | Julia Collins | -589.47 | 2.08 ‰ | 2020-06-05 | 2019-12-08 | 10 | 180 |  |  |  |  | -9.66‰ | -9.66‰ | 2.00‰ | 2.00‰ | -9.66 ‰ | 2.00 ‰ | AMS | -8.50‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -589.47 ‰ | -589.47 ‰ | 2.08 | 2.27 | -584.4 | 2.10 ‰ | -584.4 | 2.10 ‰ |  |  | 2020-06-06 | Jocelyn Turnbull | 9:28 | 0.41053 | 0.41053 | 0.00208 | 0.00227 | 0.00201 | 0.00201 | 2.8 |  |  | 0.41421 | 0.00124 | 20379 | 2020-06-06 | 18:32 | albert zondervan | 214981 | 0.00585 | 0.0059 | 0.00263 | 0.0026 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.05 | 0.21% | 20 | 5 | 4 | ... | 0.39836 | 0.00107 |  |  |  |  | 0.98777 | 0.98777 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.4766480753458344. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00585 | 0.00514 | 0.00207 | 0.39482 | 0.39‰ | 0.00193 | 0.00‰ | Yes | 1 | 0.99999 | 0.00069 | 0.00069 | 0.41053 | 0 | 0.00208 | 0 | 0 | 0.00585 | 0 | 0.00263 | 0 | 0.39482 | 0 | 0.00193 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-04 | Julia Collins | 14:02 |  | Completed |  |  |  |  | 79121 | 2000000 | 79103 79099 79054 78922 78893 78782 78739 78621 | 79103~79099~79054~78922~78893~78782~78739~78621 | 79176 79177 79178 79179 79195 79196 | 79176~79177~79178~79179~79195~79196 | 3286 | 3286 3284 3282 3280 3278 3275 3274 3272 | 3286 3286 3286 3286 3286 3286 | 2020-06-06 | 18:43 | albert zondervan |  | 0.68 mg | 0.13 | 0.68 | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 0.99861 | 0.00167 | 79176 | EA | XCAMS | ST | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | -8.5 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-06-24 | Water |  | JOB-214981 |  | C13 is just within range- please double check results at AMS | 69608 | Normal | 14047/12 |  |  | 2020-05-22 | 4642 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 0.9978 | 0.00169 | 79177 | EA | [25 empty] | 2020-05-26 | 19131 | Graphite | Completed
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 0.99889 | 0.00168 | 79178 | EA | [25 empty] | 2020-05-27 | 7283 | C13 Measurement | Completed
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 1.0023 | 0.00174 | 79179 | ST | [25 empty] | 2020-06-06 | 20379 | AMS Submission Results | Completed
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 1.0018 | 0.00169 | 79195 | EA | [25 empty] | 2020-06-08 | 24274 | Close Job | Completed
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 | 1.0007 | 0.00168 | 79196
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 |  |  |  | ST
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 |  |  |  | EA
ST
[7 empty] | 97.3 | 87.7964 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43184 | 43184 | 43184 | 199 | 198 | 198 |  |  | Julia Collins | -995.37 | 0.11 ‰ | 2020-06-05 | 2019-12-08 | 14 | 180 |  |  |  |  | 1.20‰ | 1.20‰ | 2.00‰ | 2.00‰ | 1.20 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -995.37 ‰ | -995.37 ‰ | 0.11 | 0.11 | -995.32 | 0.12 ‰ | -995.32 | 0.12 ‰ |  |  | 2020-06-06 | Jocelyn Turnbull | 9:29 | 0.0046271 | 0.0046271 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3.2 |  |  | 0.00463 | 0.00011 | 20380 | 2020-06-06 | 18:32 | albert zondervan | 214983 |  |  |  |  | 45% | 5 | 6/06/20 | 09:29:16 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 1 | 5 | 6 | ... | 0.00445 | 0.00011 |  |  |  |  | 0.98777 | 0.98777 | Yes | Yes | [5 empty] | 0.00445 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 0.99999 | 0.00069 | 0.00069 | 0.00463 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00445 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-04 | Julia Collins | 14:02 |  | Completed |  |  |  |  | 79103 | 2000000 |  |  | 79176 79177 79178 79179 79195 79196 | 79176~79177~79178~79179~79195~79196 | 3286 |  | 3286 3286 3286 3286 3286 3286 | 2020-06-06 | 18:43 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99861 | 0.00167 | 79176 |  | XCAMS | ST | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-06-24 | Water |  | JOB-214983 |  |  | 69609 | Normal | 14047/11 |  |  | 2020-05-21 | 4640 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9978 | 0.00169 | 79177 | [26 empty] | 2020-05-21 | 2815 | Gas Splitting | Completed
[168 empty] | 0.99889 | 0.00168 | 79178 | [26 empty] | 2020-05-21 | 4787 | Choose Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00174 | 79179 | [26 empty] | 2020-05-22 | 19113 | Graphite | Completed
[168 empty] | 1.0018 | 0.00169 | 79195 | [26 empty] | 2020-05-25 | 7268 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00168 | 79196 | [26 empty] | 2020-06-06 | 20380 | AMS Submission Results | Completed
[197 empty] | 2020-06-08 | 24275 | Close Job | Completed
[5 empty] | 4.1269e-05 | 0.0055388 | 95.9 | 88.534 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7162 | 7162 | 7162 | 41 | 40 | 44 | 7121.8 | 7203.2 | Margaret Norris | -590.01 | 2.08 ‰ | 2020-06-12 | 2019-12-15 | 9 | 180 |  |  |  |  | -7.39‰ | -7.39‰ | 2.00‰ | 2.00‰ | -7.39 ‰ | 2.00 ‰ | AMS | -8.35‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -590.01 ‰ | -590.01 ‰ | 2.08 | 2.27 | -584.95 | 2.10 ‰ | -584.95 | 2.10 ‰ |  |  | 2020-06-14 | Jocelyn Turnbull | 16:54 | 0.40999 | 0.40999 | 0.00208 | 0.00227 | 0.00201 | 0.00201 | 2.3 |  |  | 0.41379 | 0.00117 | 20442 | 2020-06-15 | 8:47 | albert zondervan | 215222 | 0.00603 | 0.006 | 0.00271 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41 | 0.21% | 29 | 5 | 6 | ... | 0.39795 | 0.001 |  |  |  |  | 0.98752 | 0.98752 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.5223757861299378. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00603 | 0.00533 | 0.00201 | 0.3943 | 0.39‰ | 0.00193 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.40999 | 0 | 0.00208 | 0 | 0 | 0.00603 | 0 | 0.00271 | 0 | 0.3943 | 0 | 0.00193 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-11 | Margaret Norris | 13:41 |  | Completed |  |  |  |  | 79190 | 2000000 | 79191 79103 79099 79054 78922 78893 78782 78739 78621 | 79191~79103~79099~79054~78922~78893~78782~78739~78621 | 79258 79262 79265 79277 79279 79282 | 79258~79262~79265~79277~79279~79282 | 3288 | 3288 3286 3284 3282 3280 3278 3275 3274 3272 | 3288 3288 3288 3288 3288 3288 | 2020-06-15 | 9:07 | albert zondervan |  | 0.91 mg | 0.13 | 0.91 | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 1.0011 | 0.00156 | 79258 | EA | XCAMS | EA | 0% |  | TW3288 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks wood via CE, Rafter method: 6x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-07-22 | Water |  | JOB-215222 |  |  | 69671 | Normal | 14047/12 |  |  | 2020-06-02 | 4661 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 0.9983 | 0.00156 | 79262 | EA | [25 empty] | 2020-06-02 | 2833 | Gas Splitting | Completed
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 1.0008 | 0.00157 | 79265 | EA | [25 empty] | 2020-06-02 | 4816 | Choose Gas Splitting | Completed
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 0.9983 | 0.00157 | 79277 | ST | [25 empty] | 2020-06-03 | 19200 | Graphite | Completed
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 1.0015 | 0.00161 | 79279 | EA | [25 empty] | 2020-06-04 | 7302 | C13 Measurement | Completed
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 | 1 | 0.00155 | 79282 | [26 empty] | 2020-06-14 | 20442 | AMS Submission Results | Completed
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 |  |  |  | ST | [25 empty] | 2020-06-15 | 24349 | Close Job | Completed
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 |  |  |  | EA
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 |  |  |  | ST
EA
[7 empty] | 94.4 | 87.5996 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38991 | 38990 | 38990 | 150 | 149 | 149 |  |  | Margaret Norris | -992.2 | 0.15 ‰ | 2020-06-12 | 2019-12-15 | 9 | 180 |  |  |  |  | 3.19‰ | 3.19‰ | 2.00‰ | 2.00‰ | 3.19 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.20 ‰ | -992.20 ‰ | 0.15 | 0.15 | -992.11 | 0.15 ‰ | -992.11 | 0.15 ‰ |  |  | 2020-06-14 | Jocelyn Turnbull | 16:55 | 0.0077985 | 0.0077985 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.0078 | 0.00015 | 20443 | 2020-06-15 | 8:47 | albert zondervan | 215220 |  |  |  |  | 45% | 5 | 14/06/20 | 16:54:53 | Jocelyn Turnbull | 0 |  | 0.78 | 0.01% | 8 | 5 | 4 | ... | 0.0075 | 0.00014 |  |  |  |  | 0.98752 | 0.98752 | Yes | Yes | [5 empty] | 0.0075 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.0078 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.0075 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-11 | Margaret Norris | 13:41 |  | Completed |  |  |  |  | 79191 | 2000000 |  |  | 79258 79262 79265 79277 79279 79282 | 79258~79262~79265~79277~79279~79282 | 3288 |  | 3288 3288 3288 3288 3288 3288 | 2020-06-15 | 9:07 | albert zondervan |  | 1.11 mg |  | 1.11 | [5 empty] | 1.0011 | 0.00156 | 79258 |  | XCAMS | EA | 0% |  | TW3288 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks wood via CE, Rafter method: 6x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-07-22 | Water |  | JOB-215220 |  |  | 69672 | Normal | 14047/11 |  |  | 2020-06-02 | 4660 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9983 | 0.00156 | 79262 | [26 empty] | 2020-06-03 | 19201 | Graphite | Completed
[168 empty] | 1.0008 | 0.00157 | 79265 | [26 empty] | 2020-06-04 | 7301 | C13 Measurement | Completed
[168 empty] | 0.9983 | 0.00157 | 79277 | [26 empty] | 2020-06-14 | 20443 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00161 | 79279 | [26 empty] | 2020-06-15 | 24348 | Close Job | Completed
[168 empty] | 1 | 0.00155 | 79282
[7 empty] | 94.5 | 87.779 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46011 | 46010 | 46010 | 231 | 230 | 230 |  |  | Julia Collins | -996.75 | 0.09 ‰ | 2020-06-16 | 2019-12-19 | 6 | 180 |  |  |  |  | 3.88‰ | 3.88‰ | 2.00‰ | 2.00‰ | 3.88 ‰ | 2.00 ‰ | AMS | 2.24‰ | 0.2‰ | C13 | No | -19 | -19 | 0.82 | 0.8165 | -996.75 ‰ | -996.75 ‰ | 0.09 | 0.09 | -996.71 | 0.09 ‰ | -996.71 | 0.09 ‰ |  |  | 2020-06-18 | Jocelyn Turnbull | 17:58 | 0.0032546 | 0.0032546 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 3.4 |  |  | 0.00325 | 9e-05 | 20503 | 2020-06-18 | 14:06 | albert zondervan | 215221 |  |  |  |  | 45% | 5 | 22/06/20 | 10:34:38 | albert zondervan | 0 |  | 0.33 | 0.01% | 8 | 5 | 6 | ... | 0.00313 | 9e-05 |  |  |  |  | 0.99022 | 0.99022 | Yes | Yes | [5 empty] | 0.00313 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00325 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00313 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-15 | Julia Collins | 13:35 |  | Completed |  |  |  |  | 79274 | 2000000 |  |  | 79034 79035 79048 79063 79064 79165 | 79034~79035~79048~79063~79064~79165 | 3289 |  | 3289 3289 3289 3289 3289 3289 | 2020-06-22 | 10:30 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99805 | 0.00158 | 79034 |  | XCAMS | FLASK | 0% |  | TW3289 measurement: * large+small samples * sediments, waters CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 1x EA + 5x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks organic via RP: 12x unknowns + 0x secondaries + 3x blanks + 3x OxI with blank-correction calculations outside RLIMS organic via AAA: 2x unknowns (replicates of RP unknowns) + 0x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures * Because a few cathodes gave noticably higher del13C_AMS than del13C_IRMS, these parameters were compared for all samples of which the del13C_IRMS was known. Their difference values correlate with sample mass, with ~6 permil/mgC proportionality in the 0.1-1.1 mgC range. The primary standard OxI graphites in this wheel are confined to a much narrower range: 0.40-0.63 mgC. The IAEA-C2 secondary standard (1.02 mgC) generated a result consistent with its consensus value. While this dependency needs more study (to attribute this to fractionation during graphitization or ion beam transport, it is expected that it is mostly (if not completely) accounted for through our use of 'full fractionation correction' mode 5 in CalAMS. | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-07-22 | Water | Complete | JOB-215221 |  |  | 69731 | Normal | 14047/11 |  |  | 2020-06-09 | 4693 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00168 | 79035 | [26 empty] | 2020-06-09 | 2849 | Gas Splitting | Completed
[168 empty] | 1.0021 | 0.00156 | 79048 | [26 empty] | 2020-06-09 | 4835 | Choose Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00158 | 79063 | [26 empty] | 2020-06-10 | 19284 | Graphite | Completed
[168 empty] | 1.0004 | 0.00157 | 79064 | [26 empty] | 2020-06-15 | 7332 | C13 Measurement | Completed
[168 empty] | 1.0002 | 0.00158 | 79165 | [26 empty] | 2020-06-18 | 20503 | AMS Submission Results | Completed
[197 empty] | 2020-06-22 | 24412 | Close Job | Completed
[5 empty] | 3.776e-05 | 0.0051669 | 91.2 | 88.6533 | 0.069 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7168 | 7168 | 7168 | 40 | 39 | 43 | 7128.5 | 7207.9 | Julia Collins | -590.3 | 2.03 ‰ | 2020-06-16 | 2019-12-19 | 7 | 180 |  |  |  |  | -6.02‰ | -6.02‰ | 2.00‰ | 2.00‰ | -6.02 ‰ | 2.00 ‰ | AMS | -8.26‰ | 0.2‰ | C13 | No | -19 | -19 | 0.82 | 0.8165 | -590.30 ‰ | -590.30 ‰ | 2.03 | 2.23 | -585.24 | 2.05 ‰ | -585.24 | 2.05 ‰ |  |  | 2020-06-18 | Jocelyn Turnbull | 17:59 | 0.4097 | 0.4097 | 0.00203 | 0.00223 | 0.00195 | 0.00195 | 2 |  |  | 0.41332 | 0.00119 | 20504 | 2020-06-18 | 14:06 | albert zondervan | 215223 | 0.00574 | 0.0057 | 0.00258 | 0.0026 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 40.97 | 0.2% | 22 | 5 | 4 | ... | 0.3975 | 0.00102 |  |  |  |  | 0.99022 | 0.99022 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 43.6492873680762574. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00574 | 0.00497 | 0.00211 | 0.39402 | 0.39‰ | 0.00188 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.4097 | 0 | 0.00203 | 0 | 0 | 0.00574 | 0 | 0.00258 | 0 | 0.39402 | 0 | 0.00188 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-15 | Julia Collins | 13:35 |  | Completed |  |  |  |  | 79257 | 2000000 | 79274 79191 79103 79099 79054 78922 78893 78782 78739 78621 | 79274~79191~79103~79099~79054~78922~78893~78782~78739~78621 | 79034 79035 79048 79063 79064 79165 | 79034~79035~79048~79063~79064~79165 | 3289 | 3289 3288 3286 3284 3282 3280 3278 3275 3274 3272 | 3289 3289 3289 3289 3289 3289 | 2020-06-22 | 10:30 | albert zondervan |  | 1.02 mg | 0.13 | 1.02 | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 | 0.99805 | 0.00158 | 79034 | EA | XCAMS | FLASK | 0% |  | TW3289 measurement: * large+small samples * sediments, waters CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 1x EA + 5x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks organic via RP: 12x unknowns + 0x secondaries + 3x blanks + 3x OxI with blank-correction calculations outside RLIMS organic via AAA: 2x unknowns (replicates of RP unknowns) + 0x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures * Because a few cathodes gave noticably higher del13C_AMS than del13C_IRMS, these parameters were compared for all samples of which the del13C_IRMS was known. Their difference values correlate with sample mass, with ~6 permil/mgC proportionality in the 0.1-1.1 mgC range. The primary standard OxI graphites in this wheel are confined to a much narrower range: 0.40-0.63 mgC. The IAEA-C2 secondary standard (1.02 mgC) generated a result consistent with its consensus value. While this dependency needs more study (to attribute this to fractionation during graphitization or ion beam transport, it is expected that it is mostly (if not completely) accounted for through our use of 'full fractionation correction' mode 5 in CalAMS. | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-07-22 | Water | Complete | JOB-215223 |  |  | 69732 | Normal | 14047/12 |  |  | 2020-06-08 | 4682 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 1 | 0.00168 | 79035 | EA | [25 empty] | 2020-06-09 | 19267 | Graphite | Completed
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 1.0021 | 0.00156 | 79048 | EA | [25 empty] | 2020-06-15 | 7327 | C13 Measurement | Completed
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 0.99922 | 0.00158 | 79063 | ST | [25 empty] | 2020-06-18 | 20504 | AMS Submission Results | Completed
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 1.0004 | 0.00157 | 79064 | EA | [25 empty] | 2020-06-22 | 24413 | Close Job | Completed
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 | 1.0002 | 0.00158 | 79165
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 |  |  |  | ST
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 |  |  |  | EA
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 |  |  |  | ST
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 |  |  |  | EA
FLASK
[7 empty] | 99.9 | 87.4138 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45113 | 45113 | 45113 | 207 | 206 | 206 |  |  | Julia Collins | -996.36 | 0.09 ‰ | 2020-06-19 | 2019-12-22 | 3 | 180 |  |  |  |  | 5.26‰ | 5.26‰ | 2.00‰ | 2.00‰ | 5.26 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | -996.36 ‰ | -996.36 ‰ | 0.09 | 0.09 | -996.32 | 0.09 ‰ | -996.32 | 0.09 ‰ |  |  | 2020-06-22 | Jocelyn Turnbull | 12:46 | 0.0036393 | 0.0036393 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00364 | 9e-05 | 20520 | 2020-06-20 | 18:56 | albert zondervan | 215380 |  |  |  |  | 45% | 5 | 22/06/20 | 12:46:17 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 1 | 5 | 5 | ... | 0.0035 | 9e-05 |  |  |  |  | 0.98745 | 0.98745 | Yes | Yes | [5 empty] | 0.0035 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00364 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0035 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-17 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 79317 | 2000000 |  |  | 79309 79310 79311 79312 79313 79319 | 79309~79310~79311~79312~79313~79319 | 3290 |  | 3290 3290 3290 3290 3290 3290 | 2020-06-20 | 19:29 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0032 | 0.00165 | 79309 |  | XCAMS | ST | 0% |  | TW3290 measurement: * large samples * partial wheel CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 6x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * two groups identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-08-04 | Water |  | JOB-215380 |  | new job for fresh water std prep | 69748 | Normal | 14047/11 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00159 | 79310 | [26 empty] | 2020-06-15 | 4711 | Water CO2 Evolution | Completed
[168 empty] | 0.99809 | 0.00158 | 79311 | [26 empty] | 2020-06-16 | 19327 | Graphite | Completed
[168 empty] | 1.0007 | 0.00158 | 79312 | [26 empty] | 2020-06-22 | 7351 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00161 | 79313 | [26 empty] | 2020-06-22 | 20520 | AMS Submission Results | Completed
[168 empty] | 0.99954 | 0.00159 | 79319 | [26 empty] | 2020-06-23 | 24422 | Close Job | Completed
[5 empty] | 3.5017e-05 | 0.004953 | 101.1 | 88.7114 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7099 | 7098 | 7098 | 38 | 38 | 42 | 7060.7 | 7137.2 | Julia Collins | -586.75 | 1.97 ‰ | 2020-06-19 | 2019-12-22 | 3 | 180 |  |  |  |  | -9.45‰ | -9.45‰ | 2.00‰ | 2.00‰ | -9.45 ‰ | 2.00 ‰ | AMS | -8.08‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | -586.75 ‰ | -586.75 ‰ | 1.97 | 2.18 | -581.65 | 1.99 ‰ | -581.65 | 1.99 ‰ |  |  | 2020-06-22 | Jocelyn Turnbull | 12:51 | 0.41325 | 0.41325 | 0.00197 | 0.00218 | 0.00189 | 0.00189 | 3 |  |  | 0.41672 | 0.00118 | 20529 | 2020-06-20 | 18:56 | albert zondervan | 215279 | 0.00554 | 0.0055 | 0.00249 | 0.0025 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.32 | 0.2% | 22 | 5 | 4 | ... | 0.40077 | 0.00101 |  |  |  |  | 0.98745 | 0.98745 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 42.3020503696227534. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00554 | 0.00476 | 0.00211 | 0.39743 | 0.40‰ | 0.00182 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41325 | 0 | 0.00197 | 0 | 0 | 0.00554 | 0 | 0.00249 | 0 | 0.39743 | 0 | 0.00182 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-17 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 79318 | 2000000 | 79317 79274 79191 79103 79099 79054 78922 78893 78782 78739 78621 | 79317~79274~79191~79103~79099~79054~78922~78893~78782~78739~78621 | 79309 79310 79311 79312 79313 79319 | 79309~79310~79311~79312~79313~79319 | 3290 | 3290 3289 3288 3286 3284 3282 3280 3278 3275 3274 3272 | 3290 3290 3290 3290 3290 3290 | 2020-06-20 | 19:29 | albert zondervan |  | 0.66 mg | 0.13 | 0.66 | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04 | 1.0032 | 0.00165 | 79309 | EA | XCAMS | ST | 0% |  | TW3290 measurement: * large samples * partial wheel CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 6x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * two groups identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-07-28 | Water | Complete | JOB-215279 |  |  | 69757 | Normal | 14047/12 |  |  | 2020-06-15 | 4710 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 | 0.9998 | 0.00159 | 79310 | EA | [25 empty] | 2020-06-16 | 19328 | Graphite | Completed
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 0.99809 | 0.00158 | 79311 | EA | [25 empty] | 2020-06-22 | 7352 | C13 Measurement | Completed
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 1.0007 | 0.00158 | 79312 | ST | [25 empty] | 2020-06-22 | 20529 | AMS Submission Results | Completed
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 0.9989 | 0.00161 | 79313 | EA | [25 empty] | 2020-06-23 | 24420 | Close Job | Completed
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 | 0.99954 | 0.00159 | 79319
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 |  |  |  | ST
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 |  |  |  | EA
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 |  |  |  | ST
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 |  |  |  | EA
[163 empty] | 2020-02-07 | 0.00441 | 78621 | 3272 | 1.23 |  |  |  | FLASK
ST
[7 empty] | 100.8 | 87.6471 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48556 | 48555 | 48555 | 317 | 317 | 317 |  |  | Julia Collins | -997.63 | 0.09 ‰ | 2020-07-09 | 2020-01-11 | 1 | 180 |  |  |  |  | 1.66‰ | 1.66‰ | 2.00‰ | 2.00‰ | 1.66 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -997.63 ‰ | -997.63 ‰ | 0.09 | 0.09 | -997.6 | 0.09 ‰ | -997.6 | 0.09 ‰ |  |  | 2020-07-15 | jenny dahl | 15:18 | 0.0023707 | 0.0023707 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00237 | 9e-05 | 20627 | 2020-07-10 | 12:53 | albert zondervan | 215379 |  |  |  |  | 45% | 5 | 15/07/20 | 15:18:11 | jenny dahl | 0 |  | 0.24 | 0.01% | 20 | 5 | 5 | ... | 0.00228 | 9e-05 |  |  |  |  | 0.98654 | 0.98654 | Yes | Yes | [5 empty] | 0.00228 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00237 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00228 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-07-08 | Julia Collins | 13:53 |  | Completed |  |  |  |  | 79452 | 2000000 |  |  | 79461 79462 79463 79464 79465 79466 | 79461~79462~79463~79464~79465~79466 | 3292 |  | 3292 3292 3292 3292 3292 3292 | 2020-07-10 | 12:57 | albert zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0025 | 0.0019 | 79461 |  | XCAMS | EA | 0% |  | TW3292 measurement: * large carbonate samples CalAMS error adjustments: * factor 1.04 to account for excess global run-to-run variability * factor 1.3 to account for excess cathode #21 (Ox-I) run-to-run variability primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value blanks: * one groups identified to blank-correct: carbonate: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2020-08-04 | Carbonate | Complete | JOB-215379 |  | new job for fresh marble std prep | 69905 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99887 | 0.00189 | 79462 | [26 empty] | 2020-07-07 | 4746 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99855 | 0.00188 | 79463 | [26 empty] | 2020-07-08 | 19462 | Graphite | Completed
[168 empty] | 0.99586 | 0.00252 | 79464 | [26 empty] | 2020-07-09 | 7401 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00188 | 79465 | [26 empty] | 2020-07-15 | 20627 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00189 | 79466 | [26 empty] | 2020-07-15 | 24625 | Close Job | Completed
[7 empty] | 73.7 | 87.7512 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49812 | 49811 | 49811 | 330 | 329 | 329 |  |  | Julia Collins | -997.97 | 0.08 ‰ | 2020-07-24 | 2020-01-26 | 8 | 180 |  |  |  |  | 2.27‰ | 2.27‰ | 2.00‰ | 2.00‰ | 2.27 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -997.97 ‰ | -997.97 ‰ | 0.08 | 0.08 | -997.95 | 0.08 ‰ | -997.95 | 0.08 ‰ |  |  | 2020-07-27 | Jocelyn Turnbull | 10:58 | 0.0020276 | 0.0020276 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.4 |  |  | 0.00203 | 8e-05 | 20742 | 2020-07-27 | 9:41 | albert zondervan | 215437 |  |  |  |  | 45% | 5 | 27/07/20 | 10:58:13 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 22 | 5 | 7 | ... | 0.00195 | 8e-05 |  |  |  |  | 0.98832 | 0.98832 | Yes | Yes | [5 empty] | 0.00195 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00203 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00195 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-07-23 | Julia Collins | 13:35 |  | Completed |  |  |  |  | 79534 | 2000000 |  |  | 79537 79538 79539 79598 79600 79602 | 79537~79538~79539~79598~79600~79602 | 3295 |  | 3295 3295 3295 3295 3295 3295 | 2020-07-27 | 10:32 | albert zondervan |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99938 | 0.00156 | 79537 |  | XCAMS | EA | 0% |  | TW3295 measurement: * large samples CalAMS error adjustments: * factor 1.04 to account for excess global run-to-run variability primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: carbonate: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AA: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks wood via CE-Rafter: 2x unknowns + 0x secondaries + 1x blanks 180 days: capturing 3x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2020-08-10 | Carbonate | Complete | JOB-215437 |  | C13 split taken on carbonate line under same job number | 69985 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.003 | 0.00157 | 79538 | [26 empty] | 2020-07-14 | 4782 | Carbonate CO2 Evolution | Completed
[168 empty] | 1 | 0.00157 | 79539 | [26 empty] | 2020-07-14 | 2866 | Gas Splitting | Completed
[168 empty] | 0.99759 | 0.00156 | 79598 | [26 empty] | 2020-07-14 | 4866 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00157 | 79600 | [26 empty] | 2020-07-16 | 19544 | Graphite | Completed
[168 empty] | 0.99905 | 0.00157 | 79602 | [26 empty] | 2020-07-21 | 7431 | C13 Measurement | Completed
[197 empty] | 2020-07-27 | 20742 | AMS Submission Results | Completed
[197 empty] | 2020-07-28 | 24720 | Close Job | Completed
[7 empty] | 61.5 | 88.0598 | 0.0007 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44020 | 44020 | 44020 | 200 | 200 | 200 |  |  | Julia Collins | -995.83 | 0.10 ‰ | 2020-07-27 | 2020-01-29 | 10 | 180 |  |  |  |  | -3.79‰ | -3.79‰ | 2.00‰ | 2.00‰ | -3.79 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -995.83 ‰ | -995.83 ‰ | 0.1 | 0.1 | -995.78 | 0.11 ‰ | -995.78 | 0.11 ‰ |  |  | 2020-07-29 | Jocelyn Turnbull | 12:25 | 0.0041696 | 0.0041696 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 10.8 |  |  | 0.00417 | 0.0001 | 20785 | 2020-07-29 | 9:24 | albert zondervan | 215571 |  |  |  |  | 45% | 5 | 29/07/20 | 12:25:01 | Jocelyn Turnbull | 0 |  | 0.42 | 0.01% | 13 | 5 | 4 | ... | 0.00401 | 0.0001 |  |  |  |  | 0.9858 | 0.9858 | Yes | Yes | [5 empty] | 0.00401 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00417 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00401 | 0 | 0.0001 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2020-07-27 | Julia Collins | 8:22 |  | Completed |  |  |  |  | 79550 | 2000000 |  |  | 79524 79551 79558 79563 79564 79565 | 79524~79551~79558~79563~79564~79565 | 3296 |  | 3296 3296 3296 3296 3296 3296 | 2020-07-29 | 10:22 | albert zondervan |  | 0.20 mg |  | 0.2 | [5 empty] | 0.99977 | 0.00152 | 79524 |  | XCAMS | ST | 0% |  | TW3296 measurement: * small & large samples * XCAMS: no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large carbonates: 8x unknowns + 0x secondaries + 1x blanks 180 days: capturing 7x blanks large organics via AAA: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks, all Kauri Renton Rd small organics via AAA: 6x unknowns + 1x secondaries + 1x blanks Excel file 'to blank-correct small samples.xlsm': 2x other wheels selected also, to assist with determining d, m parameters: TW3253 and TW3278 d = 1.2 ug C m = 0.77 ug C, derived from 4x Queets A blanks + 3x KRR blanks; rejected outlier KRR blank TP79570 in TW3296 small carbonates: 1x unknowns + 1x secondaries + 1x blanks m = 0.81 ug C, derived from only TP79550 in TW3296 unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-09-02 | Carbonate |  | JOB-215571 |  | SLUSHIE to match size of unknowns~ 0.22mg | 70025 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0014 | 0.00144 | 79551 | [26 empty] | 2020-07-15 | 4790 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0009 | 0.00146 | 79558 | [26 empty] | 2020-07-17 | 19560 | Graphite | Completed
[168 empty] | 0.99979 | 0.00144 | 79563 | [26 empty] | 2020-07-29 | 20785 | AMS Submission Results | Completed
[168 empty] | 0.9992 | 0.00145 | 79564 | [26 empty] | 2020-07-31 | 24751 | Close Job | Completed
[168 empty] | 0.99887 | 0.00143 | 79565
[7 empty] | 83.4 | 87.8182 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48843 | 48842 | 48842 | 256 | 255 | 255 |  |  | Julia Collins | -997.71 | 0.07 ‰ | 2020-07-27 | 2020-01-29 | 11 | 180 |  |  |  |  | -1.05‰ | -1.05‰ | 2.00‰ | 2.00‰ | -1.05 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -997.71 ‰ | -997.71 ‰ | 0.07 | 0.07 | -997.68 | 0.07 ‰ | -997.68 | 0.07 ‰ |  |  | 2020-07-29 | Jocelyn Turnbull | 12:25 | 0.0022875 | 0.0022875 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.2 |  |  | 0.00229 | 7e-05 | 20789 | 2020-07-29 | 9:24 | albert zondervan | 215581 |  |  |  |  | 45% | 5 | 29/07/20 | 12:25:05 | Jocelyn Turnbull | 0 |  | 0.23 | 0.01% | 37 | 5 | 5 | ... | 0.0022 | 7e-05 |  |  |  |  | 0.9858 | 0.9858 | Yes | Yes | [5 empty] | 0.0022 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00229 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0022 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-07-27 | Julia Collins | 8:22 |  | Completed |  |  |  |  | 79536 | 2000000 |  |  | 79524 79551 79558 79563 79564 79565 | 79524~79551~79558~79563~79564~79565 | 3296 |  | 3296 3296 3296 3296 3296 3296 | 2020-07-29 | 10:22 | albert zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99977 | 0.00152 | 79524 |  | XCAMS | ST | 0% |  | TW3296 measurement: * small & large samples * XCAMS: no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large carbonates: 8x unknowns + 0x secondaries + 1x blanks 180 days: capturing 7x blanks large organics via AAA: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks, all Kauri Renton Rd small organics via AAA: 6x unknowns + 1x secondaries + 1x blanks Excel file 'to blank-correct small samples.xlsm': 2x other wheels selected also, to assist with determining d, m parameters: TW3253 and TW3278 d = 1.2 ug C m = 0.77 ug C, derived from 4x Queets A blanks + 3x KRR blanks; rejected outlier KRR blank TP79570 in TW3296 small carbonates: 1x unknowns + 1x secondaries + 1x blanks m = 0.81 ug C, derived from only TP79550 in TW3296 unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2020-09-02 | Carbonate | Not Complete | JOB-215581 |  | regular sized blank | 70026 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0014 | 0.00144 | 79551 | [26 empty] | 2020-07-15 | 4783 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0009 | 0.00146 | 79558 | [26 empty] | 2020-07-16 | 19546 | Graphite | Completed
[168 empty] | 0.99979 | 0.00144 | 79563 | [26 empty] | 2020-07-21 | 7438 | C13 Measurement | Completed
[168 empty] | 0.9992 | 0.00145 | 79564 | [26 empty] | 2020-07-29 | 20789 | AMS Submission Results | Completed
[168 empty] | 0.99887 | 0.00143 | 79565 | [26 empty] | 2020-07-31 | 24754 | Close Job | Completed
[7 empty] | 87.8 | 87.6205 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47720 | 47720 | 47720 | 286 | 285 | 285 |  |  | Julia Collins | -997.37 | 0.09 ‰ | 2020-07-29 | 2020-01-31 | 7 | 180 |  |  |  |  | 3.90‰ | 3.90‰ | 2.00‰ | 2.00‰ | 3.90 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -997.37 ‰ | -997.37 ‰ | 0.09 | 0.09 | -997.34 | 0.09 ‰ | -997.34 | 0.09 ‰ |  |  | 2020-07-31 | Jocelyn Turnbull | 9:56 | 0.0026307 | 0.0026307 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.1 |  |  | 0.00263 | 9e-05 | 20797 | 2020-07-31 | 9:05 | albert zondervan | 215615 |  |  |  |  | 45% | 5 | 31/07/20 | 13:11:26 | jenny dahl | 0 |  | 0.26 | 0.01% | 1 | 5 | 5 | ... | 0.00253 | 9e-05 |  |  |  |  | 0.98845 | 0.98845 | Yes | Yes | [5 empty] | 0.00253 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00263 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00253 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-07-28 | Julia Collins | 14:02 |  | Completed |  |  |  |  | 79593 | 2000000 |  |  | 79596 79604 79606 79609 79611 79613 | 79596~79604~79606~79609~79611~79613 | 3297 |  | 3297 3297 3297 3297 3297 3297 | 2020-07-31 | 9:10 | albert zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0016 | 0.00169 | 79596 |  | XCAMS | EA | 0% |  | TW3297 measurement: * large carbonate samples CalAMS error adjustments: * factor 1.05 to account for excess global run-to-run variability * factor 1.4 to account for excess cathode (TP79532) run-to-run variability primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value blanks: * one groups identified to blank-correct: carbonate: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-09-04 | Carbonate | Complete | JOB-215615 |  |  | 70065 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0018 | 0.00171 | 79604 | [26 empty] | 2020-07-20 | 4806 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99993 | 0.0017 | 79606 | [26 empty] | 2020-07-22 | 19603 | Graphite | Completed
[168 empty] | 0.99952 | 0.0017 | 79609 | [26 empty] | 2020-07-28 | 7475 | C13 Measurement | Completed
[168 empty] | 0.99797 | 0.00169 | 79611 | [26 empty] | 2020-07-31 | 20797 | AMS Submission Results | Completed
[168 empty] | 0.99917 | 0.00171 | 79613 | [26 empty] | 2020-07-31 | 24756 | Close Job | Completed
[7 empty] | 127.2 | 87.3398 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45537 | 45537 | 45537 | 217 | 217 | 217 |  |  | Julia Collins | -996.55 | 0.09 ‰ | 2020-08-14 | 2020-02-16 | 14 | 180 |  |  |  |  | 2.72‰ | 2.72‰ | 2.00‰ | 2.00‰ | 2.72 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -996.55 ‰ | -996.55 ‰ | 0.09 | 0.09 | -996.51 | 0.09 ‰ | -996.51 | 0.09 ‰ |  |  | 2020-08-18 | Jocelyn Turnbull | 15:09 | 0.0034521 | 0.0034521 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00345 | 9e-05 | 20898 | 2020-08-16 | 13:00 | albert zondervan | 215602 |  |  |  |  | 45% | 5 | 18/08/20 | 15:08:37 | Jocelyn Turnbull | 0 |  | 0.35 | 0.01% | 15 | 5 | 5 | ... | 0.00332 | 9e-05 |  |  |  |  | 0.98413 | 0.98413 | Yes | Yes | [5 empty] | 0.00332 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 0.99999 | 0.00093 | 0.00093 | 0.00345 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00332 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-13 | Julia Collins | 13:49 |  | Completed |  |  |  |  | 79667 | 2000000 |  |  | 79630 79631 79632 79687 79688 79689 | 79630~79631~79632~79687~79688~79689 | 3299 |  | 3299 3299 3299 3299 3299 3299 | 2020-08-18 | 13:23 | albert zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0004 | 0.00226 | 79630 |  | XCAMS | ST | 0.17% |  | TW3299 measurement: * organics, waters CalAMS error adjustments: * global factor 1.04 to account for excess run-to-run variability * system error 0.17% to account for excess OxI-to-OxI variability primaries: * OxI cathodes: no failures 6x ST * No clear cause was identified for the excess OxI-to-OxI variability, and no specific could be identified as the offending outlier. secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: water: sample is combination of two extractions 1x unknowns + 0x secondaries + 2x blanks 180 days: capturing 2x blanks water: as per normal, i.e. through single extraction 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures * The AMS-IRMS difference in del13C is strongly associated with sample size, as has been seen in past measurements. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-09-03 | Water |  | JOB-215602 |  |  | 70145 | Normal | 14047/11 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.004 | 0.00228 | 79631 | [26 empty] | 2020-07-30 | 4839 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00228 | 79632 | [26 empty] | 2020-07-31 | 19677 | Graphite | Completed
[168 empty] | 0.99929 | 0.00227 | 79687 | [26 empty] | 2020-08-03 | 7493 | C13 Measurement | Completed
[168 empty] | 0.99742 | 0.00226 | 79688 | [26 empty] | 2020-08-18 | 20898 | AMS Submission Results | Completed
[168 empty] | 0.99849 | 0.00228 | 79689 | [26 empty] | 2020-08-25 | 24869 | Close Job | Completed
[5 empty] | 3.4551e-05 | 0.0047882 | 115.3 | 88.4727 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7111 | 7111 | 7111 | 37 | 37 | 41 | 7073.6 | 7148.5 | Julia Collins | -587.38 | 1.92 ‰ | 2020-08-14 | 2020-02-16 | 14 | 180 |  |  |  |  | -10.12‰ | -10.12‰ | 2.00‰ | 2.00‰ | -10.12 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -587.38 ‰ | -587.38 ‰ | 1.92 | 2.14 | -582.28 | 1.95 ‰ | -582.28 | 1.95 ‰ |  |  | 2020-08-18 | Jocelyn Turnbull | 15:09 | 0.41262 | 0.41262 | 0.00192 | 0.00214 | 0.00185 | 0.00185 | 2 |  |  | 0.41603 | 0.00114 | 20899 | 2020-08-16 | 13:00 | albert zondervan | 215603 | 0.00544 | 0.0054 | 0.00245 | 0.0025 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.26 | 0.19% | 29 | 5 | 4 | ... | 0.40011 | 0.00118 |  |  |  |  | 0.98413 | 0.98413 | Yes | Yes |  | result is consistent with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 41.5814743168846876. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00544 | 0.00461 | 0.00219 | 0.39683 | 0.40‰ | 0.0019 | 0.00‰ | Yes | 1 | 0.99999 | 0.00093 | 0.00093 | 0.41262 | 0 | 0.00192 | 0 | 0 | 0.00544 | 0 | 0.00245 | 0 | 0.39683 | 0 | 0.0019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-13 | Julia Collins | 13:49 |  | Completed |  |  |  |  | 79668 | 2000000 | 79667 79317 79274 79191 79103 79099 79054 78922 78893 78782 78739 | 79667~79317~79274~79191~79103~79099~79054~78922~78893~78782~78739 | 79630 79631 79632 79687 79688 79689 | 79630~79631~79632~79687~79688~79689 | 3299 | 3299 3290 3289 3288 3286 3284 3282 3280 3278 3275 3274 | 3299 3299 3299 3299 3299 3299 | 2020-08-18 | 13:23 | albert zondervan |  | 0.91 mg | 0.13 | 0.91 | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 1.0004 | 0.00226 | 79630 | EA | XCAMS | ST | 0.17% |  | TW3299 measurement: * organics, waters CalAMS error adjustments: * global factor 1.04 to account for excess run-to-run variability * system error 0.17% to account for excess OxI-to-OxI variability primaries: * OxI cathodes: no failures 6x ST * No clear cause was identified for the excess OxI-to-OxI variability, and no specific could be identified as the offending outlier. secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: water: sample is combination of two extractions 1x unknowns + 0x secondaries + 2x blanks 180 days: capturing 2x blanks water: as per normal, i.e. through single extraction 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures * The AMS-IRMS difference in del13C is strongly associated with sample size, as has been seen in past measurements. | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-09-03 | Water | Complete | JOB-215603 |  |  | 70148 | Normal | 14047/12 |  |  | 2020-07-30 | 4840 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04 | 1.004 | 0.00228 | 79631 | EA | [25 empty] | 2020-07-31 | 19678 | Graphite | Completed
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 | 1.0004 | 0.00228 | 79632 | ST | [25 empty] | 2020-11-30 | 7510 | C13 Measurement | Completed
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 0.99929 | 0.00227 | 79687 | EA | [25 empty] | 2020-08-18 | 20899 | AMS Submission Results | Completed
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 0.99742 | 0.00226 | 79688 | [26 empty] | 2020-12-09 | 25767 | Close Job | Completed
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 | 0.99849 | 0.00228 | 79689 | ST
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 |  |  |  | EA
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 |  |  |  | ST
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 |  |  |  | EA
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 |  |  |  | FLASK
[163 empty] | 2020-02-26 | 0.00502 | 78739 | 3274 | 1.09 |  |  |  | ST
ST
[7 empty] | 103.4 | 87.8632 | 0.0023 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34383 | 34382 | 34382 | 126 | 126 | 126 |  |  | Julia Collins | -986.16 | 0.22 ‰ | 2020-08-14 | 2020-02-16 | 14 | 180 |  |  |  |  | -3.25‰ | -3.25‰ | 2.00‰ | 2.00‰ | -3.25 ‰ | 2.00 ‰ | AMS | 2.21‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -986.16 ‰ | -986.16 ‰ | 0.22 | 0.22 | -985.99 | 0.22 ‰ | -985.99 | 0.22 ‰ |  |  | 2020-08-18 | Jocelyn Turnbull | 15:09 | 0.01384 | 0.01384 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 3.7 |  |  | 0.01384 | 0.00022 | 20910 | 2020-08-16 | 13:00 | albert zondervan | 215976 |  |  |  |  | 45% | 5 | 18/08/20 | 15:08:50 | Jocelyn Turnbull | 0 |  | 1.38 | 0.02% | 36 | 5 | 7 | ... | 0.01331 | 0.00021 |  |  |  |  | 0.98413 | 0.98413 | Yes | Yes | [5 empty] | 0.01331 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 0.99999 | 0.00093 | 0.00093 | 0.01384 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01331 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-13 | Julia Collins | 13:49 |  | Completed |  |  |  |  | 79669 | 2000000 |  |  | 79630 79631 79632 79687 79688 79689 | 79630~79631~79632~79687~79688~79689 | 3299 |  | 3299 3299 3299 3299 3299 3299 | 2020-08-18 | 13:23 | albert zondervan |  | 0.51 mg |  | 0.51 | [5 empty] | 1.0004 | 0.00226 | 79630 |  | XCAMS | ST | 0.17% |  | TW3299 measurement: * organics, waters CalAMS error adjustments: * global factor 1.04 to account for excess run-to-run variability * system error 0.17% to account for excess OxI-to-OxI variability primaries: * OxI cathodes: no failures 6x ST * No clear cause was identified for the excess OxI-to-OxI variability, and no specific could be identified as the offending outlier. secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: water: sample is combination of two extractions 1x unknowns + 0x secondaries + 2x blanks 180 days: capturing 2x blanks water: as per normal, i.e. through single extraction 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures * The AMS-IRMS difference in del13C is strongly associated with sample size, as has been seen in past measurements. | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-09-24 | Water | Complete | JOB-215976 |  | blank to match combined sample 41477/7 | 70146 | Normal | 14047/13 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Use this R 14047/11 for CO2 produced on water line AND COMBINED, to match with water samples that are combined from two water extractions. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. Use R14047/11 for single shot water line samples. |  |  | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C1: Carrara Marble WATER LINE COMBINED | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate. This is for samples where two water extractions are combined for graphitisation.
[168 empty] | 1.004 | 0.00228 | 79631 | [26 empty] | 2020-07-30 | 4842 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00228 | 79632 | [26 empty] | 2020-07-30 | 4843 | Water CO2 Evolution | Completed
[168 empty] | 0.99929 | 0.00227 | 79687 | [26 empty] | 2020-07-30 | 288 | Combine Water CO2 Evolution | Completed
[168 empty] | 0.99742 | 0.00226 | 79688 | [26 empty] | 2020-07-31 | 19679 | Graphite | Completed
[168 empty] | 0.99849 | 0.00228 | 79689 | [26 empty] | 2020-08-03 | 7494 | C13 Measurement | Completed
[197 empty] | 2020-08-18 | 20910 | AMS Submission Results | Completed
[197 empty] | 2020-08-25 | 24873 | Close Job | Completed
[7 empty] | 105 | 87.8588 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37808 | 37807 | 37807 | 157 | 156 | 156 |  |  | Julia Collins | -990.96 | 0.18 ‰ | 2020-08-14 | 2020-02-16 | 1 | 180 |  |  |  |  | -3.20‰ | -3.20‰ | 2.00‰ | 2.00‰ | -3.20 ‰ | 2.00 ‰ | AMS | 2.19‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -990.96 ‰ | -990.96 ‰ | 0.18 | 0.18 | -990.85 | 0.18 ‰ | -990.85 | 0.18 ‰ |  |  | 2020-08-18 | Jocelyn Turnbull | 15:09 | 0.0090358 | 0.0090358 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 4.1 |  |  | 0.00904 | 0.00018 | 20915 | 2020-08-16 | 13:00 | albert zondervan | 216033 |  |  |  |  | 45% | 5 | 19/08/20 | 11:55:13 | jenny dahl | 0 |  | 0.9 | 0.02% | 1 | 5 | 7 | ... | 0.00869 | 0.00017 |  |  |  |  | 0.98413 | 0.98413 | Yes | Yes | [5 empty] | 0.00869 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 0.99999 | 0.00093 | 0.00093 | 0.00904 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00869 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-13 | Julia Collins | 13:49 |  | Completed |  |  |  |  | 79711 | 2000000 |  |  | 79630 79631 79632 79687 79688 79689 | 79630~79631~79632~79687~79688~79689 | 3299 |  | 3299 3299 3299 3299 3299 3299 | 2020-08-18 | 13:23 | albert zondervan |  | 0.50 mg |  | 0.5 | [5 empty] | 1.0004 | 0.00226 | 79630 |  | XCAMS | ST | 0.17% |  | TW3299 measurement: * organics, waters CalAMS error adjustments: * global factor 1.04 to account for excess run-to-run variability * system error 0.17% to account for excess OxI-to-OxI variability primaries: * OxI cathodes: no failures 6x ST * No clear cause was identified for the excess OxI-to-OxI variability, and no specific could be identified as the offending outlier. secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: water: sample is combination of two extractions 1x unknowns + 0x secondaries + 2x blanks 180 days: capturing 2x blanks water: as per normal, i.e. through single extraction 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures * The AMS-IRMS difference in del13C is strongly associated with sample size, as has been seen in past measurements. | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-09-30 | Water | Not Complete | JOB-216033 |  | blank to match combined sample 41477/7 | 70147 | Normal | 14047/13 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Use this R 14047/11 for CO2 produced on water line AND COMBINED, to match with water samples that are combined from two water extractions. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. Use R14047/11 for single shot water line samples. |  |  | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C1: Carrara Marble WATER LINE COMBINED | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate. This is for samples where two water extractions are combined for graphitisation.
[168 empty] | 1.004 | 0.00228 | 79631 | [26 empty] | 2020-08-12 | 4855 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00228 | 79632 | [26 empty] | 2020-08-12 | 4858 | Water CO2 Evolution | Completed
[168 empty] | 0.99929 | 0.00227 | 79687 | [26 empty] | 2020-08-12 | 289 | Combine Water CO2 Evolution | Completed
[168 empty] | 0.99742 | 0.00226 | 79688 | [26 empty] | 2020-08-13 | 19721 | Graphite | Completed
[168 empty] | 0.99849 | 0.00228 | 79689 | [26 empty] | 2020-08-18 | 7507 | C13 Measurement | Completed
[197 empty] | 2020-08-18 | 20915 | AMS Submission Results | Completed
[197 empty] | 2020-08-25 | 24878 | Close Job | Completed
[7 empty] | 118.8 | 87.643 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42779 | 42779 | 42779 | 222 | 222 | 222 |  |  | albert zondervan | -995.13 | 0.13 ‰ | 2020-08-25 | 2020-02-27 | 6 | 180 |  |  |  |  | 2.51‰ | 2.51‰ | 2.00‰ | 2.00‰ | 2.51 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -995.13 ‰ | -995.13 ‰ | 0.13 | 0.13 | -995.07 | 0.14 ‰ | -995.07 | 0.14 ‰ |  |  | 2020-08-27 | Jocelyn Turnbull | 11:19 | 0.0048662 | 0.0048662 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2.4 |  |  | 0.00487 | 0.00013 | 20916 | 2020-08-26 | 14:42 | albert zondervan | 215278 |  |  |  |  | 45% | 5 | 27/08/20 | 11:19:06 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 37 | 5 | 6 | ... | 0.00468 | 0.00013 |  |  |  |  | 0.98733 | 0.98733 | Yes | Yes | [5 empty] | 0.00468 | 0.00‰ | 0.00013 | 0.00‰ | No | 1 | 0.99999 | 0.00125 | 0.00125 | 0.00487 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00468 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-25 | albert zondervan | 10:04 |  | Completed |  |  |  |  | 79745 | 2000000 |  |  | 79752 79753 79754 79755 79756 79757 | 79752~79753~79754~79755~79756~79757 | 3300 |  | 3300 3300 3300 3300 3300 3300 | 2020-08-26 | 14:48 | albert zondervan |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0036 | 0.00307 | 79752 |  | XCAMS | EA | 0.25% |  | TW3300 measurement: * large water samples CalAMS error adjustments: * 0.25% primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-07-28 | Water | Complete | JOB-215278 |  |  | 70185 | Normal | 14047/11 |  |  | 2020-08-18 | 4866 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0012 | 0.00307 | 79753 | [26 empty] | 2020-08-18 | 2881 | Gas Splitting | Completed
[168 empty] | 0.99451 | 0.00304 | 79754 | [26 empty] | 2020-08-18 | 4884 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00306 | 79755 | [26 empty] | 2020-08-19 | 19755 | Graphite | Completed
[168 empty] | 0.99897 | 0.00304 | 79756 | [26 empty] | 2020-08-20 | 7518 | C13 Measurement | Completed
[168 empty] | 1.001 | 0.00304 | 79757 | [26 empty] | 2020-08-27 | 20916 | AMS Submission Results | Completed
[197 empty] | 2020-08-27 | 24909 | Close Job | Completed
[5 empty] | 3.4387e-05 | 0.0047692 | 113 | 88.604 | 0.0687 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7172 | 7171 | 7171 | 39 | 39 | 43 | 7132.6 | 7211.5 | albert zondervan | -590.5 | 2.01 ‰ | 2020-08-25 | 2020-02-27 | 1 | 180 |  |  |  |  | -8.36‰ | -8.36‰ | 2.00‰ | 2.00‰ | -8.36 ‰ | 2.00 ‰ | AMS | -8.29‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -590.50 ‰ | -590.50 ‰ | 2.01 | 2.21 | -585.44 | 2.04 ‰ | -585.44 | 2.04 ‰ |  |  | 2020-08-27 | Jocelyn Turnbull | 11:19 | 0.4095 | 0.4095 | 0.00201 | 0.00221 | 0.00194 | 0.00194 | 2.6 |  |  | 0.41291 | 0.00127 | 20955 | 2020-08-26 | 14:42 | albert zondervan | 216009 | 0.00541 | 0.0054 | 0.00243 | 0.0024 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 40.95 | 0.2% | 36 | 5 | 6 | ... | 0.39711 | 0.00149 |  |  |  |  | 0.98733 | 0.98733 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 43.4185870030151316. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00541 | 0.00459 | 0.0022 | 0.39383 | 0.39‰ | 0.00211 | 0.00‰ | No | 1 | 0.99999 | 0.00125 | 0.00125 | 0.4095 | 0 | 0.00201 | 0 | 0 | 0.00541 | 0 | 0.00243 | 0 | 0.39383 | 0 | 0.00211 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-25 | albert zondervan | 10:05 |  | Completed |  |  |  |  | 79788 | 2000000 | 79745 79667 79317 79274 79191 79103 79099 79054 78922 78893 78782 | 79745~79667~79317~79274~79191~79103~79099~79054~78922~78893~78782 | 79752 79753 79754 79755 79756 79757 | 79752~79753~79754~79755~79756~79757 | 3300 | 3300 3299 3290 3289 3288 3286 3284 3282 3280 3278 3275 | 3300 3300 3300 3300 3300 3300 | 2020-08-26 | 14:48 | albert zondervan |  | 0.82 mg | 0.13 | 0.82 | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 1.0036 | 0.00307 | 79752 | EA | XCAMS | EA | 0.25% |  | TW3300 measurement: * large water samples CalAMS error adjustments: * 0.25% primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-09-25 | Water | Complete | JOB-216009 |  |  | 70186 | Normal | 14047/12 |  |  | 2020-08-21 | 4890 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 1.0012 | 0.00307 | 79753 | ST | [25 empty] | 2020-08-21 | 2889 | Gas Splitting | Completed
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04 | 0.99451 | 0.00304 | 79754 | EA | [25 empty] | 2020-08-21 | 4894 | Choose Gas Splitting | Completed
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 | 1.0009 | 0.00306 | 79755 | [26 empty] | 2020-08-24 | 19798 | Graphite | Completed
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 0.99897 | 0.00304 | 79756 | ST | [25 empty] | 2020-08-25 | 7540 | C13 Measurement | Completed
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 | 1.001 | 0.00304 | 79757 | EA | [25 empty] | 2020-08-27 | 20955 | AMS Submission Results | Completed
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 |  |  |  | ST | [25 empty] | 2020-08-27 | 24911 | Close Job | Completed
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 |  |  |  | EA
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 |  |  |  | FLASK
[163 empty] | 2020-03-12 | 0.00873 | 78893 | 3278 | 1.1 |  |  |  | ST
[163 empty] | 2020-03-03 | 0.00608 | 78782 | 3275 | 1.16 |  |  |  | ST
EA
[7 empty] | 90.7 | 87.7199 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46624 | 46623 | 46623 | 249 | 249 | 249 |  |  | Julia Collins | -996.98 | 0.09 ‰ | 2020-09-02 | 2020-03-06 | 6 | 180 |  |  |  |  | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -996.98 ‰ | -996.98 ‰ | 0.09 | 0.09 | -996.95 | 0.09 ‰ | -996.95 | 0.09 ‰ |  |  | 2020-09-04 | Jocelyn Turnbull | 9:37 | 0.0030154 | 0.0030154 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00302 | 9e-05 | 21028 | 2020-09-03 | 15:19 | albert zondervan | 216123 |  |  |  |  | 45% | 5 | 4/09/20 | 09:36:33 | Jocelyn Turnbull | 0 |  | 0.3 | 0.01% | 36 | 5 | 5 | ... | 0.0029 | 9e-05 |  |  |  |  | 0.98814 | 0.98814 | Yes | Yes | [5 empty] | 0.0029 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00302 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0029 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-09-01 | Julia Collins | 14:40 |  | Completed |  |  |  |  | 79821 | 2000000 |  |  | 79782 79808 79809 79811 79812 79813 | 79782~79808~79809~79811~79812~79813 | 3302 |  | 3302 3302 3302 3302 3302 3302 | 2020-09-03 | 15:54 | albert zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0015 | 0.00169 | 79782 |  | XCAMS | ST | 0% |  | TW3302 measurement: * large & small samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to co-determine 'small d' small OxI from 3x other wheels selected also: TW3296,3253,3278 d = 0.78 ug C, see 'to blank-correct small samples.xlsm' secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value * 3x small FIRI-I cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: large organics via AAA: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks small organics via AAA: 6x unknowns + 3x secondaries + 2x blanks small KauriRR from 3x other wheels selected also: TW3296,3253,3278 m = 0.72 ug C, see 'to blank-correct small samples.xlsm' unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-10-13 | Carbonate |  | JOB-216123 |  |  | 70265 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99926 | 0.00166 | 79808 | [26 empty] | 2020-08-25 | 4904 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00166 | 79809 | [26 empty] | 2020-08-27 | 19831 | Graphite | Completed
[168 empty] | 0.99829 | 0.00167 | 79811 | [26 empty] | 2020-08-31 | 7556 | C13 Measurement | Completed
[168 empty] | 0.99841 | 0.00167 | 79812 | [26 empty] | 2020-09-04 | 21028 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00167 | 79813 | [26 empty] | 2020-09-04 | 25014 | Close Job | Completed
[7 empty] | 100.7 | 87.4148 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46541 | 46541 | 46541 | 274 | 274 | 274 |  |  | Julia Collins | -996.95 | 0.10 ‰ | 2020-09-07 | 2020-03-11 | 7 | 180 |  |  |  |  | 4.40‰ | 4.40‰ | 2.00‰ | 2.00‰ | 4.40 ‰ | 2.00 ‰ | AMS | 2.69‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -996.95 ‰ | -996.95 ‰ | 0.1 | 0.1 | -996.92 | 0.11 ‰ | -996.92 | 0.11 ‰ |  |  | 2020-09-15 | Jocelyn Turnbull | 9:47 | 0.0030466 | 0.0030466 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.1 |  |  | 0.00305 | 0.0001 | 21113 | 2020-09-09 | 9:53 | albert zondervan | 216010 |  |  |  |  | 45% | 5 | 15/09/20 | 09:47:11 | Jocelyn Turnbull | 0 |  | 0.3 | 0.01% | 20 | 5 | 5 | ... | 0.00293 | 0.0001 |  |  |  |  | 0.98662 | 0.98662 | Yes | Yes | [5 empty] | 0.00293 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00305 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00293 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-09-04 | Julia Collins | 14:33 |  | Completed |  |  |  |  | 79838 | 2000000 |  |  | 79758 79759 79760 79761 79762 79763 | 79758~79759~79760~79761~79762~79763 | 3304 |  | 3304 3304 3304 3304 3304 3304 | 2020-09-09 | 10:06 | albert zondervan |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0001 | 0.00183 | 79758 |  | XCAMS | EA | 0% |  | TW3304 measurement: * large carbonate samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2: result agrees with consensus value blanks: * one groups identified to blank-correct: carbonates: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2020-09-25 | Carbonate |  | JOB-216010 |  |  | 70382 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.002 | 0.00183 | 79759 | [26 empty] | 2020-08-28 | 4918 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00182 | 79760 | [26 empty] | 2020-08-31 | 19848 | Graphite | Completed
[168 empty] | 0.99955 | 0.0018 | 79761 | [26 empty] | 2020-09-03 | 7578 | C13 Measurement | Completed
[168 empty] | 0.99915 | 0.00181 | 79762 | [26 empty] | 2020-09-15 | 21113 | AMS Submission Results | Completed
[168 empty] | 0.9989 | 0.0018 | 79763 | [26 empty] | 2020-09-17 | 25105 | Close Job | Completed
[7 empty] | 124.7 | 87.608 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39319 | 39318 | 39318 | 134 | 133 | 133 |  |  | jenny dahl | -992.51 | 0.12 ‰ | 2020-09-16 | 2020-03-20 | 6 | 180 |  |  |  |  | 1.03‰ | 1.03‰ | 2.00‰ | 2.00‰ | 1.03 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.51 ‰ | -992.51 ‰ | 0.12 | 0.12 | -992.42 | 0.13 ‰ | -992.42 | 0.13 ‰ |  |  | 2020-09-17 | Jocelyn Turnbull | 11:32 | 0.0074865 | 0.0074865 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00749 | 0.00012 | 21134 | 2020-09-17 | 11:02 | albert zondervan | 216131 |  |  |  |  | 45% | 5 | 17/09/20 | 11:40:26 | jenny dahl | 0 |  | 0.75 | 0.01% | 1 | 5 | 4 | ... | 0.0072 | 0.00012 |  |  |  |  | 0.98548 | 0.98548 | Yes | Yes | [5 empty] | 0.0072 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00749 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0072 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-09-10 | jenny dahl | 12:45 |  | Completed |  |  |  |  | 79961 | 2000000 |  |  | 79924 79925 79926 79927 79928 79929 | 79924~79925~79926~79927~79928~79929 | 3305 |  | 3305 3305 3305 3305 3305 3305 | 2020-09-17 | 11:12 | albert zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 0.99796 | 0.00149 | 79924 |  | XCAMS | ST | 0% |  | TW3305 measurement: * large samples * partially filled wheel CalAMS error adjustments: * TP79942: error multiplier = 1.3 to account for excess run-to-run variability primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value blanks: * two groups identified to blank-correct: large organics via AAA: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large water: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-10-14 | Water |  | JOB-216131 |  |  | 70403 | Normal | 14047/11 |  |  | 2020-09-08 | 4943 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0027 | 0.00156 | 79925 | [26 empty] | 2020-09-10 | 19971 | Graphite | Completed
[168 empty] | 0.9993 | 0.00152 | 79926 | [26 empty] | 2020-09-11 | 7618 | C13 Measurement | Completed
[168 empty] | 0.99936 | 0.00151 | 79927 | [26 empty] | 2020-09-17 | 21134 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00151 | 79928 | [26 empty] | 2020-09-17 | 25106 | Close Job | Completed
[168 empty] | 0.9996 | 0.00157 | 79929
[5 empty] | 3.4786e-05 | 0.004703 | 120.1 | 88.5634 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7134 | 7133 | 7133 | 37 | 36 | 40 | 7097.2 | 7170.4 | jenny dahl | -588.54 | 1.88 ‰ | 2020-09-16 | 2020-03-20 | 6 | 180 |  |  |  |  | -9.77‰ | -9.77‰ | 2.00‰ | 2.00‰ | -9.77 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -588.54 ‰ | -588.54 ‰ | 1.88 | 2.09 | -583.46 | 1.90 ‰ | -583.46 | 1.90 ‰ |  |  | 2020-09-17 | Jocelyn Turnbull | 11:32 | 0.41146 | 0.41146 | 0.00188 | 0.00209 | 0.0018 | 0.0018 | 2.5 |  |  | 0.41472 | 0.00115 | 21135 | 2020-09-17 | 11:02 | albert zondervan | 216132 | 0.00519 | 0.0052 | 0.00234 | 0.0023 | 45% | 5 | 5/07/22 | 14:07:05 | Jocelyn Turnbull | 0 |  | 41.15 | 0.19% | 2 | 5 | 4 | ... | 0.39885 | 0.00098 |  |  |  |  | 0.98548 | 0.98548 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.8605863652246256. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00519 | 0.00452 | 0.0021 | 0.39571 | 0.40‰ | 0.00173 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.41146 | 0 | 0.00188 | 0 | 0 | 0.00519 | 0 | 0.00234 | 0 | 0.39571 | 0 | 0.00173 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-09-10 | jenny dahl | 12:45 |  | Completed |  |  |  |  | 79962 | 2000000 | 79961 79745 79667 79317 79274 79191 79103 79099 79054 78922 | 79961~79745~79667~79317~79274~79191~79103~79099~79054~78922 | 79924 79925 79926 79927 79928 79929 | 79924~79925~79926~79927~79928~79929 | 3305 | 3305 3300 3299 3290 3289 3288 3286 3284 3282 3280 | 3305 3305 3305 3305 3305 3305 | 2020-09-17 | 11:12 | albert zondervan |  | 0.83 mg | 0.13 | 0.83 | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 0.99796 | 0.00149 | 79924 | EA | XCAMS | ST | 0% |  | TW3305 measurement: * large samples * partially filled wheel CalAMS error adjustments: * TP79942: error multiplier = 1.3 to account for excess run-to-run variability primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value blanks: * two groups identified to blank-correct: large organics via AAA: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large water: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-10-14 | Water | Not Complete | JOB-216132 |  |  | 70404 | Normal | 14047/12 |  |  | 2020-09-08 | 4944 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 1.0027 | 0.00156 | 79925 | [26 empty] | 2020-09-10 | 19972 | Graphite | Completed
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 0.9993 | 0.00152 | 79926 | ST | [25 empty] | 2020-09-11 | 7621 | C13 Measurement | Completed
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04 | 0.99936 | 0.00151 | 79927 | EA | [25 empty] | 2020-09-17 | 21135 | AMS Submission Results | Completed
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 | 1.0013 | 0.00151 | 79928 | ST | [25 empty] | 2020-09-17 | 25107 | Close Job | Completed
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 | 0.9996 | 0.00157 | 79929 | EA
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 |  |  |  | FLASK
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 |  |  |  | ST
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 |  |  |  | ST
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66 |  |  |  | EA
ST
[7 empty] | 94.3 | 87.7011 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45858 | 45858 | 45858 | 227 | 226 | 226 |  |  | Julia Collins | -996.68 | 0.09 ‰ | 2020-10-11 | 2020-04-14 | 26 | 180 |  |  |  |  | 2.27‰ | 2.27‰ | 2.00‰ | 2.00‰ | 2.27 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -996.68 ‰ | -996.68 ‰ | 0.09 | 0.09 | -996.64 | 0.09 ‰ | -996.64 | 0.09 ‰ |  |  | 2020-10-13 | Jocelyn Turnbull | 14:59 | 0.0033169 | 0.0033169 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00332 | 9e-05 | 21277 | 2020-10-13 | 13:56 | jeremy thompson | 216277 |  |  |  |  | 45% | 5 | 13/10/20 | 14:58:35 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 1 | 5 | 5 | ... | 0.00319 | 9e-05 |  |  |  |  | 0.98775 | 0.98775 | Yes | Yes | [5 empty] | 0.00319 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00332 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00319 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-10-08 | Julia Collins | 13:38 |  | Completed |  |  |  |  | 79984 | 2000000 |  |  | 80105 80108 80116 80125 80128 80133 | 80105~80108~80116~80125~80128~80133 | 3309 |  | 3309 3309 3309 3309 3309 3309 | 2020-10-13 | 14:37 | jeremy thompson |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0002 | 0.00159 | 80105 |  | XCAMS | EA | 0% |  | TW3309 measurement: * large samples CalAMS error adjustments: *no error adjustment primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value blanks: * two groups identified to blank-correct: large organics via AAA: 14x unknowns + 2x secondaries + 1x blanks 180 days: capturing 13x blanks large carbonate: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-10-28 | Carbonate |  | JOB-216277 |  |  | 70539 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0003 | 0.00158 | 80108 | [26 empty] | 2020-09-14 | 4951 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0001 | 0.00157 | 80116 | [26 empty] | 2020-09-15 | 19994 | Graphite | Completed
[168 empty] | 0.9991 | 0.00159 | 80125 | [26 empty] | 2020-09-18 | 7622 | C13 Measurement | Completed
[168 empty] | 0.99983 | 0.00158 | 80128 | [26 empty] | 2020-10-13 | 21277 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00155 | 80133 | [26 empty] | 2020-10-23 | 25366 | Close Job | Completed
[7 empty] | 98.1 | 87.3683 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41338 | 41337 | 41337 | 186 | 186 | 186 |  |  | Margaret Norris | -994.18 | 0.14 ‰ | 2020-10-13 | 2020-04-16 | 12 | 180 |  |  |  |  | 6.75‰ | 6.75‰ | 2.00‰ | 2.00‰ | 6.75 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.75593 | -994.18 ‰ | -994.18 ‰ | 0.14 | 0.14 | -994.11 | 0.14 ‰ | -994.11 | 0.14 ‰ |  |  | 2020-10-15 | Jocelyn Turnbull | 14:18 | 0.0058229 | 0.0058229 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.00582 | 0.00014 | 21335 | 2020-10-15 | 12:50 | albert zondervan | 216380 |  |  |  |  | 45% | 5 | 15/10/20 | 14:17:37 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 1 | 5 | 6 | ... | 0.0056 | 0.00013 |  |  |  |  | 0.9884 | 0.9884 | Yes | Yes | [5 empty] | 0.0056 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00582 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0056 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-10-12 | Margaret Norris | 15:40 |  | Completed |  |  |  |  | 80120 | 2000000 |  |  | 80140 80141 80142 80143 80161 80199 80205 | 80140~80141~80142~80143~80161~80199~80205 | 3310 |  | 3310 3310 3310 3310 3310 3310 3310 | 2020-10-15 | 13:07 | albert zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0011 | 0.0018 | 80140 |  | XCAMS | FLASK | 0% |  | TW3310 measurement: * no issues * atm-14CH4 & water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x from flask + 1x from EA secondaries: * IAEA-C2: 1x result agrees with consensus value blanks: * Two groups identified to blank-correct: 14x air samples (atm-CH4 from NIWA): 14x unknowns, 0x secondaries, 0x blanks window = 360 days, capturing 7x blanks Extractions were done at NIWA - Hinrich Schaefer has determined that there is no quantifiable additional blank correction required related to the extraction procedure. Therefore we use our CO2 in air blank to estimate the graphite/measurement blank 14x water samples: 13x unknowns, 1x secondaries, 1x blanks window = 180 days, capturing 13x blanks unknowns: * no failures, however: * Some cathodes filled with graphite prepared from water samples generated a ca. 15% lower than average beam current, generating a consistent pattern throughout the measurement. Post-measurement the cathodes in question were tested for electrical and thermal conductivity, showing no clear pattern match with lower-current cathodes. The cause remains unclear at this stage, but the ratio measurements do not appear to have been affected. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-11-20 | Water | Complete | JOB-216380 |  |  | 70579 | Normal | 14047/11 |  |  | 2020-09-30 | 4974 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00176 | 80141 | [26 empty] | 2020-09-30 | 2909 | Gas Splitting | Completed
[168 empty] | 0.99868 | 0.00174 | 80142 | [26 empty] | 2020-09-30 | 4924 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00172 | 80143 | [26 empty] | 2020-10-01 | 20130 | Graphite | Completed
[168 empty] | 1.0003 | 0.00169 | 80161 | [26 empty] | 2020-10-05 | 7646 | C13 Measurement | Completed
[168 empty] | 0.99829 | 0.00167 | 80199 | [26 empty] | 2020-10-15 | 21335 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00174 | 80205 | [26 empty] | 2020-10-19 | 25338 | Close Job | Completed
[5 empty] | 3.271e-05 | 0.0051256 | 91.3 | 88.5552 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7129 | 7129 | 7129 | 42 | 41 | 45 | 7087.4 | 7170.9 | Margaret Norris | -588.3 | 2.14 ‰ | 2020-10-13 | 2020-04-16 | 12 | 180 |  |  |  |  | -6.75‰ | -6.75‰ | 2.00‰ | 2.00‰ | -6.75 ‰ | 2.00 ‰ | AMS | -8.19‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.75593 | -588.30 ‰ | -588.30 ‰ | 2.14 | 2.33 | -583.22 | 2.16 ‰ | -583.22 | 2.16 ‰ |  |  | 2020-10-15 | Jocelyn Turnbull | 14:18 | 0.4117 | 0.4117 | 0.00214 | 0.00233 | 0.00207 | 0.00207 | 2.8 |  |  | 0.41553 | 0.00125 | 21336 | 2020-10-15 | 12:50 | albert zondervan | 216381 | 0.00611 | 0.0061 | 0.00275 | 0.0028 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.17 | 0.21% | 19 | 5 | 4 | ... | 0.39963 | 0.00108 |  |  |  |  | 0.9884 | 0.9884 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 45.4562884548942814. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00611 | 0.00493 | 0.00299 | 0.39594 | 0.40‰ | 0.00199 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.4117 | 0 | 0.00214 | 0 | 0 | 0.00611 | 0 | 0.00275 | 0 | 0.39594 | 0 | 0.00199 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-10-12 | Margaret Norris | 15:40 |  | Completed |  |  |  |  | 80121 | 2000000 | 80120 79961 79745 79667 79669 79711 79317 79274 79191 79103 79099 79054 78922 | 80120~79961~79745~79667~79669~79711~79317~79274~79191~79103~79099~79054~78922 | 80140 80141 80142 80143 80161 80199 80205 | 80140~80141~80142~80143~80161~80199~80205 | 3310 | 3310 3305 3300 3299 3299 3299 3290 3289 3288 3286 3284 3282 3280 | 3310 3310 3310 3310 3310 3310 3310 | 2020-10-15 | 13:07 | albert zondervan |  | 0.77 mg | 0.13 | 0.77 | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 | 1.0011 | 0.0018 | 80140 | EA | XCAMS | FLASK | 0% |  | TW3310 measurement: * no issues * atm-14CH4 & water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x from flask + 1x from EA secondaries: * IAEA-C2: 1x result agrees with consensus value blanks: * Two groups identified to blank-correct: 14x air samples (atm-CH4 from NIWA): 14x unknowns, 0x secondaries, 0x blanks window = 360 days, capturing 7x blanks Extractions were done at NIWA - Hinrich Schaefer has determined that there is no quantifiable additional blank correction required related to the extraction procedure. Therefore we use our CO2 in air blank to estimate the graphite/measurement blank 14x water samples: 13x unknowns, 1x secondaries, 1x blanks window = 180 days, capturing 13x blanks unknowns: * no failures, however: * Some cathodes filled with graphite prepared from water samples generated a ca. 15% lower than average beam current, generating a consistent pattern throughout the measurement. Post-measurement the cathodes in question were tested for electrical and thermal conductivity, showing no clear pattern match with lower-current cathodes. The cause remains unclear at this stage, but the ratio measurements do not appear to have been affected. | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-11-20 | Water | Complete | JOB-216381 |  |  | 70580 | Normal | 14047/12 |  |  | 2020-09-30 | 4975 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 1.0002 | 0.00176 | 80141 | [26 empty] | 2020-10-01 | 20131 | Graphite | Completed | 14047/13 | IAEA-C1: Carrara Marble WATER LINE COMBINED
[163 empty] | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 0.99868 | 0.00174 | 80142 | ST | [25 empty] | 2020-10-05 | 7656 | C13 Measurement | Completed
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 1.0004 | 0.00172 | 80143 | EA | [25 empty] | 2020-10-15 | 21336 | AMS Submission Results | Completed
[163 empty] | 2020-08-14 | 0.01331 | 79669 | 3299 | 0.51 | 1.0003 | 0.00169 | 80161 | ST | [25 empty] | 2020-10-19 | 25337 | Close Job | Completed
[163 empty] | 2020-08-14 | 0.00869 | 79711 | 3299 | 0.5 | 0.99829 | 0.00167 | 80199 | EA
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04 | 1.0012 | 0.00174 | 80205 | FLASK
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6 |  |  |  | ST
[163 empty] | 2020-06-12 | 0.0075 | 79191 | 3288 | 1.11 |  |  |  | ST
[163 empty] | 2020-06-05 | 0.00445 | 79103 | 3286 | 0.6 |  |  |  | EA
[163 empty] | 2020-05-29 | 0.00609 | 79099 | 3284 | 1.12 |  |  |  | ST
[163 empty] | 2020-05-27 | 0.0089 | 79054 | 3282 | 0.96 |  |  |  | FLASK
[163 empty] | 2020-05-22 | 0.0031 | 78922 | 3280 | 0.66
[7 empty] | 140.1 | 87.3306 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46820 | 46819 | 46819 | 199 | 198 | 198 |  |  | Julia Collins | -997.06 | 0.07 ‰ | 2020-11-13 | 2020-05-17 | 2 | 180 |  |  |  |  | 6.05‰ | 6.05‰ | 2.00‰ | 2.00‰ | 6.05 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0035 | 0 | 0.8165 | -997.06 ‰ | -997.06 ‰ | 0.07 | 0.07 | -997.02 | 0.07 ‰ | -997.02 | 0.07 ‰ |  |  | 2020-11-16 | Jocelyn Turnbull | 10:07 | 0.0029426 | 0.0029426 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.3 |  |  | 0.00294 | 7e-05 | 21621 | 2020-11-16 | 8:39 | jeremy thompson | 216489 |  |  |  |  | 45% | 5 | 16/11/20 | 10:07:25 | Jocelyn Turnbull | 0 |  | 0.29 | 0.01% | 1 | 5 | 5 | ... | 0.00283 | 7e-05 |  |  |  |  | 0.98729 | 0.98729 | Yes | Yes | [5 empty] | 0.00283 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.00294 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00283 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-11-12 | Julia Collins | 13:58 |  | Completed |  |  |  |  | 80411 | 2000000 |  |  | 80390 80391 80392 80424 80425 80426 | 80390~80391~80392~80424~80425~80426 | 3317 |  | 3317 3317 3317 3317 3317 3317 | 2020-11-16 | 9:12 | jeremy thompson |  | 0.98 mg |  | 0.98 | [5 empty] | 1 | 0.00136 | 80390 |  | XCAMS | EA | 0% |  | TW3317 measurement: * large samples primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value blanks: * two groups identified to blank-correct: large organics via AAA: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 17x blanks large carbonate: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-12-08 | Carbonate |  | JOB-216489 |  |  | 70857 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99743 | 0.00182 | 80391 | [26 empty] | 2020-11-09 | 5002 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99862 | 0.00139 | 80392 | [26 empty] | 2020-11-11 | 20421 | Graphite | Completed
[168 empty] | 1.0005 | 0.00145 | 80424 | [26 empty] | 2020-11-12 | 7703 | C13 Measurement | Completed
[168 empty] | 1.002 | 0.00135 | 80425 | [26 empty] | 2020-11-16 | 21621 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00138 | 80426 | [26 empty] | 2020-11-16 | 25649 | Close Job | Completed
[7 empty] | 108.7 | 87.8247 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42865 | 42865 | 42865 | 191 | 190 | 190 |  |  | Julia Collins | -995.19 | 0.11 ‰ | 2020-12-10 | 2020-06-13 | 2 | 180 |  |  |  |  | 0.53‰ | 0.53‰ | 2.00‰ | 2.00‰ | 0.53 ‰ | 2.00 ‰ | AMS | 2.14‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -995.19 ‰ | -995.19 ‰ | 0.11 | 0.11 | -995.13 | 0.12 ‰ | -995.13 | 0.12 ‰ |  |  | 2020-12-11 | Jocelyn Turnbull | 14:55 | 0.0048142 | 0.0048142 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3.2 |  |  | 0.00481 | 0.00011 | 21784 | 2020-12-11 | 13:16 | albert zondervan | 216791 |  |  |  |  | 45% | 5 | 11/12/20 | 14:54:41 | Jocelyn Turnbull | 0 |  | 0.48 | 0.01% | 1 | 5 | 6 | ... | 0.00463 | 0.00011 |  |  |  |  | 0.98743 | 0.98743 | Yes | Yes | [5 empty] | 0.00463 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00481 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00463 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-08 | Julia Collins | 14:29 |  | Completed |  |  |  |  | 80600 | 2000000 |  |  | 80533 80602 80603 80604 80605 80606 | 80533~80602~80603~80604~80605~80606 | 3321 |  | 3321 3321 3321 3321 3321 3321 | 2020-12-11 | 13:22 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99896 | 0.00161 | 80533 |  | XCAMS | ST | 0% |  | TW3321 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2021-01-19 | Water |  | JOB-216791 |  |  | 71051 | Normal | 14047/11 |  |  | 2020-12-07 | 5035 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.0016 | 80602 | [26 empty] | 2020-12-07 | 2973 | Gas Splitting | Completed
[168 empty] | 1.0007 | 0.0016 | 80603 | [26 empty] | 2020-12-07 | 5004 | Choose Gas Splitting | Completed
[168 empty] | 1.0014 | 0.00163 | 80604 | [26 empty] | 2020-12-08 | 20610 | Graphite | Completed
[168 empty] | 1 | 0.0016 | 80605 | [26 empty] | 2020-12-11 | 7743 | C13 Measurement | Completed
[168 empty] | 0.99935 | 0.0016 | 80606 | [26 empty] | 2020-12-11 | 21784 | AMS Submission Results | Completed
[197 empty] | 2020-12-14 | 25899 | Close Job | Completed
[5 empty] | 3.9312e-05 | 0.0049517 | 107.9 | 88.8 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7164 | 7164 | 7164 | 41 | 40 | 44 | 7123.2 | 7205 | Julia Collins | -590.09 | 2.09 ‰ | 2020-12-10 | 2020-06-13 | 2 | 180 |  |  |  |  | -10.46‰ | -10.46‰ | 2.00‰ | 2.00‰ | -10.46 ‰ | 2.00 ‰ | AMS | -8.03‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -590.09 ‰ | -590.09 ‰ | 2.09 | 2.28 | -585.03 | 2.11 ‰ | -585.03 | 2.11 ‰ |  |  | 2020-12-11 | Jocelyn Turnbull | 14:55 | 0.40991 | 0.40991 | 0.00209 | 0.00228 | 0.00202 | 0.00202 | 3.1 |  |  | 0.41369 | 0.0012 | 21785 | 2020-12-11 | 13:16 | albert zondervan | 216792 | 0.00601 | 0.006 | 0.0027 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 40.99 | 0.21% | 20 | 5 | 6 | ... | 0.39786 | 0.00103 |  |  |  |  | 0.98743 | 0.98743 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.7094900339092007. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00601 | 0.00476 | 0.00331 | 0.39422 | 0.39‰ | 0.00194 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.40991 | 0 | 0.00209 | 0 | 0 | 0.00601 | 0 | 0.0027 | 0 | 0.39422 | 0 | 0.00194 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-08 | Julia Collins | 14:29 |  | Completed |  |  |  |  | 80601 | 2000000 | 80600 80120 79961 79745 79667 79669 79711 79317 79274 | 80600~80120~79961~79745~79667~79669~79711~79317~79274 | 80533 80602 80603 80604 80605 80606 | 80533~80602~80603~80604~80605~80606 | 3321 | 3321 3310 3305 3300 3299 3299 3299 3290 3289 | 3321 3321 3321 3321 3321 3321 | 2020-12-11 | 13:22 | albert zondervan |  | 0.62 mg | 0.13 | 0.62 | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 | 0.99896 | 0.00161 | 80533 | FLASK | XCAMS | ST | 0% |  | TW3321 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-01-19 | Water |  | JOB-216792 |  |  | 71052 | Normal | 14047/12 |  |  | 2020-12-07 | 5036 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 | 0.99962 | 0.0016 | 80602 | ST | [25 empty] | 2020-12-07 | 2974 | Gas Splitting | Completed | 14047/13 | IAEA-C1: Carrara Marble WATER LINE COMBINED
[163 empty] | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 1.0007 | 0.0016 | 80603 | ST | [25 empty] | 2020-12-07 | 5005 | Choose Gas Splitting | Completed
[163 empty] | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 1.0014 | 0.00163 | 80604 | EA | [25 empty] | 2020-12-08 | 20611 | Graphite | Completed
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 1 | 0.0016 | 80605 | ST | [25 empty] | 2020-12-11 | 7744 | C13 Measurement | Completed
[163 empty] | 2020-08-14 | 0.01331 | 79669 | 3299 | 0.51 | 0.99935 | 0.0016 | 80606 | FLASK | [25 empty] | 2020-12-11 | 21785 | AMS Submission Results | Completed
[163 empty] | 2020-08-14 | 0.00869 | 79711 | 3299 | 0.5 |  |  |  | ST | [25 empty] | 2020-12-14 | 25900 | Close Job | Completed
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04
[163 empty] | 2020-06-16 | 0.00313 | 79274 | 3289 | 0.6
[7 empty] | 101.3 | 87.6827 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44431 | 44431 | 44431 | 210 | 209 | 209 |  |  | Julia Collins | -996.04 | 0.10 ‰ | 2020-12-16 | 2020-06-19 | 2 | 180 |  |  |  |  | 2.06‰ | 2.06‰ | 2.00‰ | 2.00‰ | 2.06 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -996.04 ‰ | -996.04 ‰ | 0.1 | 0.1 | -995.99 | 0.10 ‰ | -995.99 | 0.10 ‰ |  |  | 2020-12-22 | albert zondervan | 14:37 | 0.0039616 | 0.0039616 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.8 |  |  | 0.00396 | 0.0001 | 21856 | 2020-12-17 | 13:20 | albert zondervan | 216588 |  |  |  |  | 45% | 5 | 22/12/20 | 14:36:52 | albert zondervan | 0 |  | 0.4 | 0.01% | 11 | 5 | 4 | ... | 0.00381 | 0.0001 |  |  |  |  | 0.98734 | 0.98734 | Yes | Yes | [5 empty] | 0.00381 | 0.00‰ | 0.0001 | 0.00‰ | No | 1 | 1.0001 | 0.00124 | 0.00124 | 0.00396 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00381 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-15 | Julia Collins | 10:50 |  | Completed |  |  |  |  | 80648 | 2000000 |  |  | 80527 80529 80530 80531 80532 80607 | 80527~80529~80530~80531~80532~80607 | 3323 |  | 3323 3323 3323 3323 3323 3323 |  |  |  |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99717 | 0.003 | 80527 |  | XCAMS | ST | 0.25% |  | TW3323 measurement: * large & small samples CalAMS error adjustments: * 0.25% system error primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 2x small: no failures, used to co-determine 'small d' small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 d = 0.79 ug C, see 'to blank-correct small samples.xlsm' secondaries: * 1x large FIRI-D cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value * 2x small FIRI-D cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: large waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large organics via AAA: 6x unknowns + 0x secondaries + 0x blanks 180 days: capturing 17x blanks large organics via CE-Rafter: 3x unknowns + 1x secondaries + 1x blanks 360 days: capturing 4x blanks small organics via AAA & CE-Rafter: AAA: 4x unknowns + 0x secondaries + 1x blanks CE-Rafter: 0x unknowns + 2x secondaries + 2x blanks small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 m = 0.74 ug C, see 'to blank-correct small samples.xlsm' unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-12-30 | Water | Complete | JOB-216588 |  |  | 71097 | Normal | 14047/11 |  |  | 2020-12-11 | 5042 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.004 | 0.0031 | 80529 | [26 empty] | 2020-12-14 | 20658 | Graphite | Completed
[168 empty] | 1.0027 | 0.00305 | 80530 | [26 empty] | 2020-12-17 | 7760 | C13 Measurement | Completed
[168 empty] | 0.99867 | 0.00302 | 80531 | [26 empty] | 2020-12-22 | 21856 | AMS Submission Results | Completed
[168 empty] | 1.0016 | 0.00301 | 80532 | [26 empty] | 2020-12-22 | 25961 | Close Job | Completed
[168 empty] | 0.99656 | 0.00302 | 80607
[5 empty] | 3.9961e-05 | 0.0051136 | 102.9 | 88.7575 | 0.0691 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7132 | 7132 | 7132 | 42 | 41 | 45 | 7090.7 | 7174.5 | Julia Collins | -588.48 | 2.15 ‰ | 2020-12-16 | 2020-06-19 | 2 | 180 |  |  |  |  | -10.07‰ | -10.07‰ | 2.00‰ | 2.00‰ | -10.07 ‰ | 2.00 ‰ | AMS | -8.08‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -588.48 ‰ | -588.48 ‰ | 2.15 | 2.34 | -583.4 | 2.17 ‰ | -583.4 | 2.17 ‰ |  |  | 2020-12-22 | albert zondervan | 14:37 | 0.41152 | 0.41152 | 0.00215 | 0.00234 | 0.00208 | 0.00208 | 2.4 |  |  | 0.41534 | 0.00126 | 21857 | 2020-12-17 | 13:20 | albert zondervan | 216589 | 0.00608 | 0.0061 | 0.00274 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.15 | 0.21% | 26 | 5 | 4 | ... | 0.39944 | 0.00148 |  |  |  |  | 0.98734 | 0.98734 | Yes | Yes |  | secondary agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 45.6247220703550439. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00608 | 0.00492 | 0.00324 | 0.39577 | 0.40‰ | 0.00223 | 0.00‰ | No | 1 | 1.0001 | 0.00124 | 0.00124 | 0.41152 | 0 | 0.00215 | 0 | 0 | 0.00608 | 0 | 0.00274 | 0 | 0.39577 | 0 | 0.00223 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-15 | Julia Collins | 10:50 |  | Completed |  |  |  |  | 80649 | 2000000 | 80648 80600 80120 79961 79745 79667 79669 79711 79317 | 80648~80600~80120~79961~79745~79667~79669~79711~79317 | 80527 80529 80530 80531 80532 80607 | 80527~80529~80530~80531~80532~80607 | 3323 | 3323 3321 3310 3305 3300 3299 3299 3299 3290 | 3323 3323 3323 3323 3323 3323 | 2020-12-22 | 13:33 | Cathy Ginnane |  | 0.82 mg | 0.13 | 0.82 | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 | 0.99717 | 0.003 | 80527 | ST | XCAMS | ST | 0.25% |  | TW3323 measurement: * large & small samples CalAMS error adjustments: * 0.25% system error primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 2x small: no failures, used to co-determine 'small d' small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 d = 0.79 ug C, see 'to blank-correct small samples.xlsm' secondaries: * 1x large FIRI-D cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value * 2x small FIRI-D cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: large waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large organics via AAA: 6x unknowns + 0x secondaries + 0x blanks 180 days: capturing 17x blanks large organics via CE-Rafter: 3x unknowns + 1x secondaries + 1x blanks 360 days: capturing 4x blanks small organics via AAA & CE-Rafter: AAA: 4x unknowns + 0x secondaries + 1x blanks CE-Rafter: 0x unknowns + 2x secondaries + 2x blanks small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 m = 0.74 ug C, see 'to blank-correct small samples.xlsm' unknowns: * no failures | -8.1 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2020-12-30 | Water | Not Complete | JOB-216589 |  |  | 71098 | Normal | 14047/12 |  |  | 2020-12-11 | 5043 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 | 1.004 | 0.0031 | 80529 | ST | [25 empty] | 2020-12-14 | 20659 | Graphite | Completed | 14047/13 | IAEA-C1: Carrara Marble WATER LINE COMBINED
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 | 1.0027 | 0.00305 | 80530 | EA | [25 empty] | 2020-12-17 | 7764 | C13 Measurement | Completed
[163 empty] | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 0.99867 | 0.00302 | 80531 | ST | [25 empty] | 2020-12-22 | 21857 | AMS Submission Results | Completed
[163 empty] | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 1.0016 | 0.00301 | 80532 | FLASK | [25 empty] | 2020-12-22 | 25962 | Close Job | Completed
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 | 0.99656 | 0.00302 | 80607 | ST
[163 empty] | 2020-08-14 | 0.01331 | 79669 | 3299 | 0.51 |  |  |  | ST
[163 empty] | 2020-08-14 | 0.00869 | 79711 | 3299 | 0.5
[163 empty] | 2020-06-19 | 0.0035 | 79317 | 3290 | 1.04
[7 empty] | 125.3 | 87.0788 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44537 | 44537 | 44537 | 192 | 192 | 192 |  |  | Julia Collins | -996.09 | 0.09 ‰ | 2021-01-28 | 2020-08-01 | 8 | 180 |  |  |  |  | 6.93‰ | 6.93‰ | 2.00‰ | 2.00‰ | 6.93 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0051 | 0 | 0.8165 | -996.09 ‰ | -996.09 ‰ | 0.09 | 0.09 | -996.04 | 0.09 ‰ | -996.04 | 0.09 ‰ |  |  | 2021-02-01 | Jocelyn Turnbull | 11:14 | 0.0039096 | 0.0039096 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.9 |  |  | 0.00391 | 9e-05 | 21892 | 2021-02-01 | 8:25 | jeremy thompson | 216877 |  |  |  |  | 45% | 5 | 1/02/21 | 11:13:46 | Jocelyn Turnbull | 0 |  | 0.39 | 0.01% | 1 | 5 | 5 | ... | 0.00376 | 9e-05 |  |  |  |  | 0.9853 | 0.9853 | Yes | Yes | [5 empty] | 0.00376 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00391 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00376 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-01-26 | Julia Collins | 7:50 |  | Completed |  |  |  |  | 80688 | 2000000 |  |  | 80681 80682 80683 80684 80699 80714 | 80681~80682~80683~80684~80699~80714 | 3324 |  | 3324 3324 3324 3324 3324 3324 | 2021-02-01 | 9:46 | jeremy thompson |  | 1.08 mg |  | 1.08 | [5 empty] | 0.99764 | 0.00159 | 80681 |  | XCAMS | ST | 0% |  | TW3324 measurement: * large samples CalAMS error adjustments: Zero error adjustment primaries: * OxI cathodes: 6x large: no failures secondaries: * * 1x large IAEA-C2: Freshwater Travertine: result agrees with consensus value blanks: * Three groups identified to blank-correct: large organics via client pretreatment: 23x unknowns + 0x secondaries + 3x blanks client sent 3 blanks: large Carbonate via Carbonate CO2 Evolution 1x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks large Water via Water CO2 Evolution 5x unknowns + 0x secondaries + 0x blanks 180 days: capturing 8x blanks(none in this measurement) unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2021-02-05 | Carbonate |  | JOB-216877 |  |  | 71137 | Normal | 14047/1 |  |  | 2020-06-10 | 6604 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0004 | 0.00161 | 80682 | [26 empty] | 2021-01-18 | 5058 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0002 | 0.00163 | 80683 | [26 empty] | 2021-01-20 | 20698 | Graphite | Completed
[168 empty] | 0.9991 | 0.00155 | 80684 | [26 empty] | 2021-02-02 | 7780 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00162 | 80699 | [26 empty] | 2021-02-01 | 21892 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00201 | 80714 | [26 empty] | 2021-02-09 | 26096 | Close Job | Completed
[7 empty] | 82.1 | 87.77 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48843 | 48842 | 48842 | 256 | 255 | 255 |  |  | Julia Collins | -997.71 | 0.07 ‰ | 2021-02-09 | 2020-08-13 | 5 | 180 |  |  |  |  | 3.64‰ | 3.64‰ | 2.00‰ | 2.00‰ | 3.64 ‰ | 2.00 ‰ | AMS | 2.00‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -997.71 ‰ | -997.71 ‰ | 0.07 | 0.07 | -997.68 | 0.07 ‰ | -997.68 | 0.07 ‰ |  |  | 2021-02-12 | Cathy Ginnane | 8:15 | 0.0022875 | 0.0022875 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00229 | 7e-05 | 21998 | 2021-02-11 | 11:57 | albert zondervan | 216938 |  |  |  |  | 45% | 5 | 12/02/21 | 08:14:31 | Cathy Ginnane | 0 |  | 0.23 | 0.01% | 1 | 5 | 4 | ... | 0.0022 | 7e-05 |  |  |  |  | 0.98988 | 0.98988 | Yes | Yes |  | The result for this blank is very low, much like a Kauri wood combustion blank. |  |  |  | 0.0022 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00229 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0022 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-02-09 | Julia Collins | 8:29 |  | Completed |  |  |  |  | 80827 | 2000000 |  |  | 80685 80686 80696 80697 80698 80743 | 80685~80686~80696~80697~80698~80743 | 3327 |  | 3327 3327 3327 3327 3327 3327 | 2021-02-11 | 12:10 | albert zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 0.99874 | 0.00152 | 80685 |  | XCAMS | ST | 0% |  | TW3327 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 29x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks The result for the blank in this wheel is very low, much like a Kauri wood combustion blank. unknowns: * no failures | 2 | RRL | [5 empty] | Background Inorganic |  | 2021-02-12 | Water | Complete | JOB-216938 |  |  | 71224 | Normal | 14047/11 |  |  | 2021-02-03 | 5080 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00157 | 80686 | [26 empty] | 2021-02-04 | 20837 | Graphite | Completed
[168 empty] | 1.0009 | 0.00159 | 80696 | [26 empty] | 2021-02-10 | 7803 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00156 | 80697 | [26 empty] | 2021-02-12 | 21998 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00156 | 80698 | [26 empty] | 2021-02-12 | 26132 | Close Job | Completed
[168 empty] | 0.99949 | 0.00151 | 80743
[5 empty] | 3.7775e-05 | 0.0045926 | 80.5 | 88.7593 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7096 | 7096 | 7096 | 40 | 39 | 43 | 7056.4 | 7136.3 | Julia Collins | -586.62 | 2.06 ‰ | 2021-02-09 | 2020-08-13 | 5 | 180 |  |  |  |  | -7.55‰ | -7.55‰ | 2.00‰ | 2.00‰ | -7.55 ‰ | 2.00 ‰ | AMS | -8.36‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -586.62 ‰ | -586.62 ‰ | 2.06 | 2.26 | -581.52 | 2.08 ‰ | -581.52 | 2.08 ‰ |  |  | 2021-02-12 | Cathy Ginnane | 8:15 | 0.41338 | 0.41338 | 0.00206 | 0.00226 | 0.00199 | 0.00199 | 1.9 |  |  | 0.4171 | 0.00118 | 21999 | 2021-02-11 | 11:57 | albert zondervan | 216939 | 0.00594 | 0.0059 | 0.00267 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.34 | 0.21% | 20 | 5 | 4 | ... | 0.40114 | 0.00101 |  |  |  |  | 0.98988 | 0.98988 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 43.8820657291812288. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00594 | 0.00441 | 0.0034 | 0.39756 | 0.40‰ | 0.00191 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.41338 | 0 | 0.00206 | 0 | 0 | 0.00594 | 0 | 0.00267 | 0 | 0.39756 | 0 | 0.00191 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-02-09 | Julia Collins | 8:29 |  | Completed |  |  |  |  | 80828 | 2000000 | 80827 80648 80600 80120 79961 79745 79667 79669 79711 | 80827~80648~80600~80120~79961~79745~79667~79669~79711 | 80685 80686 80696 80697 80698 80743 | 80685~80686~80696~80697~80698~80743 | 3327 | 3327 3323 3321 3310 3305 3300 3299 3299 3299 | 3327 3327 3327 3327 3327 3327 | 2021-02-11 | 12:10 | albert zondervan |  | 1.14 mg | 0.13 | 1.14 | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 | 0.99874 | 0.00152 | 80685 | ST | XCAMS | ST | 0% |  | TW3327 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 29x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks The result for the blank in this wheel is very low, much like a Kauri wood combustion blank. unknowns: * no failures | -8.4 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-02-12 | Water | Complete | JOB-216939 |  |  | 71226 | Normal | 14047/12 |  |  | 2021-02-03 | 5081 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 | 1 | 0.00157 | 80686 | EA | [25 empty] | 2021-02-04 | 20838 | Graphite | Completed | 14047/13 | IAEA-C1: Carrara Marble WATER LINE COMBINED
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 | 1.0009 | 0.00159 | 80696 | ST | [25 empty] | 2021-02-10 | 7804 | C13 Measurement | Completed
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 | 1.0006 | 0.00156 | 80697 | FLASK | [25 empty] | 2021-02-12 | 21999 | AMS Submission Results | Completed
[163 empty] | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 1.0004 | 0.00156 | 80698 | ST | [25 empty] | 2021-02-12 | 26133 | Close Job | Completed
[163 empty] | 2020-08-25 | 0.00468 | 79745 | 3300 | 0.91 | 0.99949 | 0.00151 | 80743 | ST
[163 empty] | 2020-08-14 | 0.00332 | 79667 | 3299 | 1.05 |  |  |  | ST
[163 empty] | 2020-08-14 | 0.01331 | 79669 | 3299 | 0.51
[163 empty] | 2020-08-14 | 0.00869 | 79711 | 3299 | 0.5
[5 empty] | 8.4455e-05 | 0.01003 | 85.9 | 89.0694 | 0.0704 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7054 | 7053 | 7053 | 61 | 60 | 63 | 6993.3 | 7114.5 | Julia Collins | -584.43 | 3.14 ‰ | 2021-03-03 | 2020-09-04 | 1 | 180 |  |  |  |  | -8.63‰ | -8.63‰ | 2.00‰ | 2.00‰ | -8.63 ‰ | 2.00 ‰ | AMS | -8.00‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -584.43 ‰ | -584.43 ‰ | 3.14 | 3.27 | -579.29 | 3.17 ‰ | -579.29 | 3.17 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:25 | 0.41557 | 0.41557 | 0.00314 | 0.00327 | 0.00309 | 0.00309 | 2.8 |  |  | 0.42202 | 0.00111 | 22158 | 2021-03-05 | 9:41 | Jeremy Thompson | 216997 | 0.01032 | 0.0103 | 0.00464 | 0.0046 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.56 | 0.31% | 15 | 5 | 6 | ... | 0.40587 | 0.00093 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes |  | result is 1-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 63.2425714513634317. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.01032 | 0.00964 | 0.00262 | 0.39967 | 0.40‰ | 0.00297 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.41557 | 0 | 0.00314 | 0 | 0 | 0.01032 | 0 | 0.00464 | 0 | 0.39967 | 0 | 0.00297 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:12 |  | Completed |  |  |  |  | 81048 | 2000000 | 81047 81049 79669 79711 | 81047~81049~79669~79711 | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 | 3331 3331 3299 3299 | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 0.76 mg | 0.13 | 0.76 | 2021-03-03 | 0.00763 | 81047 | 3331 | 0.75 | 0.9993 | 0.00144 | 80926 | ST | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-03-08 | Water |  | JOB-216997 |  | special secondary. Combine 2 small portions of CO2 to make secondary for check on special water blanks to match R41552/4-5 | 71425 | Normal | 14047/12 |  |  | 2021-03-01 | 5122 | Water CO2 Evolution | Completed | 14047/13 | IAEA-C1: Carrara Marble WATER LINE COMBINED | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-03-03 | 0.01163 | 81049 | 3331 | 0.64 | 0.99767 | 0.0015 | 80933 | EA | [25 empty] | 2021-03-01 | 5126 | Water CO2 Evolution | Completed
[163 empty] | 2020-08-14 | 0.01331 | 79669 | 3299 | 0.51 | 1.0018 | 0.00143 | 80935 | [26 empty] | 2021-03-01 | 292 | Combine Water CO2 Evolution | Completed
[163 empty] | 2020-08-14 | 0.00869 | 79711 | 3299 | 0.5 | 1.0012 | 0.00164 | 80965 | [26 empty] | 2021-03-02 | 21058 | Graphite | Completed
[168 empty] | 0.99925 | 0.00145 | 80977 | [26 empty] | 2021-03-05 | 7843 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00146 | 80981 | [26 empty] | 2021-03-05 | 22158 | AMS Submission Results | Completed
[197 empty] | 2021-03-08 | 26343 | Close Job | Completed
[7 empty] | 91 | 87.9221 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45934 | 45933 | 45933 | 229 | 228 | 228 |  |  | Julia Collins | -996.71 | 0.09 ‰ | 2021-03-03 | 2020-09-04 | 0 | 180 |  |  |  |  | 4.30‰ | 4.30‰ | 2.00‰ | 2.00‰ | 4.30 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -996.71 ‰ | -996.71 ‰ | 0.09 | 0.09 | -996.67 | 0.09 ‰ | -996.67 | 0.09 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:26 | 0.0032858 | 0.0032858 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00329 | 9e-05 | 22178 | 2021-03-05 | 9:41 | Jeremy Thompson | 217129 |  |  |  |  | 45% | 5 | 5/03/21 | 11:25:55 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 22 | 5 | 4 | ... | 0.00316 | 9e-05 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes | [5 empty] | 0.00316 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00329 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00316 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 81063 | 2000000 |  |  | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 |  | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 1.05 mg |  | 1.05 | [5 empty] | 0.9993 | 0.00144 | 80926 |  | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-03-23 | Water | Complete | JOB-217129 |  | reg sized blank | 71445 | Normal | 14047/11 |  |  | 2021-03-02 | 5128 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99767 | 0.0015 | 80933 | [26 empty] | 2021-03-03 | 21073 | Graphite | Completed
[168 empty] | 1.0018 | 0.00143 | 80935 | [26 empty] | 2021-03-22 | 7846 | C13 Measurement | Completed
[168 empty] | 1.0012 | 0.00164 | 80965 | [26 empty] | 2021-03-05 | 22178 | AMS Submission Results | Completed
[168 empty] | 0.99925 | 0.00145 | 80977 | [26 empty] | 2021-03-22 | 26533 | Close Job | Completed
[168 empty] | 1.0009 | 0.00146 | 80981
[5 empty] | 4.1111e-05 | 0.0039621 | 90.2 | 88.7775 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7099 | 7098 | 7098 | 32 | 32 | 37 | 7066.5 | 7131.4 | Julia Collins | -586.75 | 1.67 ‰ | 2021-03-03 | 2020-09-04 | 0 | 180 |  |  |  |  | -5.38‰ | -5.38‰ | 2.00‰ | 2.00‰ | -5.38 ‰ | 2.00 ‰ | AMS | -8.15‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -586.75 ‰ | -586.75 ‰ | 1.67 | 1.91 | -581.65 | 1.69 ‰ | -581.65 | 1.69 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:26 | 0.41325 | 0.41325 | 0.00167 | 0.00191 | 0.00158 | 0.00158 | 2.1 |  |  | 0.41602 | 0.0011 | 22179 | 2021-03-05 | 9:41 | Jeremy Thompson | 217130 | 0.00443 | 0.0044 | 0.00199 | 0.002 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.32 | 0.17% | 29 | 5 | 4 | ... | 0.4001 | 0.00092 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 37.1497874799299667. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00443 | 0.00381 | 0.00179 | 0.39743 | 0.40‰ | 0.00152 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.41325 | 0 | 0.00167 | 0 | 0 | 0.00443 | 0 | 0.00199 | 0 | 0.39743 | 0 | 0.00152 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 81050 | 2000000 | 81063 80827 80648 80600 80120 79961 | 81063~80827~80648~80600~80120~79961 | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 | 3331 3327 3323 3321 3310 3305 | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 1.07 mg | 0.13 | 1.07 | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 | 0.9993 | 0.00144 | 80926 | ST | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-03-23 | Water | Complete | JOB-217130 |  | reg sized secondary | 71446 | Normal | 14047/12 |  |  | 2021-03-02 | 5129 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 | 0.99767 | 0.0015 | 80933 | FLASK | [25 empty] | 2021-03-03 | 21060 | Graphite | Completed
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 | 1.0018 | 0.00143 | 80935 | ST | [25 empty] | 2021-03-05 | 7842 | C13 Measurement | Completed
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 | 1.0012 | 0.00164 | 80965 | ST | [25 empty] | 2021-03-05 | 22179 | AMS Submission Results | Completed
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 | 0.99925 | 0.00145 | 80977 | ST | [25 empty] | 2021-03-08 | 26349 | Close Job | Completed
[163 empty] | 2020-09-16 | 0.0072 | 79961 | 3305 | 1.05 | 1.0009 | 0.00146 | 80981 | EA
[7 empty] | 87.9 | 88.0441 | 0.002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35467 | 35466 | 35466 | 117 | 117 | 117 |  |  | Julia Collins | -987.91 | 0.18 ‰ | 2021-03-03 | 2020-09-04 | 1 | 180 |  |  |  |  | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -987.91 ‰ | -987.91 ‰ | 0.18 | 0.18 | -987.76 | 0.18 ‰ | -987.76 | 0.18 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:26 | 0.012093 | 0.012093 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 3.4 |  |  | 0.01209 | 0.00018 | 22188 | 2021-03-05 | 9:41 | Jeremy Thompson | 217312 |  |  |  |  | 45% | 5 | 5/03/21 | 11:26:29 | Jocelyn Turnbull | 0 |  | 1.21 | 0.02% | 1 | 5 | 6 | ... | 0.01163 | 0.00017 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes | [5 empty] | 0.01163 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.01209 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.01163 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 81049 | 2000000 |  |  | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 |  | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 0.64 mg |  | 0.64 | [5 empty] | 0.9993 | 0.00144 | 80926 |  | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-04-21 | Water | Complete | JOB-217312 |  | aim for 0.61mg (combined CO2) special blank to match size of R41552/5 | 71455 | Normal | 14047/13 |  |  | 2021-03-01 | 5123 | Water CO2 Evolution | Completed |  |  | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Use this R 14047/11 for CO2 produced on water line AND COMBINED, to match with water samples that are combined from two water extractions. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. Use R14047/11 for single shot water line samples. |  |  | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C1: Carrara Marble WATER LINE COMBINED | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate. This is for samples where two water extractions are combined for graphitisation.
[168 empty] | 0.99767 | 0.0015 | 80933 | [26 empty] | 2021-03-01 | 5127 | Water CO2 Evolution | Completed
[168 empty] | 1.0018 | 0.00143 | 80935 | [26 empty] | 2021-03-01 | 293 | Combine Water CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00164 | 80965 | [26 empty] | 2021-03-02 | 21059 | Graphite | Completed
[168 empty] | 0.99925 | 0.00145 | 80977 | [26 empty] | 2021-03-05 | 7844 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00146 | 80981 | [26 empty] | 2021-03-05 | 22188 | AMS Submission Results | Completed
[197 empty] | 2021-03-08 | 26351 | Close Job | Completed
[7 empty] | 89.1 | 88.0395 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38853 | 38852 | 38852 | 137 | 136 | 136 |  |  | Julia Collins | -992.07 | 0.14 ‰ | 2021-03-03 | 2020-09-04 | 1 | 180 |  |  |  |  | 2.96‰ | 2.96‰ | 2.00‰ | 2.00‰ | 2.96 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -992.07 ‰ | -992.07 ‰ | 0.14 | 0.14 | -991.97 | 0.14 ‰ | -991.97 | 0.14 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:27 | 0.0079336 | 0.0079336 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.8 |  |  | 0.00793 | 0.00014 | 22189 | 2021-03-05 | 9:41 | Jeremy Thompson | 217313 |  |  |  |  | 45% | 5 | 5/03/21 | 11:26:31 | Jocelyn Turnbull | 0 |  | 0.79 | 0.01% | 8 | 5 | 6 | ... | 0.00763 | 0.00013 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes | [5 empty] | 0.00763 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00793 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00763 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 81047 | 2000000 |  |  | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 |  | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 0.75 mg |  | 0.75 | [5 empty] | 0.9993 | 0.00144 | 80926 |  | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-04-21 | Water | Not Complete | JOB-217313 |  | aim for 0.69mg (combined CO2) special blank to match size R41552/4 | 71456 | Normal | 14047/13 |  |  | 2021-03-01 | 5124 | Water CO2 Evolution | Completed |  |  | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Use this R 14047/11 for CO2 produced on water line AND COMBINED, to match with water samples that are combined from two water extractions. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. Use R14047/11 for single shot water line samples. |  |  | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C1: Carrara Marble WATER LINE COMBINED | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate. This is for samples where two water extractions are combined for graphitisation.
[168 empty] | 0.99767 | 0.0015 | 80933 | [26 empty] | 2021-03-01 | 5125 | Water CO2 Evolution | Completed
[168 empty] | 1.0018 | 0.00143 | 80935 | [26 empty] | 2021-03-01 | 294 | Combine Water CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00164 | 80965 | [26 empty] | 2021-03-02 | 21057 | Graphite | Completed
[168 empty] | 0.99925 | 0.00145 | 80977 | [26 empty] | 2021-03-05 | 7845 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00146 | 80981 | [26 empty] | 2021-03-05 | 22189 | AMS Submission Results | Completed
[197 empty] | 2021-03-08 | 26352 | Close Job | Completed
[7 empty] | 89.2 | 87.7381 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40341 | 40340 | 40340 | 152 | 152 | 152 |  |  | Julia Collins | -993.41 | 0.12 ‰ | 2021-03-08 | 2020-09-09 | 3 | 180 |  |  |  |  | 4.62‰ | 4.62‰ | 2.00‰ | 2.00‰ | 4.62 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.41 ‰ | -993.41 ‰ | 0.12 | 0.12 | -993.33 | 0.13 ‰ | -993.33 | 0.13 ‰ |  |  | 2021-03-10 | Jocelyn Turnbull | 9:55 | 0.0065923 | 0.0065923 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00659 | 0.00012 | 22231 | 2021-03-10 | 9:04 | Jeremy Thompson | 216996 |  |  |  |  | 45% | 5 | 10/03/21 | 09:54:54 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 1 | 5 | 4 | ... | 0.00634 | 0.00012 |  |  |  |  | 0.99048 | 0.99048 | Yes | Yes | [5 empty] | 0.00634 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00659 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00634 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-08 | Julia Collins | 8:44 |  | Completed |  |  |  |  | 81078 | 2000000 |  |  | 81022 81061 81062 81090 81091 81092 | 81022~81061~81062~81090~81091~81092 | 3333 |  | 3333 3333 3333 3333 3333 3333 | 2021-03-10 | 9:38 | Jeremy Thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99767 | 0.00162 | 81022 |  | XCAMS | EA | 0% |  | TW3333 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * FIRI-D: cellulose 1x result agrees with consensus value blanks: * 2 groups identified to blank-correct: Water: 13x unknowns + 0x secondaries + 1x blanks(terrible blank) 180 days: capturing 7x blanks AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-03-08 | Water |  | JOB-216996 |  |  | 71497 | Normal | 14047/11 |  |  | 2021-03-04 | 5142 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00158 | 81061 | [26 empty] | 2021-03-05 | 21088 | Graphite | Completed
[168 empty] | 0.99992 | 0.00157 | 81062 | [26 empty] | 2021-03-22 | 7863 | C13 Measurement | Completed
[168 empty] | 0.99984 | 0.00159 | 81090 | [26 empty] | 2021-03-10 | 22231 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00156 | 81091 | [26 empty] | 2021-03-30 | 26548 | Close Job | Completed
[168 empty] | 1.0016 | 0.00158 | 81092
[7 empty] | 101.1 | 87.8528 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43257 | 43256 | 43256 | 200 | 200 | 200 |  |  | Julia Collins | -995.41 | 0.11 ‰ | 2021-03-16 | 2020-09-17 | 5 | 180 |  |  |  |  | 2.41‰ | 2.41‰ | 2.00‰ | 2.00‰ | 2.41 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -995.41 ‰ | -995.41 ‰ | 0.11 | 0.11 | -995.36 | 0.12 ‰ | -995.36 | 0.12 ‰ |  |  | 2021-03-18 | Jocelyn Turnbull | 14:11 | 0.0045855 | 0.0045855 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.7 |  |  | 0.00459 | 0.00011 | 22270 | 2021-03-18 | 13:44 | albert zondervan | 214980 |  |  |  |  | 45% | 5 | 22/03/21 | 10:14:27 | jenny dahl | 0 |  | 0.46 | 0.01% | 1 | 5 | 7 | ... | 0.00441 | 0.00011 |  |  |  |  | 0.9896 | 0.9896 | Yes | Yes | [5 empty] | 0.00441 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00459 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00441 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-15 | Julia Collins | 15:10 |  | Completed |  |  |  |  | 81139 | 2000000 |  |  | 81013 81017 81020 81053 81057 81059 | 81013~81017~81020~81053~81057~81059 | 3334 |  | 3334 3334 3334 3334 3334 3334 | 2021-03-18 | 13:52 | albert zondervan |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99968 | 0.00164 | 81013 |  | XCAMS | EA | 0% |  | TW3334 measurement: * large waters + large carbonates CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on carbonate line: 1x result agree with consensus value * IAEA-C2 on water line: 1x result agree with consensus value blanks: * 2 groups identified to blank-correct: water: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks carbonate: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-06-24 | Carbonate | Complete | JOB-214980 |  | C13 split sealed as SIL could not run. | 71537 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0007 | 0.00166 | 81017 | [26 empty] | 2021-03-10 | 5161 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99929 | 0.00161 | 81020 | [26 empty] | 2021-03-10 | 3036 | Gas Splitting | Completed
[168 empty] | 0.99856 | 0.00161 | 81053 | [26 empty] | 2021-03-10 | 5077 | Choose Gas Splitting | Completed
[168 empty] | 0.99921 | 0.0016 | 81057 | [26 empty] | 2021-03-11 | 21149 | Graphite | Completed
[168 empty] | 1.0026 | 0.00163 | 81059 | [26 empty] | 2021-03-22 | 7871 | C13 Measurement | Completed
[197 empty] | 2021-03-18 | 22270 | AMS Submission Results | Completed
[197 empty] | 2021-03-30 | 26545 | Close Job | Completed
[7 empty] | 102.3 | 87.6554 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40391 | 40391 | 40391 | 153 | 153 | 153 |  |  | Julia Collins | -993.45 | 0.12 ‰ | 2021-03-16 | 2020-09-17 | 4 | 180 |  |  |  |  | 4.67‰ | 4.67‰ | 2.00‰ | 2.00‰ | 4.67 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.45 ‰ | -993.45 ‰ | 0.12 | 0.12 | -993.37 | 0.13 ‰ | -993.37 | 0.13 ‰ |  |  | 2021-03-18 | Jocelyn Turnbull | 14:12 | 0.0065507 | 0.0065507 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00655 | 0.00012 | 22300 | 2021-03-18 | 13:44 | albert zondervan | 217217 |  |  |  |  | 45% | 5 | 22/03/21 | 10:14:27 | jenny dahl | 0 |  | 0.66 | 0.01% | 27 | 5 | 6 | ... | 0.0063 | 0.00012 |  |  |  |  | 0.9896 | 0.9896 | Yes | Yes | [5 empty] | 0.0063 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00655 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0063 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-15 | Julia Collins | 15:10 |  | Completed |  |  |  |  | 81152 | 2000000 |  |  | 81013 81017 81020 81053 81057 81059 | 81013~81017~81020~81053~81057~81059 | 3334 |  | 3334 3334 3334 3334 3334 3334 | 2021-03-18 | 13:52 | albert zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99968 | 0.00164 | 81013 |  | XCAMS | EA | 0% |  | TW3334 measurement: * large waters + large carbonates CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on carbonate line: 1x result agree with consensus value * IAEA-C2 on water line: 1x result agree with consensus value blanks: * 2 groups identified to blank-correct: water: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks carbonate: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-04-13 | Water | Complete | JOB-217217 |  |  | 71567 | Normal | 14047/11 |  |  | 2021-03-11 | 5166 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00166 | 81017 | [26 empty] | 2021-03-11 | 3038 | Gas Splitting | Completed
[168 empty] | 0.99929 | 0.00161 | 81020 | [26 empty] | 2021-03-11 | 5080 | Choose Gas Splitting | Completed
[168 empty] | 0.99856 | 0.00161 | 81053 | [26 empty] | 2021-03-12 | 21162 | Graphite | Completed
[168 empty] | 0.99921 | 0.0016 | 81057 | [26 empty] | 2021-07-16 | 7874 | C13 Measurement | Completed
[168 empty] | 1.0026 | 0.00163 | 81059 | [26 empty] | 2021-03-18 | 22300 | AMS Submission Results | Completed
[197 empty] | 2021-07-28 | 27521 | Close Job | Completed
[5 empty] | 3.9046e-05 | 0.004128 | 101.9 | 88.7977 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7071 | 7071 | 7071 | 34 | 34 | 38 | 7037 | 7105.6 | Julia Collins | -585.33 | 1.77 ‰ | 2021-03-16 | 2020-09-17 | 4 | 180 |  |  |  |  | -8.25‰ | -8.25‰ | 2.00‰ | 2.00‰ | -8.25 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -585.33 ‰ | -585.33 ‰ | 1.77 | 2 | -580.21 | 1.79 ‰ | -580.21 | 1.79 ‰ |  |  | 2021-03-18 | Jocelyn Turnbull | 14:12 | 0.41467 | 0.41467 | 0.00177 | 0.002 | 0.00168 | 0.00168 | 3.1 |  |  | 0.41753 | 0.0012 | 22301 | 2021-03-18 | 13:44 | albert zondervan | 217218 | 0.00458 | 0.0046 | 0.00206 | 0.0021 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.47 | 0.18% | 36 | 5 | 6 | ... | 0.40155 | 0.00103 |  |  |  |  | 0.9896 | 0.9896 | Yes | Yes |  | result agree with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 38.7431845210782707. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00458 | 0.00397 | 0.0016 | 0.3988 | 0.40‰ | 0.00162 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.41467 | 0 | 0.00177 | 0 | 0 | 0.00458 | 0 | 0.00206 | 0 | 0.3988 | 0 | 0.00162 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-15 | Julia Collins | 15:10 |  | Completed |  |  |  |  | 81153 | 2000000 | 81152 81078 81063 80827 80648 80600 80120 | 81152~81078~81063~80827~80648~80600~80120 | 81013 81017 81020 81053 81057 81059 | 81013~81017~81020~81053~81057~81059 | 3334 | 3334 3333 3331 3327 3323 3321 3310 | 3334 3334 3334 3334 3334 3334 | 2021-03-18 | 13:52 | albert zondervan |  | 0.68 mg | 0.13 | 0.68 | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99968 | 0.00164 | 81013 | FLASK | XCAMS | EA | 0% |  | TW3334 measurement: * large waters + large carbonates CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on carbonate line: 1x result agree with consensus value * IAEA-C2 on water line: 1x result agree with consensus value blanks: * 2 groups identified to blank-correct: water: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks carbonate: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-04-13 | Water | Complete | JOB-217218 |  |  | 71568 | Normal | 14047/12 |  |  | 2021-03-11 | 5167 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 | 1.0007 | 0.00166 | 81017 | ST | [25 empty] | 2021-03-11 | 3039 | Gas Splitting | Completed
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 | 0.99929 | 0.00161 | 81020 | ST | [25 empty] | 2021-03-11 | 5081 | Choose Gas Splitting | Completed
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 | 0.99856 | 0.00161 | 81053 | ST | [25 empty] | 2021-03-12 | 21163 | Graphite | Completed
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 | 0.99921 | 0.0016 | 81057 | EA | [25 empty] | 2021-05-31 | 7875 | C13 Measurement | Completed
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 | 1.0026 | 0.00163 | 81059 | EA | [25 empty] | 2021-03-18 | 22301 | AMS Submission Results | Completed
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 |  |  |  | EA | [25 empty] | 2021-06-14 | 27275 | Close Job | Completed
[7 empty] | 99.9 | 87.8211 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41097 | 41097 | 41097 | 153 | 152 | 152 |  |  | taylor ferrick | -994 | 0.11 ‰ | 2021-03-31 | 2020-10-02 | 7 | 180 |  |  |  |  | 3.39‰ | 3.39‰ | 2.00‰ | 2.00‰ | 3.39 ‰ | 2.00 ‰ | AMS | 2.60‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -994.00 ‰ | -994.00 ‰ | 0.11 | 0.11 | -993.93 | 0.12 ‰ | -993.93 | 0.12 ‰ |  |  | 2021-04-08 | Jocelyn Turnbull | 11:42 | 0.0059996 | 0.0059996 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.006 | 0.00011 | 22390 | 2021-04-02 | 17:20 | albert zondervan | 217395 |  |  |  |  | 45% | 5 | 8/04/21 | 11:41:50 | Jocelyn Turnbull | 0 |  | 0.6 | 0.01% | 1 | 5 | 5 | ... | 0.00577 | 0.00011 |  |  |  |  | 0.99021 | 0.99021 | Yes | Yes |  | an anomalously high blank |  |  |  | 0.00577 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.006 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00577 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-30 | taylor ferrick | 8:33 |  | Completed |  |  |  |  | 81212 | 2000000 |  |  | 81198 81217 81227 81229 81234 81235 | 81198~81217~81227~81229~81234~81235 | 3337 |  | 3337 3337 3337 3337 3337 3337 |  |  |  |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99911 | 0.00197 | 81198 |  | XCAMS | ST | 0.13% |  | TW3337 measurement: * large AAA-ST samples * small carbonate samples CalAMS error adjustments: * 0.13% system error primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I cellulose 1x result agrees with consensus value * IAEA-C2 travertine large: 1x result agrees with consensus value small: 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: AAA-ST: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks large carbonates: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks IAEA-C1 (TP81212): an anomalously high blank small carbonates: 8x unknowns + 1x secondaries + 2x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheel selected also, to assist with determining d, m parameters: TW3296 d = 0.97 ug C m = 0.99 ug C unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-05-05 | Carbonate | Not Complete | JOB-217395 |  | regular sized blank | 71656 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0036 | 0.00195 | 81217 | [26 empty] | 2021-03-22 | 5174 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99997 | 0.00194 | 81227 | [26 empty] | 2021-03-24 | 21222 | Graphite | Completed
[168 empty] | 0.99784 | 0.00192 | 81229 | [26 empty] | 2021-03-25 | 7889 | C13 Measurement | Completed
[168 empty] | 0.9998 | 0.00198 | 81234 | [26 empty] | 2021-04-08 | 22390 | AMS Submission Results | Completed
[168 empty] | 0.99969 | 0.00192 | 81235 | [26 empty] | 2021-04-08 | 26588 | Close Job | Completed
[7 empty] | 67.8 | 88.5078 | 0.001 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41352 | 41351 | 41351 | 187 | 186 | 186 |  |  | taylor ferrick | -994.19 | 0.13 ‰ | 2021-03-31 | 2020-10-02 | 7 | 180 |  |  |  |  | -4.39‰ | -4.39‰ | 2.00‰ | 2.00‰ | -4.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0011 | 0 | 0.8165 | -994.19 ‰ | -994.19 ‰ | 0.13 | 0.13 | -994.12 | 0.14 ‰ | -994.12 | 0.14 ‰ |  |  | 2021-04-08 | Jocelyn Turnbull | 11:42 | 0.0058125 | 0.0058125 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 14.4 |  |  | 0.00581 | 0.00013 | 22407 | 2021-04-02 | 17:20 | albert zondervan | 217396 |  |  |  |  | 45% | 5 | 8/04/21 | 11:41:51 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 32 | 5 | 4 | ... | 0.00559 | 0.00013 |  |  |  |  | 0.99021 | 0.99021 | Yes | Yes | [5 empty] | 0.00559 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.00581 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00559 | 0 | 0.00013 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-03-30 | taylor ferrick | 8:33 |  | Completed |  |  |  |  | 81203 | 2000000 |  |  | 81198 81217 81227 81229 81234 81235 | 81198~81217~81227~81229~81234~81235 | 3337 |  | 3337 3337 3337 3337 3337 3337 |  |  |  |  | 0.16 mg |  | 0.16 | [5 empty] | 0.99911 | 0.00197 | 81198 |  | XCAMS | ST | 0.13% |  | TW3337 measurement: * large AAA-ST samples * small carbonate samples CalAMS error adjustments: * 0.13% system error primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I cellulose 1x result agrees with consensus value * IAEA-C2 travertine large: 1x result agrees with consensus value small: 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: AAA-ST: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks large carbonates: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks IAEA-C1 (TP81212): an anomalously high blank small carbonates: 8x unknowns + 1x secondaries + 2x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheel selected also, to assist with determining d, m parameters: TW3296 d = 0.97 ug C m = 0.99 ug C unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2021-05-05 | Carbonate |  | JOB-217396 |  | SLUSHIE small blank to match sample | 71657 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0036 | 0.00195 | 81217 | [26 empty] | 2021-03-22 | 5175 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99997 | 0.00194 | 81227 | [26 empty] | 2021-03-24 | 21213 | Graphite | Completed
[168 empty] | 0.99784 | 0.00192 | 81229 | [26 empty] | 2021-04-08 | 22407 | AMS Submission Results | Completed
[168 empty] | 0.9998 | 0.00198 | 81234 | [26 empty] | 2021-04-08 | 26589 | Close Job | Completed
[168 empty] | 0.99969 | 0.00192 | 81235
[7 empty] | 90.7 | 88.2909 | 0.001 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41097 | 41097 | 41097 | 153 | 152 | 152 |  |  | taylor ferrick | -994 | 0.11 ‰ | 2021-03-31 | 2020-10-02 | 5 | 180 |  |  |  |  | -1.95‰ | -1.95‰ | 2.00‰ | 2.00‰ | -1.95 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0011 | 0 | 0.8165 | -994.00 ‰ | -994.00 ‰ | 0.11 | 0.11 | -993.93 | 0.12 ‰ | -993.93 | 0.12 ‰ |  |  | 2021-04-08 | Jocelyn Turnbull | 11:42 | 0.0059996 | 0.0059996 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 9.3 |  |  | 0.006 | 0.00011 | 22411 | 2021-04-02 | 17:20 | albert zondervan | 217530 |  |  |  |  | 45% | 5 | 8/04/21 | 11:41:52 | Jocelyn Turnbull | 0 |  | 0.6 | 0.01% | 10 | 5 | 4 | ... | 0.00577 | 0.00011 |  |  |  |  | 0.99021 | 0.99021 | Yes | Yes | [5 empty] | 0.00577 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.006 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00577 | 0 | 0.00011 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-03-30 | taylor ferrick | 8:33 |  | Completed |  |  |  |  | 81241 | 2000000 |  |  | 81198 81217 81227 81229 81234 81235 | 81198~81217~81227~81229~81234~81235 | 3337 |  | 3337 3337 3337 3337 3337 3337 |  |  |  |  | 0.26 mg |  | 0.26 | [5 empty] | 0.99911 | 0.00197 | 81198 |  | XCAMS | ST | 0.13% |  | TW3337 measurement: * large AAA-ST samples * small carbonate samples CalAMS error adjustments: * 0.13% system error primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I cellulose 1x result agrees with consensus value * IAEA-C2 travertine large: 1x result agrees with consensus value small: 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: AAA-ST: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks large carbonates: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks IAEA-C1 (TP81212): an anomalously high blank small carbonates: 8x unknowns + 1x secondaries + 2x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheel selected also, to assist with determining d, m parameters: TW3296 d = 0.97 ug C m = 0.99 ug C unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2021-05-18 | Carbonate | Complete | JOB-217530 |  | one more small blank for luck make ~0.22-mg C ish | 71658 | Normal | 14047/1 |  |  | 2020-01-10 | 6249 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0036 | 0.00195 | 81217 | [26 empty] | 2021-03-25 | 5181 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99997 | 0.00194 | 81227 | [26 empty] | 2021-03-26 | 21251 | Graphite | Completed
[168 empty] | 0.99784 | 0.00192 | 81229 | [26 empty] | 2021-04-08 | 22411 | AMS Submission Results | Completed
[168 empty] | 0.9998 | 0.00198 | 81234 | [26 empty] | 2021-04-08 | 26596 | Close Job | Completed
[168 empty] | 0.99969 | 0.00192 | 81235
[5 empty] | 3.5346e-05 | 0.0043157 | 95.7 | 88.7759 | 0.0697 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7060 | 7059 | 7059 | 38 | 37 | 41 | 7022.4 | 7097.7 | Julia Collins | -584.75 | 1.95 ‰ | 2021-04-04 | 2020-10-06 | 3 | 180 |  |  |  |  | -4.56‰ | -4.56‰ | 2.00‰ | 2.00‰ | -4.56 ‰ | 2.00 ‰ | AMS | -8.18‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.89443 | -584.75 ‰ | -584.75 ‰ | 1.95 | 2.16 | -579.62 | 1.97 ‰ | -579.62 | 1.97 ‰ |  |  | 2021-04-19 | Jocelyn Turnbull | 15:08 | 0.41525 | 0.41525 | 0.00195 | 0.00216 | 0.00187 | 0.00153 | 1.9 |  |  | 0.41814 | 0.00145 | 22466 | 2021-04-08 | 9:20 | albert zondervan | 217315 | 0.00463 | 0.0046 | 0.00208 | 0.0021 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.53 | 0.19% | 20 | 5 | 4 | ... | 0.40214 | 0.00129 |  |  |  |  | 0.99305 | 0.99305 | Yes |  |  | result is ca. 2-sigma younger than consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 41.7985913144596719. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00463 | 0.00415 | 0.00149 | 0.39936 | 0.40‰ | 0.0018 | 0.00‰ | No | 1 | 1 | 0.00128 | 0.00128 | 0.41525 | 0 | 0.00195 | 0 | 0 | 0.00463 | 0 | 0.00208 | 0 | 0.39936 | 0 | 0.0018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-01 | Julia Collins | 14:19 |  | Completed |  |  |  |  | 81315 | 2000000 | 81316 81152 81078 81063 80827 80648 80600 80120 | 81316~81152~81078~81063~80827~80648~80600~80120 | 81099 81100 81111 81290 81291 | 81099~81100~81111~81290~81291 | 3338 | 3338 3334 3333 3331 3327 3323 3321 3310 | 3338 3338 3338 3338 3338 | 2021-04-08 | 9:40 | albert zondervan |  | 1.00 mg | 0.13 | 1 | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 0.99969 | 0.00286 | 81099 | FLASK | XCAMS | FLASK | 0.26% |  | *** JCT/AZ results of further tests show that the flask Ox is behaving normally again. Our best guess (but only a guess) is room air contamination on the graphite for hte flask 1 Ox in this wheel. Also not sure about the EA Ox in this wheel, likely a separate issue. System error is 0.26% using all 5 flask Oxs, which is the same as the wheel to wheel error we currently assign to carbonates. Therefore have reported with wheel to wheel error. TW3338 measurement: * large waters * complete system retune during measurement did not resolve OxI repeatability issues primaries: * OxI cathodes: 5x accepted, from flask 1x rejected, via EA (J216823), due to being 4-sigma outlier CalAMS error adjustments: * 1.05 global error multiplier due to excess run-to-run variability * 0.26% system error: determined from excess variability in 5x accepted flask OxI-CO2 cause of this exceptionally large system error is not identifiable in XCAMS raw data variability appears associated with OxI-CO2 flask number: taking means, flask 2 has ca. 0.5% higher rts than flask 1 system error not removed in RLIMS because of its size secondaries: * IAEA-C2 on water line: 1x result differs from consensus value by ca. 2-sigma blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days capturing 8x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-04-23 | Water | Complete | JOB-217315 |  | D bank stuck at ~0.0028 for prebake | 71697 | Normal | 14047/12 |  |  | 2021-03-31 | 5195 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99725 | 0.00286 | 81100 | ST | [25 empty] | 2021-04-01 | 21325 | Graphite | Completed
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 | 0.99743 | 0.00286 | 81111 | ST | [25 empty] | 2021-04-06 | 7919 | C13 Measurement | Completed
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 | 1.0021 | 0.00288 | 81290 | ST | [25 empty] | 2021-04-19 | 22466 | AMS Submission Results | Completed
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 | 1.0038 | 0.00289 | 81291 | EA | [25 empty] | 2021-04-22 | 26670 | Close Job | Completed
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 |  |  |  | EA
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 |  |  |  | EA
[163 empty] | 2020-10-13 | 0.0056 | 80120 | 3310 | 0.95 |  |  |  | FLASK
[7 empty] | 92.3 | 87.9469 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42296 | 42296 | 42296 | 129 | 129 | 129 |  |  | Julia Collins | -994.83 | 0.08 ‰ | 2021-04-04 | 2020-10-06 | 3 | 180 |  |  |  |  | 4.83‰ | 4.83‰ | 2.00‰ | 2.00‰ | 4.83 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.89443 | -994.83 ‰ | -994.83 ‰ | 0.08 | 0.08 | -994.77 | 0.08 ‰ | -994.77 | 0.08 ‰ |  |  | 2021-04-19 | Jocelyn Turnbull | 15:08 | 0.0051678 | 0.0051678 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00517 | 8e-05 | 22467 | 2021-04-08 | 9:20 | albert zondervan | 217370 |  |  |  |  | 45% | 5 | 19/04/21 | 15:07:30 | Jocelyn Turnbull | 0 |  | 0.52 | 0.01% | 1 | 5 | 4 | ... | 0.00497 | 8e-05 |  |  |  |  | 0.99305 | 0.99305 | Yes | [6 empty] | 0.00497 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 1 | 0.00128 | 0.00128 | 0.00517 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00497 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-01 | Julia Collins | 14:19 |  | Completed |  |  |  |  | 81316 | 2000000 |  |  | 81099 81100 81111 81290 81291 | 81099~81100~81111~81290~81291 | 3338 |  | 3338 3338 3338 3338 3338 | 2021-04-08 | 9:40 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 0.99969 | 0.00286 | 81099 |  | XCAMS | FLASK | 0.26% |  | *** JCT/AZ results of further tests show that the flask Ox is behaving normally again. Our best guess (but only a guess) is room air contamination on the graphite for hte flask 1 Ox in this wheel. Also not sure about the EA Ox in this wheel, likely a separate issue. System error is 0.26% using all 5 flask Oxs, which is the same as the wheel to wheel error we currently assign to carbonates. Therefore have reported with wheel to wheel error. TW3338 measurement: * large waters * complete system retune during measurement did not resolve OxI repeatability issues primaries: * OxI cathodes: 5x accepted, from flask 1x rejected, via EA (J216823), due to being 4-sigma outlier CalAMS error adjustments: * 1.05 global error multiplier due to excess run-to-run variability * 0.26% system error: determined from excess variability in 5x accepted flask OxI-CO2 cause of this exceptionally large system error is not identifiable in XCAMS raw data variability appears associated with OxI-CO2 flask number: taking means, flask 2 has ca. 0.5% higher rts than flask 1 system error not removed in RLIMS because of its size secondaries: * IAEA-C2 on water line: 1x result differs from consensus value by ca. 2-sigma blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days capturing 8x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-05-03 | Water | Complete | JOB-217370 |  | D bank stuck at ~0.0028 for prebake | 71696 | Normal | 14047/11 |  |  | 2021-03-31 | 5196 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99725 | 0.00286 | 81100 | [26 empty] | 2021-04-01 | 21326 | Graphite | Completed
[168 empty] | 0.99743 | 0.00286 | 81111 | [26 empty] | 2021-04-06 | 7920 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00288 | 81290 | [26 empty] | 2021-04-19 | 22467 | AMS Submission Results | Completed
[168 empty] | 1.0038 | 0.00289 | 81291 | [26 empty] | 2021-04-22 | 26672 | Close Job | Completed
[5 empty] | 3.5134e-05 | 0.0043032 | 105.5 | 88.722 | 0.0695 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7088 | 7087 | 7087 | 35 | 34 | 39 | 7052.9 | 7122.4 | Julia Collins | -586.17 | 1.79 ‰ | 2021-04-21 | 2020-10-23 | 7 | 180 |  |  |  |  | -7.51‰ | -7.51‰ | 2.00‰ | 2.00‰ | -7.51 ‰ | 2.00 ‰ | AMS | -8.23‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -586.17 ‰ | -586.17 ‰ | 1.79 | 2.02 | -581.06 | 1.81 ‰ | -581.06 | 1.81 ‰ |  |  | 2021-04-22 | Cathy Ginnane | 22:53 | 0.41383 | 0.41383 | 0.00179 | 0.00202 | 0.00171 | 0.00171 | 2 |  |  | 0.4167 | 0.00122 | 22529 | 2021-04-22 | 12:35 | albert zondervan | 217314 | 0.00458 | 0.0046 | 0.00206 | 0.0021 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.38 | 0.18% | 1 | 5 | 4 | ... | 0.40075 | 0.00116 |  |  |  |  | 0.9895 | 0.9895 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 39.1732840422617579. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00458 | 0.00414 | 0.00146 | 0.39799 | 0.40‰ | 0.00171 | 0.00‰ | Yes | 1 | 0.99999 | 0.00084 | 0.00084 | 0.41383 | 0 | 0.00179 | 0 | 0 | 0.00458 | 0 | 0.00206 | 0 | 0.39799 | 0 | 0.00171 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-20 | Julia Collins | 12:58 |  | Completed |  |  |  |  | 81330 | 2000000 | 81343 81316 81152 81078 81063 80827 80648 80600 | 81343~81316~81152~81078~81063~80827~80648~80600 | 81348 81349 81350 81351 81407 81408 | 81348~81349~81350~81351~81407~81408 | 3341 | 3341 3338 3334 3333 3331 3327 3323 3321 | 3341 3341 3341 3341 3341 3341 | 2021-04-22 | 13:26 | albert zondervan |  | 1.05 mg | 0.13 | 1.05 | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 1.0029 | 0.00202 | 81348 | ST | XCAMS | ST | 0.12% |  | TW3341 measurement: * large samples CalAMS error adjustments: * 1.03 global error multiplier to account for excess run-to-run variability * 0.12% system error to account for excess Ox-I cathode-to-cathode variability primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value * FIRI-I cellulose: 2x result agrees with consensus value: R41572/4 prepared by Otlet, assuming this is a FIRI-I R24889/9 prepared by RRL blanks: * 3 groups identified to blank-correct: water: 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks RRL org: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Otlet org: 6x unknowns + 1x secondaries + 2x blanks also picking up TW3324, capturing 5x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-04-23 | Water |  | JOB-217314 |  |  | 71768 | Normal | 14047/12 |  |  | 2021-04-12 | 5198 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 1.0008 | 0.00203 | 81349 | ST | [25 empty] | 2021-04-14 | 21340 | Graphite | Completed
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 1.0013 | 0.00207 | 81350 | ST | [25 empty] | 2021-04-16 | 7921 | C13 Measurement | Completed
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 | 0.99811 | 0.00202 | 81351 | EA | [25 empty] | 2021-04-22 | 22529 | AMS Submission Results | Completed
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 | 0.99895 | 0.00212 | 81407 | EA | [25 empty] | 2021-04-23 | 26695 | Close Job | Completed
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 | 0.99781 | 0.00203 | 81408 | EA
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 |  |  |  | FLASK
[163 empty] | 2020-12-10 | 0.00463 | 80600 | 3321 | 0.6 |  |  |  | ST
[7 empty] | 104.2 | 87.8537 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41887 | 41886 | 41886 | 184 | 184 | 184 |  |  | Julia Collins | -994.56 | 0.12 ‰ | 2021-04-21 | 2020-10-23 | 7 | 180 |  |  |  |  | 2.30‰ | 2.30‰ | 2.00‰ | 2.00‰ | 2.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -994.56 ‰ | -994.56 ‰ | 0.12 | 0.12 | -994.49 | 0.13 ‰ | -994.49 | 0.13 ‰ |  |  | 2021-04-22 | Cathy Ginnane | 22:53 | 0.0054381 | 0.0054381 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00544 | 0.00012 | 22530 | 2021-04-22 | 12:35 | albert zondervan | 217316 |  |  |  |  | 45% | 5 | 22/04/21 | 22:52:34 | Cathy Ginnane | 0 |  | 0.54 | 0.01% | 15 | 5 | 6 | ... | 0.00523 | 0.00012 |  |  |  |  | 0.9895 | 0.9895 | Yes | Yes | [5 empty] | 0.00523 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 0.99999 | 0.00084 | 0.00084 | 0.00544 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00523 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-20 | Julia Collins | 12:58 |  | Completed |  |  |  |  | 81343 | 2000000 |  |  | 81348 81349 81350 81351 81407 81408 | 81348~81349~81350~81351~81407~81408 | 3341 |  | 3341 3341 3341 3341 3341 3341 | 2021-04-22 | 13:26 | albert zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 1.0029 | 0.00202 | 81348 |  | XCAMS | ST | 0.12% |  | TW3341 measurement: * large samples CalAMS error adjustments: * 1.03 global error multiplier to account for excess run-to-run variability * 0.12% system error to account for excess Ox-I cathode-to-cathode variability primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value * FIRI-I cellulose: 2x result agrees with consensus value: R41572/4 prepared by Otlet, assuming this is a FIRI-I R24889/9 prepared by RRL blanks: * 3 groups identified to blank-correct: water: 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks RRL org: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Otlet org: 6x unknowns + 1x secondaries + 2x blanks also picking up TW3324, capturing 5x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-04-23 | Water | Not Complete | JOB-217316 |  | D bank only went down to 0.0030 mbar | 71763 | Normal | 14047/11 |  |  | 2021-04-13 | 5199 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00203 | 81349 | [26 empty] | 2021-04-13 | 3075 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00207 | 81350 | [26 empty] | 2021-04-13 | 5121 | Choose Gas Splitting | Completed
[168 empty] | 0.99811 | 0.00202 | 81351 | [26 empty] | 2021-04-14 | 21353 | Graphite | Completed
[168 empty] | 0.99895 | 0.00212 | 81407 | [26 empty] | 2021-04-16 | 7929 | C13 Measurement | Completed
[168 empty] | 0.99781 | 0.00203 | 81408 | [26 empty] | 2021-04-22 | 22530 | AMS Submission Results | Completed
[197 empty] | 2021-04-23 | 26696 | Close Job | Completed
[7 empty] | 87.1 | 87.9343 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49332 | 49332 | 49332 | 272 | 271 | 271 |  |  | taylor ferrick | -997.85 | 0.07 ‰ | 2021-05-09 | 2020-11-10 | 16 | 180 |  |  |  |  | 3.76‰ | 3.76‰ | 2.00‰ | 2.00‰ | 3.76 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0041 | 0 | 0.8165 | -997.85 ‰ | -997.85 ‰ | 0.07 | 0.07 | -997.82 | 0.07 ‰ | -997.82 | 0.07 ‰ |  |  | 2021-05-11 | Jocelyn Turnbull | 12:57 | 0.0021524 | 0.0021524 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.1 |  |  | 0.00215 | 7e-05 | 22648 | 2021-05-11 | 9:05 | albert zondervan | 217607 |  |  |  |  | 45% | 5 | 11/05/21 | 12:57:05 | Jocelyn Turnbull | 0 |  | 0.22 | 0.01% | 1 | 5 | 5 | ... | 0.00207 | 7e-05 |  |  |  |  | 0.99185 | 0.99185 | Yes | Yes | [5 empty] | 0.00207 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00215 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00207 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-05-07 | taylor ferrick | 8:47 |  | Completed |  |  |  |  | 81446 | 2000000 |  |  | 81488 81494 81500 81501 81502 81522 | 81488~81494~81500~81501~81502~81522 | 3344 |  | 3344 3344 3344 3344 3344 3344 | 2021-05-11 | 9:19 | albert zondervan |  | 1.00 mg |  | 1 | [5 empty] | 0.99962 | 0.00147 | 81488 |  | XCAMS | ST | 0% |  | TW3344 measurement CalAMS error adjustments: * factor 1.3 for TP81486,81484 due to excess run-to-run variability primaries: * large OxI cathodes: 6x accepted * 1x RP OxI check (TP81486) not used for calibration secondaries: * FIRI-I cellulose 1x result agrees with consensus value * IAEA-C2 travertine 1x result agrees with consensus value blanks: * 2 groups identified to blank-correct: AAA-ST: 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks large carbonates: 20x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-06-07 | Carbonate | Complete | JOB-217607 |  | delete physical if not needed | 71875 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0031 | 0.00154 | 81494 | [26 empty] | 2021-04-21 | 5211 | Carbonate CO2 Evolution | Completed
[168 empty] | 1 | 0.0015 | 81500 | [26 empty] | 2021-04-23 | 21456 | Graphite | Completed
[168 empty] | 0.99969 | 0.00146 | 81501 | [26 empty] | 2021-04-27 | 7947 | C13 Measurement | Completed
[168 empty] | 0.99943 | 0.00145 | 81502 | [26 empty] | 2021-05-11 | 22648 | AMS Submission Results | Completed
[168 empty] | 0.99808 | 0.0016 | 81522 | [26 empty] | 2021-05-12 | 26873 | Close Job | Completed
[7 empty] | 115.4 | 87.6671 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51154 | 51153 | 51153 | 487 | 486 | 486 |  |  | taylor ferrick | -998.28 | 0.10 ‰ | 2021-05-17 | 2020-11-18 | 6 | 180 |  |  |  |  | 4.78‰ | 4.78‰ | 2.00‰ | 2.00‰ | 4.78 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -998.28 ‰ | -998.28 ‰ | 0.1 | 0.1 | -998.26 | 0.11 ‰ | -998.26 | 0.11 ‰ |  |  | 2021-05-18 | Jocelyn Turnbull | 14:14 | 0.0017157 | 0.0017157 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.9 |  |  | 0.00172 | 0.0001 | 22768 | 2021-05-18 | 13:06 | albert zondervan | 217767 |  |  |  |  | 45% | 5 | 18/05/21 | 14:14:03 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 8 | 5 | 5 | ... | 0.00165 | 0.0001 |  |  |  |  | 0.98985 | 0.98985 | Yes | Yes | [5 empty] | 0.00165 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00172 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00165 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-05-14 | taylor ferrick | 11:42 |  | Completed |  |  |  |  | 81572 | 2000000 |  |  | 81564 81565 81591 81592 81594 81596 | 81564~81565~81591~81592~81594~81596 | 3347 |  | 3347 3347 3347 3347 3347 3347 | 2021-05-18 | 13:42 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0029 | 0.00158 | 81564 |  | XCAMS | ST | 0% |  | TW3347 measurement CalAMS error adjustments: * none primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-D wood 1x result agrees with consensus value * IAEA-C2 travertine 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: CE_Rafter: 5x unknowns + 1x secondaries + 1x blanks 180 days, capturing 3x blanks AAA: 18x unknowns + 1x secondaries + 0x blanks 180 days, capturing 6x blanks carbonates: 5x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-06-23 | Carbonate |  | JOB-217767 |  |  | 71954 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99851 | 0.00159 | 81565 | [26 empty] | 2021-05-10 | 5234 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99869 | 0.00159 | 81591 | [26 empty] | 2021-05-11 | 21582 | Graphite | Completed
[168 empty] | 1.0003 | 0.00159 | 81592 | [26 empty] | 2021-05-13 | 7994 | C13 Measurement | Completed
[168 empty] | 0.99934 | 0.00159 | 81594 | [26 empty] | 2021-05-18 | 22768 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.0016 | 81596 | [26 empty] | 2021-05-18 | 26965 | Close Job | Completed
[7 empty] | 85.5 | 87.86 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50021 | 50020 | 50020 | 254 | 253 | 253 |  |  | Julia Collins | -998.02 | 0.06 ‰ | 2021-06-03 | 2020-12-05 | 15 | 180 |  |  |  |  | 3.47‰ | 3.47‰ | 2.00‰ | 2.00‰ | 3.47 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.002 | 0 | 0.8165 | -998.02 ‰ | -998.02 ‰ | 0.06 | 0.06 | -998 | 0.06 ‰ | -998 | 0.06 ‰ |  |  | 2021-06-09 | Jocelyn Turnbull | 10:49 | 0.0019756 | 0.0019756 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.1 |  |  | 0.00198 | 6e-05 | 22928 | 2021-06-08 | 11:03 | Albert Zondervan | 217908 |  |  |  |  | 45% | 5 | 9/06/21 | 10:49:17 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 1 | 5 | 5 | ... | 0.0019 | 6e-05 |  |  |  |  | 0.99073 | 0.99073 | Yes | Yes | [5 empty] | 0.0019 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00054 | 0.00054 | 0.00198 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0019 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-02 | Julia Collins | 14:55 |  | Completed |  |  |  |  | 81644 | 2000000 |  |  | 81772 81778 81779 81780 81781 81813 | 81772~81778~81779~81780~81781~81813 | 3351 |  | 3351 3351 3351 3351 3351 3351 | 2021-06-08 | 11:10 | Albert Zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 1.0005 | 0.00138 | 81772 |  | XCAMS | ST | 0% |  | TW3351 measurement: * large carbonate samples CalAMS error adjustments: * none primaries: * large OxI cathodes: 6x accepted secondaries: * IAEA-C2 travertine large: 1x result agrees with consensus value blanks: * 1 groups identified to blank-correct: large carbonates: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-07-07 | Carbonate | Complete | JOB-217908 |  |  | 72114 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00132 | 81778 | [26 empty] | 2021-05-18 | 5241 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.0013 | 81779 | [26 empty] | 2021-05-19 | 21654 | Graphite | Completed
[168 empty] | 1.0005 | 0.00132 | 81780 | [26 empty] | 2021-05-24 | 8016 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.00128 | 81781 | [26 empty] | 2021-06-09 | 22928 | AMS Submission Results | Completed
[168 empty] | 0.99701 | 0.0013 | 81813 | [26 empty] | 2021-06-09 | 27247 | Close Job | Completed
[7 empty] | 130.5 | 87.5549 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44182 | 44182 | 44182 | 225 | 224 | 224 |  |  | Julia Collins | -995.91 | 0.11 ‰ | 2021-06-10 | 2020-12-12 | 6 | 180 |  |  |  |  | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -995.91 ‰ | -995.91 ‰ | 0.11 | 0.11 | -995.86 | 0.12 ‰ | -995.86 | 0.12 ‰ |  |  | 2021-06-11 | Jocelyn Turnbull | 12:02 | 0.0040864 | 0.0040864 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.5 |  |  | 0.00409 | 0.00011 | 22969 | 2021-06-11 | 11:47 | Albert Zondervan | 217605 |  |  |  |  | 45% | 5 | 11/06/21 | 12:02:27 | Jocelyn Turnbull | 0 |  | 0.41 | 0.01% | 1 | 5 | 5 | ... | 0.00393 | 0.00011 |  |  |  |  | 0.98682 | 0.98682 | Yes | Yes | [5 empty] | 0.00393 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00409 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00393 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-09 | Julia Collins | 11:50 |  | Completed |  |  |  |  | 81829 | 2000000 |  |  | 81782 81783 81808 81809 81828 81864 | 81782~81783~81808~81809~81828~81864 | 3352 |  | 3352 3352 3352 3352 3352 3352 | 2021-06-11 | 11:51 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99999 | 0.00172 | 81782 |  | XCAMS | ST | 0% |  | TW3352 measurement: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-06-07 | Water | Complete | JOB-217605 |  |  | 72154 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99939 | 0.00164 | 81783 | [26 empty] | 2021-06-03 | 5298 | Water CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00166 | 81808 | [26 empty] | 2021-06-04 | 21839 | Graphite | Completed
[168 empty] | 0.99924 | 0.0017 | 81809 | [26 empty] | 2021-06-08 | 8102 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00168 | 81828 | [26 empty] | 2021-06-11 | 22969 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00163 | 81864 | [26 empty] | 2021-06-11 | 27256 | Close Job | Completed
[5 empty] | 3.5134e-05 | 0.0042345 | 130.1 | 88.4328 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7095 | 7094 | 7094 | 34 | 34 | 38 | 7060.7 | 7129.1 | Julia Collins | -586.55 | 1.76 ‰ | 2021-06-10 | 2020-12-12 | 6 | 180 |  |  |  |  | -6.96‰ | -6.96‰ | 2.00‰ | 2.00‰ | -6.96 ‰ | 2.00 ‰ | AMS | -8.15‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -586.55 ‰ | -586.55 ‰ | 1.76 | 1.99 | -581.44 | 1.78 ‰ | -581.44 | 1.78 ‰ |  |  | 2021-06-11 | Jocelyn Turnbull | 12:02 | 0.41345 | 0.41345 | 0.00176 | 0.00199 | 0.00167 | 0.00167 | 2 |  |  | 0.41626 | 0.00121 | 23004 | 2021-06-11 | 11:47 | Albert Zondervan | 218137 | 0.00449 | 0.0045 | 0.00202 | 0.002 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.35 | 0.18% | 20 | 5 | 4 | ... | 0.40033 | 0.00104 |  |  |  |  | 0.98682 | 0.98682 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 38.6539430978042155. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00449 | 0.00407 | 0.00148 | 0.39763 | 0.40‰ | 0.00161 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.41345 | 0 | 0.00176 | 0 | 0 | 0.00449 | 0 | 0.00202 | 0 | 0.39763 | 0 | 0.00161 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-09 | Julia Collins | 11:50 |  | Completed |  |  |  |  | 81835 | 2000000 | 81829 81343 81316 81152 81078 81063 80827 80648 | 81829~81343~81316~81152~81078~81063~80827~80648 | 81782 81783 81808 81809 81828 81864 | 81782~81783~81808~81809~81828~81864 | 3352 | 3352 3341 3338 3334 3333 3331 3327 3323 | 3352 3352 3352 3352 3352 3352 | 2021-06-11 | 11:51 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 0.99999 | 0.00172 | 81782 | ST | XCAMS | ST | 0% |  | TW3352 measurement: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-07-28 | Water | Complete | JOB-218137 |  |  | 72155 | Normal | 14047/12 |  |  | 2021-06-03 | 5299 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 0.99939 | 0.00164 | 81783 | ST | [25 empty] | 2021-06-04 | 21845 | Graphite | Completed
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 1.0021 | 0.00166 | 81808 | EA | [25 empty] | 2021-06-08 | 8108 | C13 Measurement | Completed
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99924 | 0.0017 | 81809 | EA | [25 empty] | 2021-06-11 | 23004 | AMS Submission Results | Completed
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 | 1.0003 | 0.00168 | 81828 | EA | [25 empty] | 2021-06-11 | 27272 | Close Job | Completed
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 | 0.99904 | 0.00163 | 81864 | FLASK
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 |  |  |  | ST
[163 empty] | 2020-12-16 | 0.00381 | 80648 | 3323 | 1.12 |  |  |  | ST
[7 empty] | 142.1 | 87.569 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51154 | 51153 | 51153 | 292 | 292 | 292 |  |  | Julia Collins | -998.28 | 0.06 ‰ | 2021-06-16 | 2020-12-18 | 1 | 180 |  |  |  |  | 3.03‰ | 3.03‰ | 2.00‰ | 2.00‰ | 3.03 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0 | 0.8165 | -998.28 ‰ | -998.28 ‰ | 0.06 | 0.06 | -998.26 | 0.06 ‰ | -998.26 | 0.06 ‰ |  |  | 2021-06-18 | Jocelyn Turnbull | 10:21 | 0.0017157 | 0.0017157 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.2 |  |  | 0.00172 | 6e-05 | 23076 | 2021-06-18 | 8:04 | Albert Zondervan | 217971 |  |  |  |  | 45% | 5 | 22/06/21 | 10:40:33 | jenny dahl | 0 |  | 0.17 | 0.01% | 1 | 5 | 5 | ... | 0.00165 | 6e-05 |  |  |  |  | 0.98701 | 0.98701 | Yes | Yes | [5 empty] | 0.00165 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00172 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00165 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-15 | Julia Collins | 14:10 |  | Completed |  |  |  |  | 81905 | 2000000 |  |  | 81852 81859 81863 81866 81867 81868 | 81852~81859~81863~81866~81867~81868 | 3354 |  | 3354 3354 3354 3354 3354 3354 | 2021-06-18 | 8:13 | Albert Zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0012 | 0.00153 | 81852 |  | XCAMS | ST | 0% |  | TW3354 measurement: * large carbonate samples CalAMS error adjustments: * none primaries: * large OxI cathodes: 6x accepted secondaries: * IAEA-C2 travertine large: 1x result agrees with consensus value blanks: * 1 groups identified to blank-correct: large carbonates: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-07-12 | Carbonate |  | JOB-217971 |  |  | 72303 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0015 | 0.00148 | 81859 | [26 empty] | 2021-06-14 | 5339 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00147 | 81863 | [26 empty] | 2021-06-15 | 21915 | Graphite | Completed
[168 empty] | 0.9987 | 0.00147 | 81866 | [26 empty] | 2021-06-21 | 8144 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00142 | 81867 | [26 empty] | 2021-06-18 | 23076 | AMS Submission Results | Completed
[168 empty] | 0.99978 | 0.00146 | 81868 | [26 empty] | 2021-06-22 | 27453 | Close Job | Completed
[5 empty] | 3.5767e-05 | 0.004405 | 129.9 | 88.8207 | 0.0701 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 6981 | 6981 | 6981 | 35 | 35 | 39 | 6945.7 | 7016.7 | Julia Collins | -580.65 | 1.85 ‰ | 2021-06-22 | 2020-12-24 | 4 | 180 |  |  |  |  | -10.35‰ | -10.35‰ | 2.00‰ | 2.00‰ | -10.35 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -580.65 ‰ | -580.65 ‰ | 1.85 | 2.08 | -575.47 | 1.88 ‰ | -575.47 | 1.88 ‰ |  |  | 2021-06-24 | Cathy Ginnane | 23:58 | 0.41935 | 0.41935 | 0.00185 | 0.00208 | 0.00177 | 0.00177 | 2.3 |  |  | 0.42228 | 0.00128 | 23099 | 2021-06-23 | 12:12 | Albert Zondervan | 217907 | 0.00473 | 0.0047 | 0.00213 | 0.0021 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.93 | 0.19% | 29 | 5 | 6 | ... | 0.40612 | 0.00143 |  |  |  |  | 0.98776 | 0.98776 | Yes | Yes |  | result comes out too young by 4-sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 39.8298249247933394. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00473 | 0.00423 | 0.0015 | 0.4033 | 0.40‰ | 0.00192 | 0.00‰ | Yes | 1 | 1 | 0.00113 | 0.00113 | 0.41935 | 0 | 0.00185 | 0 | 0 | 0.00473 | 0 | 0.00213 | 0 | 0.4033 | 0 | 0.00192 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-21 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 81921 | 2000000 | 81922 81829 81343 81316 81152 81078 81063 80827 | 81922~81829~81343~81316~81152~81078~81063~80827 | 81865 81893 81894 81895 81896 81937 | 81865~81893~81894~81895~81896~81937 | 3355 | 3355 3352 3341 3338 3334 3333 3331 3327 | 3355 3355 3355 3355 3355 3355 | 2021-06-23 | 12:29 | Albert Zondervan |  | 0.94 mg | 0.13 | 0.94 | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 1.0025 | 0.00275 | 81865 | ST | XCAMS | ST | 0.22% |  | TW3355 measurement CalAMS error adjustments: * factor 1.07 due to excess global run-to-run variability * 0.22% system error due to excess OxI cat-to-cat variability primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 4-sigma blanks: * 2 groups identified to blank-correct: organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks water: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures TO BE SORTED: * The IRMS del13C values of J218004 and J217995 appear swapped, while the XCAMS data do not suggest a problem. *refer to e-log 223 on possible swap concerns | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-07-07 | Water | Not Complete | JOB-217907 |  | this secondary is 4 sigma low. Reasons for this likely include: sample non-homogeneity, insufficient N2 flushing and possible air leak. There is no likely possibility of swap in AMS relative to the other AMS measurements. Travertine is known for this behaviour occasionally, although 4 sigma out is substantial (though not without precedent). The blank performed as expected, thus other waters will be signed off. Some waters are being repeated so this can be reviewed in terms of their performance in the new wheel ELK. | 72355 | Normal | 14047/12 |  |  | 2021-06-17 | 5349 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 1.0024 | 0.0028 | 81893 | EA | [25 empty] | 2021-06-17 | 3117 | Gas Splitting | Completed
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 0.99658 | 0.00273 | 81894 | EA | [25 empty] | 2021-06-17 | 5211 | Choose Gas Splitting | Completed
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 0.99815 | 0.00279 | 81895 | EA | [25 empty] | 2021-06-18 | 21931 | Graphite | Completed
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99796 | 0.00277 | 81896 | FLASK | [25 empty] | 2021-07-15 | 8159 | C13 Measurement | Completed
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 | 1.0025 | 0.00279 | 81937 | ST | [25 empty] | 2021-06-24 | 23099 | AMS Submission Results | Completed
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 |  |  |  | ST | [25 empty] | 2021-07-28 | 27527 | Close Job | Completed
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 |  |  |  | ST
[7 empty] | 136 | 87.72 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41238 | 41237 | 41237 | 169 | 169 | 169 |  |  | Julia Collins | -994.1 | 0.12 ‰ | 2021-06-22 | 2020-12-24 | 4 | 180 |  |  |  |  | 2.07‰ | 2.07‰ | 2.00‰ | 2.00‰ | 2.07 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -994.10 ‰ | -994.10 ‰ | 0.12 | 0.12 | -994.03 | 0.13 ‰ | -994.03 | 0.13 ‰ |  |  | 2021-06-24 | Cathy Ginnane | 23:58 | 0.0058956 | 0.0058956 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.2 |  |  | 0.0059 | 0.00012 | 23100 | 2021-06-23 | 12:12 | Albert Zondervan | 217909 |  |  |  |  | 45% | 5 | 24/06/21 | 23:58:21 | Cathy Ginnane | 0 |  | 0.59 | 0.01% | 37 | 5 | 7 | ... | 0.00567 | 0.00012 |  |  |  |  | 0.98776 | 0.98776 | Yes | Yes | [5 empty] | 0.00567 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00113 | 0.00113 | 0.0059 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00567 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-21 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 81922 | 2000000 |  |  | 81865 81893 81894 81895 81896 81937 | 81865~81893~81894~81895~81896~81937 | 3355 |  | 3355 3355 3355 3355 3355 3355 | 2021-06-23 | 12:29 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0025 | 0.00275 | 81865 |  | XCAMS | ST | 0.22% |  | TW3355 measurement CalAMS error adjustments: * factor 1.07 due to excess global run-to-run variability * 0.22% system error due to excess OxI cat-to-cat variability primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 4-sigma blanks: * 2 groups identified to blank-correct: organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks water: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures TO BE SORTED: * The IRMS del13C values of J218004 and J217995 appear swapped, while the XCAMS data do not suggest a problem. *refer to e-log 223 on possible swap concerns | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-07-07 | Water | Complete | JOB-217909 |  | C13 split initially sealed then transfered to MS bottle. | 72353 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.0028 | 81893 | [26 empty] | 2021-06-17 | 5350 | Water CO2 Evolution | Completed
[168 empty] | 0.99658 | 0.00273 | 81894 | [26 empty] | 2021-06-17 | 3119 | Gas Splitting | Completed
[168 empty] | 0.99815 | 0.00279 | 81895 | [26 empty] | 2021-06-17 | 5215 | Choose Gas Splitting | Completed
[168 empty] | 0.99796 | 0.00277 | 81896 | [26 empty] | 2021-06-18 | 21932 | Graphite | Completed
[168 empty] | 1.0025 | 0.00279 | 81937 | [26 empty] | 2021-06-22 | 8167 | C13 Measurement | Completed
[197 empty] | 2021-06-24 | 23100 | AMS Submission Results | Completed
[197 empty] | 2021-06-25 | 27479 | Close Job | Completed
[7 empty] | 85.3 | 88.4172 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48953 | 48953 | 48953 | 296 | 296 | 296 |  |  | Taylor Ferrick | -997.74 | 0.08 ‰ | 2021-07-19 | 2021-01-20 | 17 | 180 |  |  |  |  | 1.22‰ | 1.22‰ | 2.00‰ | 2.00‰ | 1.22 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.75593 | -997.74 ‰ | -997.74 ‰ | 0.08 | 0.08 | -997.72 | 0.08 ‰ | -997.72 | 0.08 ‰ |  |  | 2021-07-26 | Jocelyn Turnbull | 10:17 | 0.0022564 | 0.0022564 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00226 | 8e-05 | 23184 | 2021-07-20 | 16:43 | Albert Zondervan | 218106 |  |  |  |  | 45% | 5 | 26/07/21 | 10:17:00 | Jocelyn Turnbull | 0 |  | 0.23 | 0.01% | 6 | 5 | 5 | ... | 0.00217 | 8e-05 |  |  |  |  | 0.99478 | 0.99478 | Yes | Yes | [5 empty] | 0.00217 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00226 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00217 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-07-16 | Taylor Ferrick | 12:13 |  | Completed |  |  |  |  | 81980 | 2000000 |  |  | 81990 81999 82000 82021 82034 82035 82058 | 81990~81999~82000~82021~82034~82035~82058 | 3357 |  | 3357 3357 3357 3357 3357 3357 3357 | 2021-07-23 | 11:32 | Albert Zondervan |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0002 | 0.00169 | 81990 |  | XCAMS | ST | 0% |  | TW3357 measurement: * large+small carbonate samples * large+small AAA-ST samples CalAMS error adjustments: * none blank-correction small samples: * see file 'to blank-correct small samples.xlsm' in folder 'I:\C14Data\Blank Correction for Small Samples' * results from previous wheels also incorporated: TW3278,3296,3302,3323,3337 primaries: * large OxI cathodes: 7x accepted, used for standardization in CalAMS * small OxI cathodes: 1x accepted, only used to co-determine 'small d' d = 0.75 ug C secondaries: * FIRI-I: cellulose large: 1x result is in rts ~0.005 (3-sigma) larger than consensus value; quality flag set to ..X small: 1x result agrees with consensus value * IAEA-C2: travertine, via carbonate evolution large: 1x result is in rts ~0.003 (2.5-sigma) larger than consensus value; quality flag left at ... small: 1x result agrees with consensus value blanks: * 4 groups identified to blank-correct: large carbonates: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks large organics via AAA: 4x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks TP82025 in this wheel excluded; quality flag set to ..X small carbonates: 3x unknowns + 1x secondaries + 1x blanks capturing 4x blanks, m = 0.95 ug C small organics via AAA: 1x unknowns + 1x secondaries + 1x blanks capturing 9x blanks, m = 0.72 ug C unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-07-22 | Carbonate | Not Complete | JOB-218106 |  | normal size blank (do this blank before running unks) | 72433 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99944 | 0.00169 | 81999 | [26 empty] | 2021-07-01 | 5363 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0009 | 0.00169 | 82000 | [26 empty] | 2021-07-02 | 21990 | Graphite | Completed
[168 empty] | 1.0003 | 0.00168 | 82021 | [26 empty] | 2021-07-02 | 8187 | C13 Measurement | Completed
[168 empty] | 0.99861 | 0.00167 | 82034 | [26 empty] | 2021-07-26 | 23184 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.0017 | 82035 | [26 empty] | 2021-07-28 | 27569 | Close Job | Completed
[168 empty] | 0.99917 | 0.00171 | 82058
[7 empty] | 80.6 | 88.5968 | 0.0006 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46062 | 46062 | 46062 | 258 | 258 | 258 |  |  | Taylor Ferrick | -996.77 | 0.10 ‰ | 2021-07-19 | 2021-01-20 | 10 | 180 |  |  |  |  | -0.81‰ | -0.81‰ | 2.00‰ | 2.00‰ | -0.81 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0022 | 0 | 0.75593 | -996.77 ‰ | -996.77 ‰ | 0.1 | 0.1 | -996.73 | 0.11 ‰ | -996.73 | 0.11 ‰ |  |  | 2021-07-26 | Jocelyn Turnbull | 10:17 | 0.0032338 | 0.0032338 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 8.2 |  |  | 0.00323 | 0.0001 | 23205 | 2021-07-20 | 16:43 | Albert Zondervan | 218407 |  |  |  |  | 45% | 5 | 26/07/21 | 10:17:02 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 1 | 5 | 4 | ... | 0.00311 | 0.0001 |  |  |  |  | 0.99478 | 0.99478 | Yes | Yes | [5 empty] | 0.00311 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00323 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00311 | 0 | 0.0001 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-07-16 | Taylor Ferrick | 12:13 |  | Completed |  |  |  |  | 82051 | 2000000 |  |  | 81990 81999 82000 82021 82034 82035 82058 | 81990~81999~82000~82021~82034~82035~82058 | 3357 |  | 3357 3357 3357 3357 3357 3357 3357 | 2021-07-23 | 11:32 | Albert Zondervan |  | 0.25 mg |  | 0.25 | [5 empty] | 1.0002 | 0.00169 | 81990 |  | XCAMS | ST | 0% |  | TW3357 measurement: * large+small carbonate samples * large+small AAA-ST samples CalAMS error adjustments: * none blank-correction small samples: * see file 'to blank-correct small samples.xlsm' in folder 'I:\C14Data\Blank Correction for Small Samples' * results from previous wheels also incorporated: TW3278,3296,3302,3323,3337 primaries: * large OxI cathodes: 7x accepted, used for standardization in CalAMS * small OxI cathodes: 1x accepted, only used to co-determine 'small d' d = 0.75 ug C secondaries: * FIRI-I: cellulose large: 1x result is in rts ~0.005 (3-sigma) larger than consensus value; quality flag set to ..X small: 1x result agrees with consensus value * IAEA-C2: travertine, via carbonate evolution large: 1x result is in rts ~0.003 (2.5-sigma) larger than consensus value; quality flag left at ... small: 1x result agrees with consensus value blanks: * 4 groups identified to blank-correct: large carbonates: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks large organics via AAA: 4x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks TP82025 in this wheel excluded; quality flag set to ..X small carbonates: 3x unknowns + 1x secondaries + 1x blanks capturing 4x blanks, m = 0.95 ug C small organics via AAA: 1x unknowns + 1x secondaries + 1x blanks capturing 9x blanks, m = 0.72 ug C unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2021-09-01 | Carbonate |  | JOB-218407 |  | small size blank to match size of foram samplesUsing liquid nitrogen and ethanol slushies. | 72434 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99944 | 0.00169 | 81999 | [26 empty] | 2021-07-08 | 5384 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0009 | 0.00169 | 82000 | [26 empty] | 2021-07-09 | 22061 | Graphite | Completed
[168 empty] | 1.0003 | 0.00168 | 82021 | [26 empty] | 2021-07-26 | 23205 | AMS Submission Results | Completed
[168 empty] | 0.99861 | 0.00167 | 82034 | [26 empty] | 2021-07-28 | 27595 | Close Job | Completed
[168 empty] | 1.0015 | 0.0017 | 82035
[168 empty] | 0.99917 | 0.00171 | 82058
[7 empty] | 106.5 | 87.9491 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 185 | 185 | 185 |  |  | Taylor Ferrick | -995.04 | 0.11 ‰ | 2021-07-22 | 2021-01-23 | 6 | 180 |  |  |  |  | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0059 | 0 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.11 | 0.11 | -994.98 | 0.12 ‰ | -994.98 | 0.12 ‰ |  |  | 2021-07-26 | Jocelyn Turnbull | 10:23 | 0.0049598 | 0.0049598 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00496 | 0.00011 | 23228 | 2021-07-23 | 15:27 | Jeremy Thompson | 218145 |  |  |  |  | 45% | 5 | 26/07/21 | 10:22:52 | Jocelyn Turnbull | 0 |  | 0.5 | 0.01% | 1 | 5 | 5 | ... | 0.00477 | 0.00011 |  |  |  |  | 0.99118 | 0.99118 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00496 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-07-21 | Taylor Ferrick | 13:09 |  | Completed |  |  |  |  | 82086 | 2000000 |  |  | 82082 82083 82084 82085 82096 82101 | 82082~82083~82084~82085~82096~82101 | 3358 |  | 3358 3358 3358 3358 3358 3358 | 2021-07-23 | 15:46 | Jeremy Thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0032 | 0.00169 | 82082 |  | XCAMS | ST | 0% |  | TW3358 measurement CalAMS error adjustments: * no error adjustment needed primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 2-sigma blanks: * 3 groups identified to blank-correct: Organics via AAA: 13x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Water: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Carbonate 1x unknowns + 0x secondaries + 0x blanks 180 days, capturing 8x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-08-04 | Water |  | JOB-218145 |  | Using liquid nitrogen and ethanol slushies. | 72473 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00166 | 82083 | [26 empty] | 2021-07-15 | 5402 | Water CO2 Evolution | Completed
[168 empty] | 0.99757 | 0.00165 | 82084 | [26 empty] | 2021-07-16 | 22096 | Graphite | Completed
[168 empty] | 0.99887 | 0.00166 | 82085 | [26 empty] | 2021-07-19 | 8223 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00174 | 82096 | [26 empty] | 2021-07-26 | 23228 | AMS Submission Results | Completed
[168 empty] | 0.99909 | 0.00168 | 82101 | [26 empty] | 2021-07-28 | 27572 | Close Job | Completed
[5 empty] | 3.4137e-05 | 0.0044545 | 102.8 | 89.0743 | 0.0698 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7046 | 7045 | 7045 | 36 | 35 | 39 | 7010.3 | 7081.3 | Taylor Ferrick | -584.01 | 1.84 ‰ | 2021-07-22 | 2021-01-23 | 6 | 180 |  |  |  |  | -9.76‰ | -9.76‰ | 2.00‰ | 2.00‰ | -9.76 ‰ | 2.00 ‰ | AMS | -8.17‰ | 0.2‰ | C13 | No | -19 | -19.0059 | 0 | 0.8165 | -584.01 ‰ | -584.01 ‰ | 1.84 | 2.06 | -578.87 | 1.86 ‰ | -578.87 | 1.86 ‰ |  |  | 2021-07-26 | Jocelyn Turnbull | 10:23 | 0.41599 | 0.41599 | 0.00184 | 0.00206 | 0.00176 | 0.00176 | 2.1 |  |  | 0.41894 | 0.00127 | 23229 | 2021-07-23 | 15:27 | Jeremy Thompson | 218146 | 0.00473 | 0.0047 | 0.00213 | 0.0021 | 45% | 5 | 5/07/22 | 14:07:18 | Jocelyn Turnbull | 0 |  | 41.6 | 0.18% | 8 | 5 | 4 | ... | 0.40291 | 0.0011 |  |  |  |  | 0.99118 | 0.99118 | Yes | Yes |  | 2 sigma too young wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 39.8459302460874298. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00473 | 0.00428 | 0.0014 | 0.40007 | 0.40‰ | 0.00169 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.41599 | 0 | 0.00184 | 0 | 0 | 0.00473 | 0 | 0.00213 | 0 | 0.40007 | 0 | 0.00169 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-07-21 | Taylor Ferrick | 13:09 |  | Completed |  |  |  |  | 82087 | 2000000 | 82086 81922 81829 81343 81316 81152 81078 81063 80827 | 82086~81922~81829~81343~81316~81152~81078~81063~80827 | 82082 82083 82084 82085 82096 82101 | 82082~82083~82084~82085~82096~82101 | 3358 | 3358 3355 3352 3341 3338 3334 3333 3331 3327 | 3358 3358 3358 3358 3358 3358 | 2021-07-23 | 15:46 | Jeremy Thompson |  | 0.89 mg | 0.13 | 0.89 | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 1.0032 | 0.00169 | 82082 | ST | XCAMS | ST | 0% |  | TW3358 measurement CalAMS error adjustments: * no error adjustment needed primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 2-sigma blanks: * 3 groups identified to blank-correct: Organics via AAA: 13x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Water: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Carbonate 1x unknowns + 0x secondaries + 0x blanks 180 days, capturing 8x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-08-04 | Water | Complete | JOB-218146 |  | Using liquid nitrogen and ethanol slushies. | 72474 | Normal | 14047/12 |  |  | 2021-07-15 | 5403 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 1.0001 | 0.00166 | 82083 | EA | [25 empty] | 2021-07-16 | 22097 | Graphite | Completed
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 0.99757 | 0.00165 | 82084 | EA | [25 empty] | 2021-07-19 | 8224 | C13 Measurement | Completed
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 0.99887 | 0.00166 | 82085 | EA | [25 empty] | 2021-07-26 | 23229 | AMS Submission Results | Completed
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 1.0014 | 0.00174 | 82096 | FLASK | [25 empty] | 2021-07-28 | 27573 | Close Job | Completed
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99909 | 0.00168 | 82101 | ST
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 |  |  |  | ST
[163 empty] | 2021-03-03 | 0.00316 | 81063 | 3331 | 1.05 |  |  |  | ST
[163 empty] | 2021-02-09 | 0.0022 | 80827 | 3327 | 1.1 |  |  |  | ST
[7 empty] | 62.8 | 89.0767 | 0.0007 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43881 | 43881 | 43881 | 295 | 295 | 295 |  |  | Taylor Ferrick | -995.76 | 0.16 ‰ | 2021-08-09 | 2021-02-10 | 4 | 180 |  |  |  |  | -11.20‰ | -11.20‰ | 2.00‰ | 2.00‰ | -11.20 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0034 | 0 | 0.8165 | -995.76 ‰ | -995.76 ‰ | 0.16 | 0.16 | -995.71 | 0.16 ‰ | -995.71 | 0.16 ‰ |  |  | 2021-08-11 | Jocelyn Turnbull | 14:02 | 0.0042424 | 0.0042424 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 18.5 |  |  | 0.00424 | 0.00016 | 23359 | 2021-08-11 | 9:31 | Jeremy Thompson | 218422 |  |  |  |  | 45% | 5 | 11/08/21 | 14:02:27 | Jocelyn Turnbull | 0 |  | 0.42 | 0.02% | 1 | 5 | 4 | ... | 0.00408 | 0.00015 |  |  |  |  | 0.98976 | 0.98976 | Yes | Yes | [5 empty] | 0.00408 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00424 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00408 | 0 | 0.00015 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-08-06 | Taylor Ferrick | 11:07 |  | Completed |  |  |  |  | 82215 | 2000000 |  |  | 82102 82103 82186 82187 82188 82189 | 82102~82103~82186~82187~82188~82189 | 3362 |  | 3362 3362 3362 3362 3362 3362 | 2021-08-11 | 12:41 | Cathy Ginnane |  | 0.12 mg |  | 0.12 | [5 empty] | 0.99916 | 0.00179 | 82102 |  | XCAMS | ST | 0% |  |  |  | RRL | [5 empty] | Background Inorganic |  | 2021-09-06 | Carbonate |  | JOB-218422 |  | SLUSHIE small blank | 72624 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99829 | 0.00171 | 82103 | [26 empty] | 2021-08-03 | 5409 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00174 | 82186 | [26 empty] | 2021-08-05 | 22225 | Graphite | Completed
[168 empty] | 1.0004 | 0.0017 | 82187 | [26 empty] | 2021-08-11 | 23359 | AMS Submission Results | Completed
[168 empty] | 0.99913 | 0.0018 | 82188 | [26 empty] | 2021-08-12 | 27784 | Close Job | Completed
[168 empty] | 1.0008 | 0.00172 | 82189
[7 empty] | 113.4 | 87.7707 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46088 | 46087 | 46087 | 233 | 233 | 233 |  |  | Taylor Ferrick | -996.78 | 0.09 ‰ | 2021-08-09 | 2021-02-10 | 27 | 180 |  |  |  |  | 3.51‰ | 3.51‰ | 2.00‰ | 2.00‰ | 3.51 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0034 | 0 | 0.8165 | -996.78 ‰ | -996.78 ‰ | 0.09 | 0.09 | -996.74 | 0.09 ‰ | -996.74 | 0.09 ‰ |  |  | 2021-08-11 | Jocelyn Turnbull | 14:02 | 0.0032234 | 0.0032234 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.2 |  |  | 0.00322 | 9e-05 | 23360 | 2021-08-11 | 9:31 | Jeremy Thompson | 218245 |  |  |  |  | 45% | 5 | 11/08/21 | 14:02:29 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 33 | 5 | 7 | ... | 0.0031 | 9e-05 |  |  |  |  | 0.98976 | 0.98976 | Yes | Yes | [5 empty] | 0.0031 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00322 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0031 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-08-06 | Taylor Ferrick | 11:07 |  | Completed |  |  |  |  | 82065 | 2000000 |  |  | 82102 82103 82186 82187 82188 82189 | 82102~82103~82186~82187~82188~82189 | 3362 |  | 3362 3362 3362 3362 3362 3362 | 2021-08-11 | 12:41 | Cathy Ginnane |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99916 | 0.00179 | 82102 |  | XCAMS | ST | 0% |  |  | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-08-19 | Carbonate |  | JOB-218245 |  | accidentally made normal sized sample - seal for later. Using liquid nitrogen and ethanol slushies. | 72625 |  | 14047/1 |  |  | 2021-04-21 | 7259 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99829 | 0.00171 | 82103 | [26 empty] | 2021-07-07 | 5382 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00174 | 82186 | [26 empty] | 2021-07-07 | 3137 | Gas Splitting | Completed
[168 empty] | 1.0004 | 0.0017 | 82187 | [26 empty] | 2021-07-07 | 5250 | Choose Gas Splitting | Completed
[168 empty] | 0.99913 | 0.0018 | 82188 | [26 empty] | 2021-07-12 | 8209 | C13 Measurement | Completed
[168 empty] | 1.0008 | 0.00172 | 82189 | [26 empty] | 2021-07-13 | 22075 | Graphite | Completed
[197 empty] | 2021-08-11 | 23360 | AMS Submission Results | Completed
[197 empty] | 2021-08-12 | 27777 | Close Job | Completed
[5 empty] | 4.244e-05 | 0.0054315 | 110.2 | 88.7532 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7007 | 7007 | 7007 | 37 | 36 | 40 | 6970.6 | 7044 | Taylor Ferrick | -582.01 | 1.91 ‰ | 2021-09-01 | 2021-03-05 | 15 | 180 |  |  |  |  | -3.66‰ | -3.66‰ | 2.00‰ | 2.00‰ | -3.66 ‰ | 2.00 ‰ | AMS | -8.20‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -582.01 ‰ | -582.01 ‰ | 1.91 | 2.13 | -576.85 | 1.93 ‰ | -576.85 | 1.93 ‰ |  |  | 2021-09-07 | Jocelyn Turnbull | 9:02 | 0.41799 | 0.41799 | 0.00191 | 0.00213 | 0.00183 | 0.00183 | 2.2 |  |  | 0.42129 | 0.00118 | 23445 | 2021-09-03 | 14:05 | Jeremy Thompson | 218718 | 0.00532 | 0.0053 | 0.00239 | 0.0024 | 45% | 5 | 29/08/22 | 11:19:03 | Christian Lewis | 0 |  | 41.8 | 0.19% | 1 | 5 | 4 | ... | 0.40517 | 0.00101 |  |  |  |  | 0.99368 | 0.99368 | Yes | Yes |  | result 2.5 simga off consensus wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.9042979075559431. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00532 | 0.00522 | 0.00086 | 0.40199 | 0.40‰ | 0.00176 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.41799 | 0 | 0.00191 | 0 | 0 | 0.00532 | 0 | 0.00239 | 0 | 0.40199 | 0 | 0.00176 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-08-30 | Taylor Ferrick | 13:42 |  | Completed |  |  |  |  | 82288 | 2000000 | 82086 81922 81829 81343 81316 81152 81078 | 82086~81922~81829~81343~81316~81152~81078 | 82244 82289 82290 82296 82297 82298 | 82244~82289~82290~82296~82297~82298 | 3364 | 3358 3355 3352 3341 3338 3334 3333 | 3364 3364 3364 3364 3364 3364 | 2021-09-03 | 15:34 | Jeremy Thompson |  | 0.94 mg | 0.13 | 0.94 | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 0.99789 | 0.00161 | 82244 | EA | XCAMS | FLASK | 0% |  | TW3364 measurement: CalAMS error adjustments: * 0% system error primaries: * OxI cathodes: 6x from flask secondaries: * IAEA-C2: Freshwater Travertine - 2.5 sigma out from consensus blanks: * two groups identified to blank-correct: CO2 from atmospheric CO: 6x unknowns only window = use of dilution blanks used for CO dilution (capturing only the 4x blanks measured in ~2015 TP; 68062, 68063, 68064, 68065 TW 3007) Water: 24x unknowns + 1x secondaries + 0x blanks 180 days, capturing 7x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-09-28 | Water |  | JOB-218718 |  |  | 72713 | Normal | 14047/12 |  |  | 2021-08-13 | 5436 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 1.0014 | 0.00169 | 82289 | EA | [25 empty] | 2021-08-17 | 22298 | Graphite | Completed
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 1.0015 | 0.00164 | 82290 | FLASK | [25 empty] | 2021-08-17 | 8287 | C13 Measurement | Completed
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 0.99829 | 0.0017 | 82296 | ST | [25 empty] | 2021-09-07 | 23445 | AMS Submission Results | Completed
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 1.0017 | 0.00167 | 82297 | ST | [25 empty] | 2021-09-07 | 27881 | Close Job | Completed
[163 empty] | 2021-03-16 | 0.0063 | 81152 | 3334 | 0.97 | 0.99932 | 0.00163 | 82298 | ST
[163 empty] | 2021-03-08 | 0.00634 | 81078 | 3333 | 1.12 |  |  |  | ST
[7 empty] | 88.3 | 88.1101 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48211 | 48211 | 48211 | 304 | 303 | 303 |  |  | Taylor Ferrick | -997.53 | 0.09 ‰ | 2021-09-10 | 2021-03-14 | 7 | 180 |  |  |  |  | -0.66‰ | -0.66‰ | 2.00‰ | 2.00‰ | -0.66 ‰ | 2.00 ‰ | AMS | 2.77‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.66667 | -997.53 ‰ | -997.53 ‰ | 0.09 | 0.09 | -997.49 | 0.09 ‰ | -997.49 | 0.09 ‰ |  |  | 2021-09-13 | Jocelyn Turnbull | 10:20 | 0.0024747 | 0.0024747 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00247 | 9e-05 | 23477 | 2021-09-13 | 9:46 | Jeremy Thompson | 218419 |  |  |  |  | 45% | 5 | 13/09/21 | 10:20:28 | Jocelyn Turnbull | 0 |  | 0.25 | 0.01% | 5 | 5 | 7 | ... | 0.00238 | 9e-05 |  |  |  |  | 0.98946 | 0.98946 | Yes | Yes | [5 empty] | 0.00238 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 0.99999 | 0.00057 | 0.00057 | 0.00247 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00238 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-09-08 | Taylor Ferrick | 12:10 |  | Completed |  |  |  |  | 82334 | 2000000 |  |  | 82304 82305 82317 82318 82319 82347 82348 82349 82350 | 82304~82305~82317~82318~82319~82347~82348~82349~82350 | 3365 |  | 3365 3365 3365 3365 3365 3365 3365 3365 3365 | 2021-09-13 | 10:00 | Jeremy Thompson |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99919 | 0.00165 | 82304 |  | XCAMS | ST | 0% |  | TW3365 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 11 large OxI cathodes: 9x accepted, 2x tests secondaries: * FIRI-I 2x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 2 groups identified to blank-correct: Organics via AAA: 13x unknowns + 1x secondaries + 0x blanks 3x test blanks 180 days, capturing 9x blanks Carbonate 18x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2021-09-02 | Carbonate | Complete | JOB-218419 |  | to refresh C1 standard material | 72745 |  | 14047/1 |  |  | 2021-07-09 | 7387 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1 | 0.00166 | 82305 | [26 empty] | 2021-09-02 | 5441 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99773 | 0.00163 | 82317 | [26 empty] | 2021-09-02 | 3180 | Gas Splitting | Completed
[168 empty] | 0.99992 | 0.00166 | 82318 | [26 empty] | 2021-09-02 | 5317 | Choose Gas Splitting | Completed
[168 empty] | 1.0015 | 0.00169 | 82319 | [26 empty] | 2021-09-03 | 22344 | Graphite | Completed
[168 empty] | 0.99925 | 0.00218 | 82347 | [26 empty] | 2021-09-08 | 8291 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00172 | 82348 | [26 empty] | 2021-09-13 | 23477 | AMS Submission Results | Completed
[168 empty] | 0.99862 | 0.00166 | 82349 | [26 empty] | 2021-09-14 | 27911 | Close Job | Completed
[168 empty] | 1.002 | 0.00168 | 82350
[7 empty] | 93.4 | 88.0575 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40252 | 40252 | 40252 | 163 | 162 | 162 |  |  | Taylor Ferrick | -993.33 | 0.14 ‰ | 2021-09-29 | 2021-04-02 | 7 | 180 |  |  |  |  | 1.15‰ | 1.15‰ | 2.00‰ | 2.00‰ | 1.15 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0051 | 0 | 0.8165 | -993.33 ‰ | -993.33 ‰ | 0.14 | 0.14 | -993.25 | 0.14 ‰ | -993.25 | 0.14 ‰ |  |  | 2021-10-01 | Jocelyn Turnbull | 11:31 | 0.0066651 | 0.0066651 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.00667 | 0.00014 | 23629 | 2021-10-01 | 8:51 | Jeremy Thompson | 218818 |  |  |  |  | 45% | 5 | 1/10/21 | 11:30:52 | Jocelyn Turnbull | 0 |  | 0.67 | 0.01% | 22 | 5 | 5 | ... | 0.00641 | 0.00013 |  |  |  |  | 0.99065 | 0.99065 | Yes | Yes | [5 empty] | 0.00641 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00667 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00641 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-09-24 | Taylor Ferrick | 8:49 |  | Completed |  |  |  |  | 82456 | 2000000 |  |  | 82438 82439 82440 82441 82458 82459 | 82438~82439~82440~82441~82458~82459 | 3368 |  | 3368 3368 3368 3368 3368 3368 | 2021-10-01 | 11:00 | Jeremy Thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0017 | 0.00148 | 82438 |  | XCAMS | ST | 0% |  | TW3368 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 0x blanks 180 days, capturing 1x blanks + 9 AAA blanks Water 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-10-20 | Water |  | JOB-218818 |  |  | 72865 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00145 | 82439 | [26 empty] | 2021-09-20 | 5464 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.00146 | 82440 | [26 empty] | 2021-09-22 | 22466 | Graphite | Completed
[168 empty] | 0.99716 | 0.00142 | 82441 | [26 empty] | 2021-09-23 | 8372 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00159 | 82458 | [26 empty] | 2021-10-01 | 23629 | AMS Submission Results | Completed
[168 empty] | 0.99945 | 0.00161 | 82459 | [26 empty] | 2021-10-01 | 28074 | Close Job | Completed
[5 empty] | 4.5574e-05 | 0.0052676 | 95.5 | 88.9456 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7101 | 7100 | 7100 | 36 | 36 | 40 | 7064.5 | 7137.1 | Taylor Ferrick | -586.85 | 1.87 ‰ | 2021-09-29 | 2021-04-02 | 7 | 180 |  |  |  |  | -8.85‰ | -8.85‰ | 2.00‰ | 2.00‰ | -8.85 ‰ | 2.00 ‰ | AMS | -8.24‰ | 0.2‰ | C13 | No | -19 | -19.0051 | 0 | 0.8165 | -586.85 ‰ | -586.85 ‰ | 1.87 | 2.09 | -581.75 | 1.89 ‰ | -581.75 | 1.89 ‰ |  |  | 2021-10-01 | Jocelyn Turnbull | 11:31 | 0.41315 | 0.41315 | 0.00187 | 0.00209 | 0.00179 | 0.00179 | 1.9 |  |  | 0.41639 | 0.00116 | 23630 | 2021-10-01 | 8:51 | Jeremy Thompson | 218819 | 0.00516 | 0.0052 | 0.00232 | 0.0023 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.32 | 0.19% | 29 | 5 | 4 | ... | 0.40045 | 0.00099 |  |  |  |  | 0.99065 | 0.99065 | Yes | Yes |  | Result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.563382548623561. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00516 | 0.00507 | 0.00084 | 0.39734 | 0.40‰ | 0.00172 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.41315 | 0 | 0.00187 | 0 | 0 | 0.00516 | 0 | 0.00232 | 0 | 0.39734 | 0 | 0.00172 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-09-24 | Taylor Ferrick | 8:49 |  | Completed |  |  |  |  | 82457 | 2000000 | 82456 82086 81922 81829 81343 81316 | 82456~82086~81922~81829~81343~81316 | 82438 82439 82440 82441 82458 82459 | 82438~82439~82440~82441~82458~82459 | 3368 | 3368 3358 3355 3352 3341 3338 | 3368 3368 3368 3368 3368 3368 | 2021-10-01 | 11:00 | Jeremy Thompson |  | 1.01 mg | 0.13 | 1.01 | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 | 1.0017 | 0.00148 | 82438 | FLASK | XCAMS | ST | 0% |  | TW3368 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 0x blanks 180 days, capturing 1x blanks + 9 AAA blanks Water 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-10-20 | Water |  | JOB-218819 |  |  | 72866 | Normal | 14047/12 |  |  | 2021-09-21 | 5470 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 1.0001 | 0.00145 | 82439 | ST | [25 empty] | 2021-09-22 | 22467 | Graphite | Completed
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 1.0016 | 0.00146 | 82440 | ST | [25 empty] | 2021-09-23 | 8373 | C13 Measurement | Completed
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 0.99716 | 0.00142 | 82441 | ST | [25 empty] | 2021-10-01 | 23630 | AMS Submission Results | Completed
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 1.0001 | 0.00159 | 82458 | ST | [25 empty] | 2021-10-01 | 28075 | Close Job | Completed
[163 empty] | 2021-04-04 | 0.00497 | 81316 | 3338 | 1.04 | 0.99945 | 0.00161 | 82459 | ST
[7 empty] | 96.8 | 87.7945 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42728 | 42727 | 42727 | 188 | 187 | 187 |  |  | Taylor Ferrick | -995.1 | 0.11 ‰ | 2021-10-04 | 2021-04-07 | 5 | 180 |  |  |  |  | 3.13‰ | 3.13‰ | 2.00‰ | 2.00‰ | 3.13 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0024 | 0 | 0.8165 | -995.10 ‰ | -995.10 ‰ | 0.11 | 0.11 | -995.04 | 0.12 ‰ | -995.04 | 0.12 ‰ |  |  | 2021-10-13 | Jocelyn Turnbull | 8:51 | 0.0048974 | 0.0048974 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.8 |  |  | 0.0049 | 0.00011 | 23638 | 2021-10-06 | 8:27 | Jeremy Thompson | 218996 |  |  |  |  | 45% | 5 | 13/10/21 | 08:50:43 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 1 | 5 | 5 | ... | 0.00471 | 0.00011 |  |  |  |  | 0.98965 | 0.98965 | Yes | Yes | [5 empty] | 0.00471 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.0049 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00471 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-01 | Taylor Ferrick | 11:09 |  | Completed |  |  |  |  | 82500 | 2000000 |  |  | 82472 82473 82506 82518 82519 82522 | 82472~82473~82506~82518~82519~82522 | 3369 |  | 3369 3369 3369 3369 3369 3369 | 2021-10-06 | 8:52 | Jeremy Thompson |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99683 | 0.00157 | 82472 |  | XCAMS | EA | 0% |  | TW3369 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 16x unknowns + 1x secondaries + 0x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 1x blanks 180 days, capturing 2x blanks + 9 AAA blanks Water 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-11-17 | Water | Complete | JOB-218996 |  | For new treatment of marble material. Please make jobs for KIT, LAG, MEH, etc | 72905 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99911 | 0.00162 | 82473 | [26 empty] | 2021-09-28 | 5483 | Water CO2 Evolution | Completed
[168 empty] | 1.0014 | 0.00159 | 82506 | [26 empty] | 2021-09-29 | 22510 | Graphite | Completed
[168 empty] | 1.0005 | 0.00154 | 82518 | [26 empty] | 2021-09-30 | 8394 | C13 Measurement | Completed
[168 empty] | 0.99896 | 0.00158 | 82519 | [26 empty] | 2021-10-13 | 23638 | AMS Submission Results | Completed
[168 empty] | 1.0034 | 0.0016 | 82522 | [26 empty] | 2021-10-13 | 28161 | Close Job | Completed
[5 empty] | 4.9184e-05 | 0.0052341 | 101.4 | 88.8054 | 0.0692 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7105 | 7105 | 7105 | 36 | 36 | 40 | 7069.1 | 7141.7 | Taylor Ferrick | -587.09 | 1.87 ‰ | 2021-10-04 | 2021-04-07 | 5 | 180 |  |  |  |  | -8.29‰ | -8.29‰ | 2.00‰ | 2.00‰ | -8.29 ‰ | 2.00 ‰ | AMS | -8.18‰ | 0.2‰ | C13 | No | -19 | -19.0024 | 0 | 0.8165 | -587.09 ‰ | -587.09 ‰ | 1.87 | 2.09 | -581.99 | 1.89 ‰ | -581.99 | 1.89 ‰ |  |  | 2021-10-13 | Jocelyn Turnbull | 8:51 | 0.41291 | 0.41291 | 0.00187 | 0.00209 | 0.00179 | 0.00179 | 2.3 |  |  | 0.41613 | 0.00118 | 23647 | 2021-10-06 | 8:27 | Jeremy Thompson | 218820 | 0.00512 | 0.0051 | 0.0023 | 0.0023 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.29 | 0.19% | 8 | 5 | 4 | ... | 0.4002 | 0.00101 |  |  |  |  | 0.98965 | 0.98965 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 40.5806490373557186. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00512 | 0.00503 | 0.00086 | 0.39711 | 0.40‰ | 0.00172 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41291 | 0 | 0.00187 | 0 | 0 | 0.00512 | 0 | 0.0023 | 0 | 0.39711 | 0 | 0.00172 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-01 | Taylor Ferrick | 11:09 |  | Completed |  |  |  |  | 82499 | 2000000 | 82500 82456 82086 81922 81829 81343 | 82500~82456~82086~81922~81829~81343 | 82472 82473 82506 82518 82519 82522 | 82472~82473~82506~82518~82519~82522 | 3369 | 3369 3368 3358 3355 3352 3341 | 3369 3369 3369 3369 3369 3369 | 2021-10-06 | 8:52 | Jeremy Thompson |  | 0.87 mg | 0.13 | 0.87 | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 | 0.99683 | 0.00157 | 82472 | ST | XCAMS | EA | 0% |  | TW3369 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 16x unknowns + 1x secondaries + 0x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 1x blanks 180 days, capturing 2x blanks + 9 AAA blanks Water 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-10-20 | Water |  | JOB-218820 |  |  | 72906 | Normal | 14047/12 |  |  | 2021-09-28 | 5479 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 | 0.99911 | 0.00162 | 82473 | ST | [25 empty] | 2021-09-29 | 22509 | Graphite | Completed
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 1.0014 | 0.00159 | 82506 | ST | [25 empty] | 2021-09-30 | 8393 | C13 Measurement | Completed
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 1.0005 | 0.00154 | 82518 | ST | [25 empty] | 2021-10-13 | 23647 | AMS Submission Results | Completed
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 | 0.99896 | 0.00158 | 82519 | ST | [25 empty] | 2021-10-13 | 28151 | Close Job | Completed
[163 empty] | 2021-04-21 | 0.00523 | 81343 | 3341 | 0.89 | 1.0034 | 0.0016 | 82522 | EA
[7 empty] | 107.4 | 87.7447 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36404 | 36403 | 36403 | 124 | 124 | 124 |  |  | Taylor Ferrick | -989.24 | 0.17 ‰ | 2021-10-10 | 2021-04-13 | 3 | 180 |  |  |  |  | 5.04‰ | 5.04‰ | 2.00‰ | 2.00‰ | 5.04 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.8165 | -989.24 ‰ | -989.24 ‰ | 0.17 | 0.17 | -989.11 | 0.17 ‰ | -989.11 | 0.17 ‰ |  |  | 2021-10-13 | Jocelyn Turnbull | 9:00 | 0.010762 | 0.010762 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2.2 |  |  | 0.01076 | 0.00017 | 23678 | 2021-10-12 | 11:06 | Albert Zondervan | 219248 |  |  |  |  | 45% | 5 | 13/10/21 | 09:41:48 | jenny dahl | 0 |  | 1.08 | 0.02% | 1 | 5 | 6 | ... | 0.01035 | 0.00016 |  |  |  |  | 0.99097 | 0.99097 | Yes | Yes |  | poor blank |  |  |  | 0.01035 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.01076 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01035 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-08 | Taylor Ferrick | 11:44 |  | Completed |  |  |  |  | 82573 | 2000000 |  |  | 82523 82559 82560 82577 82578 82600 | 82523~82559~82560~82577~82578~82600 | 3370 |  | 3370 3370 3370 3370 3370 3370 | 2021-10-12 | 11:27 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0001 | 0.00161 | 82523 |  | XCAMS | EA | 0% |  | TW3370 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C6 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 1x unknowns + 0x secondaries + 0x blanks 360 days, capturing 3x blanks Water 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks blank TP82573 is high unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Complete | JOB-219248 |  | refreshed material under J 218996 | 72945 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.0016 | 82559 | [26 empty] | 2021-10-06 | 5494 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99776 | 0.00162 | 82560 | [26 empty] | 2021-10-06 | 5498 | Water CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00161 | 82577 | [26 empty] | 2021-10-07 | 22583 | Graphite | Completed
[168 empty] | 0.99969 | 0.00156 | 82578 | [26 empty] | 2021-10-11 | 8405 | C13 Measurement | Completed
[168 empty] | 1.001 | 0.00164 | 82600 | [26 empty] | 2021-10-13 | 23678 | AMS Submission Results | Completed
[197 empty] | 2021-10-13 | 28163 | Close Job | Completed
[7 empty] | 98.1 | 88.1074 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40878 | 40877 | 40877 | 163 | 162 | 162 |  |  | Taylor Ferrick | -993.83 | 0.12 ‰ | 2021-10-22 | 2021-04-25 | 3 | 180 |  |  |  |  | 1.82‰ | 1.82‰ | 2.00‰ | 2.00‰ | 1.82 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -993.83 ‰ | -993.83 ‰ | 0.12 | 0.12 | -993.76 | 0.13 ‰ | -993.76 | 0.13 ‰ |  |  | 2021-10-26 | Jocelyn Turnbull | 9:40 | 0.006166 | 0.006166 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00617 | 0.00012 | 23798 | 2021-10-23 | 14:38 | Albert Zondervan | 219249 |  |  |  |  | 45% | 5 | 26/10/21 | 09:40:01 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 1 | 5 | 7 | ... | 0.00593 | 0.00012 |  |  |  |  | 0.99188 | 0.99188 | Yes | Yes | [5 empty] | 0.00593 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00617 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00593 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-19 | Taylor Ferrick | 11:39 |  | Completed |  |  |  |  | 82685 | 2000000 |  |  | 82609 82610 82611 82612 82613 82614 | 82609~82610~82611~82612~82613~82614 | 3373 |  | 3373 3373 3373 3373 3373 3373 | 2021-10-23 | 14:42 | Albert Zondervan |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0011 | 0.00163 | 82609 |  | XCAMS | ST | 0% |  | TW3373 measurement * A retune was necessary as measurement _2 had poor OxI repeatability. * Only _3 measurement data have been imported into RLIMS. CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * IAEA-C2 travertine 1x result 2-sigma younger than consensus age * LAA1 coral 1x result seems to be in keeping with previously measured values * LAC1 coral 1x result seems to be in keeping with previously measured values blanks: * 1 groups identified to blank-correct: Water 28x unknowns + 3x secondaries + 1x blanks 180 days, capturing 7x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water |  | JOB-219249 |  | refreshed material under J 218996 | 73066 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00158 | 82610 | [26 empty] | 2021-10-18 | 5527 | Water CO2 Evolution | Completed
[168 empty] | 1.0003 | 0.00159 | 82611 | [26 empty] | 2021-10-18 | 3229 | Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00161 | 82612 | [26 empty] | 2021-10-18 | 5385 | Choose Gas Splitting | Completed
[168 empty] | 1.0001 | 0.0016 | 82613 | [26 empty] | 2021-10-19 | 22695 | Graphite | Completed
[168 empty] | 0.99915 | 0.00159 | 82614 | [26 empty] | 2021-10-19 | 8453 | C13 Measurement | Completed
[197 empty] | 2021-10-26 | 23798 | AMS Submission Results | Completed
[197 empty] | 2021-11-01 | 28360 | Close Job | Completed
[5 empty] | 4.7175e-05 | 0.0057825 | 97.3 | 89.0259 | 0.0697 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7056 | 7055 | 7055 | 40 | 40 | 44 | 7015.4 | 7095.8 | Taylor Ferrick | -584.52 | 2.08 ‰ | 2021-10-22 | 2021-04-25 | 3 | 180 |  |  |  |  | -8.52‰ | -8.52‰ | 2.00‰ | 2.00‰ | -8.52 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -584.52 ‰ | -584.52 ‰ | 2.08 | 2.28 | -579.39 | 2.10 ‰ | -579.39 | 2.10 ‰ |  |  | 2021-10-26 | Jocelyn Turnbull | 9:40 | 0.41548 | 0.41548 | 0.00208 | 0.00228 | 0.00201 | 0.00201 | 2.6 |  |  | 0.4192 | 0.0012 | 23805 | 2021-10-23 | 14:38 | Albert Zondervan | 218837 | 0.00597 | 0.006 | 0.00269 | 0.0027 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.55 | 0.21% | 22 | 5 | 4 | ... | 0.40316 | 0.00103 |  |  |  |  | 0.99188 | 0.99188 | Yes | Yes |  | result 2-sigma younger than consensus age wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.0642022824245521. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00597 | 0.00556 | 0.00211 | 0.39958 | 0.40‰ | 0.00193 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41548 | 0 | 0.00208 | 0 | 0 | 0.00597 | 0 | 0.00269 | 0 | 0.39958 | 0 | 0.00193 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-19 | Taylor Ferrick | 11:39 |  | Completed |  |  |  |  | 82686 | 2000000 | 82685 82573 82500 82456 82086 81922 81829 | 82685~82573~82500~82456~82086~81922~81829 | 82609 82610 82611 82612 82613 82614 | 82609~82610~82611~82612~82613~82614 | 3373 | 3373 3370 3369 3368 3358 3355 3352 | 3373 3373 3373 3373 3373 3373 | 2021-10-23 | 14:42 | Albert Zondervan |  | 0.72 mg | 0.13 | 0.72 | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 1.0011 | 0.00163 | 82609 | ST | XCAMS | ST | 0% |  | TW3373 measurement * A retune was necessary as measurement _2 had poor OxI repeatability. * Only _3 measurement data have been imported into RLIMS. CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * IAEA-C2 travertine 1x result 2-sigma younger than consensus age * LAA1 coral 1x result seems to be in keeping with previously measured values * LAC1 coral 1x result seems to be in keeping with previously measured values blanks: * 1 groups identified to blank-correct: Water 28x unknowns + 3x secondaries + 1x blanks 180 days, capturing 7x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-10-27 | Water |  | JOB-218837 |  | to be run alongside new coral standards to gather information for the data quality project | 73073 | Normal | 14047/12 |  |  | 2021-10-18 | 5530 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 | 1.0002 | 0.00158 | 82610 | ST | [25 empty] | 2021-10-19 | 22696 | Graphite | Completed
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 | 1.0003 | 0.00159 | 82611 | ST | [25 empty] | 2021-10-19 | 8452 | C13 Measurement | Completed
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 | 0.99922 | 0.00161 | 82612 | ST | [25 empty] | 2021-10-26 | 23805 | AMS Submission Results | Completed
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 1.0001 | 0.0016 | 82613 | EA | [25 empty] | 2021-11-01 | 28356 | Close Job | Completed
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 | 0.99915 | 0.00159 | 82614 | EA
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | ST
[7 empty] | 92.5 | 88.0067 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50914 | 50913 | 50913 | 236 | 236 | 236 |  |  | jacob leath | -998.23 | 0.05 ‰ | 2021-11-01 | 2021-05-05 | 5 | 180 |  |  |  |  | 3.50‰ | 3.50‰ | 2.00‰ | 2.00‰ | 3.50 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -998.23 ‰ | -998.23 ‰ | 0.05 | 0.05 | -998.21 | 0.05 ‰ | -998.21 | 0.05 ‰ |  |  | 2021-11-03 | Jocelyn Turnbull | 12:13 | 0.0017677 | 0.0017677 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2.2 |  |  | 0.00177 | 5e-05 | 23910 | 2021-11-03 | 9:48 | Jeremy Thompson | 219186 |  |  |  |  | 45% | 5 | 1/12/21 | 11:07:40 | jenny dahl | 0 |  | 0.18 | 0.01% | 27 | 5 | 5 | ... | 0.0017 | 5e-05 |  |  |  |  | 0.99241 | 0.99241 | Yes | Yes |  | note- sample was not evolved from job 219186 material- either older materials was used of newer /11 |  |  |  | 0.0017 | 0.00‰ | 5e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00177 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.0017 | 0 | 5e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-29 | jacob leath | 11:02 |  | Completed |  |  |  |  | 82732 | 2000000 |  |  | 82702 82703 82704 82738 82739 82740 | 82702~82703~82704~82738~82739~82740 | 3375 |  | 3375 3375 3375 3375 3375 3375 | 2021-11-03 | 10:58 | Jeremy Thompson |  | 0.93 mg |  | 0.93 | [5 empty] | 0.99934 | 0.00136 | 82702 |  | XCAMS | EA | 0% |  | TW3375 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x result 1 sigma out * FIRI-D 1x result agrees with consensus value IAEA-C2: Freshwater Travertine 1x Result 2 sigma out blanks: * 3 groups identified to blank-correct: Organics via AAA: 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 10x blanks Organics via Cellulose Extraction: 1x unknowns + 1x secondaries + 0x blanks 360 days, capturing 3x blanks Carbonates 16x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2021-11-22 | Carbonate | Complete | JOB-219186 |  | refreshed material please make more jobs for future wheels (SMA needs a job from this parent job) | 73145 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99755 | 0.00134 | 82703 | [26 empty] | 2021-10-26 | 5536 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0002 | 0.00134 | 82704 | [26 empty] | 2021-10-27 | 22742 | Graphite | Completed
[168 empty] | 0.99981 | 0.00136 | 82738 | [26 empty] | 2021-10-28 | 8460 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00138 | 82739 | [26 empty] | 2021-11-03 | 23910 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00138 | 82740 | [26 empty] | 2021-12-09 | 28641 | Close Job | Completed
[7 empty] | 141.4 | 87.8688 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37499 | 37499 | 37499 | 107 | 106 | 106 |  |  | Taylor Ferrick | -990.61 | 0.12 ‰ | 2021-11-05 | 2021-05-09 | 1 | 180 |  |  |  |  | -0.35‰ | -0.35‰ | 2.00‰ | 2.00‰ | -0.35 ‰ | 2.00 ‰ | AMS | 2.22‰ | 0.2‰ | C13 | No | -19 | -19.0025 | 0 | 0.8165 | -990.61 ‰ | -990.61 ‰ | 0.12 | 0.12 | -990.49 | 0.13 ‰ | -990.49 | 0.13 ‰ |  |  | 2021-11-08 | Jocelyn Turnbull | 9:36 | 0.0093893 | 0.0093893 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.8 |  |  | 0.00939 | 0.00012 | 23918 | 2021-11-08 | 8:56 | Jeremy Thompson | 219250 |  |  |  |  | 45% | 5 | 9/11/21 | 12:11:33 | jenny dahl | 0 |  | 0.94 | 0.01% | 1 | 5 | 5 | ... | 0.00903 | 0.00012 |  |  |  |  | 0.98706 | 0.98706 | Yes | Yes | [5 empty] | 0.00903 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00054 | 0.00054 | 0.00939 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00903 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-04 | Taylor Ferrick | 12:16 |  | Completed |  |  |  |  | 82800 | 2000000 |  |  | 82752 82753 82758 82759 82760 82761 | 82752~82753~82758~82759~82760~82761 | 3376 |  | 3376 3376 3376 3376 3376 3376 | 2021-11-08 | 9:23 | Jeremy Thompson |  | 0.74 mg |  | 0.74 | [5 empty] | 1.0019 | 0.00135 | 82752 |  | XCAMS | ST | 0% |  | TW3376 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consenous IAEA-C2: Freshwater Travertine 1x Result 2 sigma out LAC: LAC1 coral 1x Result agrees with consenous LAA: LAA1 coral 1x Result agrees with consenous blanks: * 3 groups identified to blank-correct: Organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 17x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Not Complete | JOB-219250 |  | refreshed material under J 218996 | 73185 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9996 | 0.00132 | 82753 | [26 empty] | 2021-11-03 | 5569 | Water CO2 Evolution | Completed
[168 empty] | 0.99998 | 0.0013 | 82758 | [26 empty] | 2021-11-04 | 22810 | Graphite | Completed
[168 empty] | 0.99929 | 0.00128 | 82759 | [26 empty] | 2021-11-09 | 8508 | C13 Measurement | Completed
[168 empty] | 0.99934 | 0.00132 | 82760 | [26 empty] | 2021-11-08 | 23918 | AMS Submission Results | Completed
[168 empty] | 1 | 0.0013 | 82761 | [26 empty] | 2021-11-12 | 28423 | Close Job | Completed
[5 empty] | 4.062e-05 | 0.0055441 | 141.2 | 88.6969 | 0.0698 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7016 | 7016 | 7016 | 37 | 37 | 41 | 6978.8 | 7053.8 | Taylor Ferrick | -582.48 | 1.95 ‰ | 2021-11-05 | 2021-05-09 | 1 | 180 |  |  |  |  | -9.68‰ | -9.68‰ | 2.00‰ | 2.00‰ | -9.68 ‰ | 2.00 ‰ | AMS | -8.31‰ | 0.2‰ | C13 | No | -19 | -19.0025 | 0 | 0.8165 | -582.48 ‰ | -582.48 ‰ | 1.95 | 2.16 | -577.33 | 1.97 ‰ | -577.33 | 1.97 ‰ |  |  | 2021-11-08 | Jocelyn Turnbull | 9:36 | 0.41752 | 0.41752 | 0.00195 | 0.00216 | 0.00187 | 0.00187 | 3.2 |  |  | 0.42115 | 0.00105 | 23927 | 2021-11-08 | 8:56 | Jeremy Thompson | 218983 | 0.00583 | 0.0058 | 0.00262 | 0.0026 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.75 | 0.19% | 15 | 5 | 4 | ... | 0.40503 | 0.00086 |  |  |  |  | 0.98706 | 0.98706 | Yes | Yes |  | Result 2 sigma out from consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 41.6284374573993947. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00583 | 0.00533 | 0.00261 | 0.40154 | 0.40‰ | 0.0018 | 0.00‰ | Yes | 1 | 1 | 0.00054 | 0.00054 | 0.41752 | 0 | 0.00195 | 0 | 0 | 0.00583 | 0 | 0.00262 | 0 | 0.40154 | 0 | 0.0018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-04 | Taylor Ferrick | 12:16 |  | Completed |  |  |  |  | 82799 | 2000000 | 82800 82685 82573 82500 82456 82086 81922 81829 81572 | 82800~82685~82573~82500~82456~82086~81922~81829~81572 | 82752 82753 82758 82759 82760 82761 | 82752~82753~82758~82759~82760~82761 | 3376 | 3376 3373 3370 3369 3368 3358 3355 3352 3347 | 3376 3376 3376 3376 3376 3376 | 2021-11-08 | 9:23 | Jeremy Thompson |  | 0.64 mg | 0.13 | 0.64 | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 1.0019 | 0.00135 | 82752 | ST | XCAMS | ST | 0% |  | TW3376 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consenous IAEA-C2: Freshwater Travertine 1x Result 2 sigma out LAC: LAC1 coral 1x Result agrees with consenous LAA: LAA1 coral 1x Result agrees with consenous blanks: * 3 groups identified to blank-correct: Organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 17x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-11-12 | Water | Not Complete | JOB-218983 |  | to be run alongside new coral standards to gather information for the data quality project | 73186 | Normal | 14047/12 |  |  | 2021-11-03 | 5568 | Water CO2 Evolution | Completed | 14047/1 | IAEA-C1: Carrara Marble | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 0.9996 | 0.00132 | 82753 | ST | [25 empty] | 2021-11-04 | 22809 | Graphite | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 | 0.99998 | 0.0013 | 82758 | ST | [25 empty] | 2021-11-09 | 8502 | C13 Measurement | Completed
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 | 0.99929 | 0.00128 | 82759 | ST | [25 empty] | 2021-11-08 | 23927 | AMS Submission Results | Completed
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 | 0.99934 | 0.00132 | 82760 | ST | [25 empty] | 2021-11-12 | 28422 | Close Job | Completed
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 | 1 | 0.0013 | 82761 | EA
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 |  |  |  | EA
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | ST
[163 empty] | 2021-05-17 | 0.00165 | 81572 | 3347 | 1.04 |  |  |  | ST
[7 empty] | 115.2 | 87.8949 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37411 | 37411 | 37411 | 150 | 149 | 149 |  |  | Taylor Ferrick | -990.51 | 0.18 ‰ | 2021-11-11 | 2021-05-15 | 1 | 180 |  |  |  |  | 2.77‰ | 2.77‰ | 2.00‰ | 2.00‰ | 2.77 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0036 | 0 | 0.8165 | -990.51 ‰ | -990.51 ‰ | 0.18 | 0.18 | -990.39 | 0.18 ‰ | -990.39 | 0.18 ‰ |  |  | 2021-11-12 | Jocelyn Turnbull | 14:27 | 0.0094933 | 0.0094933 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.4 |  |  | 0.00949 | 0.00018 | 23989 | 2021-11-12 | 13:02 | Albert Zondervan | 219251 |  |  |  |  | 45% | 5 | 12/11/21 | 14:27:10 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 1 | 5 | 7 | ... | 0.00913 | 0.00017 |  |  |  |  | 0.99043 | 0.99043 | Yes | Yes | [5 empty] | 0.00913 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00949 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00913 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-10 | Taylor Ferrick | 12:36 |  | Completed |  |  |  |  | 82861 | 2000000 |  |  | 82790 82791 82792 82850 82853 82856 | 82790~82791~82792~82850~82853~82856 | 3377 |  | 3377 3377 3377 3377 3377 3377 | 2021-11-12 | 13:27 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 1.0009 | 0.00145 | 82790 |  | XCAMS | EA | 0% |  | TW3377 measurement CalAMS error adjustments: * none needed primaries: * 6 OxI cathodes: no failures secondaries: * water IAEA-C2: 1x result differs from consensus value by 2.5 sigma LAC1: 1x result agrees with expectation value LAA1: 1x Result agrees with expectation value * organic FIRI-I 1x result agrees with consensus value blanks: * water 14x unknowns + 3x secondaries + 1x blanks 180 days, capturing 9x blanks * organic via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Not Complete | JOB-219251 |  | refreshed material under J 218996 | 73225 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0005 | 0.00171 | 82791 | [26 empty] | 2021-11-08 | 5587 | Water CO2 Evolution | Completed
[168 empty] | 0.99918 | 0.00157 | 82792 | [26 empty] | 2021-11-08 | 3243 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00163 | 82850 | [26 empty] | 2021-11-08 | 5405 | Choose Gas Splitting | Completed
[168 empty] | 0.99657 | 0.00154 | 82853 | [26 empty] | 2021-11-10 | 22871 | Graphite | Completed
[168 empty] | 1.002 | 0.00155 | 82856 | [26 empty] | 2021-11-16 | 8524 | C13 Measurement | Completed
[197 empty] | 2021-11-12 | 23989 | AMS Submission Results | Completed
[197 empty] | 2021-11-29 | 28556 | Close Job | Completed
[5 empty] | 4.2813e-05 | 0.0064248 | 112.4 | 88.9586 | 0.0702 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 6995 | 6995 | 6995 | 43 | 42 | 46 | 6952.2 | 7038.1 | Taylor Ferrick | -581.38 | 2.24 ‰ | 2021-11-11 | 2021-05-15 | 1 | 180 |  |  |  |  | -9.22‰ | -9.22‰ | 2.00‰ | 2.00‰ | -9.22 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0036 | 0 | 0.8165 | -581.38 ‰ | -581.38 ‰ | 2.24 | 2.43 | -576.21 | 2.27 ‰ | -576.21 | 2.27 ‰ |  |  | 2021-11-12 | Jocelyn Turnbull | 14:27 | 0.41862 | 0.41862 | 0.00224 | 0.00243 | 0.00217 | 0.00217 | 1.8 |  |  | 0.42276 | 0.00122 | 23992 | 2021-11-12 | 13:02 | Albert Zondervan | 219314 | 0.00666 | 0.0067 | 0.003 | 0.003 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.86 | 0.22% | 15 | 5 | 4 | ... | 0.40658 | 0.00105 |  |  |  |  | 0.99043 | 0.99043 | Yes | Yes |  | result differs from consensus value by 2.5 sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 46.6392600491529135. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00666 | 0.00618 | 0.00228 | 0.4026 | 0.40‰ | 0.00209 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.41862 | 0 | 0.00224 | 0 | 0 | 0.00666 | 0 | 0.003 | 0 | 0.4026 | 0 | 0.00209 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-10 | Taylor Ferrick | 12:36 |  | Completed |  |  |  |  | 82863 | 2000000 | 82861 82800 82685 82573 82500 82456 82086 81922 81829 | 82861~82800~82685~82573~82500~82456~82086~81922~81829 | 82790 82791 82792 82850 82853 82856 | 82790~82791~82792~82850~82853~82856 | 3377 | 3377 3376 3373 3370 3369 3368 3358 3355 3352 | 3377 3377 3377 3377 3377 3377 | 2021-11-12 | 13:27 | Albert Zondervan |  | 1.09 mg | 0.13 | 1.09 | 2021-11-11 | 0.00913 | 82861 | 3377 | 0.78 | 1.0009 | 0.00145 | 82790 | ST | XCAMS | EA | 0% |  | TW3377 measurement CalAMS error adjustments: * none needed primaries: * 6 OxI cathodes: no failures secondaries: * water IAEA-C2: 1x result differs from consensus value by 2.5 sigma LAC1: 1x result agrees with expectation value LAA1: 1x Result agrees with expectation value * organic FIRI-I 1x result agrees with consensus value blanks: * water 14x unknowns + 3x secondaries + 1x blanks 180 days, capturing 9x blanks * organic via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | -7.9 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-12-02 | Water |  | JOB-219314 |  | bottles were obviously swapped in C13 with R41347/12 (J219312) originally came out as -1.17. Swapped data. NT 16/11/21 to be run alongside new coral standards to gather information for the data quality project | 73226 | Normal | 14047/12 |  |  | 2021-11-09 | 5589 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 1.0005 | 0.00171 | 82791 | ST | [25 empty] | 2021-11-10 | 22873 | Graphite | Completed
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 0.99918 | 0.00157 | 82792 | ST | [25 empty] | 2021-11-16 | 8526 | C13 Measurement | Completed
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 | 1.0009 | 0.00163 | 82850 | ST | [25 empty] | 2021-11-12 | 23992 | AMS Submission Results | Completed
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 | 0.99657 | 0.00154 | 82853 | EA | [25 empty] | 2021-11-29 | 28558 | Close Job | Completed
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 | 1.002 | 0.00155 | 82856 | EA
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 |  |  |  | ST
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 |  |  |  | ST
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | EA
[7 empty] | 79.8 | 88.3545 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41979 | 41979 | 41979 | 186 | 186 | 186 |  |  | Taylor Ferrick | -994.62 | 0.12 ‰ | 2021-11-22 | 2021-05-26 | 4 | 180 |  |  |  |  | 0.31‰ | 0.31‰ | 2.00‰ | 2.00‰ | 0.31 ‰ | 2.00 ‰ | AMS | 2.19‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -994.62 ‰ | -994.62 ‰ | 0.12 | 0.12 | -994.56 | 0.13 ‰ | -994.56 | 0.13 ‰ |  |  | 2021-11-24 | Jocelyn Turnbull | 10:35 | 0.0053757 | 0.0053757 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00538 | 0.00012 | 24038 | 2021-11-24 | 9:50 | Jeremy Thompson | 219337 |  |  |  |  | 45% | 5 | 24/11/21 | 10:35:18 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 1 | 5 | 7 | ... | 0.00517 | 0.00012 |  |  |  |  | 0.99316 | 0.99316 | Yes | Yes | [5 empty] | 0.00517 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00538 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00517 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-19 | Taylor Ferrick | 11:32 |  | Completed |  |  |  |  | 82938 | 2000000 |  |  | 82859 82891 82892 82893 82894 82895 | 82859~82891~82892~82893~82894~82895 | 3379 |  | 3379 3379 3379 3379 3379 3379 | 2021-11-24 | 10:01 | Jeremy Thompson |  | 0.80 mg |  | 0.8 | [5 empty] | 1.0013 | 0.00155 | 82859 |  | XCAMS | EA | 0% |  | TW3379 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with consensus blanks: * 1 groups identified to blank-correct: Waters 29x unknowns + 2x secondaries (1 coral) + 1x blanks 180 days, capturing 10x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2021-12-06 | Water | Complete | JOB-219337 |  |  | 73265 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99928 | 0.00161 | 82891 | [26 empty] | 2021-11-17 | 5623 | Water CO2 Evolution | Completed
[168 empty] | 0.9979 | 0.00163 | 82892 | [26 empty] | 2021-11-17 | 3253 | Gas Splitting | Completed
[168 empty] | 0.99989 | 0.00159 | 82893 | [26 empty] | 2021-11-17 | 5415 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00158 | 82894 | [26 empty] | 2021-11-18 | 22948 | Graphite | Completed
[168 empty] | 1.0001 | 0.00162 | 82895 | [26 empty] | 2021-11-22 | 8569 | C13 Measurement | Completed
[197 empty] | 2021-11-24 | 24038 | AMS Submission Results | Completed
[197 empty] | 2021-11-24 | 28490 | Close Job | Completed
[5 empty] | 4.0495e-05 | 0.0063143 | 91.2 | 89.1595 | 0.0689 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7134 | 7134 | 7134 | 43 | 43 | 47 | 7090.9 | 7177.9 | Taylor Ferrick | -588.57 | 2.23 ‰ | 2021-11-22 | 2021-05-26 | 4 | 180 |  |  |  |  | -8.72‰ | -8.72‰ | 2.00‰ | 2.00‰ | -8.72 ‰ | 2.00 ‰ | AMS | -8.26‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -588.57 ‰ | -588.57 ‰ | 2.23 | 2.41 | -583.49 | 2.26 ‰ | -583.49 | 2.26 ‰ |  |  | 2021-11-24 | Jocelyn Turnbull | 10:35 | 0.41143 | 0.41143 | 0.00223 | 0.00241 | 0.00216 | 0.00216 | 1.8 |  |  | 0.41551 | 0.00123 | 24046 | 2021-11-24 | 9:50 | Jeremy Thompson | 219315 | 0.00651 | 0.0065 | 0.00293 | 0.0029 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.14 | 0.22% | 29 | 5 | 4 | ... | 0.39961 | 0.00106 |  |  |  |  | 0.99316 | 0.99316 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 47.1104213748460815. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00651 | 0.00607 | 0.0022 | 0.39568 | 0.40‰ | 0.00208 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.41143 | 0 | 0.00223 | 0 | 0 | 0.00651 | 0 | 0.00293 | 0 | 0.39568 | 0 | 0.00208 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-19 | Taylor Ferrick | 11:32 |  | Completed |  |  |  |  | 82937 | 2000000 | 82938 82861 82800 82685 82573 82500 82456 82086 81922 81829 | 82938~82861~82800~82685~82573~82500~82456~82086~81922~81829 | 82859 82891 82892 82893 82894 82895 | 82859~82891~82892~82893~82894~82895 | 3379 | 3379 3377 3376 3373 3370 3369 3368 3358 3355 3352 | 3379 3379 3379 3379 3379 3379 | 2021-11-24 | 10:01 | Jeremy Thompson |  | 1.01 mg | 0.13 | 1.01 | 2021-11-22 | 0.00517 | 82938 | 3379 | 0.8 | 1.0013 | 0.00155 | 82859 | ST | XCAMS | EA | 0% |  | TW3379 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with consensus blanks: * 1 groups identified to blank-correct: Waters 29x unknowns + 2x secondaries (1 coral) + 1x blanks 180 days, capturing 10x blanks unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-12-02 | Water |  | JOB-219315 |  | to be run alongside new coral standards to gather information for the data quality project | 73266 | Normal | 14047/12 |  |  | 2021-11-17 | 5622 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-11-11 | 0.00913 | 82861 | 3377 | 0.78 | 0.99928 | 0.00161 | 82891 | ST | [25 empty] | 2021-11-18 | 22947 | Graphite | Completed
[163 empty] | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 0.9979 | 0.00163 | 82892 | ST | [25 empty] | 2021-11-22 | 8568 | C13 Measurement | Completed
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 0.99989 | 0.00159 | 82893 | ST | [25 empty] | 2021-11-24 | 24046 | AMS Submission Results | Completed
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 | 1.0013 | 0.00158 | 82894 | EA | [25 empty] | 2021-11-24 | 28489 | Close Job | Completed
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 | 1.0001 | 0.00162 | 82895 | EA
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 |  |  |  | ST
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 |  |  |  | ST
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 |  |  |  | EA
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | EA
[5 empty] | 3.7735e-05 | 0.006214 | 141.5 | 88.9211 | 0.0696 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7049 | 7048 | 7048 | 41 | 41 | 44 | 7007.7 | 7089.9 | jacob leath | -584.17 | 2.13 ‰ | 2021-11-25 | 2021-05-29 | 2 | 180 |  |  |  |  | -9.83‰ | -9.83‰ | 2.00‰ | 2.00‰ | -9.83 ‰ | 2.00 ‰ | AMS | -7.99‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -584.17 ‰ | -584.17 ‰ | 2.13 | 2.33 | -579.03 | 2.15 ‰ | -579.03 | 2.15 ‰ |  |  | 2021-11-26 | Jocelyn Turnbull | 9:22 | 0.41583 | 0.41583 | 0.00213 | 0.00233 | 0.00206 | 0.00206 | 2.5 |  |  | 0.41983 | 0.00112 | 24108 | 2021-11-26 | 7:59 | Jeremy Thompson | 219489 | 0.0064 | 0.0064 | 0.00288 | 0.0029 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.58 | 0.21% | 1 | 5 | 6 | ... | 0.40376 | 0.00094 |  |  |  |  | 0.9894 | 0.9894 | Yes | Yes |  | result agrees with consensus value wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 44.9218575333798762. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.0064 | 0.00598 | 0.00212 | 0.39992 | 0.40‰ | 0.00198 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.41583 | 0 | 0.00213 | 0 | 0 | 0.0064 | 0 | 0.00288 | 0 | 0.39992 | 0 | 0.00198 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-24 | jacob leath | 11:48 |  | Completed |  |  |  |  | 82979 | 2000000 | 82996 82938 82861 82800 82685 82573 82500 82456 82086 81922 81829 | 82996~82938~82861~82800~82685~82573~82500~82456~82086~81922~81829 | 82960 82961 82970 82971 82972 82973 | 82960~82961~82970~82971~82972~82973 | 3380 | 3380 3379 3377 3376 3373 3370 3369 3368 3358 3355 3352 | 3380 3380 3380 3380 3380 3380 | 2021-11-26 | 8:16 | Jeremy Thompson |  | 0.82 mg | 0.13 | 0.82 | 2021-11-25 | 0.00534 | 82996 | 3380 | 0.81 | 1.0008 | 0.00151 | 82960 | ST | XCAMS | ST | 0% |  | TW3380 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consensus IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with expected value LAA: LAA1 coral 1x Result agrees with expected value blanks: * 2 groups identified to blank-correct: Organics via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 14x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | -8 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2021-12-29 | Water |  | JOB-219489 |  |  | 73306 | Normal | 14047/12 |  |  | 2021-11-22 | 5635 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-11-22 | 0.00517 | 82938 | 3379 | 0.8 | 1.0009 | 0.00151 | 82961 | ST | [25 empty] | 2021-11-22 | 3259 | Gas Splitting | Completed
[163 empty] | 2021-11-11 | 0.00913 | 82861 | 3377 | 0.78 | 0.99808 | 0.00153 | 82970 | ST | [25 empty] | 2021-11-22 | 5424 | Choose Gas Splitting | Completed
[163 empty] | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 1.0004 | 0.0015 | 82971 | ST | [25 empty] | 2021-11-23 | 22989 | Graphite | Completed
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 1.0013 | 0.00144 | 82972 | EA | [25 empty] | 2021-11-24 | 8585 | C13 Measurement | Completed
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 | 0.99849 | 0.00146 | 82973 | EA | [25 empty] | 2021-11-26 | 24108 | AMS Submission Results | Completed
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 |  |  |  | ST | [25 empty] | 2021-11-29 | 28518 | Close Job | Completed
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 |  |  |  | ST
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 |  |  |  | EA
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 |  |  |  | EA
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | ST
[7 empty] | 151.1 | 87.8361 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41719 | 41719 | 41719 | 150 | 150 | 150 |  |  | jacob leath | -994.45 | 0.10 ‰ | 2021-11-25 | 2021-05-29 | 1 | 180 |  |  |  |  | 2.40‰ | 2.40‰ | 2.00‰ | 2.00‰ | 2.40 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -994.45 ‰ | -994.45 ‰ | 0.1 | 0.1 | -994.38 | 0.11 ‰ | -994.38 | 0.11 ‰ |  |  | 2021-11-26 | Jocelyn Turnbull | 9:22 | 0.0055525 | 0.0055525 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.00555 | 0.0001 | 24111 | 2021-11-26 | 7:59 | Jeremy Thompson | 219492 |  |  |  |  | 45% | 5 | 26/11/21 | 09:22:07 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 20 | 5 | 7 | ... | 0.00534 | 0.0001 |  |  |  |  | 0.9894 | 0.9894 | Yes | Yes | [5 empty] | 0.00534 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00555 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00534 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-24 | jacob leath | 11:48 |  | Completed |  |  |  |  | 82996 | 2000000 |  |  | 82960 82961 82970 82971 82972 82973 | 82960~82961~82970~82971~82972~82973 | 3380 |  | 3380 3380 3380 3380 3380 3380 | 2021-11-26 | 8:16 | Jeremy Thompson |  | 0.81 mg |  | 0.81 | [5 empty] | 1.0008 | 0.00151 | 82960 |  | XCAMS | ST | 0% |  | TW3380 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consensus IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with expected value LAA: LAA1 coral 1x Result agrees with expected value blanks: * 2 groups identified to blank-correct: Organics via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 14x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-12-29 | Water | Complete | JOB-219492 |  |  | 73305 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0009 | 0.00151 | 82961 | [26 empty] | 2021-11-23 | 5643 | Water CO2 Evolution | Completed
[168 empty] | 0.99808 | 0.00153 | 82970 | [26 empty] | 2021-11-23 | 3261 | Gas Splitting | Completed
[168 empty] | 1.0004 | 0.0015 | 82971 | [26 empty] | 2021-11-23 | 5426 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00144 | 82972 | [26 empty] | 2021-11-24 | 23006 | Graphite | Completed
[168 empty] | 0.99849 | 0.00146 | 82973 | [26 empty] | 2021-11-26 | 8595 | C13 Measurement | Completed
[197 empty] | 2021-11-26 | 24111 | AMS Submission Results | Completed
[197 empty] | 2021-11-29 | 28519 | Close Job | Completed
[7 empty] | 89.5 | 88.2171 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37634 | 37634 | 37634 | 136 | 135 | 135 |  |  | Taylor Ferrick | -990.77 | 0.16 ‰ | 2021-12-03 | 2021-06-06 | 3 | 180 |  |  |  |  | -0.03‰ | -0.03‰ | 2.00‰ | 2.00‰ | -0.03 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -990.77 ‰ | -990.77 ‰ | 0.16 | 0.16 | -990.65 | 0.16 ‰ | -990.65 | 0.16 ‰ |  |  | 2021-12-06 | Jocelyn Turnbull | 8:50 | 0.0092334 | 0.0092334 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.00923 | 0.00016 | 24158 | 2021-12-04 | 19:50 | Albert Zondervan | 219534 |  |  |  |  | 45% | 5 | 6/12/21 | 08:49:35 | Jocelyn Turnbull | 0 |  | 0.92 | 0.02% | 1 | 5 | 5 | ... | 0.00888 | 0.00015 |  |  |  |  | 0.99128 | 0.99128 | Yes | Yes | [5 empty] | 0.00888 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00923 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00888 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-12-02 | Taylor Ferrick | 12:52 |  | Completed |  |  |  |  | 83048 | 2000000 |  |  | 83028 83029 83030 83055 83056 83057 | 83028~83029~83030~83055~83056~83057 | 3382 |  | 3382 3382 3382 3382 3382 3382 | 2021-12-04 | 20:09 | Albert Zondervan |  | 0.93 mg |  | 0.93 | [5 empty] | 1.001 | 0.00161 | 83028 |  | XCAMS | EA | 0% |  | TW3382 measurement * partially filled wheel CalAMS error adjustments: * error increased for TP83036,83031 to account for excess run-to-run variability primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value waters: IAEA-C2 result deviates from consensus value by 2.5-sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses organics via CE: none in this wheel blanks: * 2 groups identified to blank-correct: organics via AAA: 13x unknowns + 2x secondaries + 1x blanks 180 days, capturing 13x blanks waters: 6x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks organics via CE: 2x unknowns + 0x secondaries + 0x blanks 360 days, capturing 4x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-01-04 | Water |  | JOB-219534 |  |  | 73385 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00155 | 83029 | [26 empty] | 2021-11-29 | 5663 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00155 | 83030 | [26 empty] | 2021-11-30 | 23058 | Graphite | Completed
[168 empty] | 0.9985 | 0.00152 | 83055 | [26 empty] | 2021-12-01 | 8625 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00152 | 83056 | [26 empty] | 2021-12-06 | 24158 | AMS Submission Results | Completed
[168 empty] | 0.99841 | 0.00154 | 83057 | [26 empty] | 2021-12-08 | 28614 | Close Job | Completed
[5 empty] | 3.6676e-05 | 0.0063809 | 92.2 | 89.1151 | 0.07 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7003 | 7002 | 7002 | 42 | 42 | 46 | 6960.5 | 7045.4 | Taylor Ferrick | -581.78 | 2.21 ‰ | 2021-12-03 | 2021-06-06 | 3 | 180 |  |  |  |  | -10.11‰ | -10.11‰ | 2.00‰ | 2.00‰ | -10.11 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -581.78 ‰ | -581.78 ‰ | 2.21 | 2.4 | -576.62 | 2.24 ‰ | -576.62 | 2.24 ‰ |  |  | 2021-12-06 | Jocelyn Turnbull | 8:50 | 0.41822 | 0.41822 | 0.00221 | 0.0024 | 0.00214 | 0.00214 | 1.9 |  |  | 0.42232 | 0.00119 | 24167 | 2021-12-04 | 19:50 | Albert Zondervan | 219535 | 0.00661 | 0.0066 | 0.00297 | 0.003 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.82 | 0.22% | 13 | 5 | 4 | ... | 0.40616 | 0.00102 |  |  |  |  | 0.99128 | 0.99128 | Yes | Yes |  | result deviates from consensus value by 2.5-sigma wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 46.1403642732965731. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00661 | 0.00614 | 0.00214 | 0.40221 | 0.40‰ | 0.00206 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.41822 | 0 | 0.00221 | 0 | 0 | 0.00661 | 0 | 0.00297 | 0 | 0.40221 | 0 | 0.00206 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-12-02 | Taylor Ferrick | 12:52 |  | Completed |  |  |  |  | 83049 | 2000000 | 83048 82996 82938 82861 82800 82685 82573 82500 82456 82086 81922 81829 | 83048~82996~82938~82861~82800~82685~82573~82500~82456~82086~81922~81829 | 83028 83029 83030 83055 83056 83057 | 83028~83029~83030~83055~83056~83057 | 3382 | 3382 3380 3379 3377 3376 3373 3370 3369 3368 3358 3355 3352 | 3382 3382 3382 3382 3382 3382 | 2021-12-04 | 20:09 | Albert Zondervan |  | 1.01 mg | 0.13 | 1.01 | 2021-12-03 | 0.00888 | 83048 | 3382 | 0.93 | 1.001 | 0.00161 | 83028 | ST | XCAMS | EA | 0% |  | TW3382 measurement * partially filled wheel CalAMS error adjustments: * error increased for TP83036,83031 to account for excess run-to-run variability primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value waters: IAEA-C2 result deviates from consensus value by 2.5-sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses organics via CE: none in this wheel blanks: * 2 groups identified to blank-correct: organics via AAA: 13x unknowns + 2x secondaries + 1x blanks 180 days, capturing 13x blanks waters: 6x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks organics via CE: 2x unknowns + 0x secondaries + 0x blanks 360 days, capturing 4x blanks unknowns: * no failures | -8.2 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2022-01-04 | Water | Complete | JOB-219535 |  |  | 73386 | Normal | 14047/12 |  |  | 2021-11-29 | 5662 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-11-25 | 0.00534 | 82996 | 3380 | 0.81 | 1.0013 | 0.00155 | 83029 | ST | [25 empty] | 2021-11-30 | 23059 | Graphite | Completed
[163 empty] | 2021-11-22 | 0.00517 | 82938 | 3379 | 0.8 | 1.0004 | 0.00155 | 83030 | ST | [25 empty] | 2021-12-02 | 8633 | C13 Measurement | Completed
[163 empty] | 2021-11-11 | 0.00913 | 82861 | 3377 | 0.78 | 0.9985 | 0.00152 | 83055 | ST | [25 empty] | 2021-12-06 | 24167 | AMS Submission Results | Completed
[163 empty] | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 1.0006 | 0.00152 | 83056 | EA | [25 empty] | 2021-12-08 | 28615 | Close Job | Completed
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 | 0.99841 | 0.00154 | 83057 | EA
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 |  |  |  | ST
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 |  |  |  | ST
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 |  |  |  | EA
[163 empty] | 2021-07-22 | 0.00477 | 82086 | 3358 | 1.01 |  |  |  | EA
[163 empty] | 2021-06-22 | 0.00567 | 81922 | 3355 | 0.94 |  |  |  | ST
[163 empty] | 2021-06-10 | 0.00393 | 81829 | 3352 | 0.88 |  |  |  | EA
[7 empty] | 77.9 | 88.6484 | 0.0009 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41674 | 41674 | 41674 | 224 | 224 | 224 |  |  | Taylor Ferrick | -994.42 | 0.16 ‰ | 2021-12-07 | 2021-06-10 | 0 | 180 |  |  |  |  | -5.21‰ | -5.21‰ | 2.00‰ | 2.00‰ | -5.21 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.42 ‰ | -994.42 ‰ | 0.16 | 0.16 | -994.35 | 0.16 ‰ | -994.35 | 0.16 ‰ |  |  | 2021-12-09 | Jocelyn Turnbull | 10:05 | 0.0055837 | 0.0055837 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 15.3 |  |  | 0.00558 | 0.00016 | 24231 | 2021-12-09 | 8:11 | Jeremy Thompson | 219683 |  |  |  |  | 45% | 5 | 9/12/21 | 10:04:52 | Jocelyn Turnbull | 0 |  | 0.56 | 0.02% | 1 | 5 | 5 | ... | 0.00537 | 0.00015 |  |  |  |  | 0.99097 | 0.99097 | Yes | Yes | [5 empty] | 0.00537 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00558 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00537 | 0 | 0.00015 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-12-07 | Taylor Ferrick | 10:05 |  | Completed |  |  |  |  | 83106 | 2000000 |  |  | 83014 83015 83067 83087 83088 83089 | 83014~83015~83067~83087~83088~83089 | 3384 |  | 3384 3384 3384 3384 3384 3384 | 2021-12-09 | 8:44 | Jeremy Thompson |  | 0.15 mg |  | 0.15 | [5 empty] | 1.0008 | 0.00164 | 83014 |  | XCAMS | ST | 0% |  | TW3384 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value carbonates: IAEA-C2 result deviates from consensus value by 2 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 14x blanks carbonates: 22x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2022-01-28 | Carbonate | Not Complete | JOB-219683 |  | small to match small Knaack samples in LAK4 wheel ~.15-.2mg recombust to match samples | 73421 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0014 | 0.00163 | 83015 | [26 empty] | 2021-12-06 | 5684 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99992 | 0.00163 | 83067 | [26 empty] | 2021-12-07 | 4810 | Sealed Tube Combustion | Completed
[168 empty] | 1.0007 | 0.00162 | 83087 | [26 empty] | 2021-12-07 | 23116 | Graphite | Completed
[168 empty] | 0.99796 | 0.00165 | 83088 | [26 empty] | 2021-12-09 | 24231 | AMS Submission Results | Completed
[168 empty] | 0.99921 | 0.00165 | 83089 | [26 empty] | 2021-12-09 | 28671 | Close Job | Completed
[7 empty] | 94.8 | 88.7656 | 0.0007 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43725 | 43725 | 43725 | 232 | 231 | 231 |  |  | Taylor Ferrick | -995.67 | 0.12 ‰ | 2021-12-07 | 2021-06-10 | 0 | 180 |  |  |  |  | -6.53‰ | -6.53‰ | 2.00‰ | 2.00‰ | -6.53 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.67 ‰ | -995.67 ‰ | 0.12 | 0.12 | -995.62 | 0.13 ‰ | -995.62 | 0.13 ‰ |  |  | 2021-12-09 | Jocelyn Turnbull | 10:05 | 0.0043255 | 0.0043255 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 12.2 |  |  | 0.00433 | 0.00012 | 24232 | 2021-12-09 | 8:11 | Jeremy Thompson | 219687 |  |  |  |  | 45% | 5 | 9/12/21 | 10:04:55 | Jocelyn Turnbull | 0 |  | 0.43 | 0.01% | 8 | 5 | 5 | ... | 0.00416 | 0.00012 |  |  |  |  | 0.99097 | 0.99097 | Yes | Yes | [5 empty] | 0.00416 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00433 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00416 | 0 | 0.00012 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2021-12-07 | Taylor Ferrick | 10:05 |  | Completed |  |  |  |  | 83107 | 2000000 |  |  | 83014 83015 83067 83087 83088 83089 | 83014~83015~83067~83087~83088~83089 | 3384 |  | 3384 3384 3384 3384 3384 3384 | 2021-12-09 | 8:44 | Jeremy Thompson |  | 0.19 mg |  | 0.19 | [5 empty] | 1.0008 | 0.00164 | 83014 |  | XCAMS | ST | 0% |  | TW3384 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value carbonates: IAEA-C2 result deviates from consensus value by 2 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 14x blanks carbonates: 22x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2022-01-31 | Carbonate | Complete | JOB-219687 |  | extra in case size doesn't match intended mg C recombust to match samples | 73422 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0014 | 0.00163 | 83015 | [26 empty] | 2021-12-06 | 5686 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99992 | 0.00163 | 83067 | [26 empty] | 2021-12-07 | 4812 | Sealed Tube Combustion | Completed
[168 empty] | 1.0007 | 0.00162 | 83087 | [26 empty] | 2021-12-07 | 23117 | Graphite | Completed
[168 empty] | 0.99796 | 0.00165 | 83088 | [26 empty] | 2021-12-09 | 24232 | AMS Submission Results | Completed
[168 empty] | 0.99921 | 0.00165 | 83089 | [26 empty] | 2021-12-09 | 28672 | Close Job | Completed
[7 empty] | 132.4 | 87.82 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52303 | 52303 | 52303 | 337 | 337 | 337 |  |  | Taylor Ferrick | -998.51 | 0.06 ‰ | 2021-12-07 | 2021-06-10 | 4 | 180 |  |  |  |  | 4.17‰ | 4.17‰ | 2.00‰ | 2.00‰ | 4.17 ‰ | 2.00 ‰ | AMS | 2.64‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -998.51 ‰ | -998.51 ‰ | 0.06 | 0.06 | -998.49 | 0.06 ‰ | -998.49 | 0.06 ‰ |  |  | 2021-12-09 | Jocelyn Turnbull | 10:05 | 0.0014869 | 0.0014869 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.8 |  |  | 0.00149 | 6e-05 | 24233 | 2021-12-09 | 8:11 | Jeremy Thompson | 219677 |  |  |  |  | 45% | 5 | 9/12/21 | 10:04:56 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 22 | 5 | 5 | ... | 0.00143 | 6e-05 |  |  |  |  | 0.99097 | 0.99097 | Yes | Yes | [5 empty] | 0.00143 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00149 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00143 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-12-07 | Taylor Ferrick | 10:05 |  | Completed |  |  |  |  | 83100 | 2000000 |  |  | 83014 83015 83067 83087 83088 83089 | 83014~83015~83067~83087~83088~83089 | 3384 |  | 3384 3384 3384 3384 3384 3384 | 2021-12-09 | 8:44 | Jeremy Thompson |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0008 | 0.00164 | 83014 |  | XCAMS | ST | 0% |  | TW3384 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value carbonates: IAEA-C2 result deviates from consensus value by 2 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 14x blanks carbonates: 22x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2022-01-26 | Carbonate |  | JOB-219677 |  |  | 73423 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0014 | 0.00163 | 83015 | [26 empty] | 2021-12-02 | 5682 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99992 | 0.00163 | 83067 | [26 empty] | 2021-12-03 | 23110 | Graphite | Completed
[168 empty] | 1.0007 | 0.00162 | 83087 | [26 empty] | 2021-12-06 | 8654 | C13 Measurement | Completed
[168 empty] | 0.99796 | 0.00165 | 83088 | [26 empty] | 2021-12-09 | 24233 | AMS Submission Results | Completed
[168 empty] | 0.99921 | 0.00165 | 83089 | [26 empty] | 2021-12-09 | 28670 | Close Job | Completed
[7 empty] | 110.1 | 87.9711 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39909 | 39908 | 39908 | 144 | 144 | 144 |  |  | Taylor Ferrick | -993.04 | 0.12 ‰ | 2022-02-04 | 2021-08-08 | 2 | 180 |  |  |  |  | 4.11‰ | 4.11‰ | 2.00‰ | 2.00‰ | 4.11 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -993.04 ‰ | -993.04 ‰ | 0.12 | 0.12 | -992.96 | 0.13 ‰ | -992.96 | 0.13 ‰ |  |  | 2022-02-08 | Jocelyn Turnbull | 10:18 | 0.0069562 | 0.0069562 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00696 | 0.00012 | 24295 | 2022-02-08 | 8:24 | Jeremy Thompson | 219613 |  |  |  |  | 45% | 5 | 8/02/22 | 10:17:31 | Jocelyn Turnbull | 0 |  | 0.7 | 0.01% | 1 | 5 | 5 | ... | 0.00669 | 0.00012 |  |  |  |  | 0.99262 | 0.99262 | Yes | Yes | [5 empty] | 0.00669 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00696 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00669 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-02-02 | Taylor Ferrick | 12:27 |  | Completed |  |  |  |  | 83189 | 2000000 |  |  | 83113 83114 83115 83116 83158 83169 | 83113~83114~83115~83116~83158~83169 | 3385 |  | 3385 3385 3385 3385 3385 3385 | 2022-02-08 | 8:41 | Jeremy Thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99812 | 0.00158 | 83113 |  | XCAMS | FLASK | 0% |  | TW3385 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: Waters: IAEA-C2 result deviates from consensus value by 3 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: Waters: 8x unknowns + 3x secondaries + 1x blanks 180 days, capturing 10x blanks C02 from CO: 20x unknowns + 0x secondaries + 4x blanks Speicfic blanks chosen to repeat past blank correction approach(TP: 68062, 68063, 68064, 68065 from TW3007) unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-01-13 | Water | Not Complete | JOB-219613 |  | New batch made for Marble water standard IAEA-C1. Has been crushed into powder instead of smaller chunks. | 73498 | Normal | 14047/11 |  |  | 2021-11-22 | 7706 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99944 | 0.00152 | 83114 | [26 empty] | 2022-02-01 | 5696 | Water CO2 Evolution | Completed
[168 empty] | 1.0013 | 0.0015 | 83115 | [26 empty] | 2022-02-02 | 23199 | Graphite | Completed
[168 empty] | 1.0005 | 0.00155 | 83116 | [26 empty] | 2022-02-04 | 8667 | C13 Measurement | Completed
[168 empty] | 0.9995 | 0.00155 | 83158 | [26 empty] | 2022-02-08 | 24295 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00148 | 83169 | [26 empty] | 2022-02-11 | 28819 | Close Job | Completed
[5 empty] | 4.1175e-05 | 0.0069786 | 110.2 | 88.8916 | 0.0697 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 7008 | 7008 | 7008 | 44 | 44 | 47 | 6964.1 | 7053 | Taylor Ferrick | -582.08 | 2.31 ‰ | 2022-02-04 | 2021-08-08 | 2 | 180 |  |  |  |  | -6.28‰ | -6.28‰ | 2.00‰ | 2.00‰ | -6.28 ‰ | 2.00 ‰ | AMS | -8.29‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -582.08 ‰ | -582.08 ‰ | 2.31 | 2.5 | -576.92 | 2.34 ‰ | -576.92 | 2.34 ‰ |  |  | 2022-02-08 | Jocelyn Turnbull | 10:18 | 0.41792 | 0.41792 | 0.00231 | 0.0025 | 0.00225 | 0.00225 | 2.9 |  |  | 0.42238 | 0.00112 | 24309 | 2022-02-08 | 8:24 | Jeremy Thompson | 219968 | 0.00716 | 0.0072 | 0.00322 | 0.0032 | 45% | 5 | 5/07/22 | 14:07:31 | Jocelyn Turnbull | 0 |  | 41.79 | 0.23% | 15 | 5 | 4 | ... | 0.40621 | 0.00094 |  |  |  |  | 0.99262 | 0.99262 | Yes | Yes |  | 3 sigma out wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 47.956828056635383. If Saved values show old error bar then the revised result has not been re-reported to the client. | 0.00716 | 0.00671 | 0.002 | 0.40193 | 0.40‰ | 0.00216 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.41792 | 0 | 0.00231 | 0 | 0 | 0.00716 | 0 | 0.00322 | 0 | 0.40193 | 0 | 0.00216 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-02-02 | Taylor Ferrick | 12:28 |  | Completed |  |  |  |  | 83188 | 2000000 | 83189 83048 82996 82938 82861 82800 82685 82573 82500 82456 | 83189~83048~82996~82938~82861~82800~82685~82573~82500~82456 | 83113 83114 83115 83116 83158 83169 | 83113~83114~83115~83116~83158~83169 | 3385 | 3385 3382 3380 3379 3377 3376 3373 3370 3369 3368 | 3385 3385 3385 3385 3385 3385 | 2022-02-08 | 8:41 | Jeremy Thompson |  | 0.71 mg | 0.13 | 0.71 | 2022-02-04 | 0.00669 | 83189 | 3385 | 1.12 | 0.99812 | 0.00158 | 83113 | ST | XCAMS | FLASK | 0% |  | TW3385 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: Waters: IAEA-C2 result deviates from consensus value by 3 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: Waters: 8x unknowns + 3x secondaries + 1x blanks 180 days, capturing 10x blanks C02 from CO: 20x unknowns + 0x secondaries + 4x blanks Speicfic blanks chosen to repeat past blank correction approach(TP: 68062, 68063, 68064, 68065 from TW3007) unknowns: * no failures | -8.3 | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2022-02-16 | Water | Complete | JOB-219968 |  | to be run alongside coral secondary for statistical purposes | 73499 | Normal | 14047/12 |  |  | 2022-02-01 | 5697 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Freshwater Travertine deposit,Munich, Federal Republic of Germany, GSF | Water line secondary |  | -8 | 7134 BP 41.14 ± 0.03 pMC |  | IAEA-C2: Freshwater Travertine | IAEA-C2 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2021-12-03 | 0.00888 | 83048 | 3382 | 0.93 | 0.99944 | 0.00152 | 83114 | EA | [25 empty] | 2022-02-02 | 23198 | Graphite | Completed
[163 empty] | 2021-11-25 | 0.00534 | 82996 | 3380 | 0.81 | 1.0013 | 0.0015 | 83115 | EA | [25 empty] | 2022-02-04 | 8670 | C13 Measurement | Completed
[163 empty] | 2021-11-22 | 0.00517 | 82938 | 3379 | 0.8 | 1.0005 | 0.00155 | 83116 | ST | [25 empty] | 2022-02-08 | 24309 | AMS Submission Results | Completed
[163 empty] | 2021-11-11 | 0.00913 | 82861 | 3377 | 0.78 | 0.9995 | 0.00155 | 83158 | ST | [25 empty] | 2022-02-11 | 28823 | Close Job | Completed
[163 empty] | 2021-11-05 | 0.00903 | 82800 | 3376 | 0.74 | 1.001 | 0.00148 | 83169 | EA
[163 empty] | 2021-10-22 | 0.00593 | 82685 | 3373 | 0.83 |  |  |  | EA
[163 empty] | 2021-10-10 | 0.01035 | 82573 | 3370 | 0.84 |  |  |  | ST
[163 empty] | 2021-10-04 | 0.00471 | 82500 | 3369 | 1.06 |  |  |  | EA
[163 empty] | 2021-09-29 | 0.00641 | 82456 | 3368 | 1.01 |  |  |  | FLASK
[7 empty] | 107.3 | 87.823 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51451 | 51451 | 51451 | 303 | 303 | 303 |  |  | Taylor Ferrick | -998.35 | 0.06 ‰ | 2022-02-22 | 2021-08-26 | 6 | 180 |  |  |  |  | 2.03‰ | 2.03‰ | 2.00‰ | 2.00‰ | 2.03 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -998.35 ‰ | -998.35 ‰ | 0.06 | 0.06 | -998.33 | 0.06 ‰ | -998.33 | 0.06 ‰ |  |  | 2022-02-23 | Jocelyn Turnbull | 15:23 | 0.0016533 | 0.0016533 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.2 |  |  | 0.00165 | 6e-05 | 24463 | 2022-02-23 | 12:16 | Albert Zondervan | 220121 |  |  |  |  | 45% | 5 | 23/02/22 | 15:38:27 | jenny dahl | 0 |  | 0.17 | 0.01% | 22 | 5 | 5 | ... | 0.00159 | 6e-05 |  |  |  |  | 0.98889 | 0.98889 | Yes | Yes | [5 empty] | 0.00159 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00165 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00159 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-02-21 | Taylor Ferrick | 8:07 |  | Completed |  |  |  |  | 83300 | 2000000 |  |  | 83245 83246 83299 83301 83302 83303 | 83245~83246~83299~83301~83302~83303 | 3389 |  | 3389 3389 3389 3389 3389 3389 | 2022-02-23 | 12:21 | Albert Zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99781 | 0.00156 | 83245 |  | XCAMS | ST | 0% |  | TW3389 measurement * large carbonates CalAMS error adjustments * none needed primaries: * 6 large OxI no failures secondaries: * carbonates 1x IAEA-C2 result deviates from consensus value by 2 sigma 2x LAC1 result deviates from expectation value by 2 sigma 1x LAC1 result agrees with expectation value blanks: * carbonates 27x unknowns + 4x secondaries + 1x blanks 180 days, capturing 4x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-03-08 | Carbonate |  | JOB-220121 |  | new ms bottle | 73651 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0018 | 0.00162 | 83246 | [26 empty] | 2022-02-15 | 5716 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9996 | 0.00161 | 83299 | [26 empty] | 2022-02-16 | 23310 | Graphite | Completed
[168 empty] | 0.99849 | 0.00157 | 83301 | [26 empty] | 2022-02-22 | 8716 | C13 Measurement | Completed
[168 empty] | 1.0008 | 0.00154 | 83302 | [26 empty] | 2022-02-23 | 24463 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00155 | 83303 | [26 empty] | 2022-02-23 | 28880 | Close Job | Completed
[7 empty] | 100.9 | 87.5479 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49608 | 49608 | 49608 | 361 | 361 | 361 |  |  | jacob leath | -997.92 | 0.09 ‰ | 2022-04-06 | 2021-10-08 | 9 | 180 |  |  |  |  | 2.66‰ | 2.66‰ | 2.00‰ | 2.00‰ | 2.66 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -997.92 ‰ | -997.92 ‰ | 0.09 | 0.09 | -997.89 | 0.09 ‰ | -997.89 | 0.09 ‰ |  |  | 2022-04-08 | Jocelyn Turnbull | 10:58 | 0.0020796 | 0.0020796 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00208 | 9e-05 | 24776 | 2022-04-07 | 14:43 | Albert Zondervan | 220197 |  |  |  |  | 45% | 5 | 6/05/22 | 12:08:47 | Cathy Ginnane | 0 |  | 0.21 | 0.01% | 19 | 5 | 5 | ... | 0.002 | 9e-05 |  |  |  |  | 0.98641 | 0.98641 | Yes | Yes |  | bit younger than recent blanks |  |  |  | 0.002 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00208 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.002 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-05 | jacob leath | 11:30 |  | Completed |  |  |  |  | 83598 | 2000000 |  |  | 83511 83599 83625 83627 83628 83629 | 83511~83599~83625~83627~83628~83629 | 3397 |  | 3397 3397 3397 3397 3397 3397 |  |  |  |  | 0.87 mg |  | 0.87 | [5 empty] | 0.99673 | 0.00206 | 83511 |  | XCAMS | ST | 0% |  | TW3397 measurement * small+large carbonates CalAMS error adjustments: * factor 1.08 to account for global run-to-run variability primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 4x coral results agree with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 6x unknowns + 4x secondaries + 1x blanks last 180 days, capturing 4x blanks small carbonates: 14x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. CG updated blank corrections and checked wheel results 6/5/22 for smalls in this wheel A bit of relevant history around the time of slushie changeover: 4-Aug-2021: the last OxI graphitised (TP82203) for the last smalls-wheel (TW3362) before the changeover to liqN2-Eth slushies 5-Aug: Taylor wrote the SOP for prepping liqN2-Eth slushies 13-Aug: Taylor instructed some members of the team how to do make such slushies 13-Sep: Cathy did the first RP run using liqN2-Eth slushies mid-Aug -- mid-Sep: the use of dry-ice was abandoned and the dry-ice freezer emptied 25-Nov + 3-Dec: OxI graphitised for TW3383, the first wheel of smalls prepped after the slushie changeover | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-03-17 | Carbonate | Complete | JOB-220197 |  | imported graphite results manually NT | 74018 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0016 | 0.00215 | 83599 | [26 empty] | 2022-03-22 | 5744 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00206 | 83625 | [26 empty] | 2022-03-28 | 8813 | C13 Measurement | Completed
[168 empty] | 1.0015 | 0.00206 | 83627 | [26 empty] | 2022-03-28 | 23608 | Graphite | Completed
[168 empty] | 0.99826 | 0.00207 | 83628 | [26 empty] | 2022-04-08 | 24776 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00193 | 83629 | [26 empty] | 2022-04-20 | 29378 | Close Job | Completed
[7 empty] | 76.6 | 88.2603 | 0.0007 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44516 | 44516 | 44516 | 298 | 298 | 298 |  |  | jacob leath | -996.08 | 0.15 ‰ | 2022-04-06 | 2021-10-08 | 13 | 180 |  |  |  |  | -5.43‰ | -5.43‰ | 2.00‰ | 2.00‰ | -5.43 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -996.08 ‰ | -996.08 ‰ | 0.15 | 0.15 | -996.03 | 0.15 ‰ | -996.03 | 0.15 ‰ |  |  | 2022-04-08 | Jocelyn Turnbull | 10:59 | 0.00392 | 0.00392 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 13.9 |  |  | 0.00392 | 0.00015 | 24780 | 2022-04-07 | 14:43 | Albert Zondervan | 220783 |  |  |  |  | 45% | 5 | 6/05/22 | 13:49:26 | Cathy Ginnane | 0 |  | 0.39 | 0.01% | 29 | 5 | 4 | ... | 0.00377 | 0.00014 |  |  |  |  | 0.98641 | 0.98641 | Yes | Yes | [5 empty] | 0.00377 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00392 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00377 | 0 | 0.00014 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2022-04-05 | jacob leath | 11:30 |  | Completed |  |  |  |  | 83579 | 2000000 |  |  | 83511 83599 83625 83627 83628 83629 | 83511~83599~83625~83627~83628~83629 | 3397 |  | 3397 3397 3397 3397 3397 3397 |  |  |  |  | 0.15 mg |  | 0.15 | [5 empty] | 0.99673 | 0.00206 | 83511 |  | XCAMS | ST | 0% |  | TW3397 measurement * small+large carbonates CalAMS error adjustments: * factor 1.08 to account for global run-to-run variability primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 4x coral results agree with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 6x unknowns + 4x secondaries + 1x blanks last 180 days, capturing 4x blanks small carbonates: 14x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. CG updated blank corrections and checked wheel results 6/5/22 for smalls in this wheel A bit of relevant history around the time of slushie changeover: 4-Aug-2021: the last OxI graphitised (TP82203) for the last smalls-wheel (TW3362) before the changeover to liqN2-Eth slushies 5-Aug: Taylor wrote the SOP for prepping liqN2-Eth slushies 13-Aug: Taylor instructed some members of the team how to do make such slushies 13-Sep: Cathy did the first RP run using liqN2-Eth slushies mid-Aug -- mid-Sep: the use of dry-ice was abandoned and the dry-ice freezer emptied 25-Nov + 3-Dec: OxI graphitised for TW3383, the first wheel of smalls prepped after the slushie changeover |  | RRL | [5 empty] | Background Inorganic |  | 2022-05-17 | Carbonate | Complete | JOB-220783 |  | small blank SLUSHIE | 74019 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0016 | 0.00215 | 83599 | [26 empty] | 2022-03-22 | 5748 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00206 | 83625 | [26 empty] | 2022-03-24 | 23589 | Graphite | Completed
[168 empty] | 1.0015 | 0.00206 | 83627 | [26 empty] | 2022-04-08 | 24780 | AMS Submission Results | Completed
[168 empty] | 0.99826 | 0.00207 | 83628 | [26 empty] | 2022-04-20 | 29385 | Close Job | Completed
[168 empty] | 1.0013 | 0.00193 | 83629
[7 empty] | 90.2 | 88.0074 | 0.0004 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48211 | 48211 | 48211 | 338 | 337 | 337 |  |  | jacob leath | -997.53 | 0.10 ‰ | 2022-04-06 | 2021-10-08 | 7 | 180 |  |  |  |  | -2.57‰ | -2.57‰ | 2.00‰ | 2.00‰ | -2.57 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -997.53 ‰ | -997.53 ‰ | 0.1 | 0.1 | -997.49 | 0.11 ‰ | -997.49 | 0.11 ‰ |  |  | 2022-04-08 | Jocelyn Turnbull | 10:59 | 0.0024747 | 0.0024747 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 8.9 |  |  | 0.00247 | 0.0001 | 24782 | 2022-04-07 | 14:43 | Albert Zondervan | 220785 |  |  |  |  | 45% | 5 | 6/05/22 | 13:49:26 | Cathy Ginnane | 0 |  | 0.25 | 0.01% | 6 | 5 | 4 | ... | 0.00238 | 0.0001 |  |  |  |  | 0.98641 | 0.98641 | Yes | Yes | [5 empty] | 0.00238 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00247 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00238 | 0 | 0.0001 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2022-04-05 | jacob leath | 11:30 |  | Completed |  |  |  |  | 83621 | 2000000 |  |  | 83511 83599 83625 83627 83628 83629 | 83511~83599~83625~83627~83628~83629 | 3397 |  | 3397 3397 3397 3397 3397 3397 |  |  |  |  | 0.24 mg |  | 0.24 | [5 empty] | 0.99673 | 0.00206 | 83511 |  | XCAMS | ST | 0% |  | TW3397 measurement * small+large carbonates CalAMS error adjustments: * factor 1.08 to account for global run-to-run variability primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 4x coral results agree with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 6x unknowns + 4x secondaries + 1x blanks last 180 days, capturing 4x blanks small carbonates: 14x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. CG updated blank corrections and checked wheel results 6/5/22 for smalls in this wheel A bit of relevant history around the time of slushie changeover: 4-Aug-2021: the last OxI graphitised (TP82203) for the last smalls-wheel (TW3362) before the changeover to liqN2-Eth slushies 5-Aug: Taylor wrote the SOP for prepping liqN2-Eth slushies 13-Aug: Taylor instructed some members of the team how to do make such slushies 13-Sep: Cathy did the first RP run using liqN2-Eth slushies mid-Aug -- mid-Sep: the use of dry-ice was abandoned and the dry-ice freezer emptied 25-Nov + 3-Dec: OxI graphitised for TW3383, the first wheel of smalls prepped after the slushie changeover |  | RRL | [5 empty] | Background Inorganic |  | 2022-05-17 | Carbonate |  | JOB-220785 |  | SLUSHIE small to match samples- aim for 0.22-0.25mg | 74020 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0016 | 0.00215 | 83599 | [26 empty] | 2022-03-25 | 5759 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00206 | 83625 | [26 empty] | 2022-03-30 | 23631 | Graphite | Completed
[168 empty] | 1.0015 | 0.00206 | 83627 | [26 empty] | 2022-04-08 | 24782 | AMS Submission Results | Completed
[168 empty] | 0.99826 | 0.00207 | 83628 | [26 empty] | 2022-04-20 | 29386 | Close Job | Completed
[168 empty] | 1.0013 | 0.00193 | 83629
[7 empty] | 113.7 | 87.7338 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49028 | 49027 | 49027 | 336 | 336 | 336 |  |  | Taylor Ferrick | -997.76 | 0.09 ‰ | 2022-04-11 | 2021-10-13 | 3 | 180 |  |  |  |  | 1.09‰ | 1.09‰ | 2.00‰ | 2.00‰ | 1.09 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -997.76 ‰ | -997.76 ‰ | 0.09 | 0.09 | -997.74 | 0.09 ‰ | -997.74 | 0.09 ‰ |  |  | 2022-04-13 | Cathy Ginnane | 12:24 | 0.0022356 | 0.0022356 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.6 |  |  | 0.00224 | 9e-05 | 24858 | 2022-04-12 | 13:03 | Albert Zondervan | 220198 |  |  |  |  | 45% | 5 | 13/04/22 | 12:24:26 | Cathy Ginnane | 0 |  | 0.22 | 0.01% | 8 | 5 | 5 | ... | 0.00215 | 9e-05 |  |  |  |  | 0.98696 | 0.98696 | Yes | Yes | [5 empty] | 0.00215 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00224 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00215 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-08 | Taylor Ferrick | 13:08 |  | Completed |  |  |  |  | 83706 | 2000000 |  |  | 83666 83667 83692 83693 83712 83713 | 83666~83667~83692~83693~83712~83713 | 3399 |  | 3399 3399 3399 3399 3399 3399 | 2022-04-12 | 13:10 | Albert Zondervan |  | 0.76 mg |  | 0.76 | [5 empty] | 1.0014 | 0.0016 | 83666 |  | XCAMS | ST | 0% |  | TW3399 measurement CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large, accepted secondaries: * organics: none * carbonates: 2x corals agree with expectation value blanks: * 2 groups identified to blank-correct: organics: 10x unknowns + 0x secondaries + 1x blanks 180 days, capturing 13x blanks carbonates: 18x unknowns + 2x secondaries + 1x blanks 180 days, capturing 5x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-03-17 | Carbonate | Complete | JOB-220198 |  |  | 74058 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99903 | 0.00162 | 83667 | [26 empty] | 2022-04-06 | 5780 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00162 | 83692 | [26 empty] | 2022-04-08 | 23716 | Graphite | Completed
[168 empty] | 1.0002 | 0.00165 | 83693 | [26 empty] | 2022-04-11 | 8852 | C13 Measurement | Completed
[168 empty] | 0.99869 | 0.00159 | 83712 | [26 empty] | 2022-04-13 | 24858 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00159 | 83713 | [26 empty] | 2022-04-20 | 29379 | Close Job | Completed
[7 empty] | 63.5 | 88.728 | 0.0007 | RRL-CBIn-TY | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44081 | 44080 | 44080 | 262 | 262 | 262 |  |  | Taylor Ferrick | -995.86 | 0.14 ‰ | 2022-04-22 | 2021-10-24 | 2 | 180 |  |  |  |  | -8.91‰ | -8.91‰ | 2.00‰ | 2.00‰ | -8.91 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0008 | 0 | 0.8165 | -995.86 ‰ | -995.86 ‰ | 0.14 | 0.14 | -995.81 | 0.14 ‰ | -995.81 | 0.14 ‰ |  |  | 2022-05-09 | Cathy Ginnane | 11:05 | 0.0041384 | 0.0041384 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 17.4 |  |  | 0.00414 | 0.00014 | 24871 | 2022-04-27 | 16:18 | Albert Zondervan | 220788 |  |  |  |  | 45% | 5 | 9/05/22 | 11:05:11 | Cathy Ginnane | 0 |  | 0.41 | 0.01% | 6 | 5 | 4 | ... | 0.00398 | 0.00013 |  |  |  |  | 0.98817 | 0.98817 | Yes | Yes |  | d13C 11.5 per mil light in AMS (small) |  |  |  | 0.00398 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00414 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00398 | 0 | 0.00013 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2022-04-21 | Taylor Ferrick | 10:43 |  | Completed |  |  |  |  | 83805 | 2000000 |  |  | 83765 83766 83767 83768 83769 83770 | 83765~83766~83767~83768~83769~83770 | 3400 |  | 3400 3400 3400 3400 3400 3400 | 2022-05-06 | 12:14 | Cathy Ginnane |  | 0.13 mg |  | 0.13 | [5 empty] | 1.0003 | 0.00136 | 83765 |  | XCAMS |  |  |  | TW3400 measurement * small+large carbonates CalAMS error adjustments: * factors ~1.4 to account for run-to-run variability in blanks in positions cat#1,38 primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 1x coral result agrees with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 13x unknowns + 1x secondaries + 1x blanks last 180 days, capturing 6x blanks small carbonates: 10x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. |  | RRL | [5 empty] | Background Inorganic |  | 2022-05-18 | Carbonate |  | JOB-220788 |  | SLUSHIE 0.10mg | 74178 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99915 | 0.00133 | 83766 | [26 empty] | 2022-04-19 | 5811 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0008 | 0.00137 | 83767 | [26 empty] | 2022-04-21 | 23815 | Graphite | Completed
[168 empty] | 1.0015 | 0.00136 | 83768 | [26 empty] | 2022-05-09 | 24871 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00138 | 83769 | [26 empty] | 2022-05-09 | 29452 | Close Job | Completed
[168 empty] | 0.998 | 0.00137 | 83770
[7 empty] | 80.1 | 88.669 | 0.0006 | RRL-CBIn-SM | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45985 | 45984 | 45984 | 384 | 383 | 383 |  |  | Taylor Ferrick | -996.74 | 0.16 ‰ | 2022-04-22 | 2021-10-24 | 2 | 180 |  |  |  |  | -8.25‰ | -8.25‰ | 2.00‰ | 2.00‰ | -8.25 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0008 | 0 | 0.8165 | -996.74 ‰ | -996.74 ‰ | 0.16 | 0.16 | -996.69 | 0.16 ‰ | -996.69 | 0.16 ‰ |  |  | 2022-05-09 | Cathy Ginnane | 11:05 | 0.003265 | 0.003265 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 13 |  |  | 0.00326 | 0.00016 | 24872 | 2022-04-27 | 16:18 | Albert Zondervan | 220995 |  |  |  |  | 45% | 5 | 9/05/22 | 11:05:15 | Cathy Ginnane | 0 |  | 0.33 | 0.02% | 1 | 5 | 4 | ... | 0.00314 | 0.00015 |  |  |  |  | 0.98817 | 0.98817 | Yes | Yes |  | d13C 11 per mil light in AMS (small) |  |  |  | 0.00314 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00326 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00314 | 0 | 0.00015 | 0 | Small | Yes | 0.9878 | 0.95 | 1 | 2022-04-21 | Taylor Ferrick | 10:43 |  | Completed |  |  |  |  | 83807 | 2000000 |  |  | 83765 83766 83767 83768 83769 83770 | 83765~83766~83767~83768~83769~83770 | 3400 |  | 3400 3400 3400 3400 3400 3400 | 2022-05-06 | 12:14 | Cathy Ginnane |  | 0.17 mg |  | 0.17 | [5 empty] | 1.0003 | 0.00136 | 83765 |  | XCAMS |  |  |  | TW3400 measurement * small+large carbonates CalAMS error adjustments: * factors ~1.4 to account for run-to-run variability in blanks in positions cat#1,38 primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 1x coral result agrees with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 13x unknowns + 1x secondaries + 1x blanks last 180 days, capturing 6x blanks small carbonates: 10x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. |  | RRL | [5 empty] | Background Inorganic |  | 2022-06-14 | Carbonate |  | JOB-220995 |  | 0.15 mg | 74179 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99915 | 0.00133 | 83766 | [26 empty] | 2022-04-19 | 5814 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0008 | 0.00137 | 83767 | [26 empty] | 2022-04-21 | 23817 | Graphite | Completed
[168 empty] | 1.0015 | 0.00136 | 83768 | [26 empty] | 2022-05-09 | 24872 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00138 | 83769 | [26 empty] | 2022-05-09 | 29467 | Close Job | Completed
[168 empty] | 0.998 | 0.00137 | 83770
[7 empty] | 127.6 | 87.7997 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52648 | 52647 | 52647 | 293 | 293 | 293 |  |  | Taylor Ferrick | -998.58 | 0.05 ‰ | 2022-04-22 | 2021-10-24 | 2 | 180 |  |  |  |  | 1.57‰ | 1.57‰ | 2.00‰ | 2.00‰ | 1.57 ‰ | 2.00 ‰ | AMS | 2.45‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -998.58 ‰ | -998.58 ‰ | 0.05 | 0.05 | -998.56 | 0.05 ‰ | -998.56 | 0.05 ‰ |  |  | 2022-05-09 | Cathy Ginnane | 11:05 | 0.0014245 | 0.0014245 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 2.3 |  |  | 0.00142 | 5e-05 | 24873 | 2022-04-27 | 16:18 | Albert Zondervan | 220786 |  |  |  |  | 45% | 5 | 9/05/22 | 11:05:16 | Cathy Ginnane | 0 |  | 0.14 | 0.01% | 11 | 5 | 5 | ... | 0.00137 | 5e-05 |  |  |  |  | 0.98817 | 0.98817 | Yes | Yes |  | best blank in 6 months |  |  |  | 0.00137 | 0.00‰ | 5e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00142 | 0 | 5e-05 | 0 | 0 |  |  | 0 | 0 | 0.00137 | 0 | 5e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-21 | Taylor Ferrick | 10:43 |  | Completed |  |  |  |  | 83808 | 2000000 |  |  | 83765 83766 83767 83768 83769 83770 | 83765~83766~83767~83768~83769~83770 | 3400 |  | 3400 3400 3400 3400 3400 3400 | 2022-05-06 | 12:14 | Cathy Ginnane |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0003 | 0.00136 | 83765 |  | XCAMS |  |  |  | TW3400 measurement * small+large carbonates CalAMS error adjustments: * factors ~1.4 to account for run-to-run variability in blanks in positions cat#1,38 primaries: * large OxI cathodes: 6x accepted, used for standardization in CalAMS * small OxI cathodes: 3x accepted, used for 'small d' determination secondaries: large carbonates: 1x coral result agrees with expectation value small carbonates: 2x coral results agree with expectation value blanks: * 2 groups identified to blank-correct: large carbonates: 13x unknowns + 1x secondaries + 1x blanks last 180 days, capturing 6x blanks small carbonates: 10x unknowns + 2x secondaries + 2x blanks data from 5x wheels used: TW3383,3384,3397,3398,3400 m = 0.61 ug C from 6x blanks d = 0.28 ug C from 10x primary standards unknowns: * no failures On 2-May MN+CG+AZ decided to redo the blank corrections for the smalls, by taking into account only wheels TW3383,3384,3397,3398,3400. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-05-18 | Carbonate |  | JOB-220786 |  | reg sized | 74180 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99915 | 0.00133 | 83766 | [26 empty] | 2022-04-19 | 5809 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0008 | 0.00137 | 83767 | [26 empty] | 2022-04-21 | 23818 | Graphite | Completed
[168 empty] | 1.0015 | 0.00136 | 83768 | [26 empty] | 2022-04-22 | 8873 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00138 | 83769 | [26 empty] | 2022-05-09 | 24873 | AMS Submission Results | Completed
[168 empty] | 0.998 | 0.00137 | 83770 | [26 empty] | 2022-05-09 | 29451 | Close Job | Completed
[7 empty] | 114.7 | 87.6372 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46514 | 46513 | 46513 | 219 | 218 | 218 |  |  | Taylor Ferrick | -996.94 | 0.08 ‰ | 2022-05-02 | 2021-11-03 | 10 | 180 |  |  |  |  | 3.04‰ | 3.04‰ | 2.00‰ | 2.00‰ | 3.04 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.75593 | -996.94 ‰ | -996.94 ‰ | 0.08 | 0.08 | -996.91 | 0.08 ‰ | -996.91 | 0.08 ‰ |  |  | 2022-05-04 | Cathy Ginnane | 23:02 | 0.003057 | 0.003057 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00306 | 8e-05 | 24990 | 2022-05-03 | 16:31 | Albert Zondervan | 220885 |  |  |  |  | 45% | 5 | 4/05/22 | 23:02:19 | Cathy Ginnane | 0 |  | 0.31 | 0.01% | 1 | 5 | 5 | ... | 0.00294 | 8e-05 |  |  |  |  | 0.98779 | 0.98779 | Yes | Yes |  | higher blank than recent carbonate blanks |  |  |  | 0.00294 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00306 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00294 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-29 | Taylor Ferrick | 12:03 |  | Completed |  |  |  |  | 83847 | 2000000 |  |  | 83786 83831 83832 83840 83843 83844 83845 | 83786~83831~83832~83840~83843~83844~83845 | 3403 |  | 3403 3403 3403 3403 3403 3403 3403 | 2022-05-04 | 9:10 | Albert Zondervan |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99936 | 0.00154 | 83786 |  | XCAMS | ST | 0% |  | TW3403 measurement: * regular-size samples only * EA tests, GIRI-D, and usual suite of unknowns CalAMS error adjustments: * none needed primaries: * OxI cathodes: 1x EA + 6x ST all 7x accepted secondaries: * carbonates: 2x corals agree with expectation value * waters: 2x corals agree with expectation value * organics via CE: 1x FIRI-D agrees with AMS consensus value * IAEA-C6 for EA test: 2x IAEA-C6 agree with each other blanks: * 4 groups identified to blank-correct: carbonates: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 6x blanks waters: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks organics via CE: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks organics via AAA: 1x unknowns + 2x secondaries + 0x blanks 180 days, capturing 12x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-05-27 | Carbonate | Complete | JOB-220885 |  | regular size | 74099 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99956 | 0.0015 | 83831 | [26 empty] | 2022-04-21 | 5824 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00157 | 83832 | [26 empty] | 2022-04-22 | 23857 | Graphite | Completed
[168 empty] | 0.99867 | 0.00152 | 83840 | [26 empty] | 2022-04-22 | 8875 | C13 Measurement | Completed
[168 empty] | 0.9983 | 0.00154 | 83843 | [26 empty] | 2022-05-04 | 24990 | AMS Submission Results | Completed
[168 empty] | 1.0009 | 0.00157 | 83844 | [26 empty] | 2022-05-06 | 29413 | Close Job | Completed
[168 empty] | 1.0024 | 0.00155 | 83845
[7 empty] | 111.9 | 87.5215 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42152 | 42152 | 42152 | 175 | 174 | 174 |  |  | Taylor Ferrick | -994.74 | 0.11 ‰ | 2022-05-02 | 2021-11-03 | 4 | 180 |  |  |  |  | 4.36‰ | 4.36‰ | 2.00‰ | 2.00‰ | 4.36 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.75593 | -994.74 ‰ | -994.74 ‰ | 0.11 | 0.11 | -994.67 | 0.12 ‰ | -994.67 | 0.12 ‰ |  |  | 2022-05-04 | Cathy Ginnane | 23:03 | 0.0052614 | 0.0052614 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00526 | 0.00011 | 24991 | 2022-05-03 | 16:31 | Albert Zondervan | 219488 |  |  |  |  | 45% | 5 | 4/05/22 | 23:02:43 | Cathy Ginnane | 0 |  | 0.53 | 0.01% | 15 | 5 | 5 | ... | 0.00506 | 0.00011 |  |  |  |  | 0.98779 | 0.98779 | Yes | Yes | [5 empty] | 0.00506 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00526 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00506 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-29 | Taylor Ferrick | 12:03 |  | Completed |  |  |  |  | 83859 | 2000000 |  |  | 83786 83831 83832 83840 83843 83844 83845 | 83786~83831~83832~83840~83843~83844~83845 | 3403 |  | 3403 3403 3403 3403 3403 3403 3403 | 2022-05-04 | 9:10 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99936 | 0.00154 | 83786 |  | XCAMS | ST | 0% |  | TW3403 measurement: * regular-size samples only * EA tests, GIRI-D, and usual suite of unknowns CalAMS error adjustments: * none needed primaries: * OxI cathodes: 1x EA + 6x ST all 7x accepted secondaries: * carbonates: 2x corals agree with expectation value * waters: 2x corals agree with expectation value * organics via CE: 1x FIRI-D agrees with AMS consensus value * IAEA-C6 for EA test: 2x IAEA-C6 agree with each other blanks: * 4 groups identified to blank-correct: carbonates: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 6x blanks waters: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks organics via CE: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks organics via AAA: 1x unknowns + 2x secondaries + 0x blanks 180 days, capturing 12x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-12-29 | Water | Not Complete | JOB-219488 |  | job for new treatment of marble std for water line | 74109 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.0015 | 83831 | [26 empty] | 2022-04-27 | 5836 | Water CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00157 | 83832 | [26 empty] | 2022-04-28 | 23869 | Graphite | Completed
[168 empty] | 0.99867 | 0.00152 | 83840 | [26 empty] | 2022-04-29 | 8881 | C13 Measurement | Completed
[168 empty] | 0.9983 | 0.00154 | 83843 | [26 empty] | 2022-05-04 | 24991 | AMS Submission Results | Completed
[168 empty] | 1.0009 | 0.00157 | 83844 | [26 empty] | 2022-05-06 | 29427 | Close Job | Completed
[168 empty] | 1.0024 | 0.00155 | 83845
[7 empty] | 133.3 | 87.6495 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45883 | 45883 | 45883 | 227 | 227 | 227 |  |  | jacob leath | -996.69 | 0.09 ‰ | 2022-05-09 | 2021-11-10 | 3 | 180 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.0015 | 0 | 0.8165 | -996.69 ‰ | -996.69 ‰ | 0.09 | 0.09 | -996.65 | 0.09 ‰ | -996.65 | 0.09 ‰ |  |  | 2022-05-12 | Cathy Ginnane | 10:02 | 0.0033065 | 0.0033065 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.9 |  |  | 0.00331 | 9e-05 | 25102 | 2022-05-11 | 9:35 | Albert Zondervan | 221031 |  |  |  |  | 45% | 5 | 12/05/22 | 10:02:00 | Cathy Ginnane | 0 |  | 0.33 | 0.01% | 27 | 5 | 5 | ... | 0.00318 | 9e-05 |  |  |  |  | 0.98778 | 0.98778 | Yes | Yes | [5 empty] | 0.00318 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00331 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00318 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-05-06 | jacob leath | 13:06 |  | Completed |  |  |  |  | 83954 | 2000000 |  |  | 83921 83930 83931 83932 83933 83934 | 83921~83930~83931~83932~83933~83934 | 3405 |  | 3405 3405 3405 3405 3405 3405 | 2022-05-11 | 11:05 | Albert Zondervan |  | 1.03 mg |  | 1.03 | [5 empty] | 0.99951 | 0.00134 | 83921 |  | XCAMS | ST | 0% |  | TW3405 measurement: * large waters + few RP samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value blanks: * 2 groups identified to blank-correct: waters: 26x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks sediments via RP: to be blank-corrected still unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-06-23 | Water |  | JOB-221031 |  |  | 74265 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99765 | 0.00144 | 83930 | [26 empty] | 2022-05-05 | 5871 | Water CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00144 | 83931 | [26 empty] | 2022-05-06 | 23964 | Graphite | Completed
[168 empty] | 1.0008 | 0.00139 | 83932 | [26 empty] | 2022-05-09 | 8967 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00136 | 83933 | [26 empty] | 2022-05-12 | 25102 | AMS Submission Results | Completed
[168 empty] | 0.99993 | 0.00142 | 83934 | [26 empty] | 2022-05-18 | 29581 | Close Job | Completed
[7 empty] | 117.8 | 87.7161 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50961 | 50961 | 50961 | 285 | 285 | 285 |  |  | Taylor Ferrick | -998.24 | 0.06 ‰ | 2022-05-13 | 2021-11-14 | 2 | 180 |  |  |  |  | 4.13‰ | 4.13‰ | 2.00‰ | 2.00‰ | 4.13 ‰ | 2.00 ‰ | AMS | 2.45‰ | 0.2‰ | C13 | No | -19 | -19.0037 | 0 | 0.8165 | -998.24 ‰ | -998.24 ‰ | 0.06 | 0.06 | -998.22 | 0.06 ‰ | -998.22 | 0.06 ‰ |  |  | 2022-05-17 | Cathy Ginnane | 13:39 | 0.0017573 | 0.0017573 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.1 |  |  | 0.00176 | 6e-05 | 25109 | 2022-05-16 | 9:08 | Albert Zondervan | 221019 |  |  |  |  | 45% | 5 | 17/05/22 | 13:38:36 | Cathy Ginnane | 0 |  | 0.18 | 0.01% | 1 | 5 | 5 | ... | 0.00169 | 6e-05 |  |  |  |  | 0.98976 | 0.98976 | Yes | Yes | [5 empty] | 0.00169 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00176 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00169 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-05-12 | Taylor Ferrick | 12:49 |  | Completed |  |  |  |  | 83982 | 2000000 |  |  | 83935 83936 83947 83948 83949 83950 | 83935~83936~83947~83948~83949~83950 | 3406 |  | 3406 3406 3406 3406 3406 3406 | 2022-05-16 | 9:23 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0014 | 0.00164 | 83935 |  | XCAMS | ST | 0% |  | TW3406 measurement: * regular-size samples * Cs-delivery poorly controlled, but no visible effect on carbon-isotope ratios CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * carbonates: 1x coral result agrees with expectation value 1x coral result disagrees with expectation value by 2-sigma * organics via AAA: none * organics via CE: 1x FIRI-D agrees with consensus value blanks: * 3 groups identified to blank-correct: carbonates: 11x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks organics via AAA: 16x unknowns + 0x secondaries + 1x blanks 180 days, capturing 11x blanks organics via CE: 0x unknowns + 1x secondaries + 0x blanks 180 days, capturing 6x blanks unknowns: * no failures+ | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-06-17 | Carbonate | Complete | JOB-221019 |  |  | 74295 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99901 | 0.00159 | 83936 | [26 empty] | 2022-05-10 | 5885 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.9998 | 0.00154 | 83947 | [26 empty] | 2022-05-11 | 23992 | Graphite | Completed
[168 empty] | 0.99868 | 0.00156 | 83948 | [26 empty] | 2022-05-11 | 8980 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00158 | 83949 | [26 empty] | 2022-05-17 | 25109 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00165 | 83950 | [26 empty] | 2022-05-18 | 29576 | Close Job | Completed
[7 empty] | 99.7 | 88.2946 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51401 | 51400 | 51400 | 301 | 301 | 301 |  |  | Taylor Ferrick | -998.34 | 0.06 ‰ | 2022-05-20 | 2021-11-21 | 2 | 180 |  |  |  |  | 1.85‰ | 1.85‰ | 2.00‰ | 2.00‰ | 1.85 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0 | 0.8165 | -998.34 ‰ | -998.34 ‰ | 0.06 | 0.06 | -998.32 | 0.06 ‰ | -998.32 | 0.06 ‰ |  |  | 2022-05-23 | Jocelyn Turnbull | 14:31 | 0.0016637 | 0.0016637 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2 |  |  | 0.00166 | 6e-05 | 25149 | 2022-05-23 | 11:23 | Albert Zondervan | 220886 |  |  |  |  | 45% | 5 | 23/05/22 | 14:31:27 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 1 | 5 | 5 | ... | 0.0016 | 6e-05 |  |  |  |  | 0.99402 | 0.99402 | Yes | Yes | [5 empty] | 0.0016 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00166 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0016 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-05-19 | Taylor Ferrick | 13:08 |  | Completed |  |  |  |  | 84036 | 2000000 |  |  | 84031 84033 84034 84035 84042 84043 | 84031~84033~84034~84035~84042~84043 | 3407 |  | 3407 3407 3407 3407 3407 3407 | 2022-05-23 | 12:38 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99716 | 0.00144 | 84031 |  | XCAMS | ST | 0% |  | TW3407 measurement: * regular-size samples * Cs-delivery poorly controlled, but good ratio data regardless CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * carbonates: 2x coral results agree with expectation value * organics via AAA: 1x IAEA-C6 agrees with consensus value * organics via CE: 1x FIRI-D agrees with consensus value blanks: * 3 groups identified to blank-correct: carbonates: 23x unknowns + 2x secondaries + 1x blanks 180 days, capturing 8x blanks organics via AAA: 4x unknowns + 1x secondaries + 1x blanks 180 days, capturing 12x blanks organics via CE: 1x unknowns + 1x secondaries + 0x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-05-27 | Carbonate |  | JOB-220886 |  | regular size | 74334 |  | 14047/1 |  |  | 2021-12-01 | 7532 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.9995 | 0.00143 | 84033 | [26 empty] | 2022-05-17 | 5911 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0003 | 0.00143 | 84034 | [26 empty] | 2022-05-18 | 24046 | Graphite | Completed
[168 empty] | 1.0004 | 0.00142 | 84035 | [26 empty] | 2022-05-19 | 9013 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00166 | 84042 | [26 empty] | 2022-05-23 | 25149 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00152 | 84043 | [26 empty] | 2022-05-26 | 29735 | Close Job | Completed
[7 empty] | 146.7 | 87.3751 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40918 | 40918 | 40918 | 123 | 122 | 122 |  |  | jacob leath | -993.87 | 0.09 ‰ | 2022-06-09 | 2021-12-11 | 13 | 180 |  |  |  |  | 6.06‰ | 6.06‰ | 2.00‰ | 2.00‰ | 6.06 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.82 | 0.8165 | -993.87 ‰ | -993.87 ‰ | 0.09 | 0.09 | -993.79 | 0.09 ‰ | -993.79 | 0.09 ‰ |  |  | 2022-06-16 | Jocelyn Turnbull | 9:45 | 0.0061348 | 0.0061348 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00613 | 9e-05 | 25270 | 2022-06-15 | 19:42 | Albert Zondervan | 221035 |  |  |  |  | 45% | 5 | 16/06/22 | 09:45:01 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 1 | 5 | 7 | ... | 0.0059 | 9e-05 |  |  |  |  | 0.9878 | 0.9878 | Yes | Yes | [5 empty] | 0.0059 | 0.01‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00055 | 0.00055 | 0.00613 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0059 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-06-08 | jacob leath | 11:59 |  | Completed |  |  |  |  | 84143 | 2000000 |  |  | 84044 84045 84195 84196 84197 84198 | 84044~84045~84195~84196~84197~84198 | 3410 |  | 3410 3410 3410 3410 3410 3410 | 2022-06-15 | 19:45 | Albert Zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99934 | 0.00126 | 84044 |  | XCAMS | ST | 0% |  | TW3410 measurement: * large waters * unexpected inter- and intra-cathode variability: only in beam current for unclear reason, probably related to SNICS condition no additional variability in ratio-to-standard * (del13C_AMS - del13C_IRMS) difference correlates with beam current * SNICS is being tested before next batch of sample graphites is measured CalAMS error adjustments: * none needed: good inter-OxI repeatability, reasonable intra-cathode (run-to-run) repeatability * data from first two rotations ignored, to cull most runs with low current primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value 1x LAC1 coral agrees with expectation value blanks: * 1 groups identified to blank-correct: waters: 28x unknowns + 2x secondaries + 1x blanks 180 days, capturing 4x blanks unknowns: * cat#18 (TP84174) 1st Q-flag set: failed in loading, not measured * cat#16 (TP84188) 2nd Q-flag set: foreign object near surface | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-06-23 | Water |  | JOB-221035 |  |  | 74471 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9992 | 0.00127 | 84045 | [26 empty] | 2022-05-26 | 5933 | Water CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00145 | 84195 | [26 empty] | 2022-05-26 | 3343 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00146 | 84196 | [26 empty] | 2022-05-26 | 5584 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00128 | 84197 | [26 empty] | 2022-05-27 | 24153 | Graphite | Completed
[168 empty] | 1.0016 | 0.00146 | 84198 | [26 empty] | 2022-05-27 | 9043 | C13 Measurement | Completed
[197 empty] | 2022-06-16 | 25270 | AMS Submission Results | Completed
[197 empty] | 2022-06-23 | 29822 | Close Job | Completed
[7 empty] | 113.2 | 88.0196 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38842 | 38842 | 38842 | 147 | 147 | 147 |  |  | Taylor Ferrick | -992.06 | 0.15 ‰ | 2022-06-20 | 2021-12-22 | 11 | 180 |  |  |  |  | 4.19‰ | 4.19‰ | 2.00‰ | 2.00‰ | 4.19 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0.82 | 0.8165 | -992.06 ‰ | -992.06 ‰ | 0.15 | 0.15 | -991.96 | 0.15 ‰ | -991.96 | 0.15 ‰ |  |  | 2022-06-21 | Jocelyn Turnbull | 17:01 | 0.007944 | 0.007944 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00794 | 0.00015 | 25304 | 2022-06-21 | 16:08 | Albert Zondervan | 221105 |  |  |  |  | 45% | 5 | 21/06/22 | 17:01:07 | Jocelyn Turnbull | 0 |  | 0.79 | 0.01% | 1 | 5 | 5 | ... | 0.00764 | 0.00014 |  |  |  |  | 0.99324 | 0.99324 | Yes | Yes | [5 empty] | 0.00764 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00794 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00764 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-06-17 | Taylor Ferrick | 13:55 |  | Completed |  |  |  |  | 84216 | 2000000 |  |  | 84225 84226 84227 84228 84229 84239 | 84225~84226~84227~84228~84229~84239 | 3411 |  | 3411 3411 3411 3411 3411 3411 | 2022-06-21 | 16:17 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99919 | 0.00162 | 84225 |  | XCAMS | ST | 0% |  | TW3411 measurement: * large waters + large organics via AAA * Cs delivery tube of SNICS has been replaced and position of bandheater on Cs reservoir has been raised * although SNICS better under control than during TW3410, still some undesirable inter-cathode variability: only in beam current no additional variability in ratio-to-standard * (del13C_AMS - del13C_IRMS) difference correlates with beam current * MAN (high-prec) will be run after REP1 (reg-prec) CalAMS error adjustments: * error multiplier of 1.05 to account for (small) excess run-to-run variability primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value 1x LAC1 coral agrees with expectation value * organics 1x IAEA-C6 sucrose agrees with expectation value blanks: * 1 groups identified to blank-correct: waters: 15x unknowns + 2x secondaries + 1x blanks 180 days, capturing 4x blanks organics: 13x unknowns + 1x secondaries + 1x blanks 180 days, capturing 4x blanks unknowns: * all accepted | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-07-01 | Water | Not Complete | JOB-221105 |  | Untested (new) water vessel used. | 74510 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.00163 | 84226 | [26 empty] | 2022-06-08 | 5967 | Water CO2 Evolution | Completed
[168 empty] | 0.99811 | 0.00163 | 84227 | [26 empty] | 2022-06-09 | 24226 | Graphite | Completed
[168 empty] | 0.99967 | 0.00165 | 84228 | [26 empty] | 2022-06-10 | 9082 | C13 Measurement | Combined
[168 empty] | 1.0006 | 0.00166 | 84229 | [26 empty] | 2022-06-21 | 25304 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00156 | 84239 | [26 empty] | 2022-06-23 | 29826 | Close Job | Completed
[7 empty] | 57.9 | 88.2774 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47946 | 47945 | 47945 | 261 | 261 | 261 |  |  | Taylor Ferrick | -997.44 | 0.08 ‰ | 2022-07-15 | 2022-01-16 | 2 | 180 |  |  |  |  | 4.35‰ | 4.35‰ | 2.00‰ | 2.00‰ | 4.35 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0 | 0.8165 | -997.44 ‰ | -997.44 ‰ | 0.08 | 0.08 | -997.41 | 0.08 ‰ | -997.41 | 0.08 ‰ |  |  | 2022-07-20 | Jocelyn Turnbull | 10:02 | 0.0025579 | 0.0025579 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00256 | 8e-05 | 25489 | 2022-07-18 | 14:42 | Christian Lewis | 221628 |  |  |  |  | 45% | 5 | 20/07/22 | 10:02:10 | Jocelyn Turnbull | 0 |  | 0.26 | 0.01% | 22 | 5 | 5 | ... | 0.00246 | 8e-05 |  |  |  |  | 0.99631 | 0.99631 | Yes | Yes | [5 empty] | 0.00246 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00256 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00246 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-14 | Taylor Ferrick | 12:42 |  | Completed |  |  |  |  | 84472 | 2000000 |  |  | 84462 84463 84487 84488 84489 84490 | 84462~84463~84487~84488~84489~84490 | 3416 |  | 3416 3416 3416 3416 3416 3416 | 2022-07-20 | 9:40 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 1.0001 | 0.00148 | 84462 |  | XCAMS | ST | 0% |  | TW3416 measurement: * large CalAMS error adjustments: * none needed Primaries: * OxI cathodes: 6x ST, accepted Secondaries: * organics via CE 1x FIRI-D agrees with expectation value * Waters 1x LAC1 coral agrees with expectation value 1x LAA1 coral agrees with expectation value Blanks: * 3 groups identified to blank-correct: organics via AAA: 2x unknowns + 0x secondaries + 0x blanks 180 days, capturing 12x blanks organics via CE: 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Waters CO2 extraction 17x unknowns + 2x secondaries + 1x blanks 180 days, capturing 6x blanks Unknowns: * all accepted | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2022-08-22 | Water |  | JOB-221628 |  |  | 74735 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00144 | 84463 | [26 empty] | 2022-07-12 | 6000 | Water CO2 Evolution | Completed
[168 empty] | 1.0014 | 0.00161 | 84487 | [26 empty] | 2022-07-13 | 24482 | Graphite | Completed
[168 empty] | 0.99928 | 0.00158 | 84488 | [26 empty] | 2022-07-15 | 9244 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00154 | 84489 | [26 empty] | 2022-07-20 | 25489 | AMS Submission Results | Completed
[168 empty] | 0.99829 | 0.00154 | 84490 | [26 empty] | 2022-07-21 | 30268 | Close Job | Completed
[7 empty] | 102.9 | 87.7847 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50148 | 50148 | 50148 | 344 | 343 | 343 |  |  | jacob leath | -998.06 | 0.08 ‰ | 2022-07-25 | 2022-01-26 | 7 | 180 |  |  |  |  | 4.98‰ | 4.98‰ | 2.00‰ | 2.00‰ | 4.98 ‰ | 2.00 ‰ | AMS | 2.77‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.8165 | -998.06 ‰ | -998.06 ‰ | 0.08 | 0.08 | -998.03 | 0.08 ‰ | -998.03 | 0.08 ‰ |  |  | 2022-07-27 | Jocelyn Turnbull | 9:27 | 0.0019444 | 0.0019444 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00194 | 8e-05 | 25600 | 2022-07-26 | 16:14 | Albert Zondervan | 221642 |  |  |  |  | 45% | 5 | 27/07/22 | 09:27:16 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 22 | 5 | 5 | ... | 0.00187 | 8e-05 |  |  |  |  | 0.99137 | 0.99137 | Yes | Yes | [5 empty] | 0.00187 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00194 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00187 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-21 | jacob leath | 12:39 |  | Completed |  |  |  |  | 84498 | 2000000 |  |  | 84535 84536 84547 84559 84560 84561 | 84535~84536~84547~84559~84560~84561 | 3418 |  | 3418 3418 3418 3418 3418 3418 | 2022-07-26 | 16:43 | Albert Zondervan |  | 0.93 mg |  | 0.93 | [5 empty] | 0.9995 | 0.0017 | 84535 |  | XCAMS | ST | 0% |  | TW3418 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * carbonates: 2x coral results agree with expectation value * organics via AAA: none * organics via CE: 1x FIRI-D agrees with consensus value blanks: * 3 groups identified to blank-correct: carbonates: 13x unknowns + 2x secondaries + 1x blanks 180 days, capturing 8x blanks organics via AAA: 14x unknowns + 0x secondaries + 0x blanks 180 days, capturing 13x blanks organics via CE: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 10x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2022-08-22 | Carbonate | Not Complete | JOB-221642 |  |  | 74815 |  | 14047/1 |  |  | 2022-05-23 | 8236 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 1.0005 | 0.00168 | 84536 | [26 empty] | 2022-07-15 | 6013 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0028 | 0.00162 | 84547 | [26 empty] | 2022-07-18 | 24508 | Graphite | Completed
[168 empty] | 0.99909 | 0.00159 | 84559 | [26 empty] | 2022-07-19 | 9268 | C13 Measurement | Completed
[168 empty] | 0.99756 | 0.00166 | 84560 | [26 empty] | 2022-07-27 | 25600 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00154 | 84561 | [26 empty] | 2022-07-28 | 30311 | Close Job | Completed
[7 empty] | 94 | 88.107 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45229 | 45228 | 45228 | 233 | 232 | 232 |  |  | Taylor Ferrick | -996.41 | 0.10 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.06‰ | 2.06‰ | 2.00‰ | 2.00‰ | 2.06 ‰ | 2.00 ‰ | AMS | 2.84‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.41 ‰ | -996.41 ‰ | 0.1 | 0.1 | -996.37 | 0.11 ‰ | -996.37 | 0.11 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0035873 | 0.0035873 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00359 | 0.0001 | 25689 | 2022-08-05 | 10:56 | Christian Lewis | 221951 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:44 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 8 | 5 | 5 | ... | 0.00345 | 0.0001 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00345 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00359 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00345 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84630 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2022-09-14 | Water |  | JOB-221951 |  | Normal water blank | 74938 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-27 | 6046 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-28 | 24640 | Graphite | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-29 | 9352 | C13 Measurement | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-08-05 | 25689 | AMS Submission Results | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 30460 | Close Job | Completed
[7 empty] | 87.3 | 88.0659 | 0.0005 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 46651 | 46651 | 46651 | 278 | 277 | 277 | 46378.2 | 46934.3 | Taylor Ferrick | -996.99 | 0.10 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.53‰ | 2.53‰ | 2.00‰ | 2.00‰ | 2.53 ‰ | 2.00 ‰ | AMS | 2.87‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.99 ‰ | -996.99 ‰ | 0.1 | 0.1 | -996.96 | 0.11 ‰ | -996.96 | 0.11 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.003005 | 0.003005 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.5 |  |  | 0.00301 | 0.0001 | 25690 | 2022-08-05 | 10:56 | Christian Lewis | 221627 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:32 | Jocelyn Turnbull | 0 |  | 0.3 | 0.01% | 1 | 5 | 6 | ... | 0.00289 | 0.0001 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00289 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00301 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00289 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84629 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.84 mg |  | 0.84 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.9 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Carbonate |  | JOB-221627 |  | 200.26 mg CuO. RECOMBUST! test of recombusted carbonate blank | 74935 |  | 14047/1 |  |  | 2022-05-23 | 8236 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-22 | 6034 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5417 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24639 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9351 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25690 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30449 | Close Job | Completed
[198 empty] | 13982 | Invoice | Started
[7 empty] | 93.8 | 88.0143 | 0.0003 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 50322 | 50322 | 50322 | 307 | 307 | 307 | 50020.5 | 50635.4 | Taylor Ferrick | -998.1 | 0.07 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 3.11‰ | 3.11‰ | 2.00‰ | 2.00‰ | 3.11 ‰ | 2.00 ‰ | AMS | 2.90‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -998.10 ‰ | -998.10 ‰ | 0.07 | 0.07 | -998.07 | 0.07 ‰ | -998.07 | 0.07 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0019028 | 0.0019028 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.4 |  |  | 0.0019 | 7e-05 | 25691 | 2022-08-05 | 10:56 | Christian Lewis | 221626 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:39 | Jocelyn Turnbull | 0 |  | 0.19 | 0.01% | 22 | 5 | 6 | ... | 0.00183 | 7e-05 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00183 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.0019 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00183 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84628 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.9 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Carbonate |  | JOB-221626 |  | 197.27 mg CuO. RECOMBUST! test of recombusted carbonate blank | 74936 |  | 14047/1 |  |  | 2022-05-23 | 8236 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-22 | 6033 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5416 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24638 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9350 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25691 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30448 | Close Job | Completed
[198 empty] | 13981 | Invoice | Started
[7 empty] | 95.5 | 88.0351 | 0.0003 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 49978 | 49978 | 49978 | 294 | 294 | 294 | 49689.2 | 50278.3 | Taylor Ferrick | -998.01 | 0.07 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.88‰ | 2.88‰ | 2.00‰ | 2.00‰ | 2.88 ‰ | 2.00 ‰ | AMS | 2.78‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -998.01 ‰ | -998.01 ‰ | 0.07 | 0.07 | -997.99 | 0.07 ‰ | -997.99 | 0.07 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.001986 | 0.001986 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2.2 |  |  | 0.00199 | 7e-05 | 25692 | 2022-08-05 | 10:56 | Christian Lewis | 221248 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:40 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 11 | 5 | 6 | ... | 0.00191 | 7e-05 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00191 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00199 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00191 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84624 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-07-13 | Carbonate |  | JOB-221248 |  | 198.44 mg CuO. RECOMBUST! Test of recombusted carbonate blank- Refreshed (Parent Job) material please make more jobs for future wheels | 74937 |  | 14047/1 |  |  | 2022-05-23 | 8236 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-22 | 6032 | Carbonate CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5415 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24634 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9341 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25692 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30426 | Close Job | Completed
[198 empty] | 13977 | Invoice | Started
[7 empty] | 94.6 | 88.0729 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40190 | 40189 | 40189 | 174 | 174 | 174 | 40017.7 | 40365.9 | Taylor Ferrick | -993.28 | 0.15 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.76‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.28 ‰ | -993.28 ‰ | 0.15 | 0.15 | -993.2 | 0.15 ‰ | -993.2 | 0.15 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0067171 | 0.0067171 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00672 | 0.00015 | 25693 | 2022-08-05 | 10:56 | Christian Lewis | 221623 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:45 | Jocelyn Turnbull | 0 |  | 0.67 | 0.01% | 15 | 5 | 6 | ... | 0.00646 | 0.00014 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00646 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00672 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00646 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84625 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water |  | JOB-221623 |  | RECOMBUST!- Test of recombusted water blank MN not completed ST process to make additional jobs for future wheels | 74939 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-12 | 6003 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5297 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24635 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9347 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25693 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30445 | Close Job | Completed
[198 empty] | 13978 | Invoice | Started
[7 empty] | 96.5 | 88.0947 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44327 | 44326 | 44326 | 229 | 228 | 228 | 44100.9 | 44558.9 | Taylor Ferrick | -995.99 | 0.11 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.20‰ | 2.20‰ | 2.00‰ | 2.00‰ | 2.20 ‰ | 2.00 ‰ | AMS | 2.79‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -995.99 ‰ | -995.99 ‰ | 0.11 | 0.11 | -995.94 | 0.12 ‰ | -995.94 | 0.12 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0040136 | 0.0040136 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00401 | 0.00011 | 25694 | 2022-08-05 | 10:56 | Christian Lewis | 221625 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:48 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 25 | 5 | 6 | ... | 0.00386 | 0.00011 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00386 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00401 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00386 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84627 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water |  | JOB-221625 |  | RECOMBUST!- Test of recombusted water blank | 74940 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-12 | 6005 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5299 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24637 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9349 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25694 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30447 | Close Job | Completed
[198 empty] | 13980 | Invoice | Started
[7 empty] | 97.8 | 88.0621 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 202 | 202 | 202 | 42426.6 | 42830.9 | Taylor Ferrick | -995.04 | 0.12 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.57‰ | 2.57‰ | 2.00‰ | 2.00‰ | 2.57 ‰ | 2.00 ‰ | AMS | 2.81‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.12 | 0.12 | -994.98 | 0.13 ‰ | -994.98 | 0.13 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0049598 | 0.0049598 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00496 | 0.00012 | 25695 | 2022-08-05 | 10:56 | Christian Lewis | 221624 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:49 | Jocelyn Turnbull | 0 |  | 0.5 | 0.01% | 34 | 5 | 6 | ... | 0.00477 | 0.00012 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00496 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84626 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water | Complete | JOB-221624 |  | RECOMBUST!- Test of recombusted water blank | 74941 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [26 empty] | 2022-07-12 | 6004 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [26 empty] | 2022-07-27 | 5298 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [26 empty] | 2022-07-28 | 24636 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [26 empty] | 2022-07-29 | 9348 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [26 empty] | 2022-08-05 | 25695 | AMS Submission Results | Completed
[197 empty] | 2022-08-05 | 30446 | Close Job | Completed
[198 empty] | 13979 | Invoice | Started
[7 empty] | 138.2 | 87.8261 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40341 | 40340 | 40340 | 152 | 152 | 152 |  |  | Taylor Ferrick | -993.41 | 0.12 ‰ | 2022-08-16 | 2022-02-17 | 13 | 180 |  |  |  |  | 2.35‰ | 2.35‰ | 2.00‰ | 2.00‰ | 2.35 ‰ | 2.00 ‰ | AMS | 3.20‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -993.41 ‰ | -993.41 ‰ | 0.12 | 0.12 | -993.33 | 0.13 ‰ | -993.33 | 0.13 ‰ |  |  | 2022-08-18 | Cathy Ginnane | 13:32 | 0.0065923 | 0.0065923 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00659 | 0.00012 | 25803 | 2022-08-17 | 10:13 | Christian Lewis | 221896 |  |  |  |  | 45% | 5 | 18/08/22 | 13:31:53 | Cathy Ginnane | 0 |  | 0.66 | 0.01% | 1 | 5 | 5 | ... | 0.00634 | 0.00012 |  |  |  |  | 0.98924 | 0.98924 | Yes | Yes | [5 empty] | 0.00634 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00659 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00634 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-08-15 | Taylor Ferrick | 8:17 |  | Completed |  |  |  |  | 84676 | 2000000 |  |  | 84735 84736 84748 84752 84753 84754 | 84735~84736~84748~84752~84753~84754 | 3424 |  | 3424 3424 3424 3424 3424 3424 | 2022-08-17 | 10:49 | Christian Lewis |  | 0.81 mg |  | 0.81 | [5 empty] | 1.0014 | 0.00149 | 84735 |  | XCAMS | ST | 0% |  | TW3424 measurement: * large CalAMS error adjustments: * none needed Primaries: * OxI cathodes: 6x ST, accepted Secondaries: * organics via CE 1x FIRI-D agrees with expectation value * Waters 1x LAC1 coral agrees with expectation value 1x LAA1 coral agrees with expectation value Blanks: * 3 groups identified to blank-correct: organics via AAA: 7x unknowns + 0x secondaries + 1x blanks 180 days, capturing 16x blanks organics via CE: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 10x blanks (this correction has been updated from 11 to 10 blanks, to eliminate TP83997 (TW3408) which had a quality flag to mark it as an outlier; CG; 23/8/22) Waters CO2 extraction 10x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks Unknowns: * all accepted | 3.2 | RRL | [5 empty] | Background Inorganic |  | 2022-09-13 | Water | Complete | JOB-221896 |  | C13 0.8 per mil out | 75023 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99971 | 0.00151 | 84736 | [26 empty] | 2022-08-02 | 6063 | Water CO2 Evolution | Completed
[168 empty] | 0.9999 | 0.00149 | 84748 | [26 empty] | 2022-08-03 | 24686 | Graphite | Completed
[168 empty] | 1.002 | 0.00149 | 84752 | [26 empty] | 2022-08-05 | 9393 | C13 Measurement | Completed
[168 empty] | 0.99879 | 0.00151 | 84753 | [26 empty] | 2022-08-18 | 25803 | AMS Submission Results | Completed
[168 empty] | 0.99821 | 0.00148 | 84754 | [28 empty] | Close Job
[7 empty] | 118.5 | 87.7108 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40328 | 40327 | 40327 | 114 | 113 | 113 |  |  | Taylor Ferrick | -993.4 | 0.09 ‰ | 2022-08-26 | 2022-02-27 | 3 | 180 |  |  |  |  | 5.07‰ | 5.07‰ | 2.00‰ | 2.00‰ | 5.07 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -993.40 ‰ | -993.40 ‰ | 0.09 | 0.09 | -993.32 | 0.09 ‰ | -993.32 | 0.09 ‰ |  |  | 2022-08-30 | Jocelyn Turnbull | 17:07 | 0.0066027 | 0.0066027 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.0066 | 9e-05 | 25963 | 2022-08-29 | 11:03 | Christian Lewis | 221989 |  |  |  |  | 45% | 5 | 30/08/22 | 17:07:09 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 31 | 5 | 5 | ... | 0.00635 | 9e-05 |  |  |  |  | 0.99062 | 0.99062 | Yes | Yes | [5 empty] | 0.00635 | 0.01‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.0066 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00635 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-08-25 | Taylor Ferrick | 12:48 |  | Completed |  |  |  |  | 84897 | 2000000 |  |  | 84694 84695 84709 84717 84718 84908 84920 | 84694~84695~84709~84717~84718~84908~84920 | 3428 |  | 3428 3428 3428 3428 3428 3428 3428 | 2022-08-29 | 12:00 | Christian Lewis |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0015 | 0.0013 | 84694 |  | XCAMS | FLASK | 0% |  | TW3428 measurement: Water samples Air samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x ST All accepted; however, one is very low. TP 84709 has RTS of 0.99591, which is on the cusp of 3-sigma for this set of standards. secondaries: * Inorganics via water evolution: Both agree with expectation values blanks: * 1 groups identified to blank-correct: Inorganics via CO2 evolution on waterline unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-09-22 | Water |  | JOB-221989 |  | New job with refreshed IAEA-C1 material for waterline blanks. | 75173 | Normal | 14047/11 |  |  | 2022-08-03 | 8437 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99988 | 0.00128 | 84695 | [26 empty] | 2022-08-22 | 6078 | Water CO2 Evolution | Completed
[168 empty] | 0.99591 | 0.00123 | 84709 | [26 empty] | 2022-08-23 | 24907 | Graphite | Completed
[168 empty] | 1.0014 | 0.00129 | 84717 | [26 empty] | 2022-08-26 | 9489 | C13 Measurement | Completed
[168 empty] | 0.99936 | 0.00127 | 84718 | [26 empty] | 2022-08-30 | 25963 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00126 | 84908 | [28 empty] | Close Job
[168 empty] | 0.99947 | 0.00124 | 84920
[10 empty] | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | [11 empty] | Taylor Ferrick |  |  |  | ? |  | 180 | [8 empty] | 2.44 ‰ | 0.20 ‰ | C13 | 2.44‰ | 0.2‰ | C13 | [5 empty] | -1000.00 ‰ |  |  |  | -1000 | [6 empty] | Taylor Ferrick | [7 empty] | 2.7 |  |  |  |  | 26000 |  |  |  | 221895 |  |  |  |  | 45% |  | 30/08/22 | 11:15:17 | Taylor Ferrick | [5 empty] | 5 | 5 | ... | [9 empty] | Yes | [21 empty] | 0 | 0 |  |  | 0 |  | Regular | Yes |  |  |  | 2022-08-30 | Taylor Ferrick | 11:09 |  | Started |  |  |  |  | 84945 | 2000000 |  |  |  |  | 3429 | [6 empty] | 0.78 mg |  | 0.78 | [9 empty] | XCAMS |  |  |  |  | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-09-13 | Carbonate | Complete | JOB-221895 |  |  |  |  | 14047/1 |  |  | 2022-05-23 | 8236 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/1 for CO2 produced on carbonate line, to match with carbonate samples. Use R14047/11 for CO2 produced on water line for water samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble | IAEA-C1
[197 empty] | 2022-08-25 | 6104 | Carbonate CO2 Evolution | Completed
[197 empty] | 2022-08-26 | 24955 | Graphite | Completed
[197 empty] | 2022-08-30 | 9515 | C13 Measurement | Completed
[198 empty] | 26000 | AMS Submission Results | Started
Close Job