## bar11
Biobased content | Biobased content error | Biobased content multiplier | Blank rts assigned | Blank rts err assigned | Blanks fm av err Calc | Blanks fm average Calc | C12 | C12C13 ratio | C14C13 ratio | Category Field | Category In Calculation | Collection Date | Collection Date Saved | Colour Description | Comment | CRA | CRA Report | CRA Report Saved | CRA_Error | CRA_Error Report | CRA_Error Report Saved | CRA_OneSigmaMinus | CRA_OneSigmaPlus | Creation User | D14C | D14C_Error | Date Run | DateLastRun | DaysAsGraphite | DaysFromCurrentRun | DCC | DCC_error | DCCstd | DCCstd_error | delta13C_AMS | delta13C_AMS Saved | delta13C_AMS_Error | delta13C_AMS_Error Stored | delta13C_In_Calculation | delta13C_In_Calculation_Error | delta13C_In_Calculation_Source | delta13C_IRMS | delta13C_IRMS_Error | delta13C_IRMS_Source | delta13C_std_av_Match | delta13C_stds_av | delta13C_stds_av_Calculation | delta13C_stds_av_error | delta13C_stds_av_error_Calculation | DELTA14C | DELTA14C Saved | DELTA14C_Error | DELTA14C_Error Saved | deltalower14C | deltalower14C_Error | dlower14C | dlower14C_Error | Email Body | Email sender | End Date | End Operator | End Time | F_corrected_normed | F_corrected_normed Saved | F_corrected_normed_error | F_corrected_normed_error Saved | F_corrected_normed_error_no_WTW_Error | F_corrected_normed_error_no_WTW_noSYS_Error | FeC | Find_TW Global | Flake Description | Fm_normalized | Fm_normalized_error | ID | Import XCAMS Results Date | Import XCAMS Results Time | Import XCAMS Results User | Job | MCC | MCC Saved | MCC_error | MCC_error Saved | MCC_fractional_error | Mode | Modification Date | Modification Time | Modification User | Normalization_exp_factor | Packing Comments | pM | pM_Error | Position | Process Group | Processes Order | Quality Flag | Ratio to standard | Ratio to standard error | Ratio_13C_standard | Ratio13C_denominator | Ratio13C_numerator | Ratio13C_standard_error | Ratio13C_std_av | Ratio13C_std_av Calc | Ratio13C_std_av_Match | Remove System Error | Result Comment for Submitter | Result Comment Internal | rts_bl_av_Calculation | rts_bl_av_Calculation Old | rts_bl_av_error_Calculation | RTS_corrected | RTS_corrected Saved | RTS_corrected_error | RTS_corrected_error Saved | rts_std_av_Match | rts_stds_av | rts_stds_av_Calculation | rts_stds_av_error | rts_stds_av_error_Calculation | Silvia_F_corrected_normed | Silvia_F_corrected_normed_Check | Silvia_F_corrected_normed_error | Silvia_F_corrected_normed_error Check | Silvia_F_corrected_normed_error_no_WTW_Error Check | Silvia_MCC | Silvia_MCC Check | Silvia_MCC_error | Silvia_MCC_error Check | Silvia_RTS_corrected | Silvia_RTS_corrected Check | Silvia_RTS_corrected_error | Silvia_RTS_corrected_error Check | Size Category | SmallSamples values calculated via Excel | Standard 13C Value Constant | Standard Specific Activity Constant | Standard Type Number | Start Date | Start Operator | Start Time | Static Description | Status | tmp_jct_F_new | tmp_jct_F_old | tmp_jct_rts_corr_new | tmp_jct_rts_corr_old | TP | TP Graphite Submitted Minimum Value | TPs Blanks | TPs Blanks Report | TPs Standards | TPs Standards Report | TW | TWs Blanks | TWs Standards | Update Blanks Average Date | Update Blanks Average Time | Update Blanks Average User | Use Assigned Blank | Weight Initial | Wheel To Wheel Error | wtgraph | Date Run | Ratio to standard | TP | TW | wtgraph | Ratio to standard | Ratio to standard error | TP | Standard Prep Type | AMS System | Standard Prep Type | System Error | Wheel Result Comment For Submitter | Wheel Result Comment Internal | delta13C | Surname | delta13C | Position Number | Run Number | delta13C | Run Number | AMS Category | delta13C Estimate | External Due Date | Fraction to date | GlobalStatus | Job Invoicing Notes | Job notes | NZA | Precision | R | Virtual TW | Contaminant observed suspected | Process End Date | Process ID | Process Name | Process Status | R Composed | Sample Description | Comments From Submitter | Comments On Sample Lab Internal | Contaminants | delta13C Submitted | Expected age | Pretreatment by Submitter | Sample Description | Sample ID | Submitter Description | Submitter Requested Pretreatment
[7 empty] | 122.9 | 88.9666 | 0.0003 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49608 | 49608 | 49607 | 321 | 321 | 309 | 49293.4 | 49936.4 | jenny dahl | -997.92 | 0.08 ‰ | 2013-07-23 | 2013-07-23 | 13 | [5 empty] | 2.59‰ | 2.59‰ | 2.00‰ | 2.00‰ | 2.59 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0032 | 0 | 0.75593 | -997.92 ‰ | -997.92 ‰ | 0.08 | 0.08 | -997.89 | 0.08 ‰ | -997.89 | 0.08 ‰ |  |  | 2013-07-24 | christine prior | 11:16 | 0.0020796 | 0.00208 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00208 | 8e-05 | 4662 |  |  |  | 196195 | [5 empty] | 5 | 5/07/22 | 15:24:48 | Albert Zondervan | 0 |  | 0.21 | 0.01% | 38 | 5 | 4 | .X. | 0.002 | 8e-05 |  |  |  |  | 1.0023 | 1.0023 | Yes |  |  | Good to see that blanks produced on larger water line are similar to blanks on smaller carbonate line. .X. to indicate that this is not really a water blank |  |  |  | 0.002 |  | 8e-05 |  | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00208 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.002 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2013-07-19 | jenny dahl | 9:42 |  | Completed |  |  |  |  | 62688 | 2000000 |  |  | 62322 62631 62632 62642 62644 62645 62694 | 62322~62631~62632~62642~62644~62645~62694 | 2876 |  | 2876 2876 2876 2876 2876 2876 2876 |  |  |  |  | 0.94 mg |  | 0.94 | [5 empty] | 0.9968 | 0.00186 | 62322 |  | XCAMS | Both | 0% |  | - first measurement after the big earthquake on Sunday - after doing a re-tune we got good results with no system error or other additional errors - Ox-I in pos#20 (TP62643) was not used for calibration as we used it for the re-tune - C1 in pos#22 (TP6265 | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2013-08-30 | Water | Complete |  | Water line evacuated like carbonate. | 54101 | Normal | 14047/11 |  |  | 2013-07-09 | 718 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00181 | 62631 | [25 empty] | 2013-07-10 | 2698 | Graphite | Completed
[168 empty] | 1.0018 | 0.0018 | 62632 | [25 empty] | 2013-07-11 | 1669 | C13 Measurement | Completed
[168 empty] | 0.99995 | 0.00172 | 62642 | [25 empty] | 2013-07-24 | 4662 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.00178 | 62644 | [25 empty] | 2013-07-26 | 3329 | Review Final Report | Completed
[168 empty] | 0.99895 | 0.00176 | 62645 | [25 empty] | 2013-07-30 | 5363 | Close Job | Completed
[168 empty] | 1.0006 | 0.00175 | 62694
[7 empty] | 112.2 | 89.0852 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50366 | 50366 | 50376 | 397 | 397 | 383 |  |  | jenny dahl | -998.11 | 0.09 ‰ | 2014-02-28 | 2014-02-28 | 31 | [5 empty] | 3.05‰ | 3.05‰ | 2.00‰ | 2.00‰ | 3.05 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.89443 | -998.11 ‰ | -998.11 ‰ | 0.09 | 0.09 | -998.08 | 0.09 ‰ | -998.08 | 0.09 ‰ |  |  | 2014-03-04 | Jocelyn Turnbull | 16:06 | 0.0018924 | 0.00189 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00189 | 9e-05 | 6321 | 2014-03-03 | 8:56 | Albert Zondervan | 198067 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.19 | 0.01% | 10 | 5 | 4 | .X. | 0.00182 | 9e-05 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.00182 |  | 9e-05 |  | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00189 | 0 | 9e-05 | 0 | 0 |  |  |  |  | 0.00182 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-02-19 | jenny dahl | 10:13 |  | Completed |  |  |  |  | 64303 | 2000000 |  |  | 64336 64340 64439 64444 64445 | 64336~64340~64439~64444~64445 | 2917 |  | 2917 2917 2917 2917 2917 | 2014-03-03 | 10:33 | Christine Prior |  | 0.97 mg |  | 0.97 | [5 empty] | 1.002 | 0.00163 | 64336 |  | XCAMS | ST | 0% |  | * Measurement was delayed due to fluctuations in cooling water. Although not fully resolved, control has been improved such that measurements could be started. * The wheel was measured twice, with a retune in between, because of poor repeatability amongst Ox-I standards. While the retune has improved things, Ox-I in cathode position 0 is still too low by 2.5 sigma and was therefore rejected as standard. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-03-24 | Water | Complete | 54050101 $450 | Water line evacuated like carbonate. | 55729 | Normal | 14047/11 |  |  | 2014-01-27 | 982 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99862 | 0.00165 | 64340 | [25 empty] | 2014-01-28 | 4313 | Graphite | Completed
[168 empty] | 0.99995 | 0.00166 | 64439 | [25 empty] | 2014-01-31 | 2247 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00166 | 64444 | [25 empty] | 2014-03-04 | 6321 | AMS Submission Results | Completed
[168 empty] | 0.99943 | 0.00164 | 64445 | [25 empty] | 2014-04-30 | 7780 | Close Job | Completed
[196 empty] | 2014-04-30 | 387 | Invoice | Completed
[7 empty] | 108.1 | 88.9283 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48382 | 48381 | 48391 | 379 | 379 | 365 |  |  | jenny dahl | -997.58 | 0.11 ‰ | 2014-03-17 | 2014-03-17 | 10 | [5 empty] | 4.51‰ | 4.51‰ | 2.00‰ | 2.00‰ | 4.51 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.89443 | -997.58 ‰ | -997.58 ‰ | 0.11 | 0.11 | -997.55 | 0.12 ‰ | -997.55 | 0.12 ‰ |  |  | 2014-03-20 | Jocelyn Turnbull | 8:41 | 0.0024227 | 0.00242 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.3 |  |  | 0.00242 | 0.00011 | 6570 | 2014-03-18 | 12:55 | John West | 198394 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.24 | 0.01% | 3 | 5 | 4 | .X. | 0.00233 | 0.00011 |  |  |  |  | 1.0038 | 1.0038 | Yes | [6 empty] | 0.00233 |  | 0.00011 |  | Yes | 1 | 1 | 0.00102 | 0.00102 | 0.00242 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00233 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-10 | jenny dahl | 13:59 |  | Completed |  |  |  |  | 64666 | 2000000 |  |  | 64615 64625 64627 64634 64645 | 64615~64625~64627~64634~64645 | 2921 |  | 2921 2921 2921 2921 2921 | 2014-03-19 | 16:52 | Christine Prior |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99782 | 0.00227 | 64615 |  | XCAMS | ST | 0.14% |  | * one calibration Ox-I (TP64616) rejected as its rts is a 2.5 sigma outlier * further error adjustments made: - global factor 1.01 to account for excess run-to-run variance - factor 1.3 on cat#3 (TP64666) to account for excess run-to-run variance - global system error 0.14% to account for excess cat-to-cat variance on remaining 5 calibration Ox-Is * The two EA-combusted Ox-Is agree with these 5 calibration Ox-Is, in terms of rts. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-04-30 | Water | Complete |  | Water line evacuated like carbonate. | 55888 | Normal | 14047/11 |  |  | 2014-03-05 | 1152 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00235 | 64625 | [25 empty] | 2014-03-07 | 4676 | Graphite | Completed
[168 empty] | 1.0031 | 0.0023 | 64627 | [25 empty] | 2014-03-13 | 2490 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00225 | 64634 | [25 empty] | 2014-03-20 | 6570 | AMS Submission Results | Completed
[168 empty] | 0.99797 | 0.00229 | 64645 | [25 empty] | 2014-05-06 | 7949 | Close Job | Completed
[7 empty] | 79.6 | 90.1402 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46651 | 46651 | 46638 | 306 | 305 | 294 |  |  | jenny dahl | -996.99 | 0.11 ‰ | 2014-04-03 | 2014-04-03 | 24 | [5 empty] | 1.51‰ | 1.51‰ | 2.00‰ | 2.00‰ | 1.51 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.001 | 0 | 0.8165 | -996.99 ‰ | -996.99 ‰ | 0.11 | 0.11 | -996.96 | 0.12 ‰ | -996.96 | 0.12 ‰ |  |  | 2014-04-07 | Jocelyn Turnbull | 13:29 | 0.003005 | 0.00301 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.2 |  |  | 0.00301 | 0.00011 | 6641 | 2014-04-07 | 11:05 | John West | 198407 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.3 | 0.01% | 25 | 5 | 3 | .X. | 0.00289 | 0.00011 |  |  |  |  | 1.0145 | 1.0145 | Yes |  |  | water line blanks have been much better, |  |  |  | 0.00289 |  | 0.00011 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00301 | 0 | 0.00011 | 0 | 0 |  |  |  |  | 0.00289 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-17 | jenny dahl | 14:18 |  | Completed |  |  |  |  | 64671 | 2000000 |  |  | 64705 64738 64739 64740 64741 64742 | 64705~64738~64739~64740~64741~64742 | 2924 |  | 2924 2924 2924 2924 2924 2924 | 2014-04-07 | 12:03 | Christine Prior |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0016 | 0.00206 | 64705 |  | XCAMS | ST | 0% |  | * good measurement, despite low currents * all runs of 1st rotation rejected as these were associated with voltage-clicks in SNICS |  | RRL | [5 empty] | Background Inorganic |  | 2014-05-01 | Water | Complete |  | Water line evacuated like carbonate. | 56011 | Normal | 14047/11 |  |  | 2014-03-06 | 1163 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99893 | 0.00203 | 64738 | [25 empty] | 2014-03-10 | 4681 | Graphite | Completed
[168 empty] | 0.99977 | 0.00202 | 64739 | [25 empty] | 2014-04-07 | 6641 | AMS Submission Results | Completed
[168 empty] | 0.99929 | 0.00204 | 64740 | [25 empty] | 2014-05-06 | 7920 | Close Job | Completed
[168 empty] | 1.0004 | 0.00203 | 64741
[168 empty] | 1.0001 | 0.00205 | 64742
[7 empty] | 103.6 | 89.1055 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48144 | 48143 | 48129 | 335 | 334 | 321 |  |  | Kelly lyons | -997.5 | 0.10 ‰ | 2014-04-01 | 2014-04-01 | 14 | [5 empty] | 2.76‰ | 2.76‰ | 2.00‰ | 2.00‰ | 2.76 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0015 | 0 | 0.8165 | -997.50 ‰ | -997.50 ‰ | 0.1 | 0.1 | -997.47 | 0.11 ‰ | -997.47 | 0.11 ‰ |  |  | 2014-04-02 | Jocelyn Turnbull | 14:24 | 0.0024955 | 0.0025 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.0025 | 0.0001 | 6686 | 2014-04-02 | 9:58 | Johannes Kaiser | 198501 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.25 | 0.01% | 24 | 5 | 4 | .X. | 0.0024 | 0.0001 |  |  |  |  | 1.0041 | 1.0041 | Yes | [6 empty] | 0.0024 |  | 0.0001 |  | No | 1 | 1 | 0.00121 | 0.00121 | 0.0025 | 0 | 0.0001 | 0 | 0 |  |  |  |  | 0.0024 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-21 | Kelly lyons | 13:24 |  | Completed |  |  |  |  | 64761 | 2000000 |  |  | 64706 64707 64723 64797 64803 64804 | 64706~64707~64723~64797~64803~64804 | 2925 |  | 2925 2925 2925 2925 2925 2925 | 2014-04-02 | 14:05 | Christine Prior |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0014 | 0.00291 | 64706 |  | XCAMS | ST | 0.22% |  | -first measurement after SNICS problems - needed to add a global run-to-run variance factor of 1.06 - Ox-I (TP64797) in pos#0 was high by 4 per mill -Ox-I (TP64707) in pos#14 was low by 4 per mill | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-05-12 | Carbonate |  |  | Water line evacuated like carbonate. | 55966 | Normal | 14047/11 |  |  | 2014-03-17 | 1188 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99591 | 0.00291 | 64707 | [25 empty] | 2014-03-18 | 4771 | Graphite | Completed
[168 empty] | 0.99915 | 0.00297 | 64723 | [25 empty] | 2014-03-19 | 2548 | C13 Measurement | Completed
[168 empty] | 1.0041 | 0.003 | 64797 | [25 empty] | 2014-04-02 | 6686 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00301 | 64803 | [25 empty] | 2014-05-06 | 7911 | Close Job | Completed
[168 empty] | 0.99802 | 0.00295 | 64804
[7 empty] | 78.5 | 89.5634 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50961 | 50961 | 50949 | 380 | 380 | 365 |  |  | Kelly lyons | -998.24 | 0.08 ‰ | 2014-04-07 | 2014-04-07 | 13 | [5 empty] | 2.14‰ | 2.14‰ | 2.00‰ | 2.00‰ | 2.14 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.24 ‰ | -998.24 ‰ | 0.08 | 0.08 | -998.22 | 0.08 ‰ | -998.22 | 0.08 ‰ |  |  | 2014-04-08 | Jocelyn Turnbull | 13:54 | 0.0017573 | 0.00176 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.7 |  |  | 0.00176 | 8e-05 | 6732 | 2014-04-08 | 12:17 | John West | 198577 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.18 | 0.01% | 4 | 5 | 4 | .X. | 0.00169 | 8e-05 |  |  |  |  | 1.0086 | 1.0086 | Yes | [6 empty] | 0.00169 |  | 8e-05 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00176 | 0 | 8e-05 | 0 | 0 |  |  |  |  | 0.00169 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-03-31 | Kelly lyons | 13:47 |  | Completed |  |  |  |  | 64857 | 2000000 |  |  | 64827 64832 64833 64834 64835 64889 | 64827~64832~64833~64834~64835~64889 | 2926 |  | 2926 2926 2926 2926 2926 2926 | 2014-04-08 | 13:38 | Christine Prior |  | 0.77 mg |  | 0.77 | [5 empty] | 0.99984 | 0.00203 | 64827 |  | XCAMS | ST | 0% |  | * good-quality measurement * not quite reached the 250k counts per Ox-I cathode | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-05-19 | Carbonate |  |  | Water line evacuated like carbonate. | 56045 | Normal | 14047/11 |  |  | 2014-03-24 | 1207 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99897 | 0.00201 | 64832 | [25 empty] | 2014-03-25 | 4867 | Graphite | Completed
[168 empty] | 1.001 | 0.00202 | 64833 | [25 empty] | 2014-03-26 | 2608 | C13 Measurement | Completed
[168 empty] | 0.99982 | 0.00204 | 64834 | [25 empty] | 2014-04-08 | 6732 | AMS Submission Results | Completed
[168 empty] | 0.99985 | 0.00203 | 64835 | [25 empty] | 2014-05-06 | 7899 | Close Job | Completed
[168 empty] | 1.0005 | 0.00201 | 64889
[7 empty] | 81.4 | 90.1071 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51401 | 51400 | 51419 | 251 | 251 | 242 |  |  | Kelly lyons | -998.34 | 0.05 ‰ | 2014-04-14 | 2014-04-14 | 3 | [5 empty] | 2.04‰ | 2.04‰ | 2.00‰ | 2.00‰ | 2.04 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.70711 | -998.34 ‰ | -998.34 ‰ | 0.05 | 0.05 | -998.32 | 0.05 ‰ | -998.32 | 0.05 ‰ |  | Christine Prior | 2014-04-17 | Jocelyn Turnbull | 13:02 | 0.0016637 | 0.00166 | 5e-05 | 5e-05 | 5e-05 | 5e-05 | 1.9 |  |  | 0.00166 | 5e-05 | 6929 | 2014-04-17 | 10:21 | John West | 198718 | [5 empty] | 5 | 5/07/22 | 15:24:50 | Albert Zondervan | 0 |  | 0.17 | 0.01% | 22 | 5 | 4 | .X. | 0.0016 | 5e-05 |  |  |  |  | 1.0146 | 1.0146 | Yes | [6 empty] | 0.0016 |  | 5e-05 |  | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00166 | 0 | 5e-05 | 0 | 0 |  |  |  |  | 0.0016 | 0 | 5e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-04-11 | Kelly lyons | 13:39 |  | Completed |  |  |  |  | 65070 | 2000000 |  |  | 64901 64903 64918 64996 64999 65034 65037 65038 | 64901~64903~64918~64996~64999~65034~65037~65038 | 2928 |  | 2928 2928 2928 2928 2928 2928 2928 2928 | 2014-04-17 | 11:27 | Christine Prior |  | 0.98 mg |  | 0.98 | [5 empty] | 1.001 | 0.00215 | 64901 |  | XCAMS | ST | 0.18% |  | * high-prec not achieved due to one Ox-I cathode (TP64903, cat#5) coming out low by 3*sigma_Poisson; found no reason to reject it; system error = 0.18% * global run-to-run variability correction factor = 1.03 * one cathode (TP65036, cat#2) corrected for excess run-to-run variability | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-06-04 | Carbonate | Complete |  | Water line evacuated like carbonate. | 56127 | Normal | 14047/11 |  |  | 2014-04-10 | 1242 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99584 | 0.00214 | 64903 | [25 empty] | 2014-04-11 | 5080 | Graphite | Completed
[168 empty] | 1.0024 | 0.00214 | 64918 | [25 empty] | 2014-04-15 | 2725 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00217 | 64996 | [25 empty] | 2014-04-17 | 6929 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.00216 | 64999 | [25 empty] | 2014-05-06 | 7864 | Close Job | Completed
[168 empty] | 1.0001 | 0.00215 | 65034
[168 empty] | 0.99805 | 0.00214 | 65037
[168 empty] | 0.99986 | 0.00214 | 65038
[7 empty] | 76.1 | 88.9932 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49178 | 49178 | 49193 | 381 | 380 | 367 |  |  | Christine Prior | -997.81 | 0.10 ‰ | 2014-05-05 | 2014-05-05 | 21 | [5 empty] | 2.28‰ | 2.28‰ | 2.00‰ | 2.00‰ | 2.28 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -997.81 ‰ | -997.81 ‰ | 0.1 | 0.1 | -997.78 | 0.11 ‰ | -997.78 | 0.11 ‰ |  |  | 2014-05-07 | Jocelyn Turnbull | 10:43 | 0.002194 | 0.00219 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.9 |  |  | 0.00219 | 0.0001 | 7037 | 2014-05-06 | 12:20 | Albert Zondervan | 198737 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.22 | 0.01% | 22 | 5 | 4 | .X. | 0.00211 | 0.0001 |  |  |  |  | 1.0023 | 1.0023 | Yes | [6 empty] | 0.00211 |  | 0.0001 |  | Yes | 1 | 0.99999 | 0.0012 | 0.0012 | 0.00219 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00211 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-02 | Christine Prior | 11:59 |  | Completed |  |  |  |  | 65085 | 2000000 |  |  | 65095 65101 65102 65103 65104 65159 | 65095~65101~65102~65103~65104~65159 | 2934 |  | 2934 2934 2934 2934 2934 2934 | 2014-05-07 | 10:36 | Christine Prior |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99777 | 0.00295 | 65095 |  | XCAMS | ST | 0.22% |  | * somewhat low C-beam intensities * error adjustments: cathode pos#2 (TP65027) and system error * some early runs rejected because of likely surface contamination * despite the good tune, it remains to be sorted why we haven't achieved a lower system error | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-06-06 | Carbonate | Complete |  | Water line evacuated like carbonate. | 56362 | Normal | 14047/11 |  |  | 2014-04-11 | 1249 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0033 | 0.00294 | 65101 | [25 empty] | 2014-04-14 | 5095 | Graphite | Completed
[168 empty] | 0.99954 | 0.00293 | 65102 | [25 empty] | 2014-04-15 | 2736 | C13 Measurement | Completed
[168 empty] | 0.99988 | 0.00292 | 65103 | [25 empty] | 2014-05-07 | 7037 | AMS Submission Results | Completed
[168 empty] | 0.99609 | 0.00292 | 65104 | [25 empty] | 2014-06-12 | 8543 | Close Job | Completed
[168 empty] | 1.0034 | 0.00295 | 65159
[7 empty] | 90.8 | 89.1722 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52082 | 52081 | 52074 | 437 | 437 | 420 |  |  | jenny dahl | -998.47 | 0.08 ‰ | 2014-05-13 | 2014-05-13 | 13 | [5 empty] | -0.68‰ | -0.68‰ | 2.00‰ | 2.00‰ | -0.68 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -998.47 ‰ | -998.47 ‰ | 0.08 | 0.08 | -998.45 | 0.08 ‰ | -998.45 | 0.08 ‰ |  |  | 2014-05-14 | Jocelyn Turnbull | 16:25 | 0.0015285 | 0.00153 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2 |  |  | 0.00153 | 8e-05 | 7113 | 2014-05-14 | 9:58 | Albert Zondervan | 198814 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 32 | 5 | 4 | .X. | 0.00147 | 8e-05 |  |  |  |  | 1.0014 | 1.0014 | Yes | [6 empty] | 0.00147 |  | 8e-05 |  | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00153 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00147 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-08 | jenny dahl | 16:16 |  | Completed |  |  |  |  | 65169 | 2000000 |  |  | 65188 65189 65190 65216 65217 65218 | 65188~65189~65190~65216~65217~65218 | 2936 |  | 2936 2936 2936 2936 2936 2936 | 2014-05-14 | 12:37 | Christine Prior |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0007 | 0.00185 | 65188 |  | XCAMS | ST | 0% |  | * excellent 1stC-14 measurement after Be-10 * no error adjustments * some first runs on blanks rejected: usual degree of surface contamination | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-06-23 | Carbonate |  |  | Water line evacuated like carbonate. | 56441 | Normal | 14047/11 |  |  | 2014-04-28 | 1264 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99958 | 0.00188 | 65189 | [25 empty] | 2014-04-30 | 5179 | Graphite | Completed
[168 empty] | 1.0005 | 0.00185 | 65190 | [25 empty] | 2014-05-02 | 2771 | C13 Measurement | Completed
[168 empty] | 0.99707 | 0.00185 | 65216 | [25 empty] | 2014-05-14 | 7113 | AMS Submission Results | Completed
[168 empty] | 0.99982 | 0.00189 | 65217 | [25 empty] | 2014-06-12 | 8558 | Close Job | Completed
[168 empty] | 1.0024 | 0.00187 | 65218
|  |  | 0 | 0 |  |  | 100.2 | 88.7845 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41266 | 41266 | 41273 | 199 | 199 | 205 | 41069.5 | 41467.7 | jenny dahl | -994.13 | 0.15 ‰ | 2014-05-19 | 2014-05-19 | 3 | [5 empty] | 3.74‰ | 3.74‰ | 2.00‰ | 2.00‰ | 3.74 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.82 | 0.8165 | -994.13 ‰ | -994.13 ‰ | 0.15 | 0.15 | -994.05 | 0.15 ‰ | -994.05 | 0.15 ‰ |  |  | 2014-05-21 | Jocelyn Turnbull | 16:21 | 0.0058748 | 0.00587 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.3 |  |  | 0.00587 | 0.00015 | 7203 | 2014-05-20 | 14:19 | Johannes Kaiser | 198966 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.59 | 0.01% | 24 | 5 | 5 | .X. | 0.00565 | 0.00014 |  |  |  |  | 1.0014 | 1.0014 | Yes |  |  | really terrible water blank |  |  |  | 0.00565 |  | 0.00014 |  | No | 1 | 0.99998 | 0.00096 | 0.00096 | 0.00587 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00565 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-05-16 | jenny dahl | 16:20 |  | Completed |  |  |  |  | 65331 | 2000000 |  |  | 65272 65273 65274 65306 65307 65308 | 65272~65273~65274~65306~65307~65308 | 2938 |  | 2938 2938 2938 2938 2938 2938 | 2014-05-22 | 9:33 | jenny dahl |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99978 | 0.00233 | 65272 |  | XCAMS | EA | 0.14% |  | - normal measurement water blank (TP 65331) was a strong outlier and was not included in the blank correction | 2.6 | RRL | [5 empty] | Background Test |  | 2014-07-07 | Carbonate |  |  | Water line evacuated like carbonate. | 56484 | Normal | 14047/11 |  |  | 2014-05-13 | 1287 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0021 | 0.00231 | 65273 | [25 empty] | 2014-05-13 | 1288 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99836 | 0.00231 | 65274 | [25 empty] | 2014-05-16 | 5341 | Graphite | Completed
[168 empty] | 0.99986 | 0.00233 | 65306 | [25 empty] | 2014-05-19 | 2827 | C13 Measurement | Completed
[168 empty] | 0.99639 | 0.00249 | 65307 | [25 empty] | 2014-05-21 | 7203 | AMS Submission Results | Completed
[168 empty] | 1.0031 | 0.00241 | 65308 | [25 empty] | 2014-06-12 | 8568 | Close Job | Completed
[7 empty] | 77.7 | 89.2551 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47913 | 47912 | 47908 | 260 | 260 | 250 |  |  | Kelly lyons | -997.43 | 0.08 ‰ | 2014-06-23 | 2014-06-23 | 24 | [5 empty] | 3.24‰ | 3.24‰ | 2.00‰ | 2.00‰ | 3.24 ‰ | 2.00 ‰ | AMS | 2.61‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.82 | 0.8165 | -997.43 ‰ | -997.43 ‰ | 0.08 | 0.08 | -997.4 | 0.08 ‰ | -997.4 | 0.08 ‰ |  |  | 2014-06-25 | Jocelyn Turnbull | 15:10 | 0.0025683 | 0.00257 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.2 |  |  | 0.00257 | 8e-05 | 7491 | 2014-06-25 | 10:13 | Johannes Kaiser | 199226 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.26 | 0.01% | 17 | 5 | 4 | .X. | 0.00247 | 8e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00247 |  | 8e-05 |  | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00257 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00247 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-19 | Kelly lyons | 9:36 |  | Completed |  |  |  |  | 65481 | 2000000 |  |  | 65525 65526 65527 65569 65570 65571 | 65525~65526~65527~65569~65570~65571 | 2947 |  | 2947 2947 2947 2947 2947 2947 | 2014-06-25 | 11:00 | Christine Prior |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0013 | 0.00162 | 65525 |  | XCAMS | ST | 0% |  | - overall a good measurement - needed to add a small factor of 1.03 to account for the global run-to-run variance - had to add an _2 measurement as 15 rotations of _1 did not give 250k counts on Ox-I std due to low beam intensity - 29/01/2015: filled Sys | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-07-23 | Carbonate |  |  | Water line evacuated like carbonate. | 56854 | Normal | 14047/11 |  |  | 2014-05-28 | 1313 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99924 | 0.00156 | 65526 | [25 empty] | 2014-05-30 | 5491 | Graphite | Completed
[168 empty] | 0.99881 | 0.00155 | 65527 | [25 empty] | 2014-06-09 | 2883 | C13 Measurement | Completed
[168 empty] | 1.0025 | 0.00167 | 65569 | [25 empty] | 2014-06-25 | 7491 | AMS Submission Results | Completed
[168 empty] | 0.99871 | 0.00165 | 65570 | [25 empty] | 2014-06-25 | 5143 | Review Final Report | Completed
[168 empty] | 0.99973 | 0.00162 | 65571 | [25 empty] | 2014-09-16 | 9688 | Close Job | Completed
[7 empty] | 72 | 89.2508 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49489 | 49488 | 49492 | 277 | 276 | 266 |  |  | Kelly lyons | -997.89 | 0.07 ‰ | 2014-06-23 | 2014-06-23 | 17 | [5 empty] | 3.29‰ | 3.29‰ | 2.00‰ | 2.00‰ | 3.29 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.82 | 0.8165 | -997.89 ‰ | -997.89 ‰ | 0.07 | 0.07 | -997.86 | 0.07 ‰ | -997.86 | 0.07 ‰ |  |  | 2014-06-25 | Jocelyn Turnbull | 15:11 | 0.0021108 | 0.00211 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 3 |  |  | 0.00211 | 7e-05 | 7497 | 2014-06-25 | 10:13 | Johannes Kaiser | 199492 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:51 | Albert Zondervan | 0 |  | 0.21 | 0.01% | 10 | 5 | 4 | .X. | 0.00203 | 7e-05 |  |  |  |  | 1.0062 | 1.0062 | Yes | [6 empty] | 0.00203 |  | 7e-05 |  | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00211 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00203 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-19 | Kelly lyons | 9:36 |  | Completed |  |  |  |  | 65567 | 2000000 |  |  | 65525 65526 65527 65569 65570 65571 | 65525~65526~65527~65569~65570~65571 | 2947 |  | 2947 2947 2947 2947 2947 2947 | 2014-06-25 | 11:00 | Christine Prior |  | 0.65 mg |  | 0.65 | [5 empty] | 1.0013 | 0.00162 | 65525 |  | XCAMS | ST | 0% |  | - overall a good measurement - needed to add a small factor of 1.03 to account for the global run-to-run variance - had to add an _2 measurement as 15 rotations of _1 did not give 250k counts on Ox-I std due to low beam intensity - 29/01/2015: filled Sys | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-07-30 | Carbonate | Complete |  | Water nitrogen flush but carbonate vessel | 56855 | Normal | 14047/11 |  |  | 2014-06-04 | 1319 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99924 | 0.00156 | 65526 | [25 empty] | 2014-06-06 | 5577 | Graphite | Completed
[168 empty] | 0.99881 | 0.00155 | 65527 | [25 empty] | 2014-06-10 | 2925 | C13 Measurement | Completed
[168 empty] | 1.0025 | 0.00167 | 65569 | [25 empty] | 2014-06-25 | 7497 | AMS Submission Results | Completed
[168 empty] | 0.99871 | 0.00165 | 65570 | [25 empty] | 2014-06-25 | 5145 | Review Final Report | Completed
[168 empty] | 0.99973 | 0.00162 | 65571 | [25 empty] | 2014-09-16 | 9689 | Close Job | Completed
[7 empty] | 97.3 | 88.7754 | 0.0002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52474 | 52473 | Background | 344 | 344 |  |  |  | Kelly lyons | -998.54 | 0.06 ‰ | 2014-07-04 | 2014-07-04 | 11 | [5 empty] | 7.89‰ | 7.89‰ | 2.00‰ | 2.00‰ | 7.89 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0015 | 0 | 0.8165 | -998.54 ‰ | -1000.00 ‰ | 0.06 |  | -998.53 | 0.06 ‰ | -998.53 | 0.06 ‰ |  |  | 2014-07-07 | Jocelyn Turnbull | 9:50 | 0.0014557 |  | 6e-05 | 0 | 6e-05 | 6e-05 | 1.6 |  |  | 0.00146 | 6e-05 | 7658 | 2014-07-07 | 8:50 | Johannes Kaiser | 199596 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.15 | 0.01% | 30 | 5 | 4 | .X. | 0.0014 | 6e-05 |  |  |  |  | 1.0055 | 1.0055 | Yes | [6 empty] | 0.0014 |  | 6e-05 |  | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00146 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.0014 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-06-30 | Kelly lyons | 15:34 |  | Completed |  |  |  |  | 65737 | 2000000 |  |  | 65572 65573 65656 65659 65660 65662 | 65572~65573~65656~65659~65660~65662 | 2950 |  | 2950 2950 2950 2950 2950 2950 | 2014-07-07 | 9:27 | Christine Prior |  | 1.21 mg |  | 1.21 | [5 empty] | 0.9998 | 0.00173 | 65572 |  | XCAMS | ST | 0% |  | - good measurement - no system error |  | RRL | [5 empty] | Background Inorganic |  | 2014-08-13 | Carbonate | Complete |  | Water line evacuated like carbonate. | 56973 | Normal | 14047/11 |  |  | 2014-06-18 | 1354 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00169 | 65573 | [25 empty] | 2014-06-19 | 1358 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99868 | 0.00171 | 65656 | [25 empty] | 2014-06-23 | 5747 | Graphite | Completed
[168 empty] | 0.99966 | 0.0017 | 65659 | [25 empty] | 2014-07-07 | 7658 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00188 | 65660 | [25 empty] | 2014-07-10 | 9001 | Close Job | Completed
[168 empty] | 0.99914 | 0.00186 | 65662
[7 empty] | 82.7 | 89.2642 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 |  |  |  | 49255 | 49254 |  | 346 | 345 |  |  |  | Kelly lyons | -997.83 | 0.09 ‰ | 2014-07-07 | 2014-07-07 | 5 | [5 empty] | 4.94‰ | 4.94‰ | 2.00‰ | 2.00‰ | 4.94 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -997.83 ‰ |  | 0.09 |  | -997.8 | 0.09 ‰ | -997.8 | 0.09 ‰ |  |  | 2014-07-09 | Jocelyn Turnbull | 8:32 | 0.0021732 |  | 9e-05 |  | 9e-05 | 9e-05 | 1.7 |  |  | 0.00217 | 9e-05 | 7693 | 2014-07-08 | 16:43 | Johannes Kaiser | 199750 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.22 | 0.01% | 30 | 5 | 4 | .X. | 0.00209 | 9e-05 |  |  |  |  | 1.008 | 1.008 | Yes | [6 empty] | 0.00209 |  | 9e-05 |  | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.00217 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00209 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-03 | Kelly lyons | 15:25 |  | Completed |  |  |  |  | 65834 | 2000000 |  |  | 65826 65835 65836 65837 65838 65840 | 65826~65835~65836~65837~65838~65840 | 2951 |  | 2951 2951 2951 2951 2951 2951 | 2014-07-09 | 6:22 | Christine Prior |  | 1.12 mg |  | 1.12 | [5 empty] | 1.0012 | 0.00204 | 65826 |  | XCAMS | ST | 0% |  | - had to let it go for 15 rotations in order to get at least 250k counts on Ox-I standards - needed a small error adjustment of 1.04 to explain the global run-to-run variance Both secondaries in this wheel are old - Oxs are biased young, perhaps? | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-08-26 | Carbonate |  |  | Water line evacuated like carbonate. | 57014 | Normal | 14047/11 |  |  | 2014-07-01 | 1373 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0026 | 0.00208 | 65835 | [25 empty] | 2014-07-02 | 5844 | Graphite | Completed
[168 empty] | 0.99672 | 0.00201 | 65836 | [25 empty] | 2014-07-07 | 3020 | C13 Measurement | Completed
[168 empty] | 0.99962 | 0.00197 | 65837 | [25 empty] | 2014-07-09 | 7693 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.002 | 65838 | [25 empty] | 2014-07-10 | 9009 | Close Job | Completed
[168 empty] | 0.99983 | 0.00206 | 65840
[7 empty] | 88 | 89.485 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47287 | 47287 | 47277 | 301 | 300 | 289 |  |  | Kelly lyons | -997.22 | 0.10 ‰ | 2014-07-30 | 2014-07-30 | 6 | [5 empty] | 4.48‰ | 4.48‰ | 2.00‰ | 2.00‰ | 4.48 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -997.22 ‰ | -997.22 ‰ | 0.1 | 0.1 | -997.19 | 0.11 ‰ | -997.19 | 0.11 ‰ |  |  | 2014-08-01 | Jocelyn Turnbull | 13:27 | 0.0027763 | 0.00278 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.1 |  |  | 0.00278 | 0.0001 | 7856 | 2014-07-31 | 16:35 | Johannes Kaiser | 199812 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.28 | 0.01% | 38 | 5 | 4 | .X. | 0.00267 | 0.0001 |  |  |  |  | 1.0101 | 1.0101 | Yes |  |  | poorer than average of blanks for water line in last 3 months |  |  |  | 0.00267 |  | 0.0001 |  | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00278 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00267 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-07-24 | Kelly lyons | 15:34 |  | Completed |  |  |  |  | 66072 | 2000000 |  |  | 65265 65661 65852 65853 65855 65856 | 65265~65661~65852~65853~65855~65856 | 2954 |  | 2954 2954 2954 2954 2954 2954 | 2014-08-01 | 13:07 | Christine Prior |  | 0.86 mg |  | 0.86 | [5 empty] | 1.0006 | 0.00193 | 65265 |  | XCAMS | ST | 0% |  | - good measurement - no error adjustments needed | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2014-09-02 | Carbonate | Not Complete |  | Water nitrogen flush but carbonate vessel | 57134 | Normal | 14047/11 |  |  | 2014-07-23 | 1449 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99917 | 0.00199 | 65661 | [25 empty] | 2014-07-24 | 6082 | Graphite | Completed
[168 empty] | 1.001 | 0.00201 | 65852 | [25 empty] | 2014-07-28 | 3107 | C13 Measurement | Completed
[168 empty] | 0.99869 | 0.00189 | 65853 | [25 empty] | 2014-08-01 | 7856 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00191 | 65855 | [25 empty] | 2014-08-01 | 9239 | Close Job | Completed
[168 empty] | 0.99926 | 0.00209 | 65856
[7 empty] | 97.2 | 89.2327 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51553 | 51553 | 51565 | 358 | 358 | 345 |  |  | jenny dahl | -998.37 | 0.07 ‰ | 2014-08-11 | 2014-08-11 | 17 | [5 empty] | 2.46‰ | 2.46‰ | 2.00‰ | 2.00‰ | 2.46 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.37 ‰ | -998.37 ‰ | 0.07 | 0.07 | -998.35 | 0.07 ‰ | -998.35 | 0.07 ‰ |  |  | 2014-08-14 | Christine Prior | 14:55 | 0.0016325 | 0.00163 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 2 |  |  | 0.00163 | 7e-05 | 7957 | 2014-08-12 | 11:06 | Johannes Kaiser | 199811 |  |  |  |  | 40% | 5 | 5/07/22 | 15:24:52 | Albert Zondervan | 0 |  | 0.16 | 0.01% | 17 | 5 | 4 | .X. | 0.00157 | 7e-05 |  |  |  |  | 1.0052 | 1.0052 | Yes | [6 empty] | 0.00157 |  | 7e-05 |  | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00163 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00157 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-08-06 | jenny dahl | 13:56 |  | Completed |  |  |  |  | 66087 | 2000000 |  |  | 66010 66022 66023 66027 66028 66142 | 66010~66022~66023~66027~66028~66142 | 2958 |  | 2958 2958 2958 2958 2958 2958 | 2014-08-12 | 11:37 | Christine Prior |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99714 | 0.0019 | 66010 |  | XCAMS | ST | 0% |  | - pos# 22 and 33 needed an additional factor of 1.3 to explain its individual run-to-run variance - 29/01/2015: filled System Error (%) box | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-09-02 | Carbonate |  |  | Water line evacuated like carbonate. | 57254 | Normal | 14047/11 |  |  | 2014-07-23 | 1448 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00191 | 66022 | [25 empty] | 2014-07-25 | 6097 | Graphite | Completed
[168 empty] | 1.0002 | 0.00187 | 66023 | [25 empty] | 2014-07-28 | 3112 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00187 | 66027 | [25 empty] | 2014-08-14 | 7957 | AMS Submission Results | Completed
[168 empty] | 1.0026 | 0.00189 | 66028 | [25 empty] | 2014-08-14 | 9360 | Close Job | Completed
[168 empty] | 0.99756 | 0.00192 | 66142
[7 empty] | 117.3 | 89.0334 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 51252 | 51251 | 51275 | 345 | 344 | 333 |  |  | jenny dahl | -998.31 | 0.07 ‰ | 2014-09-02 | 2014-09-02 | 7 | [5 empty] | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -998.31 ‰ | -998.31 ‰ | 0.07 | 0.07 | -998.28 | 0.07 ‰ | -998.28 | 0.07 ‰ |  |  | 2014-09-04 | Jocelyn Turnbull | 13:14 | 0.0016949 | 0.00169 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00169 | 7e-05 | 8184 | 2014-09-03 | 16:36 | Johannes Kaiser | 200154 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.17 | 0.01% | 23 | 5 | 6 | .X. | 0.00163 | 7e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00163 |  | 7e-05 |  | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.00169 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00163 | 0 | 7e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-01 | jenny dahl | 16:09 |  | Completed |  |  |  |  | 66332 | 2000000 |  |  | 66289 66295 66296 66297 66298 66299 | 66289~66295~66296~66297~66298~66299 | 2964 |  | 2964 2964 2964 2964 2964 2964 | 2014-09-04 | 10:05 | Christine Prior |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0018 | 0.00274 | 66289 |  | XCAMS | ST | 0.21% |  | - needed a system error of 0.21% - OX-I with TP 66297 and 66298 were ~ 0.3% low and TP 66299 was high by ~ 0.3% | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-10-20 | Carbonate |  | charge to 540GCT21 - test of waterline extraction method | Water nitrogen flush but carbonate vessel | 57494 | Normal | 14047/11 |  |  | 2014-08-25 | 1483 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00274 | 66295 | [25 empty] | 2014-08-25 | 2777 | Physical Pretreatment | Completed
[168 empty] | 0.99916 | 0.00276 | 66296 | [25 empty] | 2014-08-25 | 1485 | Carbonate CO2 Evolution | Completed
[168 empty] | 0.99732 | 0.00272 | 66297 | [25 empty] | 2014-08-26 | 6342 | Graphite | Completed
[168 empty] | 0.9966 | 0.00272 | 66298 | [25 empty] | 2014-08-27 | 3193 | C13 Measurement | Completed
[168 empty] | 1.0033 | 0.00275 | 66299 | [25 empty] | 2014-09-04 | 8184 | AMS Submission Results | Completed
[196 empty] | 2014-09-10 | 9619 | Close Job | Completed
[196 empty] | 2014-11-10 | 2088 | Invoice | Completed
[7 empty] | 115.7 | 88.8515 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49529 | 49528 | 49530 | 318 | 318 | 306 |  |  | jenny dahl | -997.9 | 0.08 ‰ | 2014-09-02 | 2014-09-02 | 7 | [5 empty] | 5.05‰ | 5.05‰ | 2.00‰ | 2.00‰ | 5.05 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -997.90 ‰ | -997.90 ‰ | 0.08 | 0.08 | -997.87 | 0.08 ‰ | -997.87 | 0.08 ‰ |  |  | 2014-09-04 | Jocelyn Turnbull | 13:14 | 0.0021004 | 0.0021 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.8 |  |  | 0.0021 | 8e-05 | 8185 | 2014-09-03 | 16:36 | Johannes Kaiser | 200155 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.21 | 0.01% | 3 | 5 | 4 | .X. | 0.00202 | 8e-05 |  |  |  |  | 1.0035 | 1.0035 | Yes | [6 empty] | 0.00202 |  | 8e-05 |  | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.0021 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00202 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-01 | jenny dahl | 16:09 |  | Completed |  |  |  |  | 66333 | 2000000 |  |  | 66289 66295 66296 66297 66298 66299 | 66289~66295~66296~66297~66298~66299 | 2964 |  | 2964 2964 2964 2964 2964 2964 | 2014-09-04 | 10:05 | Christine Prior |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0018 | 0.00274 | 66289 |  | XCAMS | ST | 0.21% |  | - needed a system error of 0.21% - OX-I with TP 66297 and 66298 were ~ 0.3% low and TP 66299 was high by ~ 0.3% | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-10-20 | Carbonate | Complete | charge to 540GCT21 - test of waterline extraction method | Water nitrogen flush but carbonate vessel | 57495 | Normal | 14047/11 |  |  | 2014-08-25 | 1484 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00274 | 66295 | [25 empty] | 2014-08-26 | 6343 | Graphite | Completed
[168 empty] | 0.99916 | 0.00276 | 66296 | [25 empty] | 2014-08-27 | 3194 | C13 Measurement | Completed
[168 empty] | 0.99732 | 0.00272 | 66297 | [25 empty] | 2014-09-04 | 8185 | AMS Submission Results | Completed
[168 empty] | 0.9966 | 0.00272 | 66298 | [25 empty] | 2014-09-10 | 9620 | Close Job | Completed
[168 empty] | 1.0033 | 0.00275 | 66299 | [25 empty] | 2014-11-10 | 2089 | Invoice | Completed
[7 empty] | 115.7 | 88.956 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38727 | 38727 | 38726 | 124 | 124 | 120 | 38604 | 38852.8 | Kelly Lyons | -991.94 | 0.12 ‰ | 2014-09-18 | 2014-09-18 | 9 | [5 empty] | 2.56‰ | 2.56‰ | 2.00‰ | 2.00‰ | 2.56 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -991.94 ‰ | -991.94 ‰ | 0.12 | 0.12 | -991.84 | 0.13 ‰ | -991.84 | 0.13 ‰ |  | Christine Prior | 2014-09-23 | Jocelyn Turnbull | 10:27 | 0.0080584 | 0.00806 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00806 | 0.00012 | 8345 | 2014-09-22 | 9:48 | John West | 200243 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.81 | 0.01% | 18 | 5 | 4 | .X. | 0.00775 | 0.00012 |  |  |  |  | 1.0022 | 1.0022 | Yes |  |  | Horrible blank. Unacceptable - need to check conditions. Going to add marble blanks from recent wheels at BG correction for this wheel. |  |  |  | 0.00775 |  | 0.00012 |  | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00806 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00775 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-09-17 | Kelly Lyons | 8:09 |  | Completed |  |  |  |  | 66448 | 2000000 |  |  | 66301 66370 66371 66372 66373 66454 | 66301~66370~66371~66372~66373~66454 | 2968 |  | 2968 2968 2968 2968 2968 2968 | 2014-09-22 | 14:40 | Christine Prior |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99874 | 0.0014 | 66301 |  | XCAMS | ST | 0% |  | - 2 measurements done on this wheel - 1st measurement Judge=ON - all rotations with Judge=ON and cathodes skipped have been rejected due to unexplained run-to-run variance - no error corrections were needed - later rotations for cathode 36 (TP66338) rejec | 2.5 | RRL | [5 empty] | Background Test |  | 2014-11-03 | Carbonate | Complete |  | Water nitrogen flush but carbonate vessel forgot to write down total extraction time | 57662 | Normal | 14047/11 |  |  | 2014-09-08 | 1491 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00139 | 66370 | [25 empty] | 2014-09-09 | 6458 | Graphite | Completed
[168 empty] | 0.99934 | 0.00142 | 66371 | [25 empty] | 2014-09-11 | 3245 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00139 | 66372 | [25 empty] | 2014-09-23 | 8345 | AMS Submission Results | Completed
[168 empty] | 0.99961 | 0.00138 | 66373 | [25 empty] | 2014-09-26 | 9970 | Close Job | Completed
[168 empty] | 1.0007 | 0.00139 | 66454
[7 empty] | 104.4 | 88.9946 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41556 | 41555 | 41551 | 192 | 191 | 198 |  |  | Kelly Lyons | -994.33 | 0.14 ‰ | 2014-10-08 | 2014-10-08 | 20 | [5 empty] | 2.68‰ | 2.68‰ | 2.00‰ | 2.00‰ | 2.68 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | -994.33 ‰ | -994.33 ‰ | 0.14 | 0.14 | -994.26 | 0.14 ‰ | -994.26 | 0.14 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.0056669 | 0.00567 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 3.2 |  |  | 0.00567 | 0.00014 | 8476 | 2014-10-09 | 13:55 | Johannes Kaiser | 200296 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 32 | 5 | 4 | .X. | 0.00545 | 0.00013 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | not a great result since waters are fairly old in wheel and job notes indicate this sample had a low yield, I have changed category to Background test and will use TW 2964 TP 66333 for correction with other InO blank in TW 2972 |  |  |  | 0.00545 |  | 0.00013 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00567 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00545 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed |  |  |  |  | 66536 | 2000000 |  |  | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 |  | 2972 2972 2972 2972 2972 2972 | 2014-10-10 | 10:47 | jenny dahl |  | 0.59 mg |  | 0.59 | [5 empty] | 1.0011 | 0.00186 | 66592 |  | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-11-10 | Carbonate |  |  | Water nitrogen flush but carbonate vessel | 57814 | Normal | 14047/11 |  |  | 2014-09-17 | 1512 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00179 | 66639 | [25 empty] | 2014-09-18 | 6546 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [25 empty] | 2014-09-19 | 3271 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [25 empty] | 2014-10-10 | 8476 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [25 empty] | 2014-10-10 | 10134 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[7 empty] | 116.6 | 88.9177 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43311 | 43311 | 43319 | 257 | 256 | 265 |  |  | Kelly Lyons | -995.45 | 0.15 ‰ | 2014-10-08 | 2014-10-08 | 16 | [5 empty] | 3.55‰ | 3.55‰ | 2.00‰ | 2.00‰ | 3.55 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0.82 | 0.8165 | -995.45 ‰ | -995.45 ‰ | 0.15 | 0.15 | -995.39 | 0.15 ‰ | -995.39 | 0.15 ‰ |  |  | 2014-10-10 | Jocelyn Turnbull | 11:15 | 0.0045543 | 0.00455 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00455 | 0.00015 | 8482 | 2014-10-09 | 13:55 | Johannes Kaiser | 200316 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 18 | 5 | 4 | .X. | 0.00438 | 0.00014 |  |  |  |  | 1.0027 | 1.0027 | Yes |  |  | not super great |  |  |  | 0.00438 |  | 0.00014 |  | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00455 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00438 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-06 | Kelly Lyons | 11:34 |  | Completed |  |  |  |  | 66541 | 2000000 |  |  | 66592 66639 66642 66645 66648 66661 | 66592~66639~66642~66645~66648~66661 | 2972 |  | 2972 2972 2972 2972 2972 2972 | 2014-10-10 | 10:47 | jenny dahl |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0011 | 0.00186 | 66592 |  | XCAMS | EA | 0% |  | - only reporting _2 measurement as _1 has poor repeatability of calibration standards - Cat#18 (TP66541) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-11-12 | Carbonate |  |  | Water line not sure how std prepared marked as test MN | 57820 | Normal | 14047/11 |  |  | 2014-09-18 | 1515 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00179 | 66639 | [25 empty] | 2014-09-22 | 6551 | Graphite | Completed
[168 empty] | 0.99605 | 0.00185 | 66642 | [25 empty] | 2014-09-24 | 3277 | C13 Measurement | Completed
[168 empty] | 0.99993 | 0.00176 | 66645 | [25 empty] | 2014-10-10 | 8482 | AMS Submission Results | Completed
[168 empty] | 0.99977 | 0.00177 | 66648 | [25 empty] | 2014-10-10 | 10136 | Close Job | Completed
[168 empty] | 1.0014 | 0.00176 | 66661
[7 empty] | 126.4 | 88.7293 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46820 | 46819 | 46827 | 255 | 255 | 246 |  |  | Kelly Lyons | -997.06 | 0.09 ‰ | 2014-10-15 | 2014-10-15 | 12 | [5 empty] | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.57‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0.82 | 0.8165 | -997.06 ‰ | -997.06 ‰ | 0.09 | 0.09 | -997.02 | 0.09 ‰ | -997.02 | 0.09 ‰ |  |  | 2014-10-20 | Jocelyn Turnbull | 14:38 | 0.0029426 | 0.00294 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00294 | 9e-05 | 8568 | 2014-10-16 | 13:49 | Johannes Kaiser | 200398 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.29 | 0.01% | 20 | 5 | 4 | .X. | 0.00283 | 9e-05 |  |  |  |  | 1.0016 | 1.0016 | Yes | [6 empty] | 0.00283 |  | 9e-05 |  | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00294 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00283 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-10-14 | Kelly Lyons | 10:28 |  | Completed |  |  |  |  | 66667 | 2000000 |  |  | 66694 66697 66699 66700 66703 66704 | 66694~66697~66699~66700~66703~66704 | 2974 |  | 2974 2974 2974 2974 2974 2974 | 2014-10-17 | 12:15 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 1.0033 | 0.00166 | 66694 |  | XCAMS | EA | 0% |  | - good measurement - needed to apply a factor of 1.04 to explain the global run-to-run variance - cathode 14 (TP66712) needed a factor of 1.4 to explain its individual run-to-run variance | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2014-11-25 | Carbonate | Complete |  | Water line not sure how std prepared | 57899 | Normal | 14047/11 |  |  | 2014-10-01 | 1545 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00168 | 66697 | [25 empty] | 2014-10-03 | 6677 | Graphite | Completed
[168 empty] | 0.99853 | 0.00165 | 66699 | [25 empty] | 2014-10-07 | 3306 | C13 Measurement | Completed
[168 empty] | 0.99903 | 0.00164 | 66700 | [25 empty] | 2014-10-20 | 8568 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00169 | 66703 | [25 empty] | 2014-10-22 | 10191 | Close Job | Completed
[168 empty] | 1 | 0.00162 | 66704
[7 empty] | 113.8 | 89.0032 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 50063 | 50062 | 50043 | 340 | 340 | 326 |  |  | jenny dahl | -998.03 | 0.08 ‰ | 2014-11-24 | 2014-11-24 | 13 | [5 empty] | 0.94‰ | 0.94‰ | 2.00‰ | 2.00‰ | 0.94 ‰ | 2.00 ‰ | AMS | 2.51‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -998.03 ‰ | -998.03 ‰ | 0.08 | 0.08 | -998.01 | 0.08 ‰ | -998.01 | 0.08 ‰ |  |  | 2014-12-01 | Jocelyn Turnbull | 15:49 | 0.0019652 | 0.00197 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00197 | 8e-05 | 8765 | 2014-11-27 | 9:56 | Johannes Kaiser | 200651 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 31 | 5 | 4 | .X. | 0.00189 | 8e-05 |  |  |  |  | 1.0011 | 1.0011 | Yes | [6 empty] | 0.00189 |  | 8e-05 |  | Yes | 1 | 1 | 0.00118 | 0.00118 | 0.00197 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00189 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-21 | jenny dahl | 14:29 |  | Completed |  |  |  |  | 66887 | 2000000 |  |  | 66817 66854 66908 66911 66914 66917 | 66817~66854~66908~66911~66914~66917 | 2979 |  | 2979 2979 2979 2979 2979 2979 | 2014-11-27 | 11:10 | Christine Prior |  | 0.90 mg |  | 0.9 | [5 empty] | 0.99761 | 0.00285 | 66817 |  | XCAMS | EA | 0.22% |  | - stopped _1 measurement after ~4.5 rotations due to poor repeatability amongst the calibration standards - only _2 measurement data were used for reporting - system error = 0.22% - cat#5 (TP66822) needed an additional factor of 1.4 to explain its individ | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2014-12-31 | Carbonate |  | standard, no charge | Water nitrogen flush but carbonate vessel | 58050 | Normal | 14047/11 |  |  | 2014-11-10 | 1587 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00287 | 66854 | [25 empty] | 2014-11-11 | 6897 | Graphite | Completed
[168 empty] | 1.0024 | 0.00291 | 66908 | [25 empty] | 2014-12-01 | 3371 | C13 Measurement | Completed
[168 empty] | 0.99609 | 0.00289 | 66911 | [25 empty] | 2014-12-01 | 8765 | AMS Submission Results | Completed
[168 empty] | 0.99882 | 0.00287 | 66914 | [25 empty] | 2014-12-01 | 10404 | Close Job | Completed
[168 empty] | 1.0028 | 0.00289 | 66917
[7 empty] | 120.5 | 88.8482 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45252 | 45251 | 45245 | 257 | 256 | 247 |  |  | jenny dahl | -996.42 | 0.11 ‰ | 2014-11-24 | 2014-11-24 | 11 | [5 empty] | 2.69‰ | 2.69‰ | 2.00‰ | 2.00‰ | 2.69 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.42 ‰ | -996.42 ‰ | 0.11 | 0.11 | -996.38 | 0.12 ‰ | -996.38 | 0.12 ‰ |  |  | 2014-12-01 | Jocelyn Turnbull | 15:49 | 0.0035769 | 0.00358 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3 |  |  | 0.00358 | 0.00011 | 8766 | 2014-11-27 | 9:56 | Johannes Kaiser | 200652 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 19 | 5 | 6 | ... | 0.00344 | 0.00011 |  |  |  |  | 1.0011 | 1.0011 | Yes |  |  | Interesting difference in blank run through water bottle versus carbonate vessel, even though on the same vacuum line. |  |  |  | 0.00344 |  | 0.00011 |  | Yes | 1 | 1 | 0.00118 | 0.00118 | 0.00358 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00344 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-11-21 | jenny dahl | 14:29 |  | Completed |  |  |  |  | 66921 | 2000000 |  |  | 66817 66854 66908 66911 66914 66917 | 66817~66854~66908~66911~66914~66917 | 2979 |  | 2979 2979 2979 2979 2979 2979 | 2014-11-27 | 11:10 | Christine Prior |  | 0.62 mg |  | 0.62 | [5 empty] | 0.99761 | 0.00285 | 66817 |  | XCAMS | EA | 0.22% |  | - stopped _1 measurement after ~4.5 rotations due to poor repeatability amongst the calibration standards - only _2 measurement data were used for reporting - system error = 0.22% - cat#5 (TP66822) needed an additional factor of 1.4 to explain its individ | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2014-12-31 | Carbonate |  | standard, no charge | water line nitrogen flush water vessel | 58051 | Normal | 14047/11 |  |  | 2014-11-11 | 1595 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00287 | 66854 | [25 empty] | 2014-11-11 | 1139 | Gas Splitting | Completed
[168 empty] | 1.0024 | 0.00291 | 66908 | [25 empty] | 2014-11-11 | 2299 | Choose Gas Splitting | Completed
[168 empty] | 0.99609 | 0.00289 | 66911 | [25 empty] | 2014-11-13 | 6931 | Graphite | Completed
[168 empty] | 0.99882 | 0.00287 | 66914 | [25 empty] | 2014-12-01 | 3376 | C13 Measurement | Completed
[168 empty] | 1.0028 | 0.00289 | 66917 | [25 empty] | 2014-12-01 | 8766 | AMS Submission Results | Completed
[196 empty] | 2014-12-01 | 10405 | Close Job | Completed
[7 empty] | 122.8 | 88.8922 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43516 | 43515 | 43516 | 207 | 206 | 199 |  |  | jenny dahl | -995.56 | 0.11 ‰ | 2014-12-11 | 2014-12-11 | 20 | [5 empty] | 2.83‰ | 2.83‰ | 2.00‰ | 2.00‰ | 2.83 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | -995.56 ‰ | -995.56 ‰ | 0.11 | 0.11 | -995.51 | 0.12 ‰ | -995.51 | 0.12 ‰ |  | Christine Prior | 2014-12-12 | Jocelyn Turnbull | 11:55 | 0.0044399 | 0.00444 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00444 | 0.00011 | 8912 | 2014-12-12 | 10:49 | Johannes Kaiser | 200767 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:41 | Jocelyn Turnbull | 0 |  | 0.44 | 0.01% | 3 | 5 | 4 | ... | 0.00427 | 0.00011 |  |  |  |  | 1.0017 | 1.0017 | Yes |  |  | hmmm, not so good. Blank from water bottle with N2 flush. Cf job 200652 TP66921; may be realistic when use of water vessels compared with carbonate vessels. |  |  |  | 0.00427 |  | 0.00011 |  | Yes | 1 | 1 | 0.00103 | 0.00103 | 0.00444 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00427 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2014-12-10 | jenny dahl | 14:19 |  | Completed |  |  |  |  | 66980 | 2000000 |  |  | 66903 66922 67039 67040 67041 67042 | 66903~66922~67039~67040~67041~67042 | 2983 |  | 2983 2983 2983 2983 2983 2983 | 2014-12-12 | 11:43 | Christine Prior |  | 1.08 mg |  | 1.08 | [5 empty] | 0.99785 | 0.00249 | 66903 |  | XCAMS | EA | 0.18% |  | #NAME? | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-01-14 | Carbonate | Complete | standard, no charge | water line nitrogen flush water vessel | 58233 | Normal | 14047/11 |  |  | 2014-11-19 | 1620 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99715 | 0.0025 | 66922 | [25 empty] | 2014-11-21 | 6990 | Graphite | Completed
[168 empty] | 1.0004 | 0.00257 | 67039 | [25 empty] | 2014-12-01 | 3385 | C13 Measurement | Completed
[168 empty] | 1.0042 | 0.00253 | 67040 | [25 empty] | 2014-12-12 | 8912 | AMS Submission Results | Completed
[168 empty] | 1.0007 | 0.00252 | 67041 | [25 empty] | 2014-12-12 | 10618 | Close Job | Completed
[168 empty] | 0.99991 | 0.00256 | 67042
[7 empty] | 123.9 | 88.9051 | 0.0018 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 36576 | 36576 | 36578 | 127 | 126 | 130 | 36450.2 | 36704 | Cathy Ginnane | -989.47 | 0.17 ‰ | 2015-01-26 | 2015-01-26 | 19 | [5 empty] | 4.40‰ | 4.40‰ | 2.00‰ | 2.00‰ | 4.40 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -989.47 ‰ | -989.47 ‰ | 0.17 | 0.17 | -989.34 | 0.17 ‰ | -989.34 | 0.17 ‰ |  |  | 2015-01-28 | Jocelyn Turnbull | 15:21 | 0.010533 | 0.01053 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.01053 | 0.00017 | 9030 | 2015-01-27 | 14:07 | Johannes Kaiser | 201062 |  |  |  |  | 40% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 24 | 5 | 4 | ... | 0.01013 | 0.00016 |  |  |  |  | 1.0034 | 1.0034 | Yes |  |  | so much younger than previous water line blanks, even considering N2 flush. Chose to use three most recent water line blanks with N2 flush instead First use of water line after holiday - maybe N2 flush not complete? |  |  |  | 0.01013 |  | 0.00016 |  | No | 1 | 1 | 0.00096 | 0.00096 | 0.01053 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01013 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-01-21 | Cathy Ginnane | 9:20 |  | Completed |  |  |  |  | 67105 | 2000000 |  |  | 67102 67103 67106 67107 67108 67109 | 67102~67103~67106~67107~67108~67109 | 2986 |  | 2986 2986 2986 2986 2986 2986 | 2015-01-28 | 15:11 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 0.99566 | 0.00235 | 67102 |  | XCAMS | ST | 0.17% |  | - rejected first rotation from analysis - needed a system error of 0.17% due to Ox-I in pos#0 (TP67102) being low by almost 3 sigma | 2.4 | RRL | [5 empty] | Background Test |  | 2015-02-16 | Carbonate |  |  | water line nitrogen flush water vessel | 58352 | Normal | 14047/11 |  |  | 2015-01-06 | 1641 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0018 | 0.00236 | 67103 | [25 empty] | 2015-01-07 | 7115 | Graphite | Completed
[168 empty] | 1.0002 | 0.00235 | 67106 | [25 empty] | 2015-01-13 | 3413 | C13 Measurement | Completed
[168 empty] | 0.99974 | 0.00235 | 67107 | [25 empty] | 2015-01-28 | 9030 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00235 | 67108 | [25 empty] | 2015-01-28 | 10788 | Close Job | Completed
[168 empty] | 1.0024 | 0.00241 | 67109
[7 empty] | 128.2 | 88.9732 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43348 | 43348 | 43354 | 203 | 202 | 195 |  |  | jenny dahl | -995.47 | 0.11 ‰ | 2015-02-11 | 2015-02-11 | 15 | [5 empty] | 3.22‰ | 3.22‰ | 2.00‰ | 2.00‰ | 3.22 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -995.47 ‰ | -995.47 ‰ | 0.11 | 0.11 | -995.41 | 0.12 ‰ | -995.41 | 0.12 ‰ |  |  | 2015-02-13 | Jocelyn Turnbull | 11:38 | 0.0045335 | 0.00453 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.2 |  |  | 0.00453 | 0.00011 | 9108 | 2015-02-12 | 16:27 | Johannes Kaiser | 201195 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.45 | 0.01% | 10 | 5 | 6 | ... | 0.00436 | 0.00011 |  |  |  |  | 1.003 | 1.003 | Yes |  |  | Blank looks like crap, but... compare this to the recent blanks-in-water-bottle-with-N2-flush. I think this is realistic blank |  |  |  | 0.00436 |  | 0.00011 |  | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00453 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00436 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-05 | jenny dahl | 16:03 |  | Completed |  |  |  |  | 67228 | 2000000 |  |  | 67240 67241 67248 67263 67300 67303 | 67240~67241~67248~67263~67300~67303 | 2988 |  | 2988 2988 2988 2988 2988 2988 | 2015-02-13 | 10:51 | Christine Prior |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00154 | 67240 |  | XCAMS | EA | 0% |  | -very good measurement - pos#26 needed an additional factor of 1.3 due to individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-03-12 | Carbonate | Not Complete |  | water line nitrogen flush water vessel | 58417 | Normal | 14047/11 |  |  | 2015-01-26 | 1689 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00154 | 67241 | [26 empty] | 1164 | Gas Splitting | Completed
[168 empty] | 0.99923 | 0.00156 | 67248 | [25 empty] | 2015-01-27 | 2333 | Choose Gas Splitting | Completed
[168 empty] | 0.99904 | 0.00168 | 67263 | [25 empty] | 2015-01-27 | 7238 | Graphite | Completed
[168 empty] | 1.0006 | 0.00155 | 67300 | [25 empty] | 2015-02-02 | 3450 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00156 | 67303 | [25 empty] | 2015-02-13 | 9108 | AMS Submission Results | Completed
[196 empty] | 2015-02-13 | 10865 | Close Job | Completed
[7 empty] | 131.5 | 88.8213 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45067 | 45067 | 45067 | 251 | 251 | 241 |  |  | Christine Prior | -996.34 | 0.11 ‰ | 2015-02-20 | 2015-02-20 | 8 | [5 empty] | 2.76‰ | 2.76‰ | 2.00‰ | 2.00‰ | 2.76 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.34 ‰ | -996.34 ‰ | 0.11 | 0.11 | -996.29 | 0.12 ‰ | -996.29 | 0.12 ‰ |  |  | 2015-02-27 | jenny dahl | 9:13 | 0.0036601 | 0.00366 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.1 |  |  | 0.00366 | 0.00011 | 9190 | 2015-02-25 | 11:18 | Johannes Kaiser | 201353 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 24 | 5 | 6 | ... | 0.00352 | 0.00011 |  |  |  |  | 1.0009 | 1.0009 | Yes |  |  | - needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance |  |  |  | 0.00352 |  | 0.00011 |  | Yes | 1 | 1 | 0.00084 | 0.00084 | 0.00366 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00352 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-02-19 | Christine Prior | 15:54 |  | Completed |  |  |  |  | 67355 | 2000000 |  |  | 67354 67387 67411 67412 67413 67414 | 67354~67387~67411~67412~67413~67414 | 2990 |  | 2990 2990 2990 2990 2990 2990 | 2015-02-25 | 14:57 | jenny dahl |  | 1.08 mg |  | 1.08 | [5 empty] | 0.9963 | 0.00206 | 67354 |  | XCAMS | EA | 0.13% |  | - system error = 0.13% and this is mainly due to Ox-I in pos# 7 (TP67354) being low by ~2.5 sigma - cat# 24 (TP67355) and cat# 27 (TP67335) needed an additional error adjustment factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP6730 | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-02 | Water | Complete |  | water line nitrogen flush water vessel | 58488 | Normal | 14047/11 |  |  | 2015-02-11 | 1763 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00209 | 67387 | [25 empty] | 2015-02-11 | 1174 | Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00206 | 67411 | [25 empty] | 2015-02-11 | 2347 | Choose Gas Splitting | Completed
[168 empty] | 0.99951 | 0.00204 | 67412 | [25 empty] | 2015-02-12 | 7365 | Graphite | Completed
[168 empty] | 0.99978 | 0.00209 | 67413 | [25 empty] | 2015-02-13 | 3492 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00205 | 67414 | [25 empty] | 2015-02-27 | 9190 | AMS Submission Results | Completed
[196 empty] | 2015-02-27 | 10904 | Close Job | Completed
[7 empty] | 94.4 | 89.4691 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46624 | 46623 | 46611 | 249 | 249 | 239 |  |  | jenny dahl | -996.98 | 0.09 ‰ | 2015-03-06 | 2015-03-06 | 29 | [5 empty] | -5.21‰ | -5.21‰ | 2.00‰ | 2.00‰ | -5.21 ‰ | 2.00 ‰ | AMS | 2.62‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | -996.98 ‰ | -996.98 ‰ | 0.09 | 0.09 | -996.95 | 0.09 ‰ | -996.95 | 0.09 ‰ |  |  | 2015-03-11 | Jocelyn Turnbull | 15:24 | 0.0030154 | 0.00302 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.5 |  |  | 0.00302 | 9e-05 | 9262 | 2015-03-09 | 11:26 | Johannes Kaiser | 201236 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.3 | 0.01% | 4 | 5 | 6 | ... | 0.0029 | 9e-05 |  |  |  |  | 1.0001 | 1.0001 | Yes | [6 empty] | 0.0029 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00302 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0029 | 0 | 9e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-05 | jenny dahl | 15:41 |  | Completed |  |  |  |  | 67304 | 2000000 |  |  | 67457 67458 67513 67514 67515 | 67457~67458~67513~67514~67515 | 2992 |  | 2992 2992 2992 2992 2992 | 2015-03-11 | 11:50 | Christine Prior |  | 1.11 mg |  | 1.11 | [5 empty] | 1.001 | 0.00154 | 67457 |  | XCAMS | EA | 0% |  | - no system error - Ox-I (TP67494) was rejected from calibration as it was low by 4.3 sigma - cat# 18 (TP67465) needed an additional factor of 1.3 to explain its individual run-to-run variance - cat# 38 (TP67377) needed an additional factor of 1.4 to expl | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-03-19 | Carbonate |  |  | water line nitrogen flush water vessel | 58588 | Normal | 14047/11 |  |  | 2015-02-03 | 1725 | Carbonate CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99995 | 0.00153 | 67458 | [25 empty] | 2015-02-03 | 1168 | Gas Splitting | Completed
[168 empty] | 0.99914 | 0.00152 | 67513 | [25 empty] | 2015-02-03 | 2338 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 67514 | [25 empty] | 2015-02-05 | 7314 | Graphite | Completed
[168 empty] | 0.99974 | 0.00155 | 67515 | [25 empty] | 2015-02-10 | 3476 | C13 Measurement | Completed
[196 empty] | 2015-03-11 | 9262 | AMS Submission Results | Completed
[196 empty] | 2015-03-12 | 10974 | Close Job | Completed
[7 empty] | 76.1 | 88.8027 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44495 | 44494 | 44495 | 254 | 254 | 254 |  |  | Christine Prior | -996.07 | 0.12 ‰ | 2015-03-18 | 2015-03-06 | 19 | 12 |  |  |  |  | 1.77‰ | 1.77‰ | 2.00‰ | 2.00‰ | 1.77 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -996.07 ‰ | -996.07 ‰ | 0.12 | 0.12 | -996.02 | 0.13 ‰ | -996.02 | 0.13 ‰ |  |  | 2015-03-20 | Jocelyn Turnbull | 12:15 | 0.0039304 | 0.0039304 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00393 | 0.00012 | 9343 | 2015-03-19 | 15:52 | Johannes Kaiser | 201402 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.39 | 0.01% | 23 | 5 | 6 | ... | 0.00378 | 0.00012 |  |  |  |  | 0.99966 | 0.99966 | Yes | Yes | [5 empty] | 0.00378 | 0.00‰ | 0.00012 | 0.00‰ | No | 1 | 0.99996 | 0.00125 | 0.00125 | 0.00393 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00378 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-17 | Christine Prior | 9:12 |  | Completed |  |  |  |  | 67468 | 2000000 |  |  | 67519 67540 67541 67543 67544 67558 | 67519~67540~67541~67543~67544~67558 | 2994 |  | 2994 2994 2994 2994 2994 2994 | 2015-03-20 | 10:11 | Christine Prior |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99664 | 0.00306 | 67519 |  | XCAMS | EA | 0.24% |  | - first measurement after major source clean - system error = 0.24% - cat in pos #25 (TP67469) needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-04-16 | Water |  |  | water line nitrogen flush water vessel | 58668 | Normal | 14047/11 |  |  | 2015-02-26 | 1807 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9989 | 0.00305 | 67540 | [25 empty] | 2015-02-26 | 1187 | Gas Splitting | Completed
[168 empty] | 0.99716 | 0.00305 | 67541 | [25 empty] | 2015-02-26 | 2369 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00305 | 67543 | [25 empty] | 2015-02-27 | 7478 | Graphite | Completed
[168 empty] | 1.0046 | 0.00305 | 67544 | [25 empty] | 2015-03-03 | 3545 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00306 | 67558 | [25 empty] | 2015-03-20 | 9343 | AMS Submission Results | Completed
[196 empty] | 2015-03-20 | 11113 | Close Job | Completed
[7 empty] | 78.1 | 88.7635 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42711 | 42710 | 42711 | 203 | 203 | 203 |  |  | Christine Prior | -995.09 | 0.12 ‰ | 2015-03-18 | 2015-03-06 | 16 | 12 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -995.09 ‰ | -995.09 ‰ | 0.12 | 0.12 | -995.03 | 0.13 ‰ | -995.03 | 0.13 ‰ |  |  | 2015-03-20 | Jocelyn Turnbull | 12:15 | 0.0049078 | 0.0049078 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00491 | 0.00012 | 9350 | 2015-03-19 | 15:52 | Johannes Kaiser | 201418 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 2 | 5 | 6 | ... | 0.00472 | 0.00012 |  |  |  |  | 0.99966 | 0.99966 | Yes | Yes | [5 empty] | 0.00472 | 0.00‰ | 0.00012 | 0.00‰ | No | 1 | 0.99996 | 0.00125 | 0.00125 | 0.00491 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00472 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-03-17 | Christine Prior | 9:12 |  | Completed |  |  |  |  | 67472 | 2000000 |  |  | 67519 67540 67541 67543 67544 67558 | 67519~67540~67541~67543~67544~67558 | 2994 |  | 2994 2994 2994 2994 2994 2994 | 2015-03-20 | 10:11 | Christine Prior |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99664 | 0.00306 | 67519 |  | XCAMS | EA | 0.24% |  | - first measurement after major source clean - system error = 0.24% - cat in pos #25 (TP67469) needed an additional uncertainty adjustment factor of 1.3 to explain its individual run-to-run variance | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2015-04-20 | Water |  |  | water line nitrogen flush water vessel | 58669 | Normal | 14047/11 |  |  | 2015-02-27 | 1812 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9989 | 0.00305 | 67540 | [25 empty] | 2015-02-27 | 1189 | Gas Splitting | Completed
[168 empty] | 0.99716 | 0.00305 | 67541 | [25 empty] | 2015-02-27 | 2373 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00305 | 67543 | [25 empty] | 2015-03-02 | 7482 | Graphite | Completed
[168 empty] | 1.0046 | 0.00305 | 67544 | [25 empty] | 2015-03-10 | 3550 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00306 | 67558 | [25 empty] | 2015-03-20 | 9350 | AMS Submission Results | Completed
[196 empty] | 2015-03-20 | 11120 | Close Job | Completed
[7 empty] | 111.2 | 89.2667 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 49853 | 49853 | 49853 | 331 | 330 | 331 |  |  | Margaret Norris | -997.98 | 0.08 ‰ | 2015-05-08 | 2015-03-18 | 21 | 51 |  |  |  |  | -1.96‰ | -1.96‰ | 2.00‰ | 2.00‰ | -1.96 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -997.98 ‰ | -997.98 ‰ | 0.08 | 0.08 | -997.96 | 0.08 ‰ | -997.96 | 0.08 ‰ |  | Christine Prior | 2015-05-12 | Jocelyn Turnbull | 10:55 | 0.0020172 | 0.0020172 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.6 |  |  | 0.00202 | 8e-05 | 9608 | 2015-05-11 | 10:15 | Johannes Kaiser | 201994 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 12 | 5 | 6 | ... | 0.00194 | 8e-05 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.00194 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 0.99998 | 0.00092 | 0.00092 | 0.00202 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00194 | 0 | 8e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-07 | Margaret Norris | 15:30 |  | Completed |  |  |  |  | 67781 | 2000000 |  |  | 67726 67727 67894 67895 67898 67899 | 67726~67727~67894~67895~67898~67899 | 3002 |  | 3002 3002 3002 3002 3002 3002 | 2015-05-11 | 15:09 | Christine Prior |  | 0.76 mg |  | 0.76 | [5 empty] | 0.99675 | 0.00227 | 67726 |  | XCAMS | EA | 0.17% |  | - interesting fact to note that Kauri (TP67855) has roughly 10x times more beam intensity than Ox-I (TP67860), Ox-I (TP657844) and FIRI-I (TP67847) and all are ~ 0.05mgC -It is noted that there is a correlation in rts differences based on the two TP grou | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-04 | Water |  |  | Took some of the acid etched and ground material from stock bottle and labelled as R14047/11 | 58945 | Normal | 14047/11 |  |  | 2015-04-14 | 1866 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99785 | 0.00227 | 67727 | [25 empty] | 2015-04-14 | 1224 | Gas Splitting | Completed
[168 empty] | 1.0019 | 0.00225 | 67894 | [25 empty] | 2015-04-14 | 2431 | Choose Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00225 | 67895 | [25 empty] | 2015-04-17 | 7791 | Graphite | Completed
[168 empty] | 1.0025 | 0.00227 | 67898 | [25 empty] | 2015-04-20 | 3622 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00224 | 67899 | [25 empty] | 2015-05-12 | 9608 | AMS Submission Results | Completed
[196 empty] | 2015-05-12 | 11479 | Close Job | Completed
[7 empty] | 111.6 | 89.2316 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40637 | 40637 | 40637 | 184 | 184 | 184 |  |  | jenny dahl | -993.65 | 0.15 ‰ | 2015-05-12 | 2015-03-18 | 11 | 55 |  |  |  |  | -2.30‰ | -2.30‰ | 2.00‰ | 2.00‰ | -2.30 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -993.65 ‰ | -993.65 ‰ | 0.15 | 0.15 | -993.57 | 0.15 ‰ | -993.57 | 0.15 ‰ |  |  | 2015-05-14 | Jocelyn Turnbull | 10:57 | 0.0063531 | 0.0063531 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.2 |  |  | 0.00635 | 0.00015 | 9681 | 2015-05-13 | 10:47 | Johannes Kaiser | 202156 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.64 | 0.01% | 17 | 5 | 6 | ... | 0.00611 | 0.00014 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes | [5 empty] | 0.00611 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00635 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00611 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-08 | jenny dahl | 16:07 |  | Completed |  |  |  |  | 67849 | 2000000 |  |  | 67885 67886 67887 67888 67914 67918 | 67885~67886~67887~67888~67914~67918 | 3003 |  | 3003 3003 3003 3003 3003 3003 | 2015-05-13 | 12:01 | Christine Prior |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99931 | 0.00167 | 67885 |  | XCAMS | ST | 0% |  | - good measurement - rejected first rotation from reporting - cat# 16 (TP67877) needed an additional factor of 1.3 to explain its individual run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-25 | Water |  |  | water line nitrogen flush water vessel | 58985 | Normal | 14047/11 |  |  | 2015-04-30 | 1912 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.00168 | 67886 | [25 empty] | 2015-04-30 | 1236 | Gas Splitting | Completed
[168 empty] | 0.99865 | 0.00168 | 67887 | [25 empty] | 2015-04-30 | 2445 | Choose Gas Splitting | Completed
[168 empty] | 0.99929 | 0.0017 | 67888 | [25 empty] | 2015-05-01 | 7859 | Graphite | Completed
[168 empty] | 1.001 | 0.0017 | 67914 | [25 empty] | 2015-05-05 | 3643 | C13 Measurement | Completed
[168 empty] | 0.99935 | 0.0017 | 67918 | [25 empty] | 2015-05-14 | 9681 | AMS Submission Results | Completed
[196 empty] | 2015-05-14 | 11527 | Close Job | Completed
[7 empty] | 123.3 | 89.1526 | 0.004 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 30014 | 30013 | 30014 | 81 | 80 | 81 | 29933.4 | 30094.6 | Margaret Norris | -976.16 | 0.24 ‰ | 2015-05-13 | 2015-05-05 | 26 | 8 |  |  |  |  | 0.84‰ | 0.84‰ | 2.00‰ | 2.00‰ | 0.84 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.0084 | 0 | 0.8165 | -976.16 ‰ | -976.16 ‰ | 0.24 | 0.24 | -975.86 | 0.24 ‰ | -975.86 | 0.24 ‰ |  |  | 2015-05-19 | Jocelyn Turnbull | 11:32 | 0.023842 | 0.023842 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.4 |  |  | 0.02384 | 0.00024 | 9716 | 2015-05-15 | 9:54 | Johannes Kaiser | 202041 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 2.38 | 0.02% | 19 | 5 | 7 | ... | 0.02293 | 0.00023 |  |  |  |  | 1.0027 | 1.0027 | Yes | Yes |  | Dreadful blank!! no obvious explanation: CO2 production and graphitization look okay. In fact, water blank graphitised on same day in next reactor as job 201994 came out great. Decided this was anomalous enough to reject as blank and change category |  |  |  | 0.02293 | 0.02‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.02384 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.02293 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-05-12 | Margaret Norris | 13:54 |  | Completed |  |  |  |  | 67782 | 2000000 |  |  | 67896 67897 67907 67908 67930 67931 | 67896~67897~67907~67908~67930~67931 | 3004 |  | 3004 3004 3004 3004 3004 3004 | 2015-05-18 | 12:33 | Christine Prior |  | 0.83 mg |  | 0.83 | [5 empty] | 0.99798 | 0.0015 | 67896 |  | XCAMS | EA | 0% |  | #NAME? | 2.3 | RRL | [5 empty] | Background Test |  | 2015-06-10 | Water |  |  | test status due to CRA of ~30014/81 | 59025 | Normal | 14047/11 |  |  | 2015-04-15 | 1882 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00148 | 67897 | [25 empty] | 2015-04-15 | 1226 | Gas Splitting | Cancelled
[168 empty] | 1.0018 | 0.00149 | 67907 | [25 empty] | 2015-04-15 | 1227 | Gas Splitting | Completed
[168 empty] | 0.9999 | 0.00152 | 67908 | [25 empty] | 2015-04-15 | 2435 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00161 | 67930 | [25 empty] | 2015-04-17 | 7792 | Graphite | Completed
[168 empty] | 1.0002 | 0.00171 | 67931 | [25 empty] | 2015-04-20 | 3624 | C13 Measurement | Completed
[196 empty] | 2015-05-19 | 9716 | AMS Submission Results | Completed
[196 empty] | 2015-06-02 | 11623 | Close Job | Completed
[7 empty] | 135.2 | 88.8502 | 0.0003 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 52082 | 52081 | 52082 | 328 | 327 | 328 |  |  | Margaret Norris | -998.47 | 0.06 ‰ | 2015-06-15 | 2015-06-02 | 5 | 13 |  |  |  |  | 2.12‰ | 2.12‰ | 2.00‰ | 2.00‰ | 2.12 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -998.47 ‰ | -998.47 ‰ | 0.06 | 0.06 | -998.45 | 0.06 ‰ | -998.45 | 0.06 ‰ |  |  | 2015-06-18 | Jocelyn Turnbull | 12:22 | 0.0015285 | 0.0015285 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 2.2 |  |  | 0.00153 | 6e-05 | 9893 | 2015-06-16 | 11:23 | Johannes Kaiser | 202157 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:42 | Jocelyn Turnbull | 0 |  | 0.15 | 0.01% | 3 | 5 | 6 | ... | 0.00147 | 6e-05 |  |  |  |  | 1.0005 | 1.0005 | Yes | Yes |  | Great blank. Probably unrealistically good. Will average other recent water blanks. |  |  |  | 0.00147 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00153 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00147 | 0 | 6e-05 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-06-11 | Margaret Norris | 16:29 |  | Completed |  |  |  |  | 68127 | 2000000 |  |  | 68120 68121 68144 68145 68146 68147 | 68120~68121~68144~68145~68146~68147 | 3009 |  | 3009 3009 3009 3009 3009 3009 |  |  |  |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99907 | 0.00162 | 68120 |  | XCAMS | EA | 0% |  | - good measurement - need to add a global factor of 1.02 to account for global run-to-run variance | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-06-25 | Water | Complete |  | water line nitrogen flush water vessel | 59231 | Normal | 14047/11 |  |  | 2015-06-09 | 1966 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99998 | 0.00161 | 68121 | [25 empty] | 2015-06-09 | 1257 | Gas Splitting | Completed
[168 empty] | 0.99981 | 0.00158 | 68144 | [25 empty] | 2015-06-09 | 2480 | Choose Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00158 | 68145 | [25 empty] | 2015-06-10 | 8137 | Graphite | Completed
[168 empty] | 0.99851 | 0.00158 | 68146 | [25 empty] | 2015-06-15 | 3716 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00157 | 68147 | [25 empty] | 2015-06-18 | 9893 | AMS Submission Results | Completed
[196 empty] | 2015-06-18 | 11783 | Close Job | Completed
[7 empty] | 80.2 | 89.1527 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43535 | 43534 | 43535 | 264 | 263 | 264 |  |  | Margaret Norris | -995.57 | 0.15 ‰ | 2015-08-11 | 2015-05-03 | 13 | 100 |  |  |  |  | 2.95‰ | 2.95‰ | 2.00‰ | 2.00‰ | 2.95 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -995.57 ‰ | -995.57 ‰ | 0.15 | 0.15 | -995.52 | 0.15 ‰ | -995.52 | 0.15 ‰ |  |  | 2015-08-13 | Jocelyn Turnbull | 15:20 | 0.0044295 | 0.0044295 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00443 | 0.00015 | 10318 | 2015-08-13 | 11:29 | Albert Zondervan | 202675 |  |  |  | 0 | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.44 | 0.01% | 32 | 5 | 6 | ... | 0.00426 | 0.00014 |  |  |  |  | 1.0048 | 1.0048 | Yes | Yes | [5 empty] | 0.00426 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00443 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00426 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-10 | Margaret Norris | 10:53 |  | Completed |  |  |  |  | 68516 | 2000000 |  |  | 68544 68545 68547 68548 68549 68550 | 68544~68545~68547~68548~68549~68550 | 3020 |  | 3020 3020 3020 3020 3020 3020 | 2015-08-13 | 14:29 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99999 | 0.0019 | 68544 |  | XCAMS | ST | 0% |  | - required quality achieved through close-up scrutiny and adjustment - very last TW measurement before SNICS source clean - error adjustments only needed for run-to-run repeatability: global factor 1.04 and cathode-specific factor 1.3 for pos#13 and #36 - | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-09-09 | Water |  |  | water line nitrogen flush water vessel | 59629 | Normal | 14047/11 |  |  | 2015-07-27 | 2027 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00203 | 68545 | [25 empty] | 2015-07-27 | 1279 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00198 | 68547 | [25 empty] | 2015-07-27 | 2516 | Choose Gas Splitting | Completed
[168 empty] | 0.99837 | 0.00197 | 68548 | [25 empty] | 2015-07-29 | 8526 | Graphite | Completed
[168 empty] | 1.0015 | 0.00196 | 68549 | [25 empty] | 2015-08-03 | 3848 | C13 Measurement | Completed
[168 empty] | 0.99767 | 0.00199 | 68550 | [25 empty] | 2015-08-13 | 10318 | AMS Submission Results | Completed
[196 empty] | 2015-08-13 | 12244 | Close Job | Completed
[7 empty] | 124.3 | 88.7803 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47109 | 47108 | 47109 | 294 | 294 | 294 |  |  | jenny dahl | -997.16 | 0.10 ‰ | 2015-08-27 | 2015-08-27 | 20 | [5 empty] | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -997.16 ‰ | -997.16 ‰ | 0.1 | 0.1 | -997.13 | 0.11 ‰ | -997.13 | 0.11 ‰ |  |  | 2015-09-01 | jenny dahl | 17:11 | 0.0028386 | 0.0028386 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00284 | 0.0001 | 10437 | 2015-08-31 | 9:18 | Albert Zondervan | 202777 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 24 | 5 | 4 | ... | 0.00273 | 0.0001 |  |  |  |  | 1.0001 | 1.0001 | Yes | Yes | [5 empty] | 0.00273 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00284 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00273 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-08-26 | jenny dahl | 16:47 |  | Completed |  |  |  |  | 68609 | 2000000 |  |  | 68551 68552 68652 68655 68656 68657 | 68551~68552~68652~68655~68656~68657 | 3023 |  | 3023 3023 3023 3023 3023 3023 | 2015-08-31 | 11:53 | Albert Zondervan |  | 1.21 mg |  | 1.21 | [5 empty] | 1.0009 | 0.00165 | 68551 |  | XCAMS | ST | 0% |  | - good measurement overall - global uncertainty multiplier is 1.05 to account for excess run-to-run variability - no system error - some questions arose regarding time-patterns of cathodes with (near-) blank 14C/C ratio values: * usual surface contaminati | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-09-25 | Water |  |  | water line nitrogen flush water vessel | 59775 | Normal | 14047/11 |  |  | 2015-08-05 | 2058 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00166 | 68552 | [25 empty] | 2015-08-07 | 8619 | Graphite | Completed
[168 empty] | 1 | 0.00164 | 68652 | [25 empty] | 2015-08-11 | 3869 | C13 Measurement | Completed
[168 empty] | 0.99921 | 0.00165 | 68655 | [25 empty] | 2015-09-01 | 10437 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00163 | 68656 | [25 empty] | 2015-09-03 | 12432 | Close Job | Completed
[168 empty] | 0.99773 | 0.00177 | 68657
[7 empty] | 124.2 | 89.008 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44327 | 44326 | 44327 | 229 | 228 | 229 |  |  | jenny dahl | -995.99 | 0.11 ‰ | 2015-09-17 | 2015-03-21 | 9 | 180 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.56‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -995.99 ‰ | -995.99 ‰ | 0.11 | 0.11 | -995.94 | 0.12 ‰ | -995.94 | 0.12 ‰ |  |  | 2015-09-21 | jenny dahl | 13:06 | 0.0040136 | 0.0040136 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.9 |  |  | 0.00401 | 0.00011 | 10565 | 2015-09-18 | 13:24 | john west | 202831 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 4 | 5 | 6 | ... | 0.00386 | 0.00011 |  |  |  |  | 1.0031 | 1.0031 | Yes | Yes | [5 empty] | 0.00386 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00401 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00386 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-11 | jenny dahl | 16:10 |  | Completed |  |  |  |  | 68799 | 2000000 |  |  | 68752 68755 68820 68821 68845 68846 | 68752~68755~68820~68821~68845~68846 | 3026 |  | 3026 3026 3026 3026 3026 3026 | 2015-09-18 | 14:41 | john west |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0015 | 0.00185 | 68752 |  | XCAMS | EA | 0% |  | - 1st measurement after tank opening - first 3 rotations rejected due to gas-stripper stabilisation - only error adjustment is on cat#7, which is an Ox-I | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-10-07 | Water |  |  | water line nitrogen flush water vessel | 59868 | Normal | 14047/11 |  |  | 2015-09-04 | 2106 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00236 | 68755 | [25 empty] | 2015-09-04 | 1304 | Gas Splitting | Completed
[168 empty] | 0.99824 | 0.00177 | 68820 | [25 empty] | 2015-09-04 | 2561 | Choose Gas Splitting | Completed
[168 empty] | 0.99697 | 0.00177 | 68821 | [25 empty] | 2015-09-08 | 8809 | Graphite | Completed
[168 empty] | 1.0009 | 0.0018 | 68845 | [25 empty] | 2015-09-14 | 3932 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.0018 | 68846 | [25 empty] | 2015-09-21 | 10565 | AMS Submission Results | Completed
[196 empty] | 2015-09-21 | 12580 | Close Job | Completed
[7 empty] | 90.8 | 89.7276 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37168 | 37168 | 37168 | 145 | 145 | 145 |  |  | jenny dahl | -990.22 | 0.18 ‰ | 2015-09-30 | 2015-04-03 | 2 | 180 |  |  |  |  | 1.23‰ | 1.23‰ | 2.00‰ | 2.00‰ | 1.23 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -990.22 ‰ | -990.22 ‰ | 0.18 | 0.18 | -990.09 | 0.18 ‰ | -990.09 | 0.18 ‰ |  |  | 2015-10-08 | jenny dahl | 11:44 | 0.0097845 | 0.0097845 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.3 |  |  | 0.00978 | 0.00018 | 10678 | 2015-10-01 | 12:54 | john west | 203037 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.98 | 0.02% | 30 | 5 | 6 | ... | 0.00941 | 0.00017 |  |  |  |  | 1.0095 | 1.0095 | Yes | Yes |  | between the horrid CRA and the notes from John about measurement we have decided to discard this sample as a blank. the secondary done same day is a bit off in CRA but still in keeping with past samples and didn't have the measurement error so believe it was not an effect in the waterline- 8/10/15-jd Measured RTS > 4x other blanks consistent results throughout measurement |  |  |  | 0.00941 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00074 | 0.00074 | 0.00978 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00941 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-09-28 | jenny dahl | 14:47 |  | Completed |  |  |  |  | 68963 | 2000000 |  |  | 68916 68917 68918 68919 68940 68946 | 68916~68917~68918~68919~68940~68946 | 3029 |  | 3029 3029 3029 3029 3029 3029 |  |  |  |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0019 | 0.00183 | 68916 |  | XCAMS | ST | 0% |  | - 1st measurement aborted after 6 rotations - 2nd measurement finished after 15 rotations - good measurement, no error adjustments required - IAEA-C1 cathode (TP68963) gave high result (>4x other blanks) consistently throughout measurement | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-11-09 | Water | Complete |  | dreadful blank, nothing unusual about its processing, have changed category so it doesn't come up in blank selection. John has alerted us that this sample is behaving badly in measurement. In looking through the evolution, splitting and graphite I can't see any issues. material was freshly etched and ground-jd From job 203101 | 59947 | Normal | 14047/11 |  |  | 2015-09-25 | 2151 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99909 | 0.00179 | 68917 | [25 empty] | 2015-09-25 | 1318 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00181 | 68918 | [25 empty] | 2015-09-25 | 2583 | Choose Gas Splitting | Completed
[168 empty] | 0.99896 | 0.00181 | 68919 | [25 empty] | 2015-09-28 | 8973 | Graphite | Completed
[168 empty] | 0.99698 | 0.00189 | 68940 | [25 empty] | 2015-10-06 | 3976 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.00181 | 68946 | [25 empty] | 2015-10-08 | 10678 | AMS Submission Results | Completed
[196 empty] | 2015-10-08 | 12635 | Close Job | Completed
[7 empty] | 92.8 | 89.2758 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39545 | 39545 | 39545 | 172 | 172 | 172 |  |  | jenny dahl | -992.72 | 0.16 ‰ | 2015-10-16 | 2015-10-16 | 8 | 0 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.8165 | -992.72 ‰ | -992.72 ‰ | 0.16 | 0.16 | -992.63 | 0.16 ‰ | -992.63 | 0.16 ‰ |  |  | 2015-10-19 | Jocelyn Turnbull | 15:16 | 0.0072786 | 0.0072786 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00728 | 0.00016 | 10767 | 2015-10-18 | 14:25 | Albert Zondervan | 203038 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.73 | 0.02% | 34 | 5 | 6 | ... | 0.007 | 0.00015 |  |  |  |  | 1.008 | 1.008 | Yes | Yes |  | Ugh! But fits with other recent waters. |  |  |  | 0.007 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0008 | 0.0008 | 0.00728 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.007 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-12 | jenny dahl | 10:11 |  | Completed |  |  |  |  | 69032 | 2000000 |  |  | 68974 68977 68988 68989 69028 69029 | 68974~68977~68988~68989~69028~69029 | 3031 |  | 3031 3031 3031 3031 3031 3031 | 2015-10-19 | 14:30 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 0.9991 | 0.00179 | 68974 |  | XCAMS | EA | 0% |  | * good measurement TW3031_2 * all cathodes ran well | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2015-11-09 | Water |  |  | water line nitrogen flush water vessel | 60052 | Normal | 14047/11 |  |  | 2015-10-07 | 2171 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99801 | 0.00187 | 68977 | [25 empty] | 2015-10-07 | 1323 | Gas Splitting | Completed
[168 empty] | 1.0039 | 0.00194 | 68988 | [25 empty] | 2015-10-07 | 2591 | Choose Gas Splitting | Completed
[168 empty] | 0.99862 | 0.00188 | 68989 | [25 empty] | 2015-10-08 | 9042 | Graphite | Completed
[168 empty] | 0.99989 | 0.00189 | 69028 | [25 empty] | 2015-10-16 | 3999 | C13 Measurement | Completed
[168 empty] | 1.0016 | 0.0026 | 69029 | [25 empty] | 2015-10-19 | 10767 | AMS Submission Results | Completed
[196 empty] | 2015-10-19 | 12761 | Close Job | Completed
[7 empty] | 84.5 | 89.4909 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39957 | 39957 | 39957 | 217 | 217 | 217 |  |  | jenny dahl | -993.09 | 0.19 ‰ | 2015-10-22 | 2015-07-24 | 6 | 90 |  |  |  |  | -0.25‰ | -0.25‰ | 2.00‰ | 2.00‰ | -0.25 ‰ | 2.00 ‰ | AMS | 2.55‰ | 0.2‰ | C13 | No | -19 | -19.0027 | 0 | 0.8165 | -993.09 ‰ | -993.09 ‰ | 0.19 | 0.19 | -993 | 0.19 ‰ | -993 | 0.19 ‰ |  |  | 2015-10-23 | Jocelyn Turnbull | 14:02 | 0.0069146 | 0.0069146 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.1 |  |  | 0.00691 | 0.00019 | 10844 | 2015-10-23 | 10:48 | Albert Zondervan | 203172 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.69 | 0.02% | 36 | 5 | 6 | ... | 0.00665 | 0.00018 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | Water blank still poor - see if improves from change in carbasorb just after this wheel |  |  |  | 0.00665 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00078 | 0.00078 | 0.00691 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00665 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-10-19 | jenny dahl | 14:19 |  | Completed |  |  |  |  | 69113 | 2000000 |  |  | 68920 68949 69087 69095 69096 69097 | 68920~68949~69087~69095~69096~69097 | 3033 |  | 3033 3033 3033 3033 3033 3033 | 2015-10-23 | 13:46 | Jocelyn Turnbull |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99947 | 0.00188 | 68920 |  | XCAMS | ST | 0% |  | * straightforward measurement * only 1 error adjustment: for TP69103 only | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-11-30 | Water |  |  | water line nitrogen flush water vessel | 60145 | Normal | 14047/11 |  |  | 2015-10-13 | 2183 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0016 | 0.0019 | 68949 | [25 empty] | 2015-10-13 | 1326 | Gas Splitting | Completed
[168 empty] | 0.99991 | 0.00199 | 69087 | [25 empty] | 2015-10-13 | 2599 | Choose Gas Splitting | Completed
[168 empty] | 0.99915 | 0.00188 | 69095 | [25 empty] | 2015-10-16 | 9123 | Graphite | Completed
[168 empty] | 1.001 | 0.00187 | 69096 | [25 empty] | 2015-10-22 | 4027 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00191 | 69097 | [25 empty] | 2015-10-23 | 10844 | AMS Submission Results | Completed
[196 empty] | 2015-10-23 | 12871 | Close Job | Completed
[7 empty] | 94.9 | 89.0656 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47348 | 47347 | 47348 | 303 | 303 | 303 |  |  | jenny dahl | -997.24 | 0.10 ‰ | 2015-11-10 | 2015-11-10 | 19 | 0 |  |  |  |  | 3.69‰ | 3.69‰ | 2.00‰ | 2.00‰ | 3.69 ‰ | 2.00 ‰ | AMS | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.0086 | 0 | 0.8165 | -997.24 ‰ | -997.24 ‰ | 0.1 | 0.1 | -997.21 | 0.11 ‰ | -997.21 | 0.11 ‰ |  |  | 2015-11-12 | Jocelyn Turnbull | 16:22 | 0.0027555 | 0.0027555 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00276 | 0.0001 | 10958 | 2015-11-12 | 13:56 | Albert Zondervan | 203173 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:43 | Jocelyn Turnbull | 0 |  | 0.28 | 0.01% | 30 | 5 | 6 | ... | 0.00265 | 0.0001 |  |  |  |  | 1.0045 | 1.0045 | Yes | Yes |  | A really good water blank. Carbasorb on N2 tank was changed day defore this blank evolved. Not certain yet that was the problem, but certainly looks promising! |  |  |  | 0.00265 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00276 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00265 | 0 | 0.0001 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-11-09 | jenny dahl | 10:09 |  | Completed |  |  |  |  | 69144 | 2000000 |  |  | 69251 69252 69253 69254 69255 69256 | 69251~69252~69253~69254~69255~69256 | 3035 |  | 3035 3035 3035 3035 3035 3035 | 2015-11-12 | 15:49 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99878 | 0.00149 | 69251 |  | XCAMS | ST | 0% |  | - 1st wheel after source clean - blanks show clear sign of surface contamination; only later runs have been retained - mode 5 calibration once again vindicates over the other modes in that returns the best possible repeatability amongst the Ox-I cathodes | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2015-11-30 | Water |  |  | A really good water blank. Carbasorb on N2 tank was changed day before this blank evolved. Not certain yet that was the problem, but certainly looks promising! | 60254 | Normal | 14047/11 |  |  | 2015-10-21 | 2205 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00162 | 69252 | [25 empty] | 2015-10-21 | 1329 | Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00178 | 69253 | [25 empty] | 2015-10-21 | 2604 | Choose Gas Splitting | Completed
[168 empty] | 0.9992 | 0.00156 | 69254 | [25 empty] | 2015-10-22 | 9154 | Graphite | Completed
[168 empty] | 1.0006 | 0.00156 | 69255 | [25 empty] | 2015-10-27 | 4045 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00157 | 69256 | [25 empty] | 2015-11-12 | 10958 | AMS Submission Results | Completed
[196 empty] | 2015-11-13 | 12998 | Close Job | Completed
[7 empty] | 114 | 89.1509 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37254 | 37254 | 37254 | 155 | 155 | 155 |  |  | jenny dahl | -990.32 | 0.19 ‰ | 2015-12-09 | 2015-12-09 | 28 | 0 |  |  |  |  | 3.63‰ | 3.63‰ | 2.00‰ | 2.00‰ | 3.63 ‰ | 2.00 ‰ | AMS | 2.52‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -990.32 ‰ | -990.32 ‰ | 0.19 | 0.19 | -990.2 | 0.19 ‰ | -990.2 | 0.19 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.0096805 | 0.0096805 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.7 |  |  | 0.00968 | 0.00019 | 11094 | 2015-12-12 | 14:02 | Albert Zondervan | 203558 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.97 | 0.02% | 5 | 5 | 6 | ... | 0.00931 | 0.00018 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes | [5 empty] | 0.00931 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00968 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00931 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:15 |  | Completed |  |  |  |  | 69274 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2015-12-22 | Water |  |  | freshly etched material water line nitrogen flush water vessel | 60429 | Normal | 14047/11 |  |  | 2015-11-09 | 2226 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [25 empty] | 2015-11-09 | 1337 | Gas Splitting | Completed
[168 empty] | 0.99975 | 0.00176 | 69370 | [25 empty] | 2015-11-09 | 2622 | Choose Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [25 empty] | 2015-11-11 | 9284 | Graphite | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [25 empty] | 2015-11-18 | 4072 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [25 empty] | 2015-12-17 | 11094 | AMS Submission Results | Completed
[196 empty] | 2015-12-21 | 13284 | Close Job | Completed
[7 empty] | 112.6 | 89.2815 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34548 | 34547 | 34548 | 123 | 123 | 123 | 34425.3 | 34671.8 | jenny dahl | -986.44 | 0.21 ‰ | 2015-12-09 | 2015-12-09 | 14 | 0 |  |  |  |  | 2.16‰ | 2.16‰ | 2.00‰ | 2.00‰ | 2.16 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -986.44 ‰ | -986.44 ‰ | 0.21 | 0.21 | -986.27 | 0.21 ‰ | -986.27 | 0.21 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.013559 | 0.013559 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2.3 |  |  | 0.01356 | 0.00021 | 11119 | 2015-12-12 | 14:02 | Albert Zondervan | 203633 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.36 | 0.02% | 29 | 5 | 8 | ... | 0.01304 | 0.0002 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | seems way too high F |  |  |  | 0.01304 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.01356 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01304 | 0 | 0.0002 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:17 |  | Completed |  |  |  |  | 69371 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.80 mg |  | 0.8 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Test |  | 2016-01-06 | Water | Complete |  | Test of new small water vessel No job created for second c13 split failed | 60454 | Normal | 14047/11 |  |  | 2015-11-23 | 2230 | Graphite | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [25 empty] | 2015-11-23 | 2241 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99975 | 0.00176 | 69370 | [25 empty] | 2015-11-24 | 2245 | Water CO2 Evolution | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [25 empty] | 2015-11-24 | 1345 | Gas Splitting | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [25 empty] | 2015-11-24 | 2635 | Choose Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [25 empty] | 2015-11-25 | 9381 | Graphite | Completed
[196 empty] | 2015-11-27 | 4092 | C13 Measurement | Completed
[196 empty] | 2015-12-17 | 11119 | AMS Submission Results | Completed
[196 empty] | 2015-12-17 | 13262 | Close Job | Completed
[7 empty] | 112.5 | 89.2114 | 0.0015 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37753 | 37752 | 37753 | 156 | 156 | 156 | 37597.9 | 37910.1 | jenny dahl | -990.9 | 0.18 ‰ | 2015-12-09 | 2015-12-09 | 15 | 0 |  |  |  |  | 2.95‰ | 2.95‰ | 2.00‰ | 2.00‰ | 2.95 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -990.90 ‰ | -990.90 ‰ | 0.18 | 0.18 | -990.79 | 0.18 ‰ | -990.79 | 0.18 ‰ |  |  | 2015-12-17 | Margaret Norris | 14:56 | 0.0090982 | 0.0090982 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.8 |  |  | 0.0091 | 0.00018 | 11127 | 2015-12-12 | 14:02 | Albert Zondervan | 203670 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.91 | 0.02% | 37 | 5 | 6 | ... | 0.00875 | 0.00017 |  |  |  |  | 1.0054 | 1.0054 | Yes | Yes |  | seems too high F |  |  |  | 0.00875 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0091 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00875 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2015-12-07 | jenny dahl | 10:17 |  | Completed |  |  |  |  | 69358 | 2000000 |  |  | 69359 69362 69370 69375 69378 69381 | 69359~69362~69370~69375~69378~69381 | 3040 |  | 3040 3040 3040 3040 3040 3040 | 2015-12-12 | 17:52 | Albert Zondervan |  | 0.71 mg |  | 0.71 | [5 empty] | 0.99931 | 0.00169 | 69359 |  | XCAMS | EA | 0% |  | - straightforward measurement - uncertainty correction terms: * system error = 0% * all Poisson errors increased by factor 1.03 due to excess run-to-run variability * additionally, Poisson error of cat#30 (TP69382) increased by factor 1.3 due to | 2.5 | RRL | [5 empty] | Background Test |  | 2016-01-15 | Water | Complete |  | test of new small water vessel | 60462 | Normal | 14047/11 |  |  | 2015-11-20 | 2237 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99933 | 0.0018 | 69362 | [25 empty] | 2015-11-20 | 1341 | Gas Splitting | Completed
[168 empty] | 0.99975 | 0.00176 | 69370 | [25 empty] | 2015-11-20 | 2629 | Choose Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00169 | 69375 | [25 empty] | 2015-11-24 | 9368 | Graphite | Completed
[168 empty] | 1.0022 | 0.00171 | 69378 | [25 empty] | 2015-11-27 | 4090 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00167 | 69381 | [25 empty] | 2015-12-17 | 11127 | AMS Submission Results | Completed
[196 empty] | 2015-12-17 | 13263 | Close Job | Completed
[7 empty] | 104.9 | 88.8485 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37826 | 37826 | 37826 | 138 | 138 | 138 |  |  | jenny dahl | -990.98 | 0.16 ‰ | 2016-01-27 | 2016-01-27 | 8 | 0 |  |  |  |  | 5.25‰ | 5.25‰ | 2.00‰ | 2.00‰ | 5.25 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.006 | 0 | 0.8165 | -990.98 ‰ | -990.98 ‰ | 0.16 | 0.16 | -990.87 | 0.16 ‰ | -990.87 | 0.16 ‰ |  |  | 2016-02-01 | Jocelyn Turnbull | 9:48 | 0.009015 | 0.009015 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00902 | 0.00016 | 11188 | 2016-01-28 | 13:18 | Albert Zondervan | 203968 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.9 | 0.02% | 36 | 5 | 7 | ... | 0.00867 | 0.00015 |  |  |  |  | 1.0036 | 1.0036 | Yes | Yes |  | clearly another poor result for the water line blank |  |  |  | 0.00867 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1.0001 | 0.00098 | 0.00098 | 0.00902 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00867 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-01-22 | jenny dahl | 12:24 |  | Completed |  |  |  |  | 69454 | 2000000 |  |  | 69441 69442 69465 69471 69472 69473 | 69441~69442~69465~69471~69472~69473 | 3042 |  | 3042 3042 3042 3042 3042 3042 | 2016-01-28 | 16:51 | Albert Zondervan |  | 1.00 mg |  | 1 | [5 empty] | 1.0038 | 0.00251 | 69441 |  | XCAMS | EA | 0.16% |  | * good measurement * non-zero system error probably incurred by somewhat differently behaving (lower beam current) Ox-I in pos#0 | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-03-10 | Water |  |  |  | 60523 | Normal | 14047/11 |  |  | 2016-01-15 | 2280 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99708 | 0.00233 | 69442 | [25 empty] | 2016-01-18 | 2284 | Water CO2 Evolution | Completed
[168 empty] | 0.99826 | 0.00234 | 69465 | [25 empty] | 2016-01-18 | 1382 | Gas Splitting | Completed
[168 empty] | 0.99991 | 0.00233 | 69471 | [25 empty] | 2016-01-18 | 2674 | Choose Gas Splitting | Completed
[168 empty] | 1 | 0.00271 | 69472 | [25 empty] | 2016-01-19 | 9464 | Graphite | Completed
[168 empty] | 1.0017 | 0.00231 | 69473 | [25 empty] | 2016-01-28 | 4135 | C13 Measurement | Completed
[196 empty] | 2016-02-01 | 11188 | AMS Submission Results | Completed
[196 empty] | 2016-02-04 | 13357 | Close Job | Completed
[7 empty] | 120.2 | 89.5649 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33552 | 33552 | 33552 | 125 | 125 | 125 |  |  | Margaret Norris | -984.65 | 0.24 ‰ | 2016-02-11 | 2016-02-11 | 6 | 0 |  |  |  |  | -2.23‰ | -2.23‰ | 2.00‰ | 2.00‰ | -2.23 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -984.65 ‰ | -984.65 ‰ | 0.24 | 0.24 | -984.46 | 0.24 ‰ | -984.46 | 0.24 ‰ |  |  | 2016-02-15 | Jocelyn Turnbull | 14:13 | 0.015347 | 0.015347 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.5 |  |  | 0.01535 | 0.00024 | 11308 | 2016-02-15 | 10:38 | Silvia Canessa | 203944 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.53 | 0.02% | 36 | 5 | 7 | ... | 0.01476 | 0.00023 |  |  |  |  | 1.0042 | 1.0042 | Yes | Yes | [5 empty] | 0.01476 | 0.01‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01535 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01476 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-05 | Margaret Norris | 15:23 |  | Completed |  |  |  |  | 69597 | 2000000 |  |  | 69499 69500 69501 69527 69528 69545 | 69499~69500~69501~69527~69528~69545 | 3045 |  | 3045 3045 3045 3045 3045 3045 | 2016-02-15 | 11:57 | Silvia Canessa |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0007 | 0.00157 | 69499 |  | XCAMS | EA | 0% |  | * good measurement: - problem with extractor resolved beforehand - high beam currents achieved by pushing source - stable ratios - no error adjustments | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-03-07 | Water |  |  | memory effect test prepared on the 3/2/16 | 60606 | Normal | 14047/11 |  |  | 2016-02-03 | 2313 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00159 | 69500 | [25 empty] | 2016-02-04 | 2315 | Water CO2 Evolution | Completed
[168 empty] | 0.99997 | 0.00159 | 69501 | [25 empty] | 2016-02-04 | 1388 | Gas Splitting | Completed
[168 empty] | 0.99838 | 0.00156 | 69527 | [25 empty] | 2016-02-04 | 2690 | Choose Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00155 | 69528 | [25 empty] | 2016-02-05 | 9607 | Graphite | Completed
[168 empty] | 0.99997 | 0.00159 | 69545 | [25 empty] | 2016-02-10 | 4178 | C13 Measurement | Completed
[196 empty] | 2016-02-15 | 11308 | AMS Submission Results | Completed
[196 empty] | 2016-02-18 | 13472 | Close Job | Completed
[7 empty] | 160.2 | 90.1592 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38871 | 38871 | 38871 | 134 | 134 | 134 |  |  | Margaret Norris | -992.08 | 0.13 ‰ | 2016-02-24 | 2015-08-28 | 8 | 180 | [8 empty] | 2.59 ‰ | 0.20 ‰ | C13 | 2.59‰ | 0.2‰ | C13 | No | -19 | -19.455 | 0.08 | 0.08165 | -992.08 ‰ | -992.08 ‰ | 0.13 | 0.13 | -991.99 | 0.13 ‰ | -991.99 | 0.13 ‰ |  |  | 2016-03-07 | Jocelyn Turnbull | 11:46 | 0.0079154 | 0.0079154 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2.3 |  |  | 0.00792 | 0.00013 | 11389 | 2016-03-05 | 13:26 | Albert Zondervan | 204024 |  |  |  |  | 45% | 1 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 1 |  | 0.79 | 0.01% | 37 | 5 | 6 | ... | 0.00778 | 0.00013 |  |  |  |  | 1.0181 | 1.0181 | Yes | Yes |  | another poor result for the water-line blank |  |  |  | 0.00778 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00792 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00778 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-19 | Margaret Norris | 16:34 |  | Completed |  |  |  |  | 69662 | 2000000 |  |  | 69476 69629 69630 69631 69688 69700 | 69476~69629~69630~69631~69688~69700 | 3046 |  | 3046 3046 3046 3046 3046 3046 | 2016-03-07 | 11:37 | Jocelyn Turnbull |  | 0.91 mg |  | 0.91 | [5 empty] | 0.99893 | 0.00156 | 69476 |  | XCAMS | EA | 0% |  | * CalAMS mode 1 was used to avoid 0.2% system error. * The full fractionation correction with mode 5 would further reduce run-to-run variability, but worsen inter-cathode variability. * The beam currents were good, but started to drop off early for the 0.5-0.6 mg C cathodes. This is another demonstration that something in the ion source blocks us to achieve high counting precision with 0.5 mg C samples. * The problems with too sharp sputter crater and lack of inter-cathode repeatability came to a head with next wheel TW3047. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-17 | Water | Complete |  | Mode 1! | 60647 | Normal | 14047/11 |  |  | 2016-02-15 | 2344 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99907 | 0.00154 | 69629 | [25 empty] | 2016-02-15 | 1407 | Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00177 | 69630 | [25 empty] | 2016-02-15 | 2712 | Choose Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00185 | 69631 | [25 empty] | 2016-02-16 | 9672 | Graphite | Completed
[168 empty] | 1.0027 | 0.00175 | 69688 | [25 empty] | 2016-02-23 | 4201 | C13 Measurement | Completed
[168 empty] | 0.99866 | 0.00159 | 69700 | [25 empty] | 2016-03-07 | 11389 | AMS Submission Results | Completed
[196 empty] | 2016-03-08 | 13541 | Close Job | Completed
[7 empty] | 100.6 | 89.0645 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39219 | 39218 | 39219 | 198 | 198 | 198 |  |  | jenny dahl | -992.42 | 0.19 ‰ | 2016-03-14 | 2016-03-14 | 17 | 0 |  |  |  |  | 2.42‰ | 2.42‰ | 2.00‰ | 2.00‰ | 2.42 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -992.42 ‰ | -992.42 ‰ | 0.19 | 0.19 | -992.33 | 0.19 ‰ | -992.33 | 0.19 ‰ |  |  | 2016-03-16 | Jocelyn Turnbull | 12:58 | 0.0075801 | 0.0075801 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.4 |  |  | 0.00758 | 0.00019 | 11420 | 2016-03-15 | 14:42 | Albert Zondervan | 204023 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.76 | 0.02% | 30 | 5 | 7 | ... | 0.00729 | 0.00018 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes | [5 empty] | 0.00729 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00758 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00729 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-02-29 | jenny dahl | 11:11 |  | Completed |  |  |  |  | 69760 | 2000000 |  |  | 69673 69674 69675 69676 69677 69754 | 69673~69674~69675~69676~69677~69754 | 3048 |  | 3048 3048 3048 3048 3048 3048 | 2016-03-15 | 16:38 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 1.0034 | 0.00214 | 69673 |  | XCAMS | ST | 0% |  | * 1st measurement after 2-week investigation into poor repeatability * good measurement * no system error * factor 1.09 multiplier for global run-to-run variability, which we think is due to long storage time of graphite | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-03-17 | Water |  |  | bad result although no obvious red flags | 60726 | Normal | 14047/11 |  |  | 2016-02-19 | 2349 | Carbonate CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99701 | 0.00222 | 69674 | [25 empty] | 2016-02-23 | 2357 | Water CO2 Evolution | Completed
[168 empty] | 0.99923 | 0.00215 | 69675 | [25 empty] | 2016-02-23 | 1420 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00215 | 69676 | [25 empty] | 2016-02-23 | 2731 | Choose Gas Splitting | Completed
[168 empty] | 0.99875 | 0.00221 | 69677 | [25 empty] | 2016-02-26 | 9770 | Graphite | Completed
[168 empty] | 1.0004 | 0.00212 | 69754 | [25 empty] | 2016-03-03 | 4231 | C13 Measurement | Completed
[196 empty] | 2016-03-16 | 11420 | AMS Submission Results | Completed
[196 empty] | 2016-03-31 | 13742 | Close Job | Completed
[7 empty] | 131.2 | 88.9136 | 0.0022 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34811 | 34810 | 34811 | 133 | 133 | 133 | 34678.6 | 34944.8 | jenny dahl | -986.88 | 0.22 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 3.86‰ | 3.86‰ | 2.00‰ | 2.00‰ | 3.86 ‰ | 2.00 ‰ | AMS | 2.45‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.88 ‰ | -986.88 ‰ | 0.22 | 0.22 | -986.72 | 0.22 ‰ | -986.72 | 0.22 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013122 | 0.013122 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.3 |  |  | 0.01312 | 0.00022 | 11644 | 2016-04-05 | 15:51 | Albert Zondervan | 204312 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.31 | 0.02% | 8 | 5 | 7 | ... | 0.01262 | 0.00021 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01262 | 0.01‰ | 0.00021 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01312 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01262 | 0 | 0.00021 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69929 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.93 mg |  | 0.93 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.5 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  |  | new batch of standard prepared. test water blank | 60943 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [25 empty] | 2016-03-17 | 2404 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [25 empty] | 2016-03-17 | 1470 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [25 empty] | 2016-03-17 | 2799 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [25 empty] | 2016-03-18 | 9939 | Graphite | Completed
[196 empty] | 2016-03-22 | 4317 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11644 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13888 | Close Job | Completed
[7 empty] | 131.7 | 88.8554 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41645 | 41644 | 41645 | 193 | 193 | 193 | 41453.5 | 41840.3 | jenny dahl | -994.4 | 0.13 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.25‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -994.40 ‰ | -994.40 ‰ | 0.13 | 0.13 | -994.33 | 0.14 ‰ | -994.33 | 0.14 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.0056045 | 0.0056045 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2 |  |  | 0.0056 | 0.00013 | 11645 | 2016-04-05 | 15:51 | Albert Zondervan | 204314 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 22 | 5 | 7 | ... | 0.00539 | 0.00013 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.00539 | 0.01‰ | 0.00013 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.0056 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00539 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69925 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 1.00 mg |  | 1 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.3 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel final weight incorrectly entered- back corrected to choose correct split. graphite yield will look odd. | 60944 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [25 empty] | 2016-03-17 | 2405 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [25 empty] | 2016-03-17 | 1469 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [25 empty] | 2016-03-17 | 2797 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [25 empty] | 2016-03-18 | 9935 | Graphite | Completed
[196 empty] | 2016-03-22 | 4318 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11645 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13889 | Close Job | Completed
[7 empty] | 126 | 89.0883 | 0.0027 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33232 | 33232 | 33232 | 114 | 114 | 114 | 33118.6 | 33347.5 | jenny dahl | -984.03 | 0.23 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 1.89‰ | 1.89‰ | 2.00‰ | 2.00‰ | 1.89 ‰ | 2.00 ‰ | AMS | 1.78‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -984.03 ‰ | -984.03 ‰ | 0.23 | 0.23 | -983.83 | 0.23 ‰ | -983.83 | 0.23 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.015971 | 0.015971 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 2.5 |  |  | 0.01597 | 0.00023 | 11646 | 2016-04-05 | 15:51 | Albert Zondervan | 204315 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.6 | 0.02% | 29 | 5 | 7 | ... | 0.01536 | 0.00022 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01536 | 0.02‰ | 0.00022 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01597 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01536 | 0 | 0.00022 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69926 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.90 mg |  | 0.9 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 1.8 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60945 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [25 empty] | 2016-03-17 | 2406 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [25 empty] | 2016-03-17 | 1468 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [25 empty] | 2016-03-17 | 2795 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [25 empty] | 2016-03-18 | 9936 | Graphite | Completed
[196 empty] | 2016-03-22 | 4319 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11646 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13890 | Close Job | Completed
[7 empty] | 123.8 | 88.9962 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34311 | 34310 | 34311 | 131 | 131 | 131 | 34180.9 | 34443 | jenny dahl | -986.04 | 0.23 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.04 ‰ | -986.04 ‰ | 0.23 | 0.23 | -985.86 | 0.23 ‰ | -985.86 | 0.23 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013964 | 0.013964 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 2.6 |  |  | 0.01396 | 0.00023 | 11647 | 2016-04-05 | 15:51 | Albert Zondervan | 204316 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.4 | 0.02% | 15 | 5 | 7 | ... | 0.01343 | 0.00022 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01343 | 0.01‰ | 0.00022 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01396 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01343 | 0 | 0.00022 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69927 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.5 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water |  |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60946 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [25 empty] | 2016-03-17 | 2407 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [25 empty] | 2016-03-17 | 1467 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [25 empty] | 2016-03-17 | 2793 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [25 empty] | 2016-03-18 | 9937 | Graphite | Completed
[196 empty] | 2016-03-22 | 4320 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11647 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13891 | Close Job | Completed
[7 empty] | 132 | 88.9452 | 0.0023 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34425 | 34425 | 34425 | 139 | 139 | 139 | 34287.5 | 34565.6 | jenny dahl | -986.23 | 0.24 ‰ | 2016-04-04 | 2016-04-04 | 17 | 0 |  |  |  |  | 3.50‰ | 3.50‰ | 2.00‰ | 2.00‰ | 3.50 ‰ | 2.00 ‰ | AMS | 2.17‰ | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.89443 | -986.23 ‰ | -986.23 ‰ | 0.24 | 0.24 | -986.06 | 0.24 ‰ | -986.06 | 0.24 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:51 | 0.013767 | 0.013767 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 2.4 |  |  | 0.01377 | 0.00024 | 11648 | 2016-04-05 | 15:51 | Albert Zondervan | 204317 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.38 | 0.02% | 1 | 5 | 7 | ... | 0.01324 | 0.00023 |  |  |  |  | 1.003 | 1.003 | Yes | Yes | [5 empty] | 0.01324 | 0.01‰ | 0.00023 | 0.00‰ | No | 1 | 0.99996 | 0.00101 | 0.00101 | 0.01377 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01324 | 0 | 0.00023 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-01 | jenny dahl | 8:10 |  | Completed |  |  |  |  | 69930 | 2000000 |  |  | 69784 69812 69813 69814 69928 | 69784~69812~69813~69814~69928 | 3055 |  | 3055 3055 3055 3055 3055 | 2016-04-05 | 16:19 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0019 | 0.00219 | 69784 |  | XCAMS | EA | 0.15% |  | * two cathodes don't match with their Ox-I description: TP69730 and TP69806 * measurement was troubled by ca. 2 hours of unstable accelerator voltage, but results data not compromised as bad runs could easily be identified and rejected * water-line blanks still not good | 2.2 | RRL | [5 empty] | Background Test |  | 2016-05-11 | Water | Complete |  | water blank test (different end points) prepare day before and run 1st All blanks to go in same wheel | 60947 | Normal | 14047/11 |  |  | 2016-03-16 | 3862 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99913 | 0.0023 | 69812 | [25 empty] | 2016-03-17 | 2408 | Water CO2 Evolution | Completed
[168 empty] | 1.0027 | 0.00228 | 69813 | [25 empty] | 2016-03-17 | 1471 | Gas Splitting | Completed
[168 empty] | 0.99729 | 0.00233 | 69814 | [25 empty] | 2016-03-17 | 2801 | Choose Gas Splitting | Completed
[168 empty] | 0.99855 | 0.00225 | 69928 | [25 empty] | 2016-03-18 | 9940 | Graphite | Completed
[196 empty] | 2016-03-22 | 4321 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11648 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13892 | Close Job | Completed
[7 empty] | 135.2 | 88.8394 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 175 | 175 | 175 |  |  | jenny dahl | -994.28 | 0.12 ‰ | 2016-04-06 | 2016-04-06 | 14 | 0 |  |  |  |  | 5.01‰ | 5.01‰ | 2.00‰ | 2.00‰ | 5.01 ‰ | 2.00 ‰ | AMS | 2.70‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.89443 | -994.28 ‰ | -994.28 ‰ | 0.12 | 0.12 | -994.21 | 0.13 ‰ | -994.21 | 0.13 ‰ |  |  | 2016-04-11 | Jocelyn Turnbull | 10:56 | 0.0057189 | 0.0057189 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00572 | 0.00012 | 11686 | 2016-04-07 | 13:39 | Albert Zondervan | 204238 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 4 | 5 | 6 | ... | 0.0055 | 0.00012 |  |  |  |  | 1.0033 | 1.0033 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00572 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-04 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 69954 | 2000000 |  |  | 69821 69916 69920 69945 69951 | 69821~69916~69920~69945~69951 | 3056 |  | 3056 3056 3056 3056 3056 | 2016-04-07 | 14:14 | Albert Zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99926 | 0.00157 | 69821 |  | XCAMS | EA | 0% |  | * good meas't, no surprises * no error adjustments * small Ox-I (TP69822) 3-sigma lower than large ones, so removed from list of calibration standards | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2016-04-26 | Water |  |  |  | 60981 | Normal | 14047/11 |  |  | 2016-03-22 | 2424 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00153 | 69916 | [25 empty] | 2016-03-22 | 1476 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00201 | 69920 | [25 empty] | 2016-03-22 | 2811 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00154 | 69945 | [25 empty] | 2016-03-23 | 9964 | Graphite | Completed
[168 empty] | 1.0001 | 0.0015 | 69951 | [25 empty] | 2016-03-24 | 4335 | C13 Measurement | Completed
[196 empty] | 2016-04-11 | 11686 | AMS Submission Results | Completed
[196 empty] | 2016-04-11 | 13847 | Close Job | Completed
[7 empty] | 89.9 | 88.822 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40987 | 40986 | 40987 | 205 | 205 | 205 |  |  | jenny dahl | -993.92 | 0.16 ‰ | 2016-04-11 | 2016-04-11 | 4 | 0 |  |  |  |  | 7.40‰ | 7.40‰ | 2.00‰ | 2.00‰ | 7.40 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0041 | 0 | 0.8165 | -993.92 ‰ | -993.92 ‰ | 0.16 | 0.16 | -993.84 | 0.16 ‰ | -993.84 | 0.16 ‰ |  |  | 2016-04-14 | Jocelyn Turnbull | 14:04 | 0.0060828 | 0.0060828 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.4 |  |  | 0.00608 | 0.00016 | 11715 | 2016-04-14 | 12:48 | Albert Zondervan | 204230 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.61 | 0.02% | 1 | 5 | 6 | ... | 0.00585 | 0.00015 |  |  |  |  | 1.0055 | 1.0055 | Yes | Yes | [5 empty] | 0.00585 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1.0001 | 0.00102 | 0.00102 | 0.00608 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00585 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-04-08 | jenny dahl | 8:44 |  | Completed |  |  |  |  | 70029 | 2000000 |  |  | 69905 69906 69935 69936 70000 70031 | 69905~69906~69935~69936~70000~70031 | 3057 |  | 3057 3057 3057 3057 3057 3057 | 2016-04-14 | 13:10 | Albert Zondervan |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99798 | 0.00252 | 69905 |  | XCAMS | ST | 0.18% |  | * good end result, after initial troubles with tuning * non-zero but acceptable system error |  | RRL | [5 empty] | Background Inorganic |  | 2016-04-25 | Water | Not Complete |  |  | 61022 | Normal | 14047/11 |  |  | 2016-04-04 | 2452 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99895 | 0.0026 | 69906 | [6 empty] | 2.5 | [18 empty] | 2016-04-04 | 1484 | Gas Splitting | Completed
[168 empty] | 0.99754 | 0.00253 | 69935 | [25 empty] | 2016-04-04 | 2822 | Choose Gas Splitting | Completed
[168 empty] | 0.99965 | 0.00225 | 69936 | [25 empty] | 2016-04-05 | 4362 | C13 Measurement | Cancelled
[168 empty] | 1.0021 | 0.0025 | 70000 | [25 empty] | 2016-04-07 | 10039 | Graphite | Completed
[168 empty] | 1.0043 | 0.00261 | 70031 | [25 empty] | 2016-04-14 | 11715 | AMS Submission Results | Completed
[196 empty] | 2016-04-13 | 4378 | C13 Measurement | Completed
[196 empty] | 2016-04-15 | 13916 | Close Job | Completed
[7 empty] | 106.7 | 89.4952 | 0.0018 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36568 | 36568 | 36568 | 151 | 150 | 151 | 36419 | 36720.1 | Margaret Norris | -989.46 | 0.20 ‰ | 2016-05-12 | 2016-05-12 | 6 | 0 |  |  |  |  | 3.47‰ | 3.47‰ | 2.00‰ | 2.00‰ | 3.47 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0039 | 0 | 0.5 | -989.46 ‰ | -989.46 ‰ | 0.2 | 0.2 | -989.33 | 0.20 ‰ | -989.33 | 0.20 ‰ |  |  | 2016-10-12 | Margaret Norris | 10:47 | 0.010544 | 0.010544 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.5 |  |  | 0.01054 | 0.0002 | 11811 | 2016-05-20 | 14:21 | Albert Zondervan | 204423 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 22 | 5 | 6 | ..T | 0.01014 | 0.00019 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.01014 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.01054 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01014 | 0 | 0.00019 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70214 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.81 mg |  | 0.81 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-03 | Water | Complete | Charged in timesheet June 2016 Jeremy 2 hours to GCT. | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61888 | Normal | 14047/11 |  |  | 2016-05-05 | 2492 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [25 empty] | 2016-05-05 | 2498 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [25 empty] | 2016-05-05 | 1513 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [25 empty] | 2016-05-05 | 2873 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [25 empty] | 2016-05-06 | 10224 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [25 empty] | 2016-10-12 | 11811 | AMS Submission Results | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [25 empty] | 2016-10-12 | 14975 | Close Job | Completed
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 100.5 | 89.1645 | 0.0011 | RRL-UNSt-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40315 | 40315 | 40315 | 202 | 202 | 202 | 40115.7 | 40519.9 | Margaret Norris | -993.39 | 0.17 ‰ | 2016-05-12 | 2016-05-12 | 6 | 0 |  |  |  |  | 7.19‰ | 7.19‰ | 2.00‰ | 2.00‰ | 7.19 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0039 | 0 | 0.5 | -993.39 ‰ | -993.39 ‰ | 0.17 | 0.17 | -993.31 | 0.17 ‰ | -993.31 | 0.17 ‰ |  |  | 2016-10-12 | Jocelyn Turnbull | 11:10 | 0.0066131 | 0.0066131 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.8 |  |  | 0.00661 | 0.00017 | 11812 | 2016-05-20 | 14:21 | Albert Zondervan | 204424 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.66 | 0.02% | 6 | 5 | 6 | ..T | 0.00636 | 0.00016 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00636 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00661 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00636 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70216 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 1.07 mg |  | 1.07 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. |  | RRL | [5 empty] | Standard as Unknown Inorganic |  | 2016-06-03 | Water | Complete | Charged in timesheet June 2016 Jeremy 2 hours to GCT. | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61889 | Normal | 14047/11 |  |  | 2016-05-05 | 2493 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [25 empty] | 2016-05-05 | 2499 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [25 empty] | 2016-05-05 | 1515 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [25 empty] | 2016-05-05 | 2877 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [25 empty] | 2016-05-06 | 10226 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [25 empty] | 2016-10-12 | 11812 | AMS Submission Results | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [25 empty] | 2016-10-12 | 14976 | Close Job | Completed
[168 empty] | 0.99865 | 0.00192 | 70185
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 102.9 | 89.1757 | 0.0012 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39802 | 39801 | 39802 | 178 | 177 | 178 | 39625.8 | 39981.3 | Margaret Norris | -992.95 | 0.16 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 7.06‰ | 7.06‰ | 2.00‰ | 2.00‰ | 7.06 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -992.95 ‰ | -992.95 ‰ | 0.16 | 0.16 | -992.86 | 0.16 ‰ | -992.86 | 0.16 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0070498 | 0.0070498 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00705 | 0.00016 | 11841 | 2016-05-20 | 14:21 | Albert Zondervan | 204507 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.7 | 0.02% | 37 | 5 | 7 | ..T | 0.00678 | 0.00015 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00678 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00705 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00678 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70222 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water | Complete | blank, no charge | Water blank test cutoff at higher pressure than normal cut off at 0.1 | 62087 | Normal | 14047/11 |  |  | 2016-05-05 | 2494 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [25 empty] | 2016-05-06 | 2503 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [25 empty] | 2016-05-06 | 1518 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [25 empty] | 2016-05-06 | 2882 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [25 empty] | 2016-05-10 | 10232 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [25 empty] | 2016-05-16 | 4421 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [25 empty] | 2016-11-02 | 11841 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [25 empty] | 2016-11-02 | 15128 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 106.5 | 89.3498 | 0.0012 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39557 | 39556 | 39557 | 184 | 183 | 184 | 39374.7 | 39742.5 | Margaret Norris | -992.73 | 0.17 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 5.10‰ | 5.10‰ | 2.00‰ | 2.00‰ | 5.10 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -992.73 ‰ | -992.73 ‰ | 0.17 | 0.17 | -992.64 | 0.17 ‰ | -992.64 | 0.17 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0072682 | 0.0072682 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2.2 |  |  | 0.00727 | 0.00017 | 11842 | 2016-05-20 | 14:21 | Albert Zondervan | 204508 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.73 | 0.02% | 29 | 5 | 7 | ..T | 0.00699 | 0.00016 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.00699 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00727 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00699 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70223 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.4 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water |  | blank, no charge | Water blank test cutoff at higher pressure than normal cut off at 1mBar | 62088 | Normal | 14047/11 |  |  | 2016-05-05 | 2495 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [25 empty] | 2016-05-06 | 2502 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [25 empty] | 2016-05-06 | 1519 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [25 empty] | 2016-05-06 | 2884 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [25 empty] | 2016-05-10 | 10233 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [25 empty] | 2016-05-16 | 4422 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [25 empty] | 2016-11-02 | 11842 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [25 empty] | 2016-11-02 | 15130 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 105.2 | 89.3472 | 0.0016 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37263 | 37262 | 37263 | 164 | 164 | 164 | 37100.4 | 37428.6 | Margaret Norris | -990.33 | 0.20 ‰ | 2016-05-12 | 2016-05-12 | 2 | 0 |  |  |  |  | 5.13‰ | 5.13‰ | 2.00‰ | 2.00‰ | 5.13 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0039 | 0 | 0.5 | -990.33 ‰ | -990.33 ‰ | 0.2 | 0.2 | -990.21 | 0.20 ‰ | -990.21 | 0.20 ‰ |  |  | 2016-11-02 | Jocelyn Turnbull | 12:52 | 0.0096701 | 0.0096701 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2 |  |  | 0.00967 | 0.0002 | 11843 | 2016-05-20 | 14:21 | Albert Zondervan | 204509 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.97 | 0.02% | 31 | 5 | 7 | ..T | 0.0093 | 0.00019 |  |  |  |  | 1.0092 | 1.0092 | Yes | Yes | [5 empty] | 0.0093 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.00967 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.0093 | 0 | 0.00019 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-10 | Margaret Norris | 16:12 |  | Completed |  |  |  |  | 70224 | 2000000 |  |  | 70139 70141 70142 70150 70175 70176 70178 70185 70186 70187 70188 70189 70190 70192 70197 70198 | 70139~70141~70142~70150~70175~70176~70178~70185~70186~70187~70188~70189~70190~70192~70197~70198 | 3060 |  | 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 3060 |  |  |  |  | 0.93 mg |  | 0.93 | [5 empty] | 0.99876 | 0.00185 | 70139 |  | XCAMS | Both | 0% |  | The main purpose of TW3060 (TST2) is to continue to identify the source of problems with our 14C analyses. It was run right after another test wheel (TW09May2016), which also focussed on XCAMS beam transport. The TW3060 results show that proper performance has not returned yet. For example, 2x ST OxIs come out 1% too low and were for that reason not used to calibrate the here reported results with. Considering other artefacts, it seems useless to take analysis of these data any further. However, the results of the IAEA-C1 blanks from the water- and carbonate lines are not compromised by the machine problems and thus may contain valuable information about the processing blank on these lines. A full AMS-technical analysis of all results obtained with this wheel can be found in I:\XCAMS\3_measurements\C-14 AMS\TW data analysis\TW3050-3099\TW3060. The conclusions drawn from it are: 1 Graphite weights > 1 mg generated on average lower 12C beam current than those < 1 mg C. This points to a problem with too low Fe:C ratio or beam emittance. 2 12C transmission and 13C/12C ratio show dependency on graphite weight: Lower weights are associated with lower ratios. 3 The run-to-run variability of these parameters is also weight dependent: Lower weights are associated with faster trending-down ratios. 4 The beam current has been trending upward while 12C transmission and 13/12 ratio trended downward. 5 Whatever the causal link between transmission and 13/12 ratio, the measurement data show correlation. The flask OxI and some other cathodes deviate from that correlation pattern. 6 Among the cathodes, the transmission values vary by 0.6% absolute. This is not 'super-tight', but by itself still acceptable. 7 After correction for del13C_irms, the 13C/12C ratios vary among the cathodes by 10 permil. This is too much. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-06-24 | Water |  | blank, no charge | Water blank test cutoff at higher pressure than normal cut off point was 0.01 | 62089 | Normal | 14047/11 |  |  | 2016-05-05 | 2496 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00187 | 70141 | [25 empty] | 2016-05-06 | 2501 | Water CO2 Evolution | Completed
[168 empty] | 0.99867 | 0.00188 | 70142 | [25 empty] | 2016-05-06 | 1517 | Gas Splitting | Completed
[168 empty] | 0.99958 | 0.00194 | 70150 | [25 empty] | 2016-05-06 | 2880 | Choose Gas Splitting | Completed
[168 empty] | 0.9996 | 0.00195 | 70175 | [25 empty] | 2016-05-10 | 10234 | Graphite | Completed
[168 empty] | 0.99708 | 0.00202 | 70176 | [25 empty] | 2016-05-16 | 4423 | C13 Measurement | Completed
[168 empty] | 0.99827 | 0.00189 | 70178 | [25 empty] | 2016-11-02 | 11843 | AMS Submission Results | Completed
[168 empty] | 0.99865 | 0.00192 | 70185 | [25 empty] | 2016-11-02 | 15129 | Close Job | Completed
[168 empty] | 0.99875 | 0.002 | 70186
[168 empty] | 1.0003 | 0.00192 | 70187
[168 empty] | 1.0032 | 0.00204 | 70188
[168 empty] | 0.99639 | 0.00199 | 70189
[168 empty] | 1.0037 | 0.00183 | 70190
[168 empty] | 1.0036 | 0.00186 | 70192
[168 empty] | 1.0021 | 0.00186 | 70197
[168 empty] | 0.99914 | 0.00191 | 70198
[7 empty] | 131 | 88.9489 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37393 | 37393 | 37393 | 140 | 140 | 140 |  |  | Margaret Norris | -990.49 | 0.17 ‰ | 2016-05-20 | 2015-11-22 | 14 | 180 |  |  |  |  | 4.60‰ | 4.60‰ | 2.00‰ | 2.00‰ | 4.60 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.003 | 0 | 0.70711 | -990.49 ‰ | -990.49 ‰ | 0.17 | 0.17 | -990.37 | 0.17 ‰ | -990.37 | 0.17 ‰ |  |  | 2016-05-25 | Jocelyn Turnbull | 16:01 | 0.0095141 | 0.0095141 | 0.00017 | 0.00017 | 0.00016 | 0.00016 | 1.7 |  |  | 0.00951 | 0.00017 | 11870 | 2016-05-25 | 14:48 | Albert Zondervan | 204422 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 6 | 5 | 6 | ... | 0.00915 | 0.00016 |  |  |  |  | 1.0041 | 1.0041 | Yes | Yes | Uncertainty on this result is larger than normally reported. The primary standards used to normalise this result exhibited a larger scatter than normal, due to a component failure during analysis. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. |  |  |  |  | 0.00915 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 1.0002 | 0.00169 | 0.00169 | 0.00951 | 0 | 0.00017 | 0 | 0 |  |  |  |  | 0.00915 | 0 | 0.00016 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-17 | Margaret Norris | 13:01 |  | Completed |  |  |  |  | 70215 | 2000000 |  |  | 70135 70136 70137 70138 70149 70174 70179 70199 | 70135~70136~70137~70138~70149~70174~70179~70199 | 3061 | 3057 3056 3048 3046 3045 3042 3040 3061 | 3061 3061 3061 3061 3061 3061 3061 3061 | 2016-05-25 | 15:52 | Jocelyn Turnbull | Yes | 1.10 mg | 0.44 | 1.1 | [5 empty] | 0.99477 | 0.00471 | 70135 |  | XCAMS | ST | 0.44% | Uncertainties on these results are larger than normally reported. The primary standards used to normalise your results exhibited a larger scatter than normal, due to a component failure during the analysis, and implying a larger uncertainty on the authentic samples. Secondary standards - known age materials treated as unknowns - measured in the same wheel did not exhibit this larger scatter and are consistent with their consensus values. We chose to be conservative and apply the larger uncertainty implied by the primary standards. | * lots to add here still JCT - despite the clear problems with scatter in OxI, the secondary standards look very good. R14047/2 water secondary was 7087±35 (vs 7134 consensus); R24889/4 FIRI-D wood was 4485±26 (vs 4508 ± 3 consensus); R24889/9 FIRI-I cellulose was 4463±26 (vs 4485±5 consensus). R14047/2 is slightly low, but we don't do very well with this material at any time. The others are within 1 sigma. Thus I chose not to increase the wheel-to-wheel error over the usual values we apply. |  | RRL | [5 empty] | Background Inorganic |  | 2016-06-03 | Water |  | really to be billed to GCT- if so at what price? | Water blank test. Normal cutoff .New valve third priani. >/=1mg | 61140 | Normal | 14047/11 |  |  | 2016-05-05 | 2491 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0044 | 0.00509 | 70136 | [25 empty] | 2016-05-05 | 2497 | Water CO2 Evolution | Completed
[168 empty] | 1.0005 | 0.00473 | 70137 | [25 empty] | 2016-05-05 | 1514 | Gas Splitting | Completed
[168 empty] | 0.99594 | 0.00473 | 70138 | [25 empty] | 2016-05-05 | 2875 | Choose Gas Splitting | Completed
[168 empty] | 1.003 | 0.00474 | 70149 | [25 empty] | 2016-05-06 | 10225 | Graphite | Completed
[168 empty] | 1.0059 | 0.00476 | 70174 | [25 empty] | 2016-05-25 | 11870 | AMS Submission Results | Completed
[168 empty] | 1.0034 | 0.00478 | 70179 | [25 empty] | 2016-05-30 | 14101 | Close Job | Completed
[168 empty] | 0.99424 | 0.00473 | 70199
[7 empty] | 76.8 | 92.029 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37571 | 37570 | 37571 | 278 | 278 | 278 |  |  | jenny dahl | -990.69 | 0.32 ‰ | 2016-06-09 | 2016-05-26 | 14 | 14 |  |  |  |  | 3.09‰ | 3.09‰ | 2.00‰ | 2.00‰ | 3.09 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -990.69 ‰ | -990.69 ‰ | 0.32 | 0.32 | -990.58 | 0.33 ‰ | -990.58 | 0.33 ‰ |  |  | 2016-06-10 | Jocelyn Turnbull | 16:52 | 0.0093062 | 0.0093062 | 0.00032 | 0.00032 | 0.00032 | 0.00032 | 2.1 |  |  | 0.00931 | 0.00032 | 11987 | 2016-06-10 | 15:56 | Albert Zondervan | 204398 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.93 | 0.03% | 36 | 5 | 6 | ..S | 0.00895 | 0.00031 |  |  |  |  | 1.0373 | 1.0373 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00895 | 0.01‰ | 0.00031 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00931 | 0 | 0.00032 | 0 | 0 |  |  | 0 | 0 | 0.00895 | 0 | 0.00031 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-05-31 | jenny dahl | 15:12 |  | Completed |  |  |  |  | 70291 | 2000000 |  |  | 70244 70246 70259 70262 70293 70309 | 70244~70246~70259~70262~70293~70309 | 3064 |  | 3064 3064 3064 3064 3064 3064 | 2016-06-10 | 16:25 | Jocelyn Turnbull |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0011 | 0.00167 | 70244 |  | SSAMS | EA | 0% |  | * 2nd wheel of unknowns measured on SSAMS at ANU * good results and FIRI-D secondaries appear OK * system error is zero * only blanks showed some excess run-to-run variance, which has been accounted for in the reported uncertainties Decided to use only blanks from this wheel for correction -this is only the second wheel from SSAMS and the pervious wheel (3061) only had air materials. Organic blanks are quite poor, but both are consistent. Water blank is also poor but not unusual. Therefore seems most appropriate to use these and not blanks run on XCAMS. FIRI-D all come out just on the low side of 1-sigma with this larger blank, but ok I think | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-06-03 | Water |  |  |  | 61197 | Normal | 14047/11 |  |  | 2016-05-24 | 2510 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0012 | 0.00166 | 70246 | [25 empty] | 2016-05-24 | 1526 | Gas Splitting | Completed
[168 empty] | 0.9984 | 0.00166 | 70259 | [25 empty] | 2016-05-24 | 2891 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00165 | 70262 | [25 empty] | 2016-05-26 | 10301 | Graphite | Completed
[168 empty] | 0.99805 | 0.00164 | 70293 | [25 empty] | 2016-05-30 | 4448 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.00165 | 70309 | [25 empty] | 2016-06-10 | 11987 | AMS Submission Results | Completed
[196 empty] | 2016-06-13 | 14157 | Close Job | Completed
[7 empty] | 83.7 | 91.8977 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36017 | 36017 | 36017 | 111 | 110 | 111 |  |  | Jocelyn Turnbull | -988.71 | 0.16 ‰ | 2016-07-02 | 2016-04-03 | 31 | 90 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -988.71 ‰ | -988.71 ‰ | 0.16 | 0.16 | -988.57 | 0.16 ‰ | -988.57 | 0.16 ‰ |  |  | 2016-07-04 | Jocelyn Turnbull | 9:54 | 0.011292 | 0.011292 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.5 |  |  | 0.01129 | 0.00016 | 12038 | 2016-07-04 | 8:52 | Albert Zondervan | 204491 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 1.13 | 0.02% | 23 | 5 | 6 | ..S | 0.01086 | 0.00015 |  |  |  |  | 1.0357 | 1.0357 | Yes | Yes | [5 empty] | 0.01086 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.01129 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.01086 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-10 | Jocelyn Turnbull | 12:00 |  | Completed |  |  |  |  | 70334 | 2000000 |  |  | 70264 70269 70346 70348 70368 70370 | 70264~70269~70346~70348~70368~70370 | 3066 |  | 3066 3066 3066 3066 3066 3066 | 2016-07-04 | 9:46 | Jocelyn Turnbull |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99725 | 0.00163 | 70264 |  | SSAMS | EA | 0% |  | * 4th wheel with unknowns run by S. Fallon on his SSAMS at ANU * physical wheel had been sent over before removing the ANU-sucrose in cathode pos#39 and moving the Ox-I from pos#0 there; this was corrected before measurement and now also in RLIMS * no error adjustments needed; good repeatability all around * again, all blanks seem a bit high Used only blanks run at ANU for blank corrections for this wheel | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-06-24 | Water |  | blank, no charge |  | 61353 | Normal | 14047/11 |  |  | 2016-05-31 | 2532 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00155 | 70269 | [25 empty] | 2016-05-31 | 1533 | Gas Splitting | Completed
[168 empty] | 0.99948 | 0.00151 | 70346 | [25 empty] | 2016-05-31 | 2900 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00149 | 70348 | [25 empty] | 2016-06-01 | 10344 | Graphite | Completed
[168 empty] | 0.99984 | 0.00148 | 70368 | [25 empty] | 2016-06-08 | 4467 | C13 Measurement | Completed
[168 empty] | 1.0022 | 0.00148 | 70370 | [25 empty] | 2016-07-04 | 12038 | AMS Submission Results | Completed
[196 empty] | 2016-07-04 | 14362 | Close Job | Completed
[7 empty] | 76.9 | 88.738 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37938 | 37938 | 37938 | 141 | 140 | 141 |  |  | jenny dahl | -991.11 | 0.16 ‰ | 2016-06-27 | 2016-03-29 | 13 | 90 |  |  |  |  | 1.03‰ | 1.03‰ | 2.00‰ | 2.00‰ | 1.03 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0 | 0.89443 | -991.11 ‰ | -991.11 ‰ | 0.16 | 0.16 | -991 | 0.16 ‰ | -991 | 0.16 ‰ |  |  | 2016-06-29 | Jocelyn Turnbull | 11:17 | 0.0088902 | 0.0088902 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00889 | 0.00016 | 12140 | 2016-06-29 | 10:23 | Albert Zondervan | 204641 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.89 | 0.02% | 10 | 5 | 6 | ... | 0.00855 | 0.00015 |  |  |  |  | 0.9982 | 0.9982 | Yes | Yes | [5 empty] | 0.00855 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00889 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00855 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-27 | jenny dahl | 8:53 |  | Completed |  |  |  |  | 70408 | 2000000 |  |  | 70376 70431 70437 70440 70441 | 70376~70431~70437~70440~70441 | 3070 |  | 3070 3070 3070 3070 3070 | 2016-06-29 | 11:07 | Jocelyn Turnbull |  | 1.07 mg |  | 1.07 | [5 empty] | 1.0006 | 0.00162 | 70376 |  | XCAMS | EA | 0% |  | * OxI in pos#7 (TP70373) is a 5-sigma high outlier and was therefore removed from the list of calibration standards. No reason for this deviation has been found yet. The lowering of Cs-oven temperature halfway during the measurement did not influence the degree of this outlying behaviour. * Excess global run-to-run variance required a 1.06 multiplication of all Poisson uncertainties. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-07-19 | Water | Complete | Blank, no charge |  | 61312 | Normal | 14047/11 |  |  | 2016-06-10 | 2562 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99896 | 0.00163 | 70431 | [25 empty] | 2016-06-10 | 1544 | Gas Splitting | Completed
[168 empty] | 0.99727 | 0.00175 | 70437 | [25 empty] | 2016-06-10 | 2913 | Choose Gas Splitting | Completed
[168 empty] | 1.0021 | 0.00164 | 70440 | [25 empty] | 2016-06-14 | 10418 | Graphite | Completed
[168 empty] | 1.0007 | 0.00157 | 70441 | [25 empty] | 2016-06-22 | 4506 | C13 Measurement | Completed
[196 empty] | 2016-06-29 | 12140 | AMS Submission Results | Completed
[196 empty] | 2016-06-29 | 14338 | Close Job | Completed
[7 empty] | 115.2 | 88.4851 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39580 | 39579 | 39580 | 127 | 126 | 127 |  |  | jenny dahl | -992.75 | 0.11 ‰ | 2016-07-04 | 2016-07-04 | 17 | 0 |  |  |  |  | 0.72‰ | 0.72‰ | 2.00‰ | 2.00‰ | 0.72 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -992.75 ‰ | -992.75 ‰ | 0.11 | 0.11 | -992.66 | 0.12 ‰ | -992.66 | 0.12 ‰ |  |  | 2016-07-11 | Jocelyn Turnbull | 11:33 | 0.0072474 | 0.0072474 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.4 |  |  | 0.00725 | 0.00011 | 12182 | 2016-07-08 | 14:51 | Albert Zondervan | 204977 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.72 | 0.01% | 30 | 5 | 6 | ... | 0.00697 | 0.00011 |  |  |  |  | 0.99504 | 0.99504 | Yes | Yes | [5 empty] | 0.00697 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00725 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00697 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-06-29 | jenny dahl | 16:06 |  | Completed |  |  |  |  | 70455 | 2000000 |  |  | 70462 70465 70529 70530 70538 70539 | 70462~70465~70529~70530~70538~70539 | 3071 |  | 3071 3071 3071 3071 3071 3071 | 2016-07-08 | 16:03 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 0.99876 | 0.00135 | 70462 |  | XCAMS | ST | 0% |  | * good data despite having to re-tune XCAMS because of power mains problems * measurement was extended beyond time needed for regular precision to get an impression how good the repeatability is now: With on average 550 k counts per OxI, we can't see a system error. * 1 cathode failed to give a reliable result: TP70453 (cat#4) * all other variances agree with their Poisson expectation: GOOD !! | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2016-08-03 | Water |  | blank, no charge |  | 61393 | Normal | 14047/11 |  |  | 2016-06-15 | 2569 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99909 | 0.00136 | 70465 | [25 empty] | 2016-06-15 | 1550 | Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00138 | 70529 | [25 empty] | 2016-06-15 | 2927 | Choose Gas Splitting | Completed
[168 empty] | 0.99974 | 0.00138 | 70530 | [25 empty] | 2016-06-17 | 10465 | Graphite | Completed
[168 empty] | 1.0029 | 0.00139 | 70538 | [25 empty] | 2016-06-22 | 4509 | C13 Measurement | Completed
[168 empty] | 1.0004 | 0.00131 | 70539 | [25 empty] | 2016-07-11 | 12182 | AMS Submission Results | Completed
[196 empty] | 2016-07-12 | 14420 | Close Job | Completed
[7 empty] | 103.6 | 88.4727 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38254 | 38254 | 38254 | 147 | 146 | 147 |  |  | jenny dahl | -991.45 | 0.16 ‰ | 2016-07-18 | 2016-07-18 | 18 | 0 |  |  |  |  | 3.03‰ | 3.03‰ | 2.00‰ | 2.00‰ | 3.03 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -991.45 ‰ | -991.45 ‰ | 0.16 | 0.16 | -991.35 | 0.16 ‰ | -991.35 | 0.16 ‰ |  |  | 2016-07-20 | jenny dahl | 11:15 | 0.0085471 | 0.0085471 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.4 |  |  | 0.00855 | 0.00016 | 12261 | 2016-07-19 | 13:01 | Albert Zondervan | 205068 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.85 | 0.02% | 10 | 5 | 6 | ... | 0.00822 | 0.00015 |  |  |  |  | 0.9972 | 0.9972 | Yes | Yes | [5 empty] | 0.00822 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00855 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00822 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-14 | jenny dahl | 10:41 |  | Completed |  |  |  |  | 70550 | 2000000 |  |  | 70419 70421 70428 70505 70506 70507 | 70419~70421~70428~70505~70506~70507 | 3073 |  | 3073 3073 3073 3073 3073 3073 | 2016-07-19 | 14:44 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 1.0003 | 0.00171 | 70419 |  | XCAMS | EA | 0% |  | * straightforward problem-free measurement | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-08-16 | Water |  | blank, no charge |  | 61491 | Normal | 14047/11 |  |  | 2016-06-29 | 2601 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9984 | 0.00173 | 70421 | [25 empty] | 2016-06-29 | 1562 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00172 | 70428 | [25 empty] | 2016-06-29 | 2959 | Choose Gas Splitting | Completed
[168 empty] | 0.99782 | 0.00173 | 70505 | [25 empty] | 2016-06-30 | 10560 | Graphite | Completed
[168 empty] | 1.0023 | 0.00173 | 70506 | [25 empty] | 2016-07-04 | 4555 | C13 Measurement | Completed
[168 empty] | 0.99914 | 0.00172 | 70507 | [25 empty] | 2016-07-20 | 12261 | AMS Submission Results | Completed
[196 empty] | 2016-07-20 | 14499 | Close Job | Completed
[7 empty] | 97.1 | 88.2885 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39330 | 39329 | 39330 | 167 | 167 | 167 |  |  | Margaret Norris | -992.52 | 0.16 ‰ | 2016-07-20 | 2016-01-22 | 5 | 180 |  |  |  |  | 3.25‰ | 3.25‰ | 2.00‰ | 2.00‰ | 3.25 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0033 | 0 | 0.8165 | -992.52 ‰ | -992.52 ‰ | 0.16 | 0.16 | -992.43 | 0.16 ‰ | -992.43 | 0.16 ‰ |  |  | 2016-07-25 | jenny dahl | 12:27 | 0.0074761 | 0.0074761 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.00748 | 0.00016 | 12293 | 2016-07-21 | 12:13 | Albert Zondervan | 205070 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.75 | 0.02% | 16 | 5 | 6 | ... | 0.00719 | 0.00015 |  |  |  |  | 0.99534 | 0.99534 | Yes | Yes | [5 empty] | 0.00719 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 0.99999 | 0.00094 | 0.00094 | 0.00748 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00719 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-07-18 | Margaret Norris | 11:31 |  | Completed |  |  |  |  | 70636 | 2000000 |  |  | 70463 70471 70473 70531 70536 70537 | 70463~70471~70473~70531~70536~70537 | 3074 |  | 3074 3074 3074 3074 3074 3074 | 2016-07-21 | 13:21 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 0.99951 | 0.00227 | 70463 |  | XCAMS | ST | 0.13% |  | * straightforward measurement | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-08-16 | Water |  |  |  | 61523 | Normal | 14047/11 |  |  | 2016-07-12 | 2642 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00228 | 70471 | [25 empty] | 2016-07-12 | 1580 | Gas Splitting | Completed
[168 empty] | 1.0017 | 0.00231 | 70473 | [25 empty] | 2016-07-12 | 2984 | Choose Gas Splitting | Completed
[168 empty] | 0.99564 | 0.00233 | 70531 | [25 empty] | 2016-07-15 | 10646 | Graphite | Completed
[168 empty] | 0.99985 | 0.00225 | 70536 | [25 empty] | 2016-07-19 | 4605 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00235 | 70537 | [25 empty] | 2016-07-25 | 12293 | AMS Submission Results | Completed
[196 empty] | 2016-07-26 | 14558 | Close Job | Completed
[7 empty] | 144.4 | 88.1432 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39230 | 39229 | 39230 | 132 | 132 | 132 |  |  | jeremy thompson | -992.43 | 0.12 ‰ | 2016-08-26 | 2016-08-26 | 17 | 0 |  |  |  |  | 1.63‰ | 1.63‰ | 2.00‰ | 2.00‰ | 1.63 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0029 | 0 | 0.8165 | -992.43 ‰ | -992.43 ‰ | 0.12 | 0.12 | -992.34 | 0.13 ‰ | -992.34 | 0.13 ‰ |  |  | 2016-08-30 | Jocelyn Turnbull | 9:51 | 0.0075697 | 0.0075697 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 3.4 |  |  | 0.00757 | 0.00012 | 12490 | 2016-08-29 | 11:25 | Albert Zondervan | 205193 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:44 | Jocelyn Turnbull | 0 |  | 0.76 | 0.01% | 10 | 5 | 7 | ... | 0.00728 | 0.00012 |  |  |  |  | 0.9921 | 0.9921 | Yes | Yes | [5 empty] | 0.00728 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00757 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00728 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-08-25 | jeremy thompson | 10:51 |  | Completed |  |  |  |  | 70748 | 2000000 |  |  | 70698 70699 70701 70721 70724 70758 | 70698~70699~70701~70721~70724~70758 | 3079 |  | 3079 3079 3079 3079 3079 3079 | 2016-08-29 | 11:46 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0026 | 0.00184 | 70698 |  | XCAMS | EA | 0.11% |  | * good measurement, with small system error and no other uncertainty adjustments * high beam currents generally * exception: cathode pos#15 (TP70745) had lower output than all others, for unknown reason | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2016-08-26 | Water | Complete |  |  | 61668 | Normal | 14047/11 |  |  | 2016-08-05 | 2682 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00183 | 70699 | [25 empty] | 2016-08-08 | 2695 | Water CO2 Evolution | Completed
[168 empty] | 0.99965 | 0.00183 | 70701 | [25 empty] | 2016-08-08 | 1599 | Gas Splitting | Completed
[168 empty] | 0.99817 | 0.00184 | 70721 | [25 empty] | 2016-08-08 | 3017 | Choose Gas Splitting | Completed
[168 empty] | 0.99783 | 0.00182 | 70724 | [25 empty] | 2016-08-09 | 10758 | Graphite | Completed
[168 empty] | 1.0011 | 0.00183 | 70758 | [25 empty] | 2016-08-11 | 4674 | C13 Measurement | Completed
[196 empty] | 2016-08-30 | 12490 | AMS Submission Results | Completed
[196 empty] | 2016-09-01 | 14748 | Close Job | Completed
[7 empty] | 124.4 | 88.2242 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40290 | 40290 | 40290 | 151 | 151 | 151 |  |  | Margaret Norris | -993.37 | 0.12 ‰ | 2016-09-20 | 2016-09-20 | 7 | 0 |  |  |  |  | 2.32‰ | 2.32‰ | 2.00‰ | 2.00‰ | 2.32 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -993.37 ‰ | -993.37 ‰ | 0.12 | 0.12 | -993.28 | 0.13 ‰ | -993.28 | 0.13 ‰ |  |  | 2016-09-21 | Jocelyn Turnbull | 12:53 | 0.0066339 | 0.0066339 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00663 | 0.00012 | 12587 | 2016-09-21 | 11:00 | Albert Zondervan | 205322 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 26 | 5 | 6 | ... | 0.00638 | 0.00012 |  |  |  |  | 0.9937 | 0.9937 | Yes | Yes | [5 empty] | 0.00638 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00663 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00638 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-19 | Margaret Norris | 10:09 |  | Completed |  |  |  |  | 70935 | 2000000 |  |  | 70820 70821 70832 70833 70834 70883 | 70820~70821~70832~70833~70834~70883 | 3082 |  | 3082 3082 3082 3082 3082 3082 | 2016-09-21 | 11:44 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0001 | 0.00156 | 70820 |  | XCAMS | ST | 0% |  | * good meas't * minor error adjustments for 3 cathodes * on basis of wildly different 14C concentration, it was found that two samples (TP70934,5) had been permutated; this was found to be caused by a swap on the water-line | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2016-09-21 | Water |  |  | JCT 21/9/16. Changed R number from 14047/12 to 14047/11 since this was obviously switched with c1 marble at evolution d13C should be - 8.2 | 61809 | Normal | 14047/11 |  |  | 2016-09-12 | 2746 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99934 | 0.00157 | 70821 | [25 empty] | 2016-09-12 | 1619 | Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00157 | 70832 | [25 empty] | 2016-09-12 | 3046 | Choose Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00157 | 70833 | [25 empty] | 2016-09-13 | 10945 | Graphite | Completed
[168 empty] | 0.99901 | 0.00157 | 70834 | [25 empty] | 2016-09-20 | 4737 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00157 | 70883 | [25 empty] | 2016-09-21 | 12587 | AMS Submission Results | Completed
[196 empty] | 2016-09-23 | 14896 | Close Job | Completed
[7 empty] | 103 | 88.0137 | 0.0094 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23244 | 23243 | 23244 | 57 | 57 | 57 | 23186.7 | 23301.4 | jenny dahl | -944.62 | 0.40 ‰ | 2016-09-30 | 2016-09-30 | 3 | 0 |  |  |  |  | 6.11‰ | 6.11‰ | 2.00‰ | 2.00‰ | 6.11 ‰ | 2.00 ‰ | AMS | 1.95‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -944.62 ‰ | -944.62 ‰ | 0.4 | 0.4 | -943.94 | 0.40 ‰ | -943.94 | 0.40 ‰ |  |  | 2016-10-03 | Jocelyn Turnbull | 9:05 | 0.055379 | 0.055379 | 0.0004 | 0.0004 | 0.0004 | 0.0004 | 2.2 |  |  | 0.05538 | 0.0004 | 12656 | 2016-10-01 | 14:01 | Albert Zondervan | 205517 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 5.54 | 0.04% | 11 | 5 | 6 | .X. | 0.05326 | 0.00038 |  |  |  |  | 0.99507 | 0.99507 | Yes | Yes | [5 empty] | 0.05326 | 0.05‰ | 0.00038 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.05538 | 0 | 0.0004 | 0 | 0 |  |  | 0 | 0 | 0.05326 | 0 | 0.00038 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-27 | jenny dahl | 14:56 |  | Completed |  |  |  |  | 71042 | 2000000 |  |  | 70889 70890 70952 70953 70959 70960 71033 | 70889~70890~70952~70953~70959~70960~71033 | 3083 |  | 3083 3083 3083 3083 3083 3083 3083 | 2016-10-01 | 14:50 | Albert Zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 1.0003 | 0.00156 | 70889 |  | XCAMS | EA | 0% |  | * solid meas't without issues | 2 | RRL | [5 empty] | Background Test |  | 2016-11-10 | Water | Complete | JCT ADDED INVOICE 540GCT21 | Water blank test. No flush with nitrogen, large vessel | 61849 | Normal | 14047/11 |  |  | 2016-09-26 | 2762 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00156 | 70890 | [25 empty] | 2016-09-26 | 1622 | Gas Splitting | Completed
[168 empty] | 0.99996 | 0.00158 | 70952 | [25 empty] | 2016-09-26 | 3054 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00155 | 70953 | [25 empty] | 2016-09-27 | 11052 | Graphite | Completed
[168 empty] | 1.0015 | 0.00156 | 70959 | [25 empty] | 2016-10-03 | 4749 | C13 Measurement | Completed
[168 empty] | 0.99895 | 0.00155 | 70960 | [25 empty] | 2016-10-03 | 12656 | AMS Submission Results | Completed
[168 empty] | 0.99894 | 0.00153 | 71033 | [25 empty] | 2016-10-06 | 14970 | Close Job | Completed
[196 empty] | 2017-03-27 | 4996 | Invoice | Completed
[7 empty] | 121.9 | 88.4284 | 0.0033 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31558 | 31557 | 31558 | 89 | 89 | 89 | 31469.1 | 31647.4 | jenny dahl | -980.33 | 0.22 ‰ | 2016-09-30 | 2016-09-30 | 3 | 0 |  |  |  |  | 1.40‰ | 1.40‰ | 2.00‰ | 2.00‰ | 1.40 ‰ | 2.00 ‰ | AMS | 2.21‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -980.33 ‰ | -980.33 ‰ | 0.22 | 0.22 | -980.08 | 0.22 ‰ | -980.08 | 0.22 ‰ |  |  | 2016-10-03 | Jocelyn Turnbull | 9:05 | 0.019673 | 0.019673 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.4 |  |  | 0.01967 | 0.00022 | 12657 | 2016-10-01 | 14:01 | Albert Zondervan | 205518 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 1.97 | 0.02% | 31 | 5 | 6 | .X. | 0.01892 | 0.00021 |  |  |  |  | 0.99507 | 0.99507 | Yes | Yes | [5 empty] | 0.01892 | 0.02‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.01967 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01892 | 0 | 0.00021 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-09-27 | jenny dahl | 14:56 |  | Completed |  |  |  |  | 71043 | 2000000 |  |  | 70889 70890 70952 70953 70959 70960 71033 | 70889~70890~70952~70953~70959~70960~71033 | 3083 |  | 3083 3083 3083 3083 3083 3083 3083 | 2016-10-01 | 14:50 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 1.0003 | 0.00156 | 70889 |  | XCAMS | EA | 0% |  | * solid meas't without issues | 2.2 | RRL | [5 empty] | Background Test |  | 2016-11-10 | Water | Not Complete | JCT ADDED INVOICE 540GCT21 | Water blank test. No flush with nitrogen, small vessel. C13 readings | 61850 | Normal | 14047/11 |  |  | 2016-09-26 | 2763 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00156 | 70890 | [25 empty] | 2016-09-26 | 1626 | Gas Splitting | Completed
[168 empty] | 0.99996 | 0.00158 | 70952 | [25 empty] | 2016-09-26 | 3062 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00155 | 70953 | [25 empty] | 2016-09-27 | 11053 | Graphite | Completed
[168 empty] | 1.0015 | 0.00156 | 70959 | [25 empty] | 2016-10-03 | 4750 | C13 Measurement | Completed
[168 empty] | 0.99895 | 0.00155 | 70960 | [25 empty] | 2016-10-03 | 12657 | AMS Submission Results | Completed
[168 empty] | 0.99894 | 0.00153 | 71033 | [25 empty] | 2016-10-06 | 14971 | Close Job | Completed
[196 empty] | 2017-03-27 | 4997 | Invoice | Completed
[7 empty] | 97.8 | 88.4872 | 0.0017 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36761 | 36760 | 36761 | 146 | 146 | 148 | 36615.5 | 36908.4 | margaret norris | -989.71 | 0.19 ‰ | 2016-11-25 | 2016-11-25 | 35 | 0 |  |  |  |  | 3.16‰ | 3.16‰ | 2.00‰ | 2.00‰ | 3.16 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.75593 | -989.71 ‰ | -989.71 ‰ | 0.19 | 0.19 | -989.58 | 0.19 ‰ | -989.58 | 0.19 ‰ |  |  | 2016-11-28 | Jocelyn Turnbull | 13:03 | 0.010294 | 0.010294 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2.7 |  |  | 0.01029 | 0.00019 | 12943 | 2016-11-28 | 8:51 | albert zondervan | 205635 |  |  |  |  | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 1.03 | 0.02% | 11 | 5 | 6 | A.. | 0.0099 | 0.00018 |  |  |  |  | 0.99749 | 0.99749 | Yes | Yes |  | 35 days old wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 147.5403229275772136. If Saved values show old error bar then the revised result has not been re-reported to the client. |  |  |  | 0.0099 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01029 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0099 | 0 | 0.00018 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-11-21 | margaret norris | 8:31 |  | Completed |  |  |  |  | 71187 | 2000000 |  |  | 71234 71246 71253 71254 71269 71314 71320 | 71234~71246~71253~71254~71269~71314~71320 | 3091 |  | 3091 3091 3091 3091 3091 3091 3091 | 2016-11-28 | 9:32 | albert zondervan |  | 0.74 mg | 0.13 | 0.74 | [5 empty] | 1.0003 | 0.00197 | 71234 |  | XCAMS | EA | 0% |  | TW3091 measurement: No tuning required as we've come off back of TW3090. error adjustments: * factor 2.7 for run-to-run variability of plant TP71277 * factor 1.3 for run-to-run variability of the small sediment TP71313 * no other adjustments warranted unknowns: * TP71277 generated an average rts that has been derived from a far too variable set of individual runs. I do not trust this average to be correct and therefore I call it a fail. * None of the other cathodes in the wheel show such a behaviour. * Small sediment TP71313 was blank-corrected along with all larges. primaries: The wheel contained 7x cathodes with OxI: all via EA. All of them performed equally well, so I changed the category for TP71269 to primary standard. blanks: Two groups have been identified in this wheel to blank-correct: * organics via AAA: 26x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied * organics via CellExtr: 0x unknowns + 1x secondaries + 0x blanks; a 360-days window was applied to find a decent number of blanks to go with it secondaries: * Both secondaries (FIRI-I, D) agree with their expectation value. | 2.5 | RRL | [5 empty] | Unknown Inorganic |  | 2016-12-09 | Water |  |  | TEST Water blank test. Flush with nitrogen before added sample as powder, C13 readings | 62189 | Normal | 14047/11 |  |  | 2016-10-20 | 2764 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.00184 | 71246 | [25 empty] | 2016-10-20 | 1637 | Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00188 | 71253 | [25 empty] | 2016-10-20 | 3077 | Choose Gas Splitting | Completed
[168 empty] | 0.99791 | 0.00192 | 71254 | [25 empty] | 2016-10-21 | 11197 | Graphite | Completed
[168 empty] | 1.0015 | 0.00193 | 71269 | [25 empty] | 2016-11-01 | 4777 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00185 | 71314 | [25 empty] | 2016-11-28 | 12943 | AMS Submission Results | Completed
[168 empty] | 1.0019 | 0.00188 | 71320 | [25 empty] | 2016-12-05 | 15339 | Close Job | Completed
[7 empty] | 98.1 | 88.1866 | 0.001 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41453 | 41453 | 41453 | 189 | 189 | 190 | 41265.9 | 41644.9 | margaret norris | -994.26 | 0.14 ‰ | 2016-11-25 | 2016-11-25 | 35 | 0 |  |  |  |  | 6.58‰ | 6.58‰ | 2.00‰ | 2.00‰ | 6.58 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0 | 0.75593 | -994.26 ‰ | -994.26 ‰ | 0.14 | 0.14 | -994.19 | 0.14 ‰ | -994.19 | 0.14 ‰ |  |  | 2016-11-28 | Jocelyn Turnbull | 13:03 | 0.0057397 | 0.0057397 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.7 |  |  | 0.00574 | 0.00014 | 12944 | 2016-11-28 | 8:51 | albert zondervan | 205636 |  |  |  |  | 45% | 5 | 5/07/22 | 14:06:22 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 31 | 5 | 5 | A.. | 0.00552 | 0.00013 |  |  |  |  | 0.99749 | 0.99749 | Yes | Yes |  | 35 days old wtwerror revised July 2022 from 0.26 to 0.13 based on new analysis of corals R41347. This resulted in a change in CRA error from 190.3323755000963192. If Saved values show old error bar then the revised result has not been re-reported to the client. |  |  |  | 0.00552 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00574 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00552 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-11-21 | margaret norris | 8:31 |  | Completed |  |  |  |  | 71188 | 2000000 |  |  | 71234 71246 71253 71254 71269 71314 71320 | 71234~71246~71253~71254~71269~71314~71320 | 3091 |  | 3091 3091 3091 3091 3091 3091 3091 | 2016-11-28 | 9:32 | albert zondervan |  | 1.24 mg | 0.13 | 1.24 | [5 empty] | 1.0003 | 0.00197 | 71234 |  | XCAMS | EA | 0% |  | TW3091 measurement: No tuning required as we've come off back of TW3090. error adjustments: * factor 2.7 for run-to-run variability of plant TP71277 * factor 1.3 for run-to-run variability of the small sediment TP71313 * no other adjustments warranted unknowns: * TP71277 generated an average rts that has been derived from a far too variable set of individual runs. I do not trust this average to be correct and therefore I call it a fail. * None of the other cathodes in the wheel show such a behaviour. * Small sediment TP71313 was blank-corrected along with all larges. primaries: The wheel contained 7x cathodes with OxI: all via EA. All of them performed equally well, so I changed the category for TP71269 to primary standard. blanks: Two groups have been identified in this wheel to blank-correct: * organics via AAA: 26x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied * organics via CellExtr: 0x unknowns + 1x secondaries + 0x blanks; a 360-days window was applied to find a decent number of blanks to go with it secondaries: * Both secondaries (FIRI-I, D) agree with their expectation value. |  | RRL | [5 empty] | Unknown Inorganic |  | 2016-12-09 | Water | Complete |  | TEST Water blank test. Flush with nitrogen before added sample as pieces , C13 readings | 62190 | Normal | 14047/11 |  |  | 2016-10-20 | 2765 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.00184 | 71246 | [25 empty] | 2016-10-20 | 1638 | Gas Splitting | Completed
[168 empty] | 0.99761 | 0.00188 | 71253 | [25 empty] | 2016-10-20 | 3079 | Choose Gas Splitting | Completed
[168 empty] | 0.99791 | 0.00192 | 71254 | [25 empty] | 2016-10-21 | 11198 | Graphite | Completed
[168 empty] | 1.0015 | 0.00193 | 71269 | [25 empty] | 2016-11-28 | 12944 | AMS Submission Results | Completed
[168 empty] | 1.0011 | 0.00185 | 71314 | [25 empty] | 2016-12-05 | 15340 | Close Job | Completed
[168 empty] | 1.0019 | 0.00188 | 71320
[7 empty] | 99.8 | 88.6789 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37836 | 37835 | 37836 | 111 | 111 | 111 |  |  | Margaret Norris | -991 | 0.12 ‰ | 2016-12-06 | 2016-12-06 | 11 | 0 |  |  |  |  | 0.91‰ | 0.91‰ | 2.00‰ | 2.00‰ | 0.91 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -991.00 ‰ | -991.00 ‰ | 0.12 | 0.12 | -990.88 | 0.13 ‰ | -990.88 | 0.13 ‰ |  |  | 2016-12-08 | Jocelyn Turnbull | 11:30 | 0.0090046 | 0.0090046 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.7 |  |  | 0.009 | 0.00012 | 13036 | 2016-12-08 | 8:33 | Albert Zondervan | 205820 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 37 | 5 | 6 | ... | 0.00866 | 0.00012 |  |  |  |  | 0.99741 | 0.99741 | Yes | Yes | [5 empty] | 0.00866 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.0005 | 0.0005 | 0.009 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00866 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2016-12-05 | Margaret Norris | 9:05 |  | Completed |  |  |  |  | 71386 | 2000000 |  |  | 71321 71359 71360 71368 71369 71455 | 71321~71359~71360~71368~71369~71455 | 3093 |  | 3093 3093 3093 3093 3093 3093 | 2016-12-08 | 9:48 | Albert Zondervan |  | 0.79 mg |  | 0.79 | [5 empty] | 0.99847 | 0.00121 | 71321 |  | XCAMS | EA | 0% |  | TW3093 measurement: * no issues CalAMS error adjustments: * none primaries: * The wheel contained 6x cathodes with OxI: all via EA. Three groups have been identified in this wheel to blank-correct: * ANSTO cellulose extraction: 2x unknowns + 1x secondaries + 1x blanks a 30-days window was applied leaving only the one blank to blank-subtract with 1x FIRI-D: result is 2.5-sigma below consensus value * other organics: a 180-days window was applied AAA pretreated: 7x unknowns + 0x secondaries + 1x blanks untreated: 17x unknowns + 1x secondaries + 0x blanks 1x FIRI-I: result is consistent with consensus value * waters: 1x unknowns + 0x secondaries + 1x blanks unknowns: * no issues | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-01-05 | Water |  |  | replaces sample lost in graphite | 62268 | Normal | 14047/11 |  |  | 2016-11-24 | 2768 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00127 | 71359 | [25 empty] | 2016-11-24 | 1645 | Gas Splitting | Completed
[168 empty] | 0.99883 | 0.0012 | 71360 | [25 empty] | 2016-11-24 | 3096 | Choose Gas Splitting | Completed
[168 empty] | 0.99959 | 0.00121 | 71368 | [25 empty] | 2016-11-25 | 11396 | Graphite | Completed
[168 empty] | 1.0013 | 0.00123 | 71369 | [25 empty] | 2016-12-09 | 4808 | C13 Measurement | Completed
[168 empty] | 1.0018 | 0.00123 | 71455 | [25 empty] | 2016-12-08 | 13036 | AMS Submission Results | Completed
[196 empty] | 2016-12-16 | 15541 | Close Job | Completed
[7 empty] | 95.3 | 88.3985 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42010 | 42010 | 42010 | 187 | 187 | 187 |  |  | margaret norris | -994.65 | 0.12 ‰ | 2017-02-10 | 2017-02-10 | 22 | 0 |  |  |  |  | 3.59‰ | 3.59‰ | 2.00‰ | 2.00‰ | 3.59 ‰ | 2.00 ‰ | AMS | 2.27‰ | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -994.65 ‰ | -994.65 ‰ | 0.12 | 0.12 | -994.58 | 0.13 ‰ | -994.58 | 0.13 ‰ |  |  | 2017-02-13 | Jocelyn Turnbull | 11:07 | 0.0053549 | 0.0053549 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.7 |  |  | 0.00535 | 0.00012 | 13133 | 2017-02-13 | 8:17 | Albert Zondervan | 206002 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 10 | 5 | 6 | ... | 0.00515 | 0.00012 |  |  |  |  | 0.99692 | 0.99692 | Yes | Yes | [5 empty] | 0.00515 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00075 | 0.00075 | 0.00535 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00515 | 0 | 0.00012 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-02 | margaret norris | 10:52 |  | Completed |  |  |  |  | 71554 | 2000000 |  |  | 71511 71512 71513 71595 71596 71597 | 71511~71512~71513~71595~71596~71597 | 3096 |  | 3096 3096 3096 3096 3096 3096 | 2017-02-13 | 10:49 | Albert Zondervan |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99691 | 0.00182 | 71511 |  | XCAMS | ST | 0% |  | TW3096 measurement: * first wheel of year 2017 AD * unknowns are all waters * large samples only * XCAMS worked fine, except for anomalously low Cs and ionizer settings to obtain normal beam intensity CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: cat#14: factor 1.3 * 0% system error primaries: * The wheel contained 6x cathodes with large OxI, combusted via ST. One group of unknowns is identified in this wheel to blank-correct: * waters: 28x unknowns + 1x secondaries (IAEA-C2, R14047/12) + 1x blanks (IAEA-C1, R14047/11) 4x blanks were found for last 180 days IAEA-C2 comes out consistent with consensus value | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-02-10 | Water | Complete |  |  | 62376 | Normal | 14047/11 |  |  | 2017-01-17 | 2789 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99979 | 0.00243 | 71512 | [25 empty] | 2017-01-17 | 1659 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00187 | 71513 | [25 empty] | 2017-01-17 | 3121 | Choose Gas Splitting | Completed
[168 empty] | 1.0025 | 0.00172 | 71595 | [25 empty] | 2017-01-19 | 11564 | Graphite | Completed
[168 empty] | 1 | 0.00172 | 71596 | [25 empty] | 2017-01-26 | 4830 | C13 Measurement | Completed
[168 empty] | 0.99951 | 0.00171 | 71597 | [25 empty] | 2017-02-13 | 13133 | AMS Submission Results | Completed
[196 empty] | 2017-02-13 | 15677 | Close Job | Completed
[7 empty] | 95.1 | 88.5976 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41780 | 41779 | 41780 | 197 | 197 | 197 |  |  | Margaret Norris | -994.49 | 0.14 ‰ | 2017-02-21 | 2017-02-21 | 6 | 0 |  |  |  |  | 2.58‰ | 2.58‰ | 2.00‰ | 2.00‰ | 2.58 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -994.49 ‰ | -994.49 ‰ | 0.14 | 0.14 | -994.42 | 0.14 ‰ | -994.42 | 0.14 ‰ |  |  | 2017-02-27 | Jocelyn Turnbull | 15:31 | 0.0055109 | 0.0055109 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 3 |  |  | 0.00551 | 0.00014 | 13220 | 2017-02-22 | 15:26 | Albert Zondervan | 206056 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 13 | 5 | 6 | ... | 0.0053 | 0.00013 |  |  |  |  | 0.99815 | 0.99815 | Yes | Yes | [5 empty] | 0.0053 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00551 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0053 | 0 | 0.00013 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-02-20 | Margaret Norris | 8:38 |  | Completed |  |  |  |  | 71702 | 2000000 |  |  | 71580 71581 71601 71602 71615 71680 | 71580~71581~71601~71602~71615~71680 | 3098 |  | 3098 3098 3098 3098 3098 3098 | 2017-02-23 | 13:10 | Albert Zondervan |  | 0.69 mg |  | 0.69 | [5 empty] | 1.0026 | 0.00164 | 71580 |  | XCAMS | EA | 0% |  | TW3098 measurement: * no issues CalAMS error adjustments: * run-to-run variability of whole dataset is ok * factor 1.3 to account for run-to-run variability of cathodes TP71601 and 71599 * 0% system error primaries: * The wheel contained 6x cathodes with OxI: all via EA. secondaries: * 1x IAEA-C2: result consistent with expectation value * 1x FIRI-I: result consistent with expectation value blanks: * Four groups have been identified to blank-correct: - waters: 6x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - carbonate: 2x unknowns + 0x secondaries + 0x blanks; a 180-days window was applied - organics via AAA and other non-CE: 14x unknowns + 1x secondaries + 1x blanks; a 180-days window was applied - organics via CE-GNSmethod: 5x unknowns + 0x secondaries + 1x blanks; a 180-days window was applied unknowns: * no issues | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-02-15 | Water |  |  |  | 62483 | Normal | 14047/11 | HAT |  | 2017-02-13 | 2831 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99774 | 0.00166 | 71581 | [25 empty] | 2017-02-13 | 1678 | Gas Splitting | Completed
[168 empty] | 0.9998 | 0.00217 | 71601 | [25 empty] | 2017-02-13 | 3148 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00169 | 71602 | [25 empty] | 2017-02-15 | 11712 | Graphite | Completed
[168 empty] | 1.0005 | 0.00169 | 71615 | [25 empty] | 2017-02-21 | 4871 | C13 Measurement | Completed
[168 empty] | 0.99868 | 0.00169 | 71680 | [25 empty] | 2017-02-27 | 13220 | AMS Submission Results | Completed
[196 empty] | 2017-03-01 | 15839 | Close Job | Completed
[7 empty] | 91.3 | 88.5952 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39352 | 39352 | 39352 | 156 | 156 | 156 |  |  | jenny dahl | -992.54 | 0.15 ‰ | 2017-03-17 | 2017-03-17 | 21 | 0 |  |  |  |  | 2.71‰ | 2.71‰ | 2.00‰ | 2.00‰ | 2.71 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -992.54 ‰ | -992.54 ‰ | 0.15 | 0.15 | -992.45 | 0.15 ‰ | -992.45 | 0.15 ‰ |  |  | 2017-03-20 | Jocelyn Turnbull | 13:02 | 0.0074553 | 0.0074553 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.2 |  |  | 0.00746 | 0.00015 | 13411 | 2017-03-20 | 8:36 | Albert Zondervan | 206559 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.75 | 0.01% | 15 | 5 | 6 | .X. | 0.00717 | 0.00014 |  |  |  |  | 0.99826 | 0.99826 | Yes | Yes | [5 empty] | 0.00717 | 0.01‰ | 0.00014 | 0.00‰ | No | 1 | 1 | 0.00087 | 0.00087 | 0.00746 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00717 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-14 | jenny dahl | 17:44 |  | Completed |  |  |  |  | 71759 | 2000000 |  |  | 71747 71762 71802 71840 71843 71844 | 71747~71762~71802~71840~71843~71844 | 3103 |  | 3103 3103 3103 3103 3103 3103 | 2017-03-20 | 9:37 | Albert Zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99774 | 0.00213 | 71747 |  | XCAMS | EA | 0.14% |  | TW3103 measurement: * virTW=ASE1 * no ST * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.14% system error primaries: * 6x large OxI groups identified to blank-correct: * water: 3x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 5x blanks IAEA-C2 (TP71760) is deviating >3-sigma from consensus value - TEST SAMPLE USING POWDER MIXED IN WITH WATER - NOT A WORRY RIGHT NOW * organic via ANSTO-method cellulose extraction: 12x unknowns + 2x secondaries + 2x blanks a 180-day window was applied, capturing 5x blanks both FIRI-D (TP71749,71798) are deviating from consensus value, towards lower F NOTE WE HAVE CHECKED THE CONSENSUS VALUE AND IT IS 4532 FOR AMS LABS WHICH MAKES THESE MUCH CLOSER TO THE EXPECTED VALUE * organic via AAA: 9x unknowns + 1x secondaries + 1x blanks a 180-day window was applied, capturing 11x blanks FIRI-I (TP71838) is 1-sigma away from consensus value | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-04-18 | Water |  | JCT ADDED INVOICE 540GCT21 | test with marble in DI water | 62622 | Normal | 14047/11 |  |  | 2017-02-22 | 2838 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00212 | 71762 | [25 empty] | 2017-02-22 | 1686 | Gas Splitting | Completed
[168 empty] | 0.99878 | 0.00213 | 71802 | [25 empty] | 2017-02-22 | 3161 | Choose Gas Splitting | Completed
[168 empty] | 0.99881 | 0.00213 | 71840 | [25 empty] | 2017-02-24 | 11769 | Graphite | Completed
[168 empty] | 1.0038 | 0.00217 | 71843 | [25 empty] | 2017-02-27 | 4897 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00216 | 71844 | [25 empty] | 2017-03-20 | 13411 | AMS Submission Results | Completed
[196 empty] | 2017-03-27 | 16035 | Close Job | Completed
[196 empty] | 2017-03-27 | 5024 | Invoice | Completed
[7 empty] | 98.9 | 88.5038 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38422 | 38422 | 38422 | 108 | 108 | 108 |  |  | jenny dahl | -991.63 | 0.11 ‰ | 2017-03-24 | 2017-03-24 | 3 | 0 |  |  |  |  | 3.66‰ | 3.66‰ | 2.00‰ | 2.00‰ | 3.66 ‰ | 2.00 ‰ | AMS | 2.58‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0.89 | 0.89443 | -991.63 ‰ | -991.63 ‰ | 0.11 | 0.11 | -991.53 | 0.11 ‰ | -991.53 | 0.11 ‰ |  |  | 2017-03-27 | Jocelyn Turnbull | 10:06 | 0.0083703 | 0.0083703 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.5 |  |  | 0.00837 | 0.00011 | 13442 | 2017-03-27 | 8:02 | Albert Zondervan | 206721 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:45 | Jocelyn Turnbull | 0 |  | 0.84 | 0.01% | 17 | 5 | 6 | ... | 0.00805 | 0.00011 |  |  |  |  | 0.99818 | 0.99818 | Yes | Yes | [5 empty] | 0.00805 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00112 | 0.00112 | 0.00837 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00805 | 0 | 0.00011 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-03-23 | jenny dahl | 16:30 |  | Completed |  |  |  |  | 71907 | 2000000 |  |  | 71885 71886 71935 71937 71938 | 71885~71886~71935~71937~71938 | 3104 |  | 3104 3104 3104 3104 3104 | 2017-03-27 | 8:31 | Albert Zondervan |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0024 | 0.00251 | 71885 |  | XCAMS | ST | 0.21% |  | TW3104 measurement: * virTW=KID * no EA * large samples * no issues with XCAMS CalAMS error adjustments: * uncertainty multiplier, due to excess run-to-run variability: none * 0.21% system error primaries: * 6x large OxI, of which 1x was rejected as calibration standard: TP71936 is an outlier in 13/12 and rts. Cathy is querying this. groups identified to blank-correct: * bone + organic : 3x EXPRESS bone unknowns + 1x organic unknown + 2x secondaries + 1x blanks a 180-day window was applied FIRI-L whalebone and IAEA-C6 sucrose secondaries come out good * carbonates: 8x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good * water: 5x unknowns + 1x secondaries + 1x blanks a 180-day window was applied IAEA-C2 travertine secondary comes out good Some IRMS del13C values are still to be reported. | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2017-05-09 | Water | Not Complete |  | to replace sample that was lost into a melted tube in graphite | 62703 | Normal | 14047/11 |  |  | 2017-03-17 | 2863 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0028 | 0.00251 | 71886 | [25 empty] | 2017-03-17 | 1701 | Gas Splitting | Completed
[168 empty] | 0.99735 | 0.0025 | 71935 | [25 empty] | 2017-03-17 | 3187 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.0025 | 71937 | [25 empty] | 2017-03-21 | 11917 | Graphite | Completed
[168 empty] | 0.99948 | 0.00253 | 71938 | [25 empty] | 2017-03-27 | 4943 | C13 Measurement | Completed
[196 empty] | 2017-03-27 | 13442 | AMS Submission Results | Completed
[196 empty] | 2017-03-29 | 16101 | Close Job | Completed
[7 empty] | 130.3 | 88.2694 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37725 | 37725 | 37725 | 137 | 137 | 137 |  |  | margaret norris | -990.87 | 0.16 ‰ | 2017-05-05 | 2017-05-05 | 24 | 0 |  |  |  |  | 5.18‰ | 5.18‰ | 2.00‰ | 2.00‰ | 5.18 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -990.87 ‰ | -990.87 ‰ | 0.16 | 0.16 | -990.76 | 0.16 ‰ | -990.76 | 0.16 ‰ |  |  | 2017-05-08 | Jocelyn Turnbull | 6:59 | 0.0091294 | 0.0091294 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00913 | 0.00016 | 13673 | 2017-05-06 | 13:53 | Albert Zondervan | 206979 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.91 | 0.02% | 20 | 5 | 6 | ... | 0.00878 | 0.00015 |  |  |  |  | 0.99704 | 0.99704 | Yes | Yes | [5 empty] | 0.00878 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00913 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00878 | 0 | 0.00015 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-04 | margaret norris | 11:18 |  | Completed |  |  |  |  | 72071 | 2000000 |  |  | 72120 72122 72125 72171 72175 72198 | 72120~72122~72125~72171~72175~72198 | 3110 |  | 3110 3110 3110 3110 3110 3110 | 2017-05-06 | 14:12 | Albert Zondervan |  | 0.92 mg |  | 0.92 | [5 empty] | 0.9985 | 0.00158 | 72120 |  | XCAMS | EA | 0% |  | TW3110 measurement: * no problems with XCAMS * normal precision * EA + waterline CalAMS error adjustments: * 0% system error * factor 1.4 for TP72112 because of excess run-to-run variability primaries: * The wheel contained 6x cathodes with OxI secondaries: * 1x IAEA-C2: results consistent with expectation value * 1x FIRI-I: results consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 13x unknowns + 1x secondaries + 1x blanks waters: 15x unknowns + 1x secondaries + 1x blanks * Waterline blank is rather high (again). unknowns: * no failures * partial loss of graphite in loading TP72124 has not affected its precision | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2017-05-30 | Water | Complete |  |  | 62930 | Normal | 14047/11 |  |  | 2017-04-10 | 2881 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0019 | 0.00155 | 72122 | [25 empty] | 2017-04-10 | 1733 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00155 | 72125 | [25 empty] | 2017-04-10 | 3228 | Choose Gas Splitting | Completed
[168 empty] | 0.99954 | 0.00158 | 72171 | [25 empty] | 2017-04-11 | 12081 | Graphite | Completed
[168 empty] | 0.99833 | 0.00157 | 72175 | [25 empty] | 2017-04-19 | 4993 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00159 | 72198 | [25 empty] | 2017-05-08 | 13673 | AMS Submission Results | Completed
[196 empty] | 2017-05-17 | 16426 | Close Job | Completed
|  |  | 0 | 0 | 5.1163e-05 | 0.0075136 | 120.1 | 88.5552 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37341 | 37340 | 37341 | 122 | 122 | 122 | 37219.7 | 37464 | margaret Norris | -990.42 | 0.15 ‰ | 2017-05-12 | 2016-11-13 | 15 | 180 |  |  |  |  | 1.55‰ | 1.55‰ | 2.00‰ | 2.00‰ | 1.55 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -990.42 ‰ | -990.42 ‰ | 0.15 | 0.15 | -990.31 | 0.15 ‰ | -990.31 | 0.15 ‰ |  |  | 2017-05-15 | Jocelyn Turnbull | 11:02 | 0.0095765 | 0.0095765 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.8 |  |  | 0.00958 | 0.00015 | 13758 | 2017-05-15 | 9:14 | Albert Zondervan | 207067 | 0 | 0 | 0 | 0 | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.96 | 0.01% | 22 | 5 | 7 | ... | 0.00921 | 0.00014 |  |  |  |  | 0.99666 | 0.99666 | Yes | Yes |  |  | 0.00728 | 0.00722 | 0.0015 | 0.00921 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 0.99999 | 0.00062 | 0.00062 | 0.00958 | 0 | 0.00015 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00921 | 0 | 0.00014 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-11 | margaret Norris | 11:39 |  | Completed |  |  |  |  | 72164 | 2000000 | 72237 72071 71907 71759 71702 71554 71386 | 72237~72071~71907~71759~71702~71554~71386 | 72069 72088 72089 72090 72153 72163 | 72069~72088~72089~72090~72153~72163 | 3112 | 3112 3110 3104 3103 3098 3096 3093 | 3112 3112 3112 3112 3112 3112 | 2017-05-15 | 10:20 | Albert Zondervan | Yes | 0.70 mg |  | 0.7 | 2017-05-12 | 0.00788 | 72237 | 3112 | 0.89 | 0.99735 | 0.00145 | 72069 | EA | XCAMS | ST | 0% |  | TW3112 measurement: * no problems with XCAMS * normal precision * ST + waterline CalAMS error adjustments: * 0% system error * large factor 3.2 for TP72253 because of excess run-to-run variability (see comment below) primaries: * The wheel contained 6x cathodes with OxI, with 450k counts each on average. secondaries: * 1x FIRI-I: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 25x unknowns + 1x secondaries + 0x blanks (+ 1x EA blank) waters: 2x unknowns + 0x secondaries + 1x blanks + 2x tests * The waterline test blanks were forced to category 'Test Background' and to receive no blank correction. unknowns: * no failures, but: * The result for TP72253 (water unknown) is suspect because of trends in the raw data suggesting ratio inhomogeneity. | 2.3 | RRL | [5 empty] | Background Test |  | 2017-06-21 | Water | Complete | 540GCT23 (test sample) | test job to compare with Kapuni bubble -flushed before addition -chunk rather than powder acid heated in water bath to melt marble chip | 63010 | Normal | 14047/11 |  |  | 2017-04-26 | 2895 | Water CO2 Evolution | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 1.0001 | 0.00149 | 72088 | ST | [24 empty] | 2017-04-26 | 1738 | Gas Splitting | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 0.99931 | 0.00147 | 72089 | EA | [24 empty] | 2017-04-26 | 3238 | Choose Gas Splitting | Cancelled
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0013 | 0.00169 | 72090 | EA | [24 empty] | 2017-04-26 | 2897 | Water CO2 Evolution | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0014 | 0.00151 | 72153 | ST | [24 empty] | 2017-04-27 | 12174 | Graphite | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 | 1.0009 | 0.0015 | 72163 | EA | [24 empty] | 2017-05-01 | 5025 | C13 Measurement | Completed
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79 |  |  |  | ST | [24 empty] | 2017-05-15 | 13758 | AMS Submission Results | Completed
[196 empty] | 2017-05-17 | 16433 | Close Job | Completed
[196 empty] | 2017-05-24 | 5180 | Invoice | Completed
|  |  | 0 | 0 | 2.927e-05 | 0.0034707 | 123.8 | 88.4731 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38594 | 38593 | 38594 | 173 | 173 | 173 |  |  | margaret Norris | -991.81 | 0.18 ‰ | 2017-05-12 | 2016-11-13 | 7 | 180 |  |  |  |  | 2.48‰ | 2.48‰ | 2.00‰ | 2.00‰ | 2.48 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0017 | 0 | 0.8165 | -991.81 ‰ | -991.81 ‰ | 0.18 | 0.18 | -991.71 | 0.18 ‰ | -991.71 | 0.18 ‰ |  |  | 2017-05-15 | Jocelyn Turnbull | 11:02 | 0.0081936 | 0.0081936 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.2 |  |  | 0.00819 | 0.00018 | 13764 | 2017-05-15 | 9:14 | Albert Zondervan | 207150 | 0 | 0 | 0 | 0 | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.82 | 0.02% | 3 | 5 | 6 | ... | 0.00788 | 0.00017 |  |  |  |  | 0.99666 | 0.99666 | Yes | Yes |  |  | 0.0054 | 0.00333 | 0.00299 | 0.00788 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 0.99999 | 0.00062 | 0.00062 | 0.00819 | 0 | 0.00018 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00788 | 0 | 0.00017 | 0 | Regular |  | 0.9878 | 0.95 | 1 | 2017-05-11 | margaret Norris | 11:39 |  | Completed |  |  |  |  | 72237 | 2000000 | 72071 72000 71898 71907 71759 71702 71554 71443 71386 | 72071~72000~71898~71907~71759~71702~71554~71443~71386 | 72069 72088 72089 72090 72153 72163 | 72069~72088~72089~72090~72153~72163 | 3112 | 3110 3106 3104 3104 3103 3098 3096 3094 3093 | 3112 3112 3112 3112 3112 3112 | 2017-05-15 | 10:20 | Albert Zondervan | Yes | 0.89 mg |  | 0.89 | 2017-05-05 | 0.00878 | 72071 | 3110 | 0.92 | 0.99735 | 0.00145 | 72069 | EA | XCAMS | ST | 0% |  | TW3112 measurement: * no problems with XCAMS * normal precision * ST + waterline CalAMS error adjustments: * 0% system error * large factor 3.2 for TP72253 because of excess run-to-run variability (see comment below) primaries: * The wheel contained 6x cathodes with OxI, with 450k counts each on average. secondaries: * 1x FIRI-I: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics via AAA: 25x unknowns + 1x secondaries + 0x blanks (+ 1x EA blank) waters: 2x unknowns + 0x secondaries + 1x blanks + 2x tests * The waterline test blanks were forced to category 'Test Background' and to receive no blank correction. unknowns: * no failures, but: * The result for TP72253 (water unknown) is suspect because of trends in the raw data suggesting ratio inhomogeneity. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-06-28 | Water | Complete |  |  | 63011 | Normal | 14047/11 |  |  | 2017-05-03 | 2900 | Water CO2 Evolution | Completed | 14047/1 | IAEA-C1: Carrara Marble | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[163 empty] | 2017-04-06 | 0.00148 | 72000 | 3106 | 0.76 | 1.0001 | 0.00149 | 72088 | EA | [24 empty] | 2017-05-03 | 1743 | Gas Splitting | Completed | 14047/11 | IAEA-C1: Carrara Marble WATER LINE
[163 empty] | 2017-03-24 | 0.0014 | 71898 | 3104 | 1.05 | 0.99931 | 0.00147 | 72089 | ST | [24 empty] | 2017-05-03 | 3245 | Choose Gas Splitting | Completed
[163 empty] | 2017-03-24 | 0.00805 | 71907 | 3104 | 0.87 | 1.0013 | 0.00169 | 72090 | EA | [24 empty] | 2017-05-05 | 12247 | Graphite | Completed
[163 empty] | 2017-03-17 | 0.00717 | 71759 | 3103 | 0.89 | 1.0014 | 0.00151 | 72153 | EA | [24 empty] | 2017-05-08 | 5057 | C13 Measurement | Completed
[163 empty] | 2017-02-21 | 0.0053 | 71702 | 3098 | 0.69 | 1.0009 | 0.0015 | 72163 | ST | [24 empty] | 2017-05-15 | 13764 | AMS Submission Results | Completed
[163 empty] | 2017-02-10 | 0.00515 | 71554 | 3096 | 1.18 |  |  |  | EA | [24 empty] | 2017-05-17 | 16436 | Close Job | Completed
[163 empty] | 2016-12-12 | 0.00258 | 71443 | 3094 | 0.99 |  |  |  | EA
[163 empty] | 2016-12-06 | 0.00866 | 71386 | 3093 | 0.79
[7 empty] | 119.9 | 88.5589 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41948 | 41948 | 41948 | 186 | 185 | 186 |  |  | jenny dahl | -994.6 | 0.12 ‰ | 2017-05-24 | 2017-05-24 | 2 | 0 |  |  |  |  | 1.63‰ | 1.63‰ | 2.00‰ | 2.00‰ | 1.63 ‰ | 2.00 ‰ | AMS | 2.11‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -994.60 ‰ | -994.60 ‰ | 0.12 | 0.12 | -994.54 | 0.13 ‰ | -994.54 | 0.13 ‰ |  |  | 2017-05-26 | Jocelyn Turnbull | 12:24 | 0.0053965 | 0.0053965 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.0054 | 0.00012 | 13835 | 2017-05-25 | 14:45 | Albert Zondervan | 207250 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 16 | 5 | 6 | ... | 0.00519 | 0.00012 |  |  |  |  | 0.99677 | 0.99677 | Yes | Yes | [5 empty] | 0.00519 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.0054 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00519 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-05-24 | jenny dahl | 8:19 |  | Completed |  |  |  |  | 72341 | 2000000 |  |  | 72240 72250 72317 72319 72326 72336 | 72240~72250~72317~72319~72326~72336 | 3114 |  | 3114 3114 3114 3114 3114 3114 | 2017-05-26 | 12:04 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 1.0007 | 0.00164 | 72240 |  | XCAMS | ST | 0% |  | TW3114 measurement: * no problems with XCAMS * normal precision * large samples, except 2 organic unknowns just under 0.3 mg C CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes secondaries: * 1x FIRI-I: result is 2-sigma deviant from expectation value * 1x FIRI-L: result consistent with expectation value * 1x IAEA-C2: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: organics via AAA: 17x unknowns + 2x secondaries + 1x blanks waters: 7x unknowns + 1x secondaries + 1x blanks shells: 3x unknowns + 0x secondaries + 0x blanks * Two organic unknowns are below the 0.3 mgC threshold: TP72206,72207. It was checked if blank-correcting them with the Santos method (using the spreadsheet for smalls) makes a difference. On the basis of one size-matching Kauri Renton Rd blank (TP71784, TW3100), MCC and DCC for these two unknowns were calculated. Changing to these MCC values makes no significant difference, but changing to these DCC values has a significant effect. This is all explained qualitatively by the fact that the unknowns are near-modern, for which we expect the blank correction to be dominated by the dead carbon component. Hence, the Santos correction has been applied for those two unknowns. unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-07-10 | Water |  |  | to be run with Troy's samples | 63089 | Normal | 14047/11 |  |  | 2017-05-18 | 2920 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99918 | 0.00164 | 72250 | [25 empty] | 2017-05-18 | 1758 | Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00166 | 72317 | [25 empty] | 2017-05-18 | 3263 | Choose Gas Splitting | Completed
[168 empty] | 0.99904 | 0.00164 | 72319 | [25 empty] | 2017-05-22 | 12351 | Graphite | Completed
[168 empty] | 0.99838 | 0.00165 | 72326 | [25 empty] | 2017-05-24 | 5102 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00165 | 72336 | [25 empty] | 2017-05-26 | 13835 | AMS Submission Results | Completed
[196 empty] | 2017-05-29 | 16537 | Close Job | Completed
[7 empty] | 133.5 | 88.3038 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41795 | 41795 | 41795 | 167 | 167 | 167 |  |  | julia collins | -994.5 | 0.11 ‰ | 2017-06-28 | 2017-06-28 | 6 | 0 |  |  |  |  | 3.06‰ | 3.06‰ | 2.00‰ | 2.00‰ | 3.06 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0008 | 0 | 0.8165 | -994.50 ‰ | -994.50 ‰ | 0.11 | 0.11 | -994.43 | 0.12 ‰ | -994.43 | 0.12 ‰ |  |  | 2017-06-29 | Jocelyn Turnbull | 10:55 | 0.0055005 | 0.0055005 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.9 |  |  | 0.0055 | 0.00011 | 14038 | 2017-06-29 | 9:51 | Albert Zondervan | 207397 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 31 | 5 | 6 | ... | 0.00529 | 0.00011 |  |  |  |  | 0.99532 | 0.99532 | Yes | Yes | [5 empty] | 0.00529 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.0055 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00529 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-06-27 | julia collins | 13:26 |  | Completed |  |  |  |  | 72525 | 2000000 |  |  | 72487 72488 72489 72515 72529 72547 | 72487~72488~72489~72515~72529~72547 | 3119 |  | 3119 3119 3119 3119 3119 3119 | 2017-06-29 | 10:10 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99902 | 0.0015 | 72487 |  | XCAMS | ST | 0% |  | TW3119 measurement: * no problems with XCAMS * large samples CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes, via ST secondaries: * IAEA-C6 (tune cathode): result consistent with expectation value * IAEA-C2 via water-line: result 3-sigma older than expectation CRA * FIRI-I cellulose: result consistent with expectation value * FIRI-L whalebone: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)water: 14x unknowns + 1x secondaries + 1x blanks bone + other organics: 7x unknowns + 3x secondaries + 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-08-16 | Water | Not Complete |  | gas too old to use for c13 | 63289 | Normal | 14047/11 |  |  | 2017-06-21 | 2955 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00148 | 72488 | [25 empty] | 2017-06-21 | 2963 | Water CO2 Evolution | Completed
[168 empty] | 1.0007 | 0.00149 | 72489 | [25 empty] | 2017-06-21 | 1787 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00154 | 72515 | [25 empty] | 2017-06-21 | 3305 | Choose Gas Splitting | Completed
[168 empty] | 0.99954 | 0.0015 | 72529 | [25 empty] | 2017-06-22 | 12535 | Graphite | Completed
[168 empty] | 0.99924 | 0.00152 | 72547 | [25 empty] | 2017-06-29 | 14038 | AMS Submission Results | Completed
[196 empty] | 2017-07-07 | 16879 | Close Job | Cancelled
[7 empty] | 121.1 | 88.4253 | 0.0025 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33790 | 33789 | 33790 | 106 | 106 | 106 |  |  | margaret norris | -985.1 | 0.20 ‰ | 2017-08-24 | 2017-08-24 | 8 | 0 |  |  |  |  | 2.58‰ | 2.58‰ | 2.00‰ | 2.00‰ | 2.58 ‰ | 2.00 ‰ | AMS | 2.14‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -985.10 ‰ | -985.10 ‰ | 0.2 | 0.2 | -984.92 | 0.20 ‰ | -984.92 | 0.20 ‰ |  |  | 2017-08-28 | Cathy Ginnane | 15:44 | 0.0149 | 0.0149 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2.3 |  |  | 0.0149 | 0.0002 | 14277 | 2017-08-28 | 9:01 | Albert Zondervan | 207533 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 1.49 | 0.02% | 8 | 5 | 6 | ... | 0.01433 | 0.00019 |  |  |  |  | 0.99622 | 0.99622 | Yes | Yes |  | very high blank- no unusual observations were made during processing evolution or graphite of this blank material. |  |  |  | 0.01433 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00083 | 0.00083 | 0.0149 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01433 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-08-23 | margaret norris | 15:47 |  | Completed |  |  |  |  | 72804 | 2000000 |  |  | 72767 72768 72846 72849 72857 72858 | 72767~72768~72846~72849~72857~72858 | 3126 |  | 3126 3126 3126 3126 3126 3126 |  |  |  |  | 0.87 mg |  | 0.87 | [5 empty] | 1.0011 | 0.00204 | 72767 |  | XCAMS | EA | 0.12% |  | TW3126 measurement: * no problems with XCAMS * regular precision * large samples CalAMS error adjustments: * 0.12% system error category line (method) unkn/blan/seco/prim window [d] ----------¦--------------¦--------------------¦---------- inorganic carbonate 8/1/1/0 180 water water 3/1/1/0 180 organic EA (AAA) 15/1/2/6 180 unknowns: * no failures blanks: * Both IAEA-C1 blanks are uncharacteristically high relative to previous C1 blanks in previous wheels. The amounts by which they are high are about the same, suggesting there is a common contamination factor at play. Querying RLIMS does not reveal an obvious connection between these two preparations and no special comments have been made for them. These C1 blanks have not been ignored for the blank-subtraction calculation. The gas for these two blanks were put into spare M/S bottles not the usual M/s bottles. * The AAA-treated Kauri Renton Rd and Kapuni CO2 blanks come out consistent with previous analyses. secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x EA IAEA-C6: result consistent with consensus value * 1x EA FIRI-I: result consistent with consensus value primaries: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-09-07 | Water |  |  |  | 63499 | Normal | 14047/11 |  |  | 2017-08-15 | 2977 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00203 | 72768 | [25 empty] | 2017-08-15 | 1829 | Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00202 | 72846 | [25 empty] | 2017-08-15 | 3381 | Choose Gas Splitting | Completed
[168 empty] | 0.99693 | 0.00203 | 72849 | [25 empty] | 2017-08-16 | 12814 | Graphite | Completed
[168 empty] | 0.99913 | 0.00203 | 72857 | [25 empty] | 2017-08-21 | 5225 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00202 | 72858 | [25 empty] | 2017-08-28 | 14277 | AMS Submission Results | Completed
[196 empty] | 2017-08-30 | 17102 | Close Job | Completed
[7 empty] | 123.9 | 88.5148 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36616 | 36615 | 36616 | 128 | 127 | 128 |  |  | Julia Collins | -989.52 | 0.17 ‰ | 2017-09-06 | 2017-09-06 | 12 | 0 |  |  |  |  | 0.20‰ | 0.20‰ | 2.00‰ | 2.00‰ | 0.20 ‰ | 2.00 ‰ | AMS | 2.07‰ | 0.2‰ | C13 | No | -19 | -19.0019 | 0 | 0.8165 | -989.52 ‰ | -989.52 ‰ | 0.17 | 0.17 | -989.39 | 0.17 ‰ | -989.39 | 0.17 ‰ |  |  | 2017-09-08 | margaret norris | 8:35 | 0.010481 | 0.010481 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.8 |  |  | 0.01048 | 0.00017 | 14357 | 2017-09-07 | 11:40 | Albert Zondervan | 207703 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 1.05 | 0.02% | 16 | 5 | 8 | ... | 0.01008 | 0.00016 |  |  |  |  | 0.99486 | 0.99486 | Yes | Yes | [5 empty] | 0.01008 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.01048 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01008 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-05 | Julia Collins | 14:57 |  | Completed |  |  |  |  | 72897 | 2000000 |  |  | 72942 72958 72959 72960 72961 72971 | 72942~72958~72959~72960~72961~72971 | 3128 |  | 3128 3128 3128 3128 3128 3128 | 2017-09-07 | 12:08 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99943 | 0.00155 | 72942 |  | XCAMS | ST | 0% |  | TW3128 measurement: * no problems with XCAMS * normal precision * large samples * no EA (except IAEA-C6) CalAMS error adjustments: * 1.3 uncertainty multiplier for cat#2 (TP72937) to account for run-to-run variability * 0% system error primaries: * OxI cathodes: 6x no failures blanks: * Two groups have been identified to blank-correct: organics: 17x unknowns + 2x secondaries + 1x blanks window = 180 days waters: 10x unknowns + 1x secondaries + 1x blanks window = 180 days water-line blank is still pretty bad * Kapuni has low rts: good secondaries: * 1x IAEA-C2: consistent with expectation value * 1x TIRI-I: consistent with expectation value * 1x TIRI-L: consistent with expectation value unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2017-10-05 | Water |  |  |  | 63572 | Normal | 14047/11 |  |  | 2017-08-21 | 2988 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.002 | 0.00147 | 72958 | [25 empty] | 2017-08-21 | 1831 | Gas Splitting | Completed
[168 empty] | 0.99835 | 0.00153 | 72959 | [25 empty] | 2017-08-21 | 3385 | Choose Gas Splitting | Cancelled
[168 empty] | 1.0005 | 0.00155 | 72960 | [25 empty] | 2017-08-21 | 1832 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00154 | 72961 | [25 empty] | 2017-08-21 | 3386 | Choose Gas Splitting | Completed
[168 empty] | 0.99829 | 0.00158 | 72971 | [25 empty] | 2017-08-25 | 12907 | Graphite | Completed
[196 empty] | 2017-08-30 | 5241 | C13 Measurement | Completed
[196 empty] | 2017-09-08 | 14357 | AMS Submission Results | Completed
[196 empty] | 2017-09-08 | 17146 | Close Job | Completed
[7 empty] | 111.8 | 88.5151 | 0.001 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS |  | 0.2‰ | Estimate | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 1.9 |  |  |  |  | 14422 | 2017-09-11 | 8:48 | Albert Zondervan | 207917 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 1 | 5 | 5 | ... | 0.00399 | 6e-05 | [7 empty] | Yes | [5 empty] | 0.00399 | 0.00‰ | 6e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.00399 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73015 | 2000000 |  |  |  |  | 3130 | [6 empty] | 0.99 mg |  | 0.99 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Unknown Inorganic |  | 2017-11-02 | Water |  |  | carbonate wheel test | 64957 | Normal | 14047/11 |  |  | 2017-09-07 | 3027 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[196 empty] | 2017-09-07 | 1855 | Gas Splitting | Completed
[196 empty] | 2017-09-07 | 3424 | Choose Gas Splitting | Completed
[196 empty] | 2017-09-08 | 13025 | Graphite | Completed
[196 empty] | 2018-05-28 | 14422 | AMS Submission Results | Completed
[196 empty] | 2018-05-28 | 18924 | Close Job | Completed
[196 empty] | 2018-05-28 | 7021 | Invoice | Completed
[7 empty] | 115.8 | 88.571 | 0.0015 | RRL-UNIn-LG | Unknown Inorganic | 1950-01-01 | 1950-01-01 | [10 empty] | margaret norris |  |  | 2017-09-10 | 2017-09-10 | 2 | 0 | [10 empty] | AMS |  | 0.2‰ | Estimate | [5 empty] | -1000.00 ‰ | -1000.00 ‰ |  |  | -1000 | [5 empty] | 2018-05-28 | Jocelyn Turnbull | 14:08 | [6 empty] | 2.6 |  |  |  |  | 14423 | 2017-09-11 | 8:48 | Albert Zondervan | 207918 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  |  |  | 36 | 5 | 5 | ... | 0.00608 | 8e-05 | [7 empty] | Yes | [5 empty] | 0.00608 | 0.01‰ | 8e-05 | 0.00‰ | [12 empty] | 0 | 0 | 0.00608 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-08 | margaret norris | 15:49 |  | Completed |  |  |  |  | 73016 | 2000000 |  |  |  |  | 3130 | [6 empty] | 0.77 mg |  | 0.77 | [9 empty] | XCAMS |  |  |  | TW3130 measurement: * purpose: check on large water-line and carbonate-line blanks * no problems with XCAMS; same tune as for TW3129 * no unknowns, no secondaries, no OxIs * To obtain rts values relative to OxI, the imported rts values, which were determined relative to IAEA-C6, need to be multiplied by 1.45. This has been done in the table at the bottom of this message. blanks: * 2x carbonate-line blanks * 2x water-line blanks * 1x Kapuni * results after re-calibration to OxI: TP rts err 71516 0.00107 0.00004 72632 1.45 0.01 73015 0.00579 0.00009 73016 0.00882 0.00012 73017 0.00152 0.00004 73018 0.00131 0.00004 |  | RRL | [5 empty] | Unknown Inorganic |  | 2017-11-02 | Water |  |  | carbonate wheel test | 64958 | Normal | 14047/11 |  |  | 2017-09-07 | 3028 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[196 empty] | 2017-09-07 | 1856 | Gas Splitting | Completed
[196 empty] | 2017-09-07 | 3426 | Choose Gas Splitting | Completed
[196 empty] | 2017-09-08 | 13026 | Graphite | Completed
[196 empty] | 2018-05-28 | 14423 | AMS Submission Results | Completed
[196 empty] | 2018-05-28 | 18925 | Close Job | Completed
[196 empty] | 2018-05-28 | 7022 | Invoice | Completed
[7 empty] | 109.8 | 88.5635 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37580 | 37579 | 37580 | 135 | 134 | 135 |  |  | helen zhang | -990.7 | 0.16 ‰ | 2017-09-15 | 2017-09-15 | 8 | 0 |  |  |  |  | 3.25‰ | 3.25‰ | 2.00‰ | 2.00‰ | 3.25 ‰ | 2.00 ‰ | AMS | 2.30‰ | 0.2‰ | C13 | No | -19 | -19.0003 | 0 | 0.8165 | -990.70 ‰ | -990.70 ‰ | 0.16 | 0.16 | -990.59 | 0.16 ‰ | -990.59 | 0.16 ‰ |  |  | 2017-09-18 | Jocelyn Turnbull | 11:00 | 0.0092958 | 0.0092958 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.7 |  |  | 0.0093 | 0.00016 | 14497 | 2017-09-18 | 9:14 | Albert Zondervan | 207819 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:46 | Jocelyn Turnbull | 0 |  | 0.93 | 0.02% | 22 | 5 | 6 | ... | 0.00894 | 0.00015 |  |  |  |  | 0.99844 | 0.99844 | Yes | Yes | [5 empty] | 0.00894 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.0093 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00894 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-14 | helen zhang | 15:10 |  | Completed |  |  |  |  | 73012 | 2000000 |  |  | 73003 73007 73040 73041 73056 73057 | 73003~73007~73040~73041~73056~73057 | 3132 |  | 3132 3132 3132 3132 3132 3132 | 2017-09-18 | 10:31 | Albert Zondervan |  | 0.74 mg |  | 0.74 | [5 empty] | 1.0002 | 0.00173 | 73003 |  | XCAMS | EA | 0% |  | TW3132 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * Four groups have been identified to blank-correct: organics via AAA: 7x unknowns + 2x secondaries + 1x blanks window = 180 days: captures 10x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks water: 8x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks carbonate: 9x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 2.5 sigma * 1x IAEA-C2 on carbonate-line: result differs from consensus value by 2.0 sigma unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-10-16 | Water |  |  | New etched material carbosorb not changed | 63724 | Normal | 14047/11 |  |  | 2017-09-04 | 3014 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00169 | 73007 | [25 empty] | 2017-09-04 | 1852 | Gas Splitting | Completed
[168 empty] | 1.0027 | 0.0017 | 73040 | [25 empty] | 2017-09-04 | 3418 | Choose Gas Splitting | Completed
[168 empty] | 0.99938 | 0.00168 | 73041 | [25 empty] | 2017-09-07 | 13022 | Graphite | Completed
[168 empty] | 0.99918 | 0.00171 | 73056 | [25 empty] | 2017-09-11 | 5290 | C13 Measurement | Completed
[168 empty] | 0.99859 | 0.0016 | 73057 | [25 empty] | 2017-09-18 | 14497 | AMS Submission Results | Completed
[196 empty] | 2017-09-19 | 17286 | Close Job | Completed
[7 empty] | 113.3 | 88.7738 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34110 | 34110 | 34110 | 111 | 110 | 111 |  |  | Helen Zhang | -985.68 | 0.20 ‰ | 2017-09-29 | 2017-09-29 | 17 | 0 |  |  |  |  | 0.89‰ | 0.89‰ | 2.00‰ | 2.00‰ | 0.89 ‰ | 2.00 ‰ | AMS | 2.28‰ | 0.2‰ | C13 | No | -19 | -19.0035 | 0 | 0.8165 | -985.68 ‰ | -985.68 ‰ | 0.2 | 0.2 | -985.51 | 0.20 ‰ | -985.51 | 0.20 ‰ |  |  | 2017-10-03 | Cathy Ginnane | 9:31 | 0.014318 | 0.014318 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 3 |  |  | 0.01432 | 0.0002 | 14617 | 2017-10-02 | 9:26 | Albert Zondervan | 207874 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.43 | 0.02% | 31 | 5 | 4 | ... | 0.01377 | 0.00019 |  |  |  |  | 0.99846 | 0.99846 | Yes | Yes | [5 empty] | 0.01377 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.01432 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01377 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-09-28 | Helen Zhang | 10:27 |  | Completed |  |  |  |  | 73029 | 2000000 |  |  | 73087 73121 73122 73131 73132 73161 | 73087~73121~73122~73131~73132~73161 | 3135 |  | 3135 3135 3135 3135 3135 3135 | 2017-10-02 | 9:40 | Albert Zondervan |  | 0.67 mg |  | 0.67 | [5 empty] | 0.99957 | 0.00154 | 73087 |  | XCAMS | ST | 0% |  | TW3135 measurement: * no problems with XCAMS * normal precision * large, (ground)water samples CalAMS error adjustments: * 0% system error * uncertainty of 3x unknowns (TP73109, TP73111, TP73099) adjusted by factor 1.3x to account for excess run-to-run variability primaries: * 6x OxI cathodes blanks: * One groups have been identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 8x blanks secondaries: * 1x IAEA-C2 via waterline: result differs from consensus value by 2.0 sigma Note that this disagreement resembles that of TP73013 in TW3132. unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2017-10-26 | Water | Complete |  | first stds after water line move | 63850 | Normal | 14047/11 |  |  | 2017-09-11 | 3035 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00154 | 73121 | [25 empty] | 2017-09-12 | 13039 | Graphite | Completed
[168 empty] | 0.9982 | 0.00157 | 73122 | [25 empty] | 2017-09-18 | 5296 | C13 Measurement | Completed
[168 empty] | 1.0017 | 0.00168 | 73131 | [25 empty] | 2017-10-03 | 14617 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00167 | 73132 | [25 empty] | 2017-10-03 | 17489 | Close Job | Completed
[168 empty] | 1.0006 | 0.00158 | 73161
[7 empty] | 117.9 | 88.4905 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39308 | 39307 | 39308 | 134 | 133 | 134 |  |  | margaret norris | -992.5 | 0.12 ‰ | 2017-10-24 | 2017-10-24 | 26 | 0 |  |  |  |  | 3.93‰ | 3.93‰ | 2.00‰ | 2.00‰ | 3.93 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.50 ‰ | -992.50 ‰ | 0.12 | 0.12 | -992.41 | 0.13 ‰ | -992.41 | 0.13 ‰ |  |  | 2017-10-30 | Cathy Ginnane | 13:48 | 0.0074969 | 0.0074969 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.5 |  |  | 0.0075 | 0.00012 | 14760 | 2017-10-26 | 15:31 | Albert Zondervan | 207963 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.75 | 0.01% | 9 | 5 | 7 | ... | 0.00721 | 0.00012 |  |  |  |  | 0.99829 | 0.99829 | Yes | Yes | [5 empty] | 0.00721 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00059 | 0.00059 | 0.0075 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00721 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-10-20 | margaret norris | 15:11 |  | Completed |  |  |  |  | 73196 | 2000000 |  |  | 73047 73061 73272 73273 73321 73324 | 73047~73061~73272~73273~73321~73324 | 3139 |  | 3139 3139 3139 3139 3139 3139 | 2017-10-26 | 16:13 | Albert Zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99809 | 0.00144 | 73047 |  | XCAMS | EA | 0% |  | TW3139 measurement: * no problems with XCAMS * normal precision * large samples * no ST CalAMS error adjustments: * 0% system error primaries: * 6x OxI cathodes blanks: * 3x groups have been identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 9x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks window = 180 days: captures 3x blanks organics not via CE: 3x unknowns + 1x secondaries + 1x blanks window = 180 days: captures 5x blanks via EA secondaries: * 1x FIRI-I: result agrees with consensus value * 1x IAEA-C2 via waterline: result disagrees from consensus value by 3 sigma unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2017-11-17 | Water | Complete |  |  | 63971 | Normal | 14047/11 |  |  | 2017-09-25 | 3076 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99996 | 0.00151 | 73061 | [25 empty] | 2017-09-27 | 3078 | Water CO2 Evolution | Completed
[168 empty] | 1.0015 | 0.00143 | 73272 | [25 empty] | 2017-09-27 | 1876 | Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00144 | 73273 | [25 empty] | 2017-09-27 | 3458 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00146 | 73321 | [25 empty] | 2017-09-28 | 13206 | Graphite | Completed
[168 empty] | 0.99995 | 0.00143 | 73324 | [25 empty] | 2017-09-29 | 5348 | C13 Measurement | Completed
[196 empty] | 2017-10-30 | 14760 | AMS Submission Results | Completed
[196 empty] | 2017-10-31 | 17643 | Close Job | Completed
[7 empty] | 139.5 | 88.2422 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40987 | 40986 | 40987 | 165 | 164 | 165 |  |  | julia collins | -993.92 | 0.12 ‰ | 2017-11-08 | 2017-11-08 | 8 | 0 |  |  |  |  | 4.63‰ | 4.63‰ | 2.00‰ | 2.00‰ | 4.63 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0005 | 0 | 0.8165 | -993.92 ‰ | -993.92 ‰ | 0.12 | 0.12 | -993.84 | 0.13 ‰ | -993.84 | 0.13 ‰ |  |  | 2017-11-13 | Jocelyn Turnbull | 16:57 | 0.0060828 | 0.0060828 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00608 | 0.00012 | 14876 | 2017-11-10 | 8:32 | Albert Zondervan | 208365 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 9 | 5 | 5 | ... | 0.00585 | 0.00012 |  |  |  |  | 0.99619 | 0.99619 | Yes | Yes | [5 empty] | 0.00585 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00608 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00585 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-08 | julia collins | 8:12 |  | Completed |  |  |  |  | 73420 | 2000000 |  |  | 73363 73364 73365 73366 73367 73384 | 73363~73364~73365~73366~73367~73384 | 3142 |  | 3142 3142 3142 3142 3142 3142 | 2017-11-10 | 10:37 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99968 | 0.00153 | 73363 |  | XCAMS | ST | 0% |  | TW3142 measurement: * no problems with XCAMS * few small samples * none processed via EA CalAMS error adjustments: * 0% system error primaries: * 6x large OxI cathodes * 3x small OxI cathodes: d = 0.00075 mgC secondaries: * 1x small IAEA-C2 travertine: result consistent with expectation value * 1x small FIRI-I cellulose: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Three groups have been identified to blank-correct: small organics (AAA): blank-corrected using data from this wheel and TW3134 3x unknowns + 1x secondaries + 2x blanks m = 0.00127 mgC large organics (AAA): 7x unknowns + 1x secondaries + 1x blanks (ground)waters: 12x unknowns + 1x secondaries + 1x blanks * machine blank: An artefact has reappeared in the 2D spectrum of the particle detector. This is most likely due to Si-28 from the stubby behind the graphite in the cathode and is visible only when the Cs sputter beam reaches the stubby. Recently we reverted back to using stubbies cut from Al welding wire, which contains some Si. Any Si-28 ions that 'leak through' fall outside the C-14 counting gate, thus not affecting the machine blank. unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-12-11 | Water | Complete |  | new carbosorb, 113% CO2 yield on graphite seems high; split issue??? | 64123 | Normal | 14047/11 |  |  | 2017-10-30 | 3118 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99883 | 0.00153 | 73364 | [25 empty] | 2017-10-30 | 1900 | Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00158 | 73365 | [25 empty] | 2017-10-30 | 3496 | Choose Gas Splitting | Completed
[168 empty] | 1.002 | 0.0016 | 73366 | [25 empty] | 2017-10-31 | 13430 | Graphite | Completed
[168 empty] | 1.0001 | 0.00157 | 73367 | [25 empty] | 2017-11-13 | 14876 | AMS Submission Results | Completed
[168 empty] | 0.99926 | 0.00153 | 73384 | [25 empty] | 2017-11-14 | 17808 | Close Job | Completed
[7 empty] | 126.4 | 88.1154 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35922 | 35921 | 35922 | 117 | 116 | 117 |  |  | julia collins | -988.57 | 0.17 ‰ | 2017-11-14 | 2017-11-14 | 6 | 0 |  |  |  |  | 3.83‰ | 3.83‰ | 2.00‰ | 2.00‰ | 3.83 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0004 | 0 | 0.8165 | -988.57 ‰ | -988.57 ‰ | 0.17 | 0.17 | -988.43 | 0.17 ‰ | -988.43 | 0.17 ‰ |  |  | 2017-11-16 | Jocelyn Turnbull | 11:58 | 0.011427 | 0.011427 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.01143 | 0.00017 | 14945 | 2017-11-15 | 10:16 | Albert Zondervan | 208538 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.14 | 0.02% | 8 | 5 | 5 | ... | 0.01099 | 0.00016 |  |  |  |  | 0.99396 | 0.99396 | Yes | Yes |  | not a great water blank |  |  |  | 0.01099 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 1.0001 | 0.00079 | 0.00079 | 0.01143 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01099 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-11-13 | julia collins | 11:40 |  | Completed |  |  |  |  | 73485 | 2000000 |  |  | 73480 73483 73484 73497 73498 73499 | 73480~73483~73484~73497~73498~73499 | 3143 |  | 3143 3143 3143 3143 3143 3143 | 2017-11-15 | 10:28 | Albert Zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 0.99775 | 0.00184 | 73480 |  | XCAMS | ST | 0.1% |  | TW3143 measurement: * no problems with XCAMS, apart from drifting stripper-gas pressure * large samples * none processed via EA CalAMS error adjustments: * 0.10% system error * error adjustments for cat#1,22,35 (TP73496,735009, 73484) factor 1.3 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 1x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: organics (bone): 20x unknowns + 1x secondaries + 1x blanks (ground)waters: 8x unknowns + 1x secondaries + 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-02 | Water | Complete |  | loaded marble chips | 64201 | Normal | 14047/11 |  |  | 2017-11-07 | 3140 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99839 | 0.00189 | 73483 | [25 empty] | 2017-11-07 | 1918 | Gas Splitting | Completed
[168 empty] | 1.0032 | 0.00232 | 73484 | [25 empty] | 2017-11-07 | 3524 | Choose Gas Splitting | Completed
[168 empty] | 1.0011 | 0.0019 | 73497 | [25 empty] | 2017-11-08 | 13495 | Graphite | Completed
[168 empty] | 1.0013 | 0.00189 | 73498 | [25 empty] | 2017-11-16 | 14945 | AMS Submission Results | Completed
[168 empty] | 0.99965 | 0.0019 | 73499 | [25 empty] | 2017-11-28 | 17922 | Close Job | Completed
[7 empty] | 98.2 | 88.4688 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38842 | 38842 | 38842 | 137 | 136 | 137 |  |  | Julia Collins | -992.06 | 0.14 ‰ | 2017-12-04 | 2017-12-04 | 13 | 0 |  |  |  |  | 1.52‰ | 1.52‰ | 2.00‰ | 2.00‰ | 1.52 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | Estimate | No | -19 | -19.0001 | 0 | 0.8165 | -992.06 ‰ | -992.06 ‰ | 0.14 | 0.14 | -991.96 | 0.14 ‰ | -991.96 | 0.14 ‰ |  |  | 2017-12-06 | Jocelyn Turnbull | 16:44 | 0.007944 | 0.007944 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.2 |  |  | 0.00794 | 0.00014 | 15062 | 2017-12-05 | 16:51 | Albert Zondervan | 208513 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.79 | 0.01% | 13 | 5 | 5 | ... | 0.00764 | 0.00013 |  |  |  |  | 0.99565 | 0.99565 | Yes | Yes | [5 empty] | 0.00764 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00794 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00764 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-01 | Julia Collins | 12:37 |  | Completed |  |  |  |  | 73570 | 2000000 |  |  | 73525 73526 73527 73637 73644 73646 | 73525~73526~73527~73637~73644~73646 | 3146 |  | 3146 3146 3146 3146 3146 3146 | 2017-12-06 | 12:27 | Albert Zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0008 | 0.00159 | 73525 |  | XCAMS | EA | 0% |  | TW3146 measurement: * no problems with XCAMS * large samples * none processed via ST CalAMS error adjustments: * 0% system error * error adjustments: none necessary primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: result consistent with expectation value * 2x large FIRI/TIRI whalebone: result consistent with expectation value * 1x large FIRI-I cellulose: result consistent with expectation value * 1x large IAEA-C6 sucrose: result consistent with expectation value blanks: * Two groups have been identified to blank-correct: (ground)waters: 11x unknowns + 1x secondaries + 1x blanks The sample swap between cat#13, 34 has been resolved. organics: CE: 3x unknowns + 0x secondaries + 1x blanks non-CE: 12x unknowns + 4x secondaries + 0x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2017-12-26 | Water |  |  | This job was originally labelled as R14047/12, but it is actually R14047/11. Changed - JCT. | 64283 | Normal | 14047/11 |  |  | 2017-11-20 | 3171 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00158 | 73526 | [25 empty] | 2017-11-20 | 1938 | Gas Splitting | Completed
[168 empty] | 0.99854 | 0.00156 | 73527 | [25 empty] | 2017-11-20 | 3547 | Choose Gas Splitting | Completed
[168 empty] | 0.99977 | 0.0016 | 73637 | [25 empty] | 2017-11-21 | 13580 | Graphite | Completed
[168 empty] | 1.0006 | 0.00163 | 73644 | [25 empty] | 2017-12-06 | 15062 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00164 | 73646 | [25 empty] | 2017-12-07 | 18019 | Close Job | Completed
[7 empty] | 100.9 | 88.3472 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37799 | 37798 | 37799 | 111 | 110 | 111 |  |  | Julia Collins | -990.95 | 0.12 ‰ | 2017-12-06 | 2017-12-06 | 19 | 0 |  |  |  |  | 3.12‰ | 3.12‰ | 2.00‰ | 2.00‰ | 3.12 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -990.95 ‰ | -990.95 ‰ | 0.12 | 0.12 | -990.84 | 0.13 ‰ | -990.84 | 0.13 ‰ |  |  | 2017-12-11 | Jocelyn Turnbull | 9:07 | 0.0090462 | 0.0090462 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00905 | 0.00012 | 15097 | 2017-12-08 | 14:10 | rachel corran | 208552 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 30 | 5 | 6 | ... | 0.0087 | 0.00012 |  |  |  |  | 0.99587 | 0.99587 | Yes | Yes | [5 empty] | 0.0087 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00905 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0087 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2017-12-05 | Julia Collins | 12:39 |  | Completed |  |  |  |  | 73551 | 2000000 |  |  | 73618 73621 73624 73639 73641 73648 | 73618~73621~73624~73639~73641~73648 | 3147 |  | 3147 3147 3147 3147 3147 3147 | 2017-12-08 | 15:08 | rachel corran |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0019 | 0.0014 | 73618 |  | XCAMS | EA | 0% |  | TW3147 measurement: * no problems with XCAMS * large samples * none processed via ST, except the C6 sucrose CalAMS error adjustments: * 0% system error * factor x1.03 across all cathodes, and 1.3 for cat#23 primaries: * 6x large OxI cathodes secondaries: * 1x IAEA-C2 travertine: !! result deviates from consensus by ~4sigma !! cause unknown blanks: * one group has been identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2018-01-10 | Water | Not Complete |  |  | 64323 | Normal | 14047/11 |  |  | 2017-11-15 | 3157 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00145 | 73621 | [25 empty] | 2017-11-15 | 1927 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00139 | 73624 | [25 empty] | 2017-11-15 | 3534 | Choose Gas Splitting | Completed
[168 empty] | 0.99773 | 0.00139 | 73639 | [25 empty] | 2017-11-17 | 13561 | Graphite | Completed
[168 empty] | 0.99885 | 0.00141 | 73641 | [25 empty] | 2017-11-21 | 5480 | C13 Measurement | Completed
[168 empty] | 0.99977 | 0.00141 | 73648 | [25 empty] | 2017-12-11 | 15097 | AMS Submission Results | Completed
[196 empty] | 2017-12-11 | 18034 | Close Job | Completed
[7 empty] | 114.9 | 88.2886 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36002 | 36002 | 36002 | 140 | 140 | 140 |  |  | margaret norris | -988.69 | 0.20 ‰ | 2018-02-01 | 2018-02-01 | 20 | 0 |  |  |  |  | 3.98‰ | 3.98‰ | 2.00‰ | 2.00‰ | 3.98 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0032 | 0 | 0.8165 | -988.69 ‰ | -988.69 ‰ | 0.2 | 0.2 | -988.55 | 0.20 ‰ | -988.55 | 0.20 ‰ |  |  | 2018-02-07 | Cathy Ginnane | 14:31 | 0.011313 | 0.011313 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 1.5 |  |  | 0.01131 | 0.0002 | 15229 |  |  |  | 208689 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.13 | 0.02% | 18 | 5 | 5 | ... | 0.01088 | 0.00019 |  |  |  |  | 0.99607 | 0.99607 | Yes | Yes |  | Not a great blank but not unusual |  |  |  | 0.01088 | 0.01‰ | 0.00019 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.01131 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01088 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-01-30 | margaret norris | 10:43 |  | Completed |  |  |  |  | 73742 | 2000000 |  |  | 73774 73775 73789 73790 73791 73792 | 73774~73775~73789~73790~73791~73792 | 3151 |  | 3151 3151 3151 3151 3151 3151 | 2018-02-05 | 10:30 | Albert Zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99939 | 0.00188 | 73774 |  | XCAMS | ST | 0% |  | TW3151 measurement: * XCAMS: ion source a bit unstable * regular precision * large samples * no EA CalAMS error adjustments: * 0% system error * 1.04 global error multiplier * 1.4 error multiplier for cat#4 (TP73715) category line (method) unkn/blan/seco window [d] ----------¦--------------¦---------------¦---------- inorganic carbonate 5/1/1 180 water water 15/1/1 180 organic AAA 4/1/2 180 organic CE-Rafter 1/1/0 180 unknowns: * no failures blanks: * results in keeping with previous analyses secondaries: * 1x carbonate-line IAEA-C2: result consistent with consensus value * 1x water-line IAEA-C2: result consistent with consensus value * 1x IAEA-C6: result consistent with consensus value * 1x FIRI-I: result consistent with consensus value primaries: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-01-31 | Water |  |  |  | 64431 | Normal | 14047/11 |  |  | 2018-01-11 | 3200 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0015 | 0.00187 | 73775 | [25 empty] | 2018-01-11 | 1945 | Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00188 | 73789 | [25 empty] | 2018-01-11 | 3558 | Choose Gas Splitting | Completed
[168 empty] | 0.99915 | 0.00189 | 73790 | [25 empty] | 2018-01-12 | 13752 | Graphite | Completed
[168 empty] | 0.99931 | 0.00179 | 73791 | [25 empty] | 2018-02-07 | 15229 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00187 | 73792 | [25 empty] | 2018-03-15 | 18378 | Close Job | Completed
[7 empty] | 77.6 | 90.2597 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42745 | 42744 | 42745 | 256 | 256 | 256 |  |  | julia collins | -995.11 | 0.16 ‰ | 2018-03-09 | 2018-03-09 | 38 | 0 |  |  |  |  | -1.39‰ | -1.39‰ | 2.00‰ | 2.00‰ | -1.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.11 ‰ | -995.11 ‰ | 0.16 | 0.16 | -995.05 | 0.16 ‰ | -995.05 | 0.16 ‰ |  |  | 2018-03-14 | margaret norris | 11:29 | 0.004887 | 0.004887 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00489 | 0.00016 | 15316 |  |  |  | 209040 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.49 | 0.02% | 13 | 5 | 5 | ... | 0.0047 | 0.00015 |  |  |  |  | 1.0129 | 1.0129 | Yes | Yes |  | d13C out by >3.5 per mil, not a great background |  |  |  | 0.0047 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00489 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.0047 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-02-14 | julia collins | 8:11 |  | Completed |  |  |  |  | 73840 | 2000000 |  |  | 73815 73817 73819 73846 73847 73893 | 73815~73817~73819~73846~73847~73893 | 3153 |  | 3153 3153 3153 3153 3153 3153 | 2018-03-12 | 13:18 | Albert Zondervan |  | 0.85 mg |  | 0.85 | [5 empty] | 0.99955 | 0.00158 | 73815 |  | SSAMS | EA | 0% |  | TW3153 measurement: * Relative to original wheel make-up, two positions are new: TP74018,TP73880. Neither one is used in RLIMS calculations for the unknowns. * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Regular precision has been achieved. * File 'context' from ANU was used to prepare the CalAMS input file. CalAMS error adjustments: * no system error * global multiplication factor 1.05 due to excess variability * individual multiplication factor 1.3 for cat#3 (TP73904) due to excess variability primaries: * OxI cathodes: 6x no failures secondaries, relative to their respective expectation values for DEL14C: * 1x IAEA-C2: result is 2-sigma away from consensus value blanks: * Four groups have been identified to blank-correct: (ground)waters: 26x unknowns + 1x secondaries + 1x blanks window = 180 days, capturing 9x blanks textiles (via AAA): 2x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 6x blanks shells: 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 5x blanks tree bark (via CE-Rafter): 1x unknowns + 0x secondaries + 0x blanks window = 180 days, capturing 8x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-03-26 | Water |  |  |  | 64550 | Normal | 14047/11 |  |  | 2018-01-29 | 3225 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00161 | 73817 | [25 empty] | 2018-01-29 | 1972 | Gas Splitting | Completed
[168 empty] | 1.0011 | 0.00159 | 73819 | [25 empty] | 2018-01-29 | 3609 | Choose Gas Splitting | Completed
[168 empty] | 0.99878 | 0.00159 | 73846 | [25 empty] | 2018-01-30 | 13850 | Graphite | Completed
[168 empty] | 0.99875 | 0.00158 | 73847 | [25 empty] | 2018-03-14 | 15316 | AMS Submission Results | Completed
[168 empty] | 1.0027 | 0.00165 | 73893 | [25 empty] | 2018-03-15 | 18395 | Close Job | Completed
[7 empty] | 71.9 | 90.5707 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34269 | 34269 | 34269 | 125 | 124 | 125 |  |  | julia collins | -985.96 | 0.22 ‰ | 2018-04-13 | 2018-04-13 | 52 | 0 |  |  |  |  | 3.39‰ | 3.39‰ | 2.00‰ | 2.00‰ | 3.39 ‰ | 2.00 ‰ | AMS | 2.22‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -985.96 ‰ | -985.96 ‰ | 0.22 | 0.22 | -985.79 | 0.22 ‰ | -985.79 | 0.22 ‰ |  |  | 2018-04-13 | Jocelyn Turnbull | 15:08 | 0.014037 | 0.014037 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.1 |  |  | 0.01404 | 0.00022 | 15427 |  |  |  | 209076 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.4 | 0.02% | 8 | 5 | 6 | ... | 0.0135 | 0.00021 |  |  |  |  | 1.0212 | 1.0212 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.0135 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01404 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.0135 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-03-14 | julia collins | 7:50 |  | Completed |  |  |  |  | 73983 | 2000000 |  |  | 73973 73985 73986 73987 73988 73989 | 73973~73985~73986~73987~73988~73989 | 3156 |  | 3156 3156 3156 3156 3156 3156 | 2018-04-13 | 14:50 | Albert Zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0009 | 0.00162 | 73973 |  | SSAMS | ST | 0% |  | TW3156 measurement: * The wheel was measured with SSAMS at ANU because of a continued problem with our ion source. * Previously (during measurement of TW3154) signal from the low-energy side Faraday cups was lost. While these signals are good for exercising quality control, they are not essential in deriving the final ratio-to-standard values. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, both agreeing with their consensus value blanks: * 2 groups have been identified to blank-correct: waters 24x unknowns + 1x secondaries + 1x blank window = 180 days organics: 4x unknowns + 2x secondaries + 1x blank window = 180 days unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2018-03-28 | Water |  |  |  | 64622 | Normal | 14047/11 |  |  | 2018-02-19 | 3270 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99806 | 0.00156 | 73985 | [25 empty] | 2018-02-19 | 2004 | Gas Splitting | Completed
[168 empty] | 0.99949 | 0.00156 | 73986 | [25 empty] | 2018-02-19 | 3648 | Choose Gas Splitting | Completed
[168 empty] | 0.99917 | 0.00154 | 73987 | [25 empty] | 2018-02-20 | 13993 | Graphite | Completed
[168 empty] | 1.0012 | 0.00155 | 73988 | [25 empty] | 2018-02-23 | 5642 | C13 Measurement | Completed
[168 empty] | 1.0011 | 0.00153 | 73989 | [25 empty] | 2018-04-13 | 15427 | AMS Submission Results | Completed
[196 empty] | 2018-04-13 | 18544 | Close Job | Completed
[7 empty] | 78.1 | 91.5912 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36736 | 36736 | 36736 | 121 | 121 | 121 |  |  | julia collins | -989.67 | 0.16 ‰ | 2018-05-13 | 2018-05-13 | 51 | 0 |  |  |  |  | 3.87‰ | 3.87‰ | 2.00‰ | 2.00‰ | 3.87 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0005 | 0 | 0.8165 | -989.67 ‰ | -989.67 ‰ | 0.16 | 0.16 | -989.55 | 0.16 ‰ | -989.55 | 0.16 ‰ |  |  | 2018-05-15 | Jocelyn Turnbull | 11:29 | 0.010325 | 0.010325 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.8 |  |  | 0.01033 | 0.00016 | 15535 |  |  |  | 209461 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1.03 | 0.02% | 20 | 5 | 5 | ... | 0.00993 | 0.00015 |  |  |  |  | 1.0332 | 1.0332 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.00993 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.01033 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00993 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-09 | julia collins | 11:04 |  | Completed |  |  |  |  | 74073 | 2000000 |  |  | 74116 74117 74120 74121 74122 74123 | 74116~74117~74120~74121~74122~74123 | 3159 |  | 3159 3159 3159 3159 3159 3159 | 2018-05-15 | 10:39 | Albert Zondervan |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0006 | 0.00152 | 74116 |  | SSAMS | ST | 0% |  | TW3159 measurement: * The wheel was measured on 12+13 May 2018 with SSAMS at ANU because of a continued problem with our ion source. NOTE: In RLIMS, the field 'Date Run' has been imported with incorrect values. JCT - corrected to 13/5/2018 data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 3 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 8x unknowns + 1x secondaries + 1x blank window = 180 days carbonates 4x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via AAA: 14x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period unknowns: * no failures JCT - blanks are not at all good, but to be expected for having sat for so long before measurement. Secondaries look fine with the corrections assigned and no really old samples, so all ok. |  | RRL | [5 empty] | Background Inorganic |  | 2018-04-18 | Water |  |  |  | 64783 | Normal | 14047/11 | KIP |  | 2018-03-12 | 3295 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99843 | 0.00153 | 74117 | [25 empty] | 2018-03-12 | 2025 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00151 | 74120 | [25 empty] | 2018-03-12 | 3670 | Choose Gas Splitting | Completed
[168 empty] | 1.0007 | 0.00151 | 74121 | [25 empty] | 2018-03-23 | 14083 | Graphite | Completed
[168 empty] | 0.99815 | 0.00151 | 74122 | [25 empty] | 2018-05-15 | 15535 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.0015 | 74123 | [25 empty] | 2018-05-15 | 18743 | Close Job | Completed
[7 empty] | 72.4 | 91.0234 | 0.0017 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37008 | 37007 | 37008 | 134 | 133 | 134 |  |  | Julia Collins | -990.02 | 0.17 ‰ | 2018-05-24 | 2018-05-24 | 37 | 0 |  |  |  |  | 6.65‰ | 6.65‰ | 2.00‰ | 2.00‰ | 6.65 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0019 | 0 | 0.8165 | -990.02 ‰ | -990.02 ‰ | 0.17 | 0.17 | -989.89 | 0.17 ‰ | -989.89 | 0.17 ‰ |  |  | 2018-05-24 | Jocelyn Turnbull | 16:18 | 0.009982 | 0.009982 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.00998 | 0.00017 | 15606 | 2018-05-24 | 15:09 | Albert Zondervan | 209652 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 1 | 0.02% | 22 | 5 | 5 | ... | 0.0096 | 0.00016 |  |  |  |  | 1.0296 | 1.0296 | Yes | Yes | AMS measurement of this sample was performed at the Australian National University AMS facility. All preparation (pretreatment, combustion, graphitisation, target packing) and data reduction was performed at the Rafter facility in our usual way. The data quality meets our usual high standard, indicated by standard materials that agree with previous measurements of the same material made in our lab within one standard deviation. |  |  |  |  | 0.0096 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.00998 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.0096 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-04-27 | Julia Collins | 10:42 |  | Completed |  |  |  |  | 74147 | 2000000 |  |  | 74152 74153 74162 74163 74171 74172 | 74152~74153~74162~74163~74171~74172 | 3160 |  | 3160 3160 3160 3160 3160 3160 | 2018-05-24 | 15:26 | Albert Zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0007 | 0.00171 | 74152 |  | SSAMS | EA | 0% |  | TW3160 measurement: * The wheel was measured on 22/23ish May 2018 with SSAMS at ANU because of a continued problem with our ion source. data analysis with CalAMS: * File 'context' recorded by the SSAMS was used to prepare the CalAMS input file. This required some reformatting in Excel. * system error = 0.0% primaries: * OxI cathodes: 6x no failures secondaries: 2 secondaries in wheel, all agreeing with their consensus value blanks: * 3 groups have been identified to blank-correct: waters 4x unknowns + 1x secondaries + 1x blank window = 180 days organics via AAA: 23x unknowns + 1x secondaries + 1x blank window = only covering ANU-SSAMS period organics via Rafter Cellulose: 1x unknowns + 0x secondaries + 0x blank window = covering beyond ANU-SSAMS period as no rafter cellulose blanks during this time unknowns: * TP74174 (cat#32) gives a low output, different trend in 13/12, and large deviation from IRMS del13C: REJECT |  | RRL | [5 empty] | Background Inorganic |  | 2018-06-06 | Water |  |  |  | 64855 | Normal | 14047/11 |  |  | 2018-04-11 | 3325 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99763 | 0.00173 | 74153 | [25 empty] | 2018-04-11 | 2055 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.0017 | 74162 | [25 empty] | 2018-04-11 | 3715 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00173 | 74163 | [25 empty] | 2018-04-20 | 14157 | Graphite | Completed
[168 empty] | 0.99872 | 0.00171 | 74171 | [25 empty] | 2018-05-24 | 15606 | AMS Submission Results | Completed
[168 empty] | 1.0016 | 0.00173 | 74172 | [25 empty] | 2018-05-28 | 18836 | Close Job | Completed
[7 empty] | 71.7 | 88.9012 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40018 | 40017 | 40018 | 122 | 121 | 122 |  |  | margaret norris | -993.14 | 0.10 ‰ | 2018-05-25 | 2017-11-26 | 21 | 180 |  |  |  |  | -1.89‰ | -1.89‰ | 2.00‰ | 2.00‰ | -1.89 ‰ | 2.00 ‰ | AMS | 1.86‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -993.14 ‰ | -993.14 ‰ | 0.1 | 0.1 | -993.05 | 0.11 ‰ | -993.05 | 0.11 ‰ |  |  | 2018-05-28 | Jocelyn Turnbull | 14:39 | 0.0068626 | 0.0068626 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00686 | 0.0001 | 15708 |  |  |  | 209758 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:47 | Jocelyn Turnbull | 0 |  | 0.69 | 0.01% | 15 | 5 | 6 | ... | 0.0066 | 0.0001 |  |  |  |  | 0.99711 | 0.99711 | Yes | Yes | [5 empty] | 0.0066 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00046 | 0.00046 | 0.00686 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.0066 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-05-24 | margaret norris | 10:37 |  | Completed |  |  |  |  | 74231 | 2000000 |  |  | 74282 74284 74285 74286 74287 74288 | 74282~74284~74285~74286~74287~74288 | 3165 |  | 3165 3165 3165 3165 3165 3165 | 2018-05-28 | 12:22 | Albert Zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0009 | 0.00114 | 74282 |  | XCAMS | ST | 0% |  | TW3165 measurement: * no problems with XCAMS * The complete tune prior measurement has resolved two issues relative to the previous wheel (TW3161): The graphite cathodes last longer (= higher efficiency) due to the a reduced sputter rate. The mode=5 ratio shows no beam current dependency: Runs with sputtering into the stubby do not have to be rejected. * Measurement was extended beyond the demand for regular precision, to prove that high precision can be achieved on 0.5 mgC graphites. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x via ST no failures secondaries: * 1x IAEA-C2 via water-line: result is consistent with expectation value * 1x IAEA-C6: result is consistent with expectation value * 1x FIRI-I: result is consistent with expectation value blanks: * One group has been identified to blank-correct: carbonates: 6x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 2x blanks waters: 1x unknowns + 1x secondaries + 1x blanks 180 days, including SSAMS runs: capturing 7x blanks organics via AAA: 20x unknowns + 2x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks organics via CE-Rafter: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures | 1.9 | RRL | [5 empty] | Background Inorganic |  | 2018-06-21 | Water |  |  | d13C lower than it should be | 64961 | Normal | 14047/11 |  |  | 2018-05-02 | 3352 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99818 | 0.00115 | 74284 | [25 empty] | 2018-05-02 | 2078 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00113 | 74285 | [25 empty] | 2018-05-02 | 3746 | Choose Gas Splitting | Completed
[168 empty] | 0.99929 | 0.00111 | 74286 | [25 empty] | 2018-05-04 | 14241 | Graphite | Completed
[168 empty] | 1.0002 | 0.00114 | 74287 | [25 empty] | 2018-05-07 | 5778 | C13 Measurement | Completed
[168 empty] | 1.0013 | 0.00114 | 74288 | [25 empty] | 2018-05-28 | 15708 | AMS Submission Results | Completed
[196 empty] | 2018-05-28 | 18941 | Close Job | Completed
[7 empty] | 107.1 | 89.1413 | 0.0022 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34913 | 34913 | 34913 | 128 | 128 | 128 |  |  | jenny dahl | -987.04 | 0.21 ‰ | 2018-07-12 | 2018-01-13 | 9 | 180 |  |  |  |  | -1.30‰ | -1.30‰ | 2.00‰ | 2.00‰ | -1.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -987.04 ‰ | -987.04 ‰ | 0.21 | 0.21 | -986.88 | 0.21 ‰ | -986.88 | 0.21 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:01 | 0.012956 | 0.012956 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2 |  |  | 0.01296 | 0.00021 | 16048 | 2018-07-13 | 14:25 | Albert Zondervan | 209984 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 1.3 | 0.02% | 37 | 5 | 5 | ... | 0.01246 | 0.0002 |  |  |  |  | 1.0004 | 1.0004 | Yes | Yes | [5 empty] | 0.01246 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 0.99999 | 0.00097 | 0.00097 | 0.01296 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01246 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-04 | jenny dahl | 8:32 |  | Completed |  |  |  |  | 74710 | 2000000 |  |  | 74587 74591 74597 74606 74607 74608 | 74587~74591~74597~74606~74607~74608 | 3174 |  | 3174 3174 3174 3174 3174 3174 | 2018-07-13 | 14:32 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 0.99795 | 0.00241 | 74587 |  | XCAMS | EA | 0.15% |  | TW3174 measurement: * A mistake was made during graphite loading: One funnel must have been used for loading two cathodes. One of the 40 funnels was left over after loading all 40 graphites. It is unknown which cathodes were possibly compromised. The resulting data do not show a cross-contamination problem. * There were a few issues with XCAMS: There is a small step in 13/12 for some cathodes halfway in the measurement. The tune for TW3173 has already shown a possible issue with the scan of ESX reg#1 for 13C. They are probably associated with the slight lack of repeatability amongst the standards. These need to be sorted before high-precision is attempted: A source clean is warranted. After his return, Albert to look into the ESX reg#1 profile of 13C+ transmission into MFC 04-2. CalAMS error adjustments: * small system error: 0.15% * small error factors x1.3 on cat#33 primaries: * OxI cathodes: 6x from flask no failures secondaries: * IAEA-C2: 1x result ~ 2 sigma younger than consensus value blanks: * One groups have been identified to blank-correct: waters: window = 180 days, capturing 7x blanks unknowns: * no failures in wheel other: * It appears the field 'Standard Prep Type' in RLIMS is no longer filled in. Is this no longer needed? |  | RRL | [5 empty] | Background Inorganic |  | 2018-07-24 | Water |  |  |  | 65279 | Normal | 14047/11 |  |  | 2018-07-02 | 3450 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00238 | 74591 | [25 empty] | 2018-07-02 | 2155 | Gas Splitting | Completed
[168 empty] | 0.99931 | 0.00237 | 74597 | [25 empty] | 2018-07-02 | 3885 | Choose Gas Splitting | Completed
[168 empty] | 0.99676 | 0.00236 | 74606 | [25 empty] | 2018-07-03 | 14720 | Graphite | Completed
[168 empty] | 1.0018 | 0.0024 | 74607 | [25 empty] | 2018-07-17 | 16048 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00236 | 74608 | [25 empty] | 2018-07-17 | 19414 | Close Job | Completed
[7 empty] | 101.6 | 88.8725 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37863 | 37863 | 37863 | 130 | 130 | 130 |  |  | Cathy Ginnane | -991.03 | 0.15 ‰ | 2018-07-14 | 2018-01-15 | 8 | 180 |  |  |  |  | 1.93‰ | 1.93‰ | 2.00‰ | 2.00‰ | 1.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -991.03 ‰ | -991.03 ‰ | 0.15 | 0.15 | -990.92 | 0.15 ‰ | -990.92 | 0.15 ‰ |  |  | 2018-07-17 | Jocelyn Turnbull | 11:21 | 0.0089734 | 0.0089734 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.8 |  |  | 0.00897 | 0.00015 | 16091 | 2018-07-17 | 11:04 | Jocelyn Turnbull | 210223 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.9 | 0.01% | 37 | 5 | 5 | ... | 0.00863 | 0.00014 |  |  |  |  | 1.0006 | 1.0006 | Yes | Yes | [5 empty] | 0.00863 | 0.01‰ | 0.00014 | 0.00‰ | No | 1 | 0.99998 | 0.0008 | 0.0008 | 0.00897 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00863 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-07-11 | Cathy Ginnane | 15:25 |  | Completed |  |  |  |  | 74758 | 2000000 |  |  | 74703 74704 74712 74713 74760 74761 | 74703~74704~74712~74713~74760~74761 | 3175 |  | 3175 3175 3175 3175 3175 3175 | 2018-07-17 | 11:16 | Jocelyn Turnbull |  | 1.09 mg |  | 1.09 | [5 empty] | 0.99969 | 0.00194 | 74703 |  | XCAMS | EA |  |  |  |  | RRL | [5 empty] | Background Inorganic |  | 2018-08-13 | Water |  |  |  | 65319 | Normal | 14047/11 |  |  | 2018-07-04 | 3453 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99996 | 0.00197 | 74704 | [25 empty] | 2018-07-04 | 2161 | Gas Splitting | Completed
[168 empty] | 0.99776 | 0.00198 | 74712 | [25 empty] | 2018-07-04 | 3896 | Choose Gas Splitting | Completed
[168 empty] | 0.99851 | 0.00197 | 74713 | [25 empty] | 2018-07-06 | 14768 | Graphite | Completed
[168 empty] | 1.0034 | 0.00195 | 74760 | [25 empty] | 2018-07-17 | 16091 | AMS Submission Results | Completed
[168 empty] | 1.0005 | 0.00194 | 74761 | [25 empty] | 2018-07-17 | 19418 | Close Job | Completed
[7 empty] | 113.3 | 88.7041 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38492 | 38492 | 38492 | 161 | 161 | 161 |  |  | jenny dahl | -991.7 | 0.17 ‰ | 2018-08-13 | 2018-02-14 | 12 | 180 |  |  |  |  | 2.26‰ | 2.26‰ | 2.00‰ | 2.00‰ | 2.26 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -991.70 ‰ | -991.70 ‰ | 0.17 | 0.17 | -991.6 | 0.17 ‰ | -991.6 | 0.17 ‰ |  |  | 2018-08-16 | Jocelyn Turnbull | 15:21 | 0.0082976 | 0.0082976 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.9 |  |  | 0.0083 | 0.00017 | 16361 |  |  |  | 210353 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 0.83 | 0.02% | 20 | 5 | 6 | ... | 0.00798 | 0.00016 |  |  |  |  | 0.99904 | 0.99904 | Yes | Yes | [5 empty] | 0.00798 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0083 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00798 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-08-09 | jenny dahl | 9:26 |  | Completed |  |  |  |  | 74967 | 2000000 |  |  | 74881 74882 74907 74912 75033 75034 | 74881~74882~74907~74912~75033~75034 | 3183 |  | 3183 3183 3183 3183 3183 3183 | 2018-08-14 | 16:23 | Albert Zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0004 | 0.00177 | 74881 |  | XCAMS | EA | 0% |  | TW3183 measurement: * no problems with XCAMS CalAMS error adjustments: * 0% system error primaries: * OxI cathodes: all 6x via EA no failures secondaries: * 1x IAEA-C2: result is consistent with expectation value * 1x IAEA-C6: result is NOT consistent with expectation value * 1x FIRI-I: result is consistent with expectation value * 1x FIRI whalebone: result is consistent with expectation value blanks: * four groups have been identified to blank-correct: organics via AAA: 14x unknowns + 3x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks waters: 8x unknowns + 1x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 4x blanks carbonates: 2x unknowns + 0x secondaries + 0x blanks 180 days, excluding SSAMS runs: capturing 4x blanks organics via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days, excluding SSAMS runs: capturing 2x blanks unknowns: * TP74905: 8 permil difference in IRMS d13C to AMS checked carbonate and graphite processes for this sample and found nothing unusual; can't see a problem in XCAMS raw data either | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2018-08-30 | Water | Complete |  | PC was left open between the left and right sides of the line during sample transfer. Consultation with JT concurs that unlikely to be any cross-contamination but be aware if sample results look odd | 65611 | Normal | 14047/11 |  |  | 2018-07-31 | 3485 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0004 | 0.00172 | 74882 | [25 empty] | 2018-07-31 | 2185 | Gas Splitting | Completed
[168 empty] | 0.99798 | 0.00163 | 74907 | [25 empty] | 2018-07-31 | 3924 | Choose Gas Splitting | Completed
[168 empty] | 1.0006 | 0.00165 | 74912 | [25 empty] | 2018-08-01 | 14977 | Graphite | Completed
[168 empty] | 1.0002 | 0.00172 | 75033 | [25 empty] | 2018-08-02 | 5981 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00177 | 75034 | [25 empty] | 2018-08-16 | 16361 | AMS Submission Results | Completed
[196 empty] | 2018-08-17 | 19769 | Close Job | Completed
[7 empty] | 97.4 | 88.622 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33531 | 33530 | 33530 | 114 | 113 | 113 |  |  | Julia Collins | -984.61 | 0.22 ‰ | 2018-10-29 | 2018-05-02 | 21 | 180 |  |  |  |  | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0017 | 0 | 0.8165 | -984.61 ‰ | -984.61 ‰ | 0.22 | 0.22 | -984.42 | 0.22 ‰ | -984.42 | 0.22 ‰ |  |  | 2018-11-05 | Jocelyn Turnbull | 13:58 | 0.015389 | 0.015389 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 2.3 |  |  | 0.01539 | 0.00022 | 16995 |  |  |  | 209937 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:48 | Jocelyn Turnbull | 0 |  | 1.54 | 0.02% | 1 | 5 | 5 | ... | 0.0148 | 0.00021 |  |  |  |  | 0.99885 | 0.99885 | Yes | Yes | [5 empty] | 0.0148 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01539 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.0148 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-10-24 | Julia Collins | 15:03 |  | Completed |  |  |  |  | 75583 | 2000000 |  |  | 75434 75435 75447 75448 75449 75605 | 75434~75435~75447~75448~75449~75605 | 3199 |  | 3199 3199 3199 3199 3199 3199 | 2018-11-02 | 9:02 | Albert Zondervan |  | 0.86 mg |  | 0.86 | [5 empty] | 0.99908 | 0.00173 | 75434 |  | XCAMS | ST | 0% |  | TW3199 measurement: * The problem with cooling-water temperature stability, causing transmission variations, was recognised easily; corresponding outlier data were rejected. * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x, all via ST 6x large: no failures, used for normalisation in CalAMS 1x small: no failures used for calculation of DCC d = 2.41 ugC using this wheel's small OxI only because reduction in rts is significantly higher than in previous wheels secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value blanks: * Four groups have been identified to blank-correct: waters, large: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, large: 9x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 11x unknowns + 0x secondaries + 2x blanks blanks used: from this wheel only because blank level is high m = 1.55 ugC, double the previous value reported in TW3185 organics via CE, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures JCT - checked small sample corrections. Modern contamination is higher because this wheel had longer days as graphite than normal. So larger m than typical is expected. d is also higher than normal, but is based on a single small OxI target that is only marginally higher than in previous wheels. Seems reasonable to use d based on this wheel, with the associated larger error bar, to be consistent with using m from this wheel alone. Even though there is no obvious reason why d would be larger than normal. |  | RRL | [5 empty] | Background Inorganic |  | 2018-07-19 | Water |  |  |  | 66211 | Normal | 14047/11 |  |  | 2018-10-05 | 3512 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99841 | 0.00176 | 75435 | [25 empty] | 2018-10-05 | 2247 | Gas Splitting | Completed
[168 empty] | 1.0018 | 0.00181 | 75447 | [25 empty] | 2018-10-05 | 4005 | Choose Gas Splitting | Completed
[168 empty] | 0.99903 | 0.00179 | 75448 | [25 empty] | 2018-10-08 | 15593 | Graphite | Completed
[168 empty] | 1.0015 | 0.00173 | 75449 | [25 empty] | 2018-11-05 | 16995 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00175 | 75605 | [25 empty] | 2018-11-06 | 20462 | Close Job | Completed
[7 empty] | 117.3 | 88.1779 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39088 | 39087 | 39087 | 152 | 151 | 151 |  |  | Julia Collins | -992.3 | 0.15 ‰ | 2018-11-15 | 2018-05-19 | 16 | 180 |  |  |  |  | 4.96‰ | 4.96‰ | 2.00‰ | 2.00‰ | 4.96 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -992.30 ‰ | -992.30 ‰ | 0.15 | 0.15 | -992.2 | 0.15 ‰ | -992.2 | 0.15 ‰ |  |  | 2018-11-20 | Jocelyn Turnbull | 9:45 | 0.0077049 | 0.0077049 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.6 |  |  | 0.0077 | 0.00015 | 17076 | 2018-11-19 | 16:44 | Albert Zondervan | 210849 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.77 | 0.01% | 1 | 5 | 5 | ... | 0.00741 | 0.00014 |  |  |  |  | 0.99579 | 0.99579 | Yes | Yes | [5 empty] | 0.00741 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.0077 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00741 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-12 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 75742 | 2000000 |  |  | 75715 75716 75717 75794 75795 75796 | 75715~75716~75717~75794~75795~75796 | 3201 |  | 3201 3201 3201 3201 3201 3201 | 2018-11-19 | 17:11 | Albert Zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 1.0003 | 0.0017 | 75715 |  | XCAMS | ST |  |  | TW3201 measurement: * XCAMS ok * raw 14C/C ratios low because of known effect in detector setup * large samples only CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I celluse: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks woods via ABBA, for Chris Moy (R41230): 4x unknowns + 0x secondaries + 1x blanks 10 days: capturing 1x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-11-27 | Water |  |  |  | 66330 | Normal | 14047/11 |  |  | 2018-10-29 | 3584 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0025 | 0.00175 | 75716 | [25 empty] | 2018-10-29 | 2270 | Gas Splitting | Completed
[168 empty] | 0.99955 | 0.00172 | 75717 | [25 empty] | 2018-10-29 | 4044 | Choose Gas Splitting | Completed
[168 empty] | 0.99913 | 0.00175 | 75794 | [25 empty] | 2018-10-30 | 15752 | Graphite | Completed
[168 empty] | 1.0001 | 0.00172 | 75795 | [25 empty] | 2018-11-20 | 17076 | AMS Submission Results | Completed
[168 empty] | 0.99856 | 0.00171 | 75796 | [25 empty] | 2018-11-21 | 20597 | Close Job | Completed
[7 empty] | 100.6 | 88.5094 | 0.0019 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35936 | 35936 | 35936 | 154 | 153 | 153 |  |  | jenny dahl | -988.59 | 0.22 ‰ | 2018-11-19 | 2018-05-23 | 13 | 180 |  |  |  |  | 3.90‰ | 3.90‰ | 2.00‰ | 2.00‰ | 3.90 ‰ | 2.00 ‰ | AMS | 2.35‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.70711 | -988.59 ‰ | -988.59 ‰ | 0.22 | 0.22 | -988.45 | 0.22 ‰ | -988.45 | 0.22 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 11:07 | 0.011407 | 0.011407 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 1.7 |  |  | 0.01141 | 0.00022 | 17176 | 2018-11-20 | 12:57 | Albert Zondervan | 211055 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.14 | 0.02% | 1 | 5 | 6 | ... | 0.01097 | 0.00021 |  |  |  |  | 0.99848 | 0.99848 | Yes | Yes | [5 empty] | 0.01097 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.0007 | 0.01141 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.01097 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:40 |  | Completed |  |  |  |  | 75841 | 2000000 |  |  | 75819 75820 75822 75825 75840 75856 75875 75876 | 75819~75820~75822~75825~75840~75856~75875~75876 | 3202 |  | 3202 3202 3202 3202 3202 3202 3202 3202 | 2018-11-20 | 13:46 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99902 | 0.00194 | 75819 |  | XCAMS | ST | 0% |  | TW3202 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.32 mgC CalAMS error adjustments: * none needed primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures AMS category changed from 'standard as unknown' to 'primary standard OxI' NOT used for calculation of DCC, because their rts equals that of large OxIs d = 2.51 ugC extracted from the one small OxI in TW3199 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * Three groups have been identified to blank-correct: waters, large: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA, large: 2x unknowns + 1x secondaries + 0x blanks 180 days: capturing 8x blanks organics via AAA, small, via Excel: 4x unknowns + 1x secondaries + 2x blanks blanks used: 2x from this wheel + 2x from TW3199 m = 1.32 ugC unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2018-12-27 | Water |  |  |  | 66406 | Normal | 14047/11 |  |  | 2018-11-05 | 3613 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.00196 | 75820 | [25 empty] | 2018-11-05 | 2293 | Gas Splitting | Completed
[168 empty] | 1.0012 | 0.00195 | 75822 | [25 empty] | 2018-11-05 | 4076 | Choose Gas Splitting | Completed
[168 empty] | 1.0031 | 0.00201 | 75825 | [25 empty] | 2018-11-06 | 15851 | Graphite | Completed
[168 empty] | 0.99842 | 0.00198 | 75840 | [25 empty] | 2018-11-07 | 6222 | C13 Measurement | Completed
[168 empty] | 0.99898 | 0.00198 | 75856 | [25 empty] | 2018-11-21 | 17176 | AMS Submission Results | Completed
[168 empty] | 1.0001 | 0.002 | 75875 | [25 empty] | 2018-11-22 | 20617 | Close Job | Completed
[168 empty] | 0.99971 | 0.002 | 75876
[7 empty] | 110.5 | 87.9367 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39969 | 39969 | 39969 | 181 | 181 | 181 |  |  | jenny dahl | -993.1 | 0.16 ‰ | 2018-11-20 | 2018-05-24 | 13 | 180 |  |  |  |  | 6.13‰ | 6.13‰ | 2.00‰ | 2.00‰ | 6.13 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0013 | 0 | 0.8165 | -993.10 ‰ | -993.10 ‰ | 0.16 | 0.16 | -993.01 | 0.16 ‰ | -993.01 | 0.16 ‰ |  |  | 2018-11-21 | Jocelyn Turnbull | 10:59 | 0.0069042 | 0.0069042 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.5 |  |  | 0.0069 | 0.00016 | 17222 | 2018-11-21 | 9:46 | Albert Zondervan | 211057 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.69 | 0.02% | 8 | 5 | 5 | ... | 0.00664 | 0.00015 |  |  |  |  | 0.99422 | 0.99422 | Yes | Yes | [5 empty] | 0.00664 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00079 | 0.00079 | 0.0069 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00664 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-15 | jenny dahl | 9:47 |  | Completed |  |  |  |  | 75854 | 2000000 |  |  | 75849 75851 75853 75904 75905 75906 | 75849~75851~75853~75904~75905~75906 | 3204 |  | 3204 3204 3204 3204 3204 3204 | 2018-11-21 | 10:25 | Albert Zondervan |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0007 | 0.00194 | 75849 |  | XCAMS | EA | 0% |  | TW3204 measurement: * large samples CalAMS error adjustments: * factor 1.3 for cat#17 (TP75850) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is 1.5 sigma older than consensus value * 1x IAEA-C6: result looks ok * 1x FIRI-I: result agrees with consensus value blanks: * Three groups have been identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days: capturing 9x blanks organics via CE-Rafter: 2x unknowns + 0x secondaries + 0x blanks 180 days: capturing 3x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2018-12-27 | Water |  |  |  | 66370 | Normal | 14047/11 |  |  | 2018-11-05 | 3615 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99999 | 0.002 | 75851 | [25 empty] | 2018-11-05 | 2294 | Gas Splitting | Completed
[168 empty] | 1.0033 | 0.00194 | 75853 | [25 empty] | 2018-11-05 | 4078 | Choose Gas Splitting | Completed
[168 empty] | 0.99807 | 0.00194 | 75904 | [25 empty] | 2018-11-07 | 15864 | Graphite | Completed
[168 empty] | 0.99968 | 0.00192 | 75905 | [25 empty] | 2018-11-21 | 17222 | AMS Submission Results | Completed
[168 empty] | 0.99838 | 0.00187 | 75906 | [25 empty] | 2018-11-22 | 20619 | Close Job | Completed
[7 empty] | 74.9 | 88.7507 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45808 | 45807 | 45807 | 275 | 275 | 275 |  |  | jenny dahl | -996.66 | 0.11 ‰ | 2018-12-02 | 2018-06-05 | 9 | 180 |  |  |  |  | 2.50‰ | 2.50‰ | 2.00‰ | 2.00‰ | 2.50 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19 | 0 | 0.8165 | -996.66 ‰ | -996.66 ‰ | 0.11 | 0.11 | -996.62 | 0.12 ‰ | -996.62 | 0.12 ‰ |  |  | 2018-12-06 | Jocelyn Turnbull | 15:28 | 0.0033377 | 0.0033377 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00334 | 0.00011 | 17354 |  |  |  | 211438 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 27 | 5 | 5 | ... | 0.00321 | 0.00011 |  |  |  |  | 0.99981 | 0.99981 | Yes | Yes |  | rather a good water blank! |  |  |  | 0.00321 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00334 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00321 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-11-28 | jenny dahl | 15:01 |  | Completed |  |  |  |  | 75971 | 2000000 |  |  | 75993 75994 75995 75996 75997 75998 | 75993~75994~75995~75996~75997~75998 | 3207 |  | 3207 3207 3207 3207 3207 3207 | 2018-12-04 | 15:43 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99754 | 0.00216 | 75993 |  | XCAMS | EA | 0.13% |  | TW3207 measurement: * large water samples * SNICS indexer stopped measurement unexpectedly, but dataset is large enough CalAMS error adjustments: * factor 1.3 for cat#16,26 (TP75931,75949) primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result looks ok blanks: * Three groups have been identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks IAEA-C1 blank in this wheel (TP75971) looks remarkably good unknowns: * no failures * TP75930: AMS-IRMS difference in del13C = 4 permil |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-07 | Water |  |  |  | 66571 | Normal | 14047/11 |  |  | 2018-11-21 | 3673 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0034 | 0.00214 | 75994 | [25 empty] | 2018-11-21 | 2321 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00217 | 75995 | [25 empty] | 2018-11-21 | 4112 | Choose Gas Splitting | Completed
[168 empty] | 0.99806 | 0.00215 | 75996 | [25 empty] | 2018-11-23 | 15981 | Graphite | Completed
[168 empty] | 1.0013 | 0.00217 | 75997 | [25 empty] | 2018-12-06 | 17354 | AMS Submission Results | Completed
[168 empty] | 0.99965 | 0.00215 | 75998 | [25 empty] | 2018-12-07 | 20845 | Close Job | Completed
[7 empty] | 99.1 | 88.7236 | 0.002 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 35648 | 35648 | 35648 | 127 | 127 | 127 |  |  | Margaret Norris | -988.18 | 0.19 ‰ | 2018-12-07 | 2018-06-10 | 14 | 180 |  |  |  |  | 1.59‰ | 1.59‰ | 2.00‰ | 2.00‰ | 1.59 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -988.18 ‰ | -988.18 ‰ | 0.19 | 0.19 | -988.03 | 0.19 ‰ | -988.03 | 0.19 ‰ |  |  | 2018-12-10 | Jocelyn Turnbull | 10:15 | 0.011822 | 0.011822 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.5 |  |  | 0.01182 | 0.00019 | 17404 |  |  |  | 211402 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.18 | 0.02% | 10 | 5 | 5 | ... | 0.01137 | 0.00018 |  |  |  |  | 0.99859 | 0.99859 | Yes | Yes | [5 empty] | 0.01137 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.01182 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.01137 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2018-12-07 | Margaret Norris | 14:06 |  | Completed |  |  |  |  | 75968 | 2000000 |  |  | 76029 76035 76036 76037 76042 76043 | 76029~76035~76036~76037~76042~76043 | 3209 |  | 3209 3209 3209 3209 3209 3209 | 2018-12-09 | 12:43 | Albert Zondervan |  | 1.28 mg |  | 1.28 | [5 empty] | 1.0005 | 0.00149 | 76029 |  | XCAMS | ST | 0% |  | TW3209 measurement: * some EXPRESS samples * all large samples CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.4 for cat#17 (TP75966), 1.3 for cat#34 (TP76040) * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x IAEA-C6: result agrees with consensus value * 1x FIRI-I: result is 1.5 sigma older than consensus value * 1x FIRI-L whalebone: result is 1.5 sigma younger than consensus value blanks: * Three groups have been identified to blank-correct: (ground)waters: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks bone via AAA: 6x unknowns + 3x secondaries + 0x blanks 180 days: capturing 8x blanks wood via CE-Rafter: 3x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-02 | Water | Complete |  |  | 66651 | Normal | 14047/11 |  |  | 2018-11-21 | 3671 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99794 | 0.00149 | 76035 | [25 empty] | 2018-11-21 | 2323 | Gas Splitting | Completed
[168 empty] | 1.0002 | 0.00151 | 76036 | [25 empty] | 2018-11-21 | 4116 | Choose Gas Splitting | Completed
[168 empty] | 1.0014 | 0.0015 | 76037 | [25 empty] | 2018-11-23 | 15978 | Graphite | Completed
[168 empty] | 0.99942 | 0.00148 | 76042 | [25 empty] | 2018-12-10 | 17404 | AMS Submission Results | Completed
[168 empty] | 1.0006 | 0.00149 | 76043 | [25 empty] | 2018-12-10 | 20895 | Close Job | Completed
[7 empty] | 77.8 | 88.7397 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 175 | 175 | 175 |  |  | Julia Collins | -994.28 | 0.12 ‰ | 2019-01-16 | 2018-07-20 | 6 | 180 |  |  |  |  | 2.84‰ | 2.84‰ | 2.00‰ | 2.00‰ | 2.84 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -994.28 ‰ | -994.28 ‰ | 0.12 | 0.12 | -994.21 | 0.13 ‰ | -994.21 | 0.13 ‰ |  |  | 2019-01-23 | Cathy Ginnane | 14:06 | 0.0057189 | 0.0057189 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.8 |  |  | 0.00572 | 0.00012 | 17466 | 2019-01-18 | 9:05 | Albert Zondervan | 211585 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.57 | 0.01% | 8 | 5 | 5 | ... | 0.0055 | 0.00012 |  |  |  |  | 1 | 1 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00572 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-15 | Julia Collins | 13:47 |  | Completed |  |  |  |  | 76104 | 2000000 |  |  | 76072 76073 76074 76075 76115 76118 | 76072~76073~76074~76075~76115~76118 | 3210 |  | 3210 3210 3210 3210 3210 3210 | 2019-01-18 | 9:21 | Albert Zondervan |  | 1.11 mg |  | 1.11 | [5 empty] | 1.001 | 0.00158 | 76072 |  | XCAMS | ST | 0% |  | TW3210 measurement: * some EXPRESS samples * all large samples * 1st wheel of 2019 CalAMS error adjustments: * all-cathodes factor 1.07 * single-cathode factors 1.3 for cat#23, 1.2 for cat#33 * system error is zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I: result is consistent with consensus value * 1x FIRI-L whalebone: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 23x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks bone&textile via AAA: 4x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-01-29 | Water | Not Complete |  |  | 66721 | Normal | 14047/11 |  |  | 2018-12-11 | 3708 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0009 | 0.00154 | 76073 | [25 empty] | 2018-12-11 | 2352 | Gas Splitting | Completed
[168 empty] | 0.99832 | 0.0015 | 76074 | [25 empty] | 2018-12-11 | 4153 | Choose Gas Splitting | Completed
[168 empty] | 0.99943 | 0.00152 | 76075 | [25 empty] | 2019-01-10 | 16114 | Graphite | Completed
[168 empty] | 0.99969 | 0.00156 | 76115 | [25 empty] | 2019-01-23 | 17466 | AMS Submission Results | Completed
[168 empty] | 1.0008 | 0.00156 | 76118 | [25 empty] | 2019-01-30 | 21006 | Close Job | Completed
[7 empty] | 112.3 | 89.0279 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38373 | 38372 | 38372 | 148 | 148 | 148 |  |  | jenny dahl | -991.58 | 0.16 ‰ | 2019-01-29 | 2018-08-02 | 6 | 180 |  |  |  |  | 0.69‰ | 0.69‰ | 2.00‰ | 2.00‰ | 0.69 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -991.58 ‰ | -991.58 ‰ | 0.16 | 0.16 | -991.47 | 0.16 ‰ | -991.47 | 0.16 ‰ |  |  | 2019-02-04 | Jocelyn Turnbull | 15:37 | 0.0084223 | 0.0084223 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.3 |  |  | 0.00842 | 0.00016 | 17541 |  |  |  | 211700 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 8 | 5 | 5 | ... | 0.0081 | 0.00015 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.0081 | 0.01‰ | 0.00015 | 0.00‰ | No | 1 | 1 | 0.00066 | 0.00066 | 0.00842 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.0081 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-01-25 | jenny dahl | 12:26 |  | Completed |  |  |  |  | 76214 | 2000000 |  |  | 76153 76154 76155 76183 76184 76185 | 76153~76154~76155~76183~76184~76185 | 3212 |  | 3212 3212 3212 3212 3212 3212 |  |  |  |  | 0.89 mg |  | 0.89 | [5 empty] | 1.0008 | 0.0016 | 76153 |  | XCAMS | EA | 0.11% |  | TW3212 measurement: * all large samples CalAMS error adjustments: * single-cathode factors: 1.3 for cat#39 * system error: 0.11% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * groups identified to blank-correct: (ground)waters: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-02-08 | Water |  |  |  | 66771 | Normal | 14047/11 |  |  | 2018-12-18 | 3722 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9996 | 0.00161 | 76154 | [25 empty] | 2018-12-18 | 2365 | Gas Splitting | Completed
[168 empty] | 1.0019 | 0.0016 | 76155 | [25 empty] | 2018-12-18 | 4166 | Choose Gas Splitting | Completed
[168 empty] | 1.0015 | 0.00166 | 76183 | [25 empty] | 2019-01-23 | 16224 | Graphite | Completed
[168 empty] | 0.9987 | 0.00158 | 76184 | [25 empty] | 2019-02-04 | 17541 | AMS Submission Results | Completed
[168 empty] | 0.99783 | 0.00159 | 76185 | [25 empty] | 2019-02-07 | 21069 | Close Job | Completed
[7 empty] | 86.9 | 88.775 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36319 | 36318 | 36318 | 121 | 120 | 120 |  |  | jenny dahl | -989.12 | 0.16 ‰ | 2019-02-22 | 2018-08-26 | 10 | 180 |  |  |  |  | 3.53‰ | 3.53‰ | 2.00‰ | 2.00‰ | 3.53 ‰ | 2.00 ‰ | AMS | 2.64‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -989.12 ‰ | -989.12 ‰ | 0.16 | 0.16 | -988.99 | 0.17 ‰ | -988.99 | 0.17 ‰ |  |  | 2019-02-25 | Jocelyn Turnbull | 11:54 | 0.010876 | 0.010876 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.01088 | 0.00016 | 17719 |  |  |  | 211586 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.09 | 0.02% | 29 | 5 | 8 | ... | 0.01046 | 0.00016 |  |  |  |  | 1.0011 | 1.0011 | Yes | Yes | [5 empty] | 0.01046 | 0.01‰ | 0.00016 | 0.00‰ | No | 1 | 0.99998 | 0.00127 | 0.00127 | 0.01088 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.01046 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-02-15 | jenny dahl | 13:57 |  | Completed |  |  |  |  | 76353 | 2000000 |  |  | 76186 76187 76188 76194 76372 76394 | 76186~76187~76188~76194~76372~76394 | 3217 |  | 3217 3217 3217 3217 3217 3217 | 2019-02-25 | 10:27 | Albert Zondervan |  | 1.01 mg |  | 1.01 | [5 empty] | 0.99924 | 0.0031 | 76186 |  | XCAMS | EA | 0.27% |  | TW3217 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * global factor: 1.03 * single-cathode factors: 1.3 for cat#8 (TP76363) * system error: 0.27% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result is 1.5-sigma younger than consensus value * 1x IAEA-C2 water: result is 2.5-sigma younger than consensus value blanks: * groups identified to blank-correct: carbonates: 26x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2019-01-29 | Water | Complete |  |  | 66972 | Normal | 14047/11 |  |  | 2018-12-14 | 3718 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00313 | 76187 | [25 empty] | 2019-01-23 | 3773 | Water CO2 Evolution | Cancelled
[168 empty] | 1.0057 | 0.00311 | 76188 | [25 empty] | 2019-01-23 | 3775 | Water CO2 Evolution | Completed
[168 empty] | 0.99656 | 0.00309 | 76194 | [25 empty] | 2019-01-23 | 2418 | Gas Splitting | Completed
[168 empty] | 0.999 | 0.0031 | 76372 | [25 empty] | 2019-01-23 | 4228 | Choose Gas Splitting | Completed
[168 empty] | 0.99911 | 0.00317 | 76394 | [25 empty] | 2019-01-28 | 6383 | C13 Measurement | Completed
[196 empty] | 2019-02-12 | 16363 | Graphite | Completed
[196 empty] | 2019-02-25 | 17719 | AMS Submission Results | Completed
[196 empty] | 2019-02-26 | 21253 | Close Job | Completed
[7 empty] | 118 | 88.6531 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38353 | 38352 | 38352 | 148 | 148 | 148 |  |  | Julia Collins | -991.56 | 0.16 ‰ | 2019-03-07 | 2018-09-08 | 15 | 180 |  |  |  |  | 3.38‰ | 3.38‰ | 2.00‰ | 2.00‰ | 3.38 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0046 | 0 | 0.8165 | -991.56 ‰ | -991.56 ‰ | 0.16 | 0.16 | -991.45 | 0.16 ‰ | -991.45 | 0.16 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:33 | 0.0084431 | 0.0084431 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00844 | 0.00016 | 17793 | 2019-03-11 | 9:42 | Albert Zondervan | 212062 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 8 | 5 | 4 | ... | 0.00812 | 0.00015 |  |  |  |  | 0.99958 | 0.99958 | Yes | Yes | [5 empty] | 0.00812 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00844 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00812 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-04 | Julia Collins | 14:12 |  | Completed |  |  |  |  | 76446 | 2000000 |  |  | 76399 76400 76401 76480 76483 76486 | 76399~76400~76401~76480~76483~76486 | 3218 |  | 3218 3218 3218 3218 3218 3218 | 2019-03-18 | 15:08 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99966 | 0.00167 | 76399 |  | XCAMS | EA | 0% |  | TW3218 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% * uncertainty multiplier for cat#36 (TP76442) = 1.4 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks atmos-CH4 as CO2 in break-seals: 13x unknowns + 0x secondaries + 0x blanks Since blank-correction must involve NIWA's process blank as well and since it has to be done separately, only the first step has been done here, by manual entering the graphitisation blank using 117x Kapuni (R40699/1) measurements. unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-04-04 | Water | Complete |  |  | 67011 | Normal | 14047/11 | CH4 |  | 2019-02-19 | 3846 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00195 | 76400 | [25 empty] | 2019-02-20 | 16456 | Graphite | Completed
[168 empty] | 0.99954 | 0.00192 | 76401 | [25 empty] | 2019-02-21 | 6464 | C13 Measurement | Completed
[168 empty] | 0.99953 | 0.00169 | 76480 | [25 empty] | 2019-03-21 | 17793 | AMS Submission Results | Completed
[168 empty] | 0.99774 | 0.00169 | 76483 | [25 empty] | 2019-03-21 | 21458 | Close Job | Completed
[168 empty] | 1.0035 | 0.0017 | 76486
[7 empty] | 94.5 | 88.7104 | 0.0025 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33678 | 33678 | 33678 | 127 | 127 | 127 |  |  | Julia Collins | -984.89 | 0.24 ‰ | 2019-03-14 | 2018-09-15 | 13 | 180 |  |  |  |  | 2.61‰ | 2.61‰ | 2.00‰ | 2.00‰ | 2.61 ‰ | 2.00 ‰ | AMS | 2.40‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -984.89 ‰ | -984.89 ‰ | 0.24 | 0.24 | -984.71 | 0.24 ‰ | -984.71 | 0.24 ‰ |  |  | 2019-03-21 | Jocelyn Turnbull | 9:41 | 0.015108 | 0.015108 | 0.00024 | 0.00024 | 0.00024 | 0.00024 | 3.5 |  |  | 0.01511 | 0.00024 | 17870 | 2019-03-15 | 16:48 | Albert Zondervan | 212111 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:49 | Jocelyn Turnbull | 0 |  | 1.51 | 0.02% | 15 | 5 | 7 | ... | 0.01453 | 0.00023 |  |  |  |  | 0.99946 | 0.99946 | Yes | Yes | [5 empty] | 0.01453 | 0.01‰ | 0.00023 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.01511 | 0 | 0.00024 | 0 | 0 |  |  | 0 | 0 | 0.01453 | 0 | 0.00023 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-03-13 | Julia Collins | 14:28 |  | Completed |  |  |  |  | 76506 | 2000000 |  |  | 76491 76524 76525 76589 76599 76600 | 76491~76524~76525~76589~76599~76600 | 3220 |  | 3220 3220 3220 3220 3220 3220 | 2019-03-18 | 14:12 | Albert Zondervan |  | 0.57 mg |  | 0.57 | [5 empty] | 0.99959 | 0.00195 | 76491 |  | XCAMS | EA | 0% |  | TW3220 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is 1.5-sigma younger than consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 16x unknowns + 2x secondaries + 1x blanks 180 days: capturing 6x blanks organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2019-04-08 | Water | Not Complete |  | C13 split taken on RG20, not waterline. Higher non-cons when opening to add hydrogen on graphite than others. | 67091 | Normal | 14047/11 |  |  | 2019-02-26 | 3858 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00197 | 76524 | [25 empty] | 2019-02-26 | 3860 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99786 | 0.00196 | 76525 | [25 empty] | 2019-02-28 | 3861 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.00193 | 76589 | [25 empty] | 2019-03-01 | 2442 | Gas Splitting | Cancelled
[168 empty] | 0.99918 | 0.00199 | 76599 | [25 empty] | 2019-03-01 | 16516 | Graphite | Completed
[168 empty] | 1.0005 | 0.00205 | 76600 | [25 empty] | 2019-03-01 | 6480 | C13 Measurement | Completed
[196 empty] | 2019-03-21 | 17870 | AMS Submission Results | Completed
[196 empty] | 2019-03-21 | 21465 | Close Job | Completed
[7 empty] | 87.9 | 88.6826 | 0.0028 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 32834 | 32834 | 32834 | 134 | 134 | 134 |  |  | Julia Collins | -983.22 | 0.28 ‰ | 2019-04-11 | 2018-10-13 | 28 | 180 |  |  |  |  | 2.21‰ | 2.21‰ | 2.00‰ | 2.00‰ | 2.21 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -983.22 ‰ | -983.22 ‰ | 0.28 | 0.28 | -983.01 | 0.28 ‰ | -983.01 | 0.28 ‰ |  |  | 2019-04-15 | Cathy Ginnane | 10:01 | 0.016782 | 0.016782 | 0.00028 | 0.00028 | 0.00028 | 0.00028 | 2.8 |  |  | 0.01678 | 0.00028 | 18077 |  |  |  | 212206 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.68 | 0.03% | 6 | 5 | 4 | ... | 0.01614 | 0.00027 |  |  |  |  | 0.99875 | 0.99875 | Yes | Yes |  | Youngest blank in a long time. Not a good blank. Sitting in graphite for three weeks--possibly contributing? |  |  |  | 0.01614 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00082 | 0.00082 | 0.01678 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01614 | 0 | 0.00027 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-10 | Julia Collins | 9:55 |  | Completed |  |  |  |  | 76607 | 2000000 |  |  | 76712 76715 76717 76739 76742 76745 | 76712~76715~76717~76739~76742~76745 | 3226 |  | 3226 3226 3226 3226 3226 3226 | 2019-04-12 | 12:53 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 0.99692 | 0.002 | 76712 |  | XCAMS | EA | 0% |  | TW3226 measurement: * no issues with XCAMS * all large samples CalAMS error adjustments: * system error = 0% primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result agrees with consensus value * 1x FIRI-I cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: waters: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 12x blanks organics via AAA: 7x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks organics via CE-Rafter (cellulose): 6x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2019-05-03 | Water | Complete |  | C13 split sealed as it was not needed at the time | 67339 | Normal | 14047/11 |  |  | 2019-03-11 | 3878 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00201 | 76715 | [25 empty] | 2019-03-19 | 16617 | Graphite | Completed
[168 empty] | 0.99958 | 0.002 | 76717 | [25 empty] | 2019-03-28 | 6505 | C13 Measurement | Completed
[168 empty] | 1.002 | 0.00205 | 76739 | [25 empty] | 2019-04-15 | 18077 | AMS Submission Results | Completed
[168 empty] | 1.0014 | 0.00201 | 76742 | [25 empty] | 2019-04-17 | 21704 | Close Job | Completed
[168 empty] | 1.0001 | 0.002 | 76745
[7 empty] | 75 | 88.8684 | 0.0031 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 32012 | 32011 | 32011 | 90 | 89 | 89 |  |  | Julia Collins | -981.41 | 0.21 ‰ | 2019-04-24 | 2018-10-26 | 9 | 180 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -981.41 ‰ | -981.41 ‰ | 0.21 | 0.21 | -981.18 | 0.21 ‰ | -981.18 | 0.21 ‰ |  |  | 2019-05-07 | Jocelyn Turnbull | 12:49 | 0.018592 | 0.018592 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 2.3 |  |  | 0.01859 | 0.00021 | 18150 | 2019-04-26 | 13:00 | Albert Zondervan | 212255 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.86 | 0.02% | 8 | 5 | 5 | ... | 0.01788 | 0.0002 |  |  |  |  | 1.0008 | 1.0008 | Yes | Yes | [5 empty] | 0.01788 | 0.02‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.01859 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.01788 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-04-18 | Julia Collins | 13:23 |  | Completed |  |  |  |  | 76841 | 2000000 |  |  | 76773 76774 76775 76869 76871 76894 | 76773~76774~76775~76869~76871~76894 | 3228 |  | 3228 3228 3228 3228 3228 3228 | 2019-05-07 | 12:32 | Albert Zondervan |  | 0.80 mg |  | 0.8 | [5 empty] | 0.99885 | 0.00159 | 76773 |  | XCAMS | ST | 0% |  | TW3228 measurement: * no issues with XCAMS * Cathodes had to be transferred to another wheel because of persistent ticking on cat#38,39. This resolved the issue completely. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1.04 * cathode run-to-run variability factor = 1.3 (cat#25), 1.2 (cat#36) primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * two groups identified to blank-correct: waters: 25x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 3x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks: 1x Queets 'A' from this wheel 7x Kauri Renton Rd from previous wheels unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-05-13 | Water |  |  |  | 67451 | Normal | 14047/11 |  |  | 2019-03-18 | 3892 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0006 | 0.00156 | 76774 | [25 empty] | 2019-03-18 | 2457 | Gas Splitting | Completed
[168 empty] | 0.99864 | 0.00161 | 76775 | [25 empty] | 2019-03-18 | 4285 | Choose Gas Splitting | Completed
[168 empty] | 0.9999 | 0.00155 | 76869 | [25 empty] | 2019-04-15 | 16851 | Graphite | Completed
[168 empty] | 0.99943 | 0.00156 | 76871 | [25 empty] | 2019-05-07 | 18150 | AMS Submission Results | Completed
[168 empty] | 1.0025 | 0.00155 | 76894 | [25 empty] | 2019-05-13 | 21849 | Close Job | Completed
[7 empty] | 62.5 | 88.6224 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38895 | 38894 | 38894 | 159 | 158 | 158 |  |  | Margaret Norris | -992.11 | 0.16 ‰ | 2019-05-14 | 2018-11-15 | 7 | 180 |  |  |  |  | 3.50‰ | 3.50‰ | 2.00‰ | 2.00‰ | 3.50 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -18.9413 | 0 | 0.8165 | -992.11 ‰ | -992.11 ‰ | 0.16 | 0.16 | -992.01 | 0.16 ‰ | -992.01 | 0.16 ‰ |  |  | 2019-05-22 | Jocelyn Turnbull | 9:35 | 0.007892 | 0.007892 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.9 |  |  | 0.00789 | 0.00016 | 18309 | 2019-05-17 | 11:00 | Albert Zondervan | 212412 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.79 | 0.02% | 8 | 5 | 6 | ... | 0.00759 | 0.00015 |  |  |  |  | 0.99936 | 0.99929 | No | Yes | [5 empty] | 0.00759 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.0007 | 0.00072 | 0.00789 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00759 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-13 | Margaret Norris | 10:44 |  | Completed |  |  |  |  | 76956 | 2000000 |  |  | 76887 76968 76969 76970 76986 76989 | 76887~76968~76969~76970~76986~76989 | 3232 |  | 3232 3232 3232 3232 3232 3232 | 2019-05-17 | 11:02 | Albert Zondervan |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99914 | 0.00175 | 76887 |  | XCAMS | EA | 0% |  | TW3232 measurement: * SNICS ion source ticky during first rotation; run data rejected. * all large samples CalAMS error adjustments: * system error = 0% * wheel run-to-run variability factor = 1 primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x IAEA-C2 water: result is consistent with consensus value * 1x FIRI-I cellulose: result is consistent with consensus value * 1x IAEA-C6 ANU sucrose result is consistent with consensus value blanks: * three groups identified to blank-correct: waters: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: organics via CE-Rafter: 4x unknowns + 0x secondaries + 0x blanks 180 days: capturing 2x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-06-03 | Water | Not Complete |  |  | 67541 | Normal | 14047/11 |  |  | 2019-05-06 | 5868 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0003 | 0.00176 | 76968 | [25 empty] | 2019-05-06 | 3943 | Water CO2 Evolution | Completed
[168 empty] | 1.0017 | 0.00176 | 76969 | [25 empty] | 2019-05-06 | 2517 | Gas Splitting | Completed
[168 empty] | 1.0008 | 0.00173 | 76970 | [25 empty] | 2019-05-06 | 4362 | Choose Gas Splitting | Completed
[168 empty] | 0.99874 | 0.00179 | 76986 | [25 empty] | 2019-05-07 | 16966 | Graphite | Completed
[168 empty] | 0.99928 | 0.00173 | 76989 | [25 empty] | 2019-05-22 | 18309 | AMS Submission Results | Completed
[196 empty] | 2019-05-22 | 21954 | Close Job | Completed
[7 empty] | 82 | 88.6192 | 0.0024 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33920 | 33919 | 33919 | 102 | 101 | 101 |  |  | jenny dahl | -985.34 | 0.19 ‰ | 2019-05-28 | 2018-11-29 | 8 | 180 |  |  |  |  | 3.62‰ | 3.62‰ | 2.00‰ | 2.00‰ | 3.62 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.89443 | -985.34 ‰ | -985.34 ‰ | 0.19 | 0.19 | -985.16 | 0.19 ‰ | -985.16 | 0.19 ‰ |  |  | 2019-06-05 | Jocelyn Turnbull | 14:03 | 0.014661 | 0.014661 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 3.1 |  |  | 0.01466 | 0.00019 | 18350 | 2019-05-30 | 10:30 | Albert Zondervan | 212238 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.47 | 0.02% | 17 | 5 | 5 | ... | 0.0141 | 0.00018 |  |  |  |  | 0.99944 | 0.99944 | Yes | Yes | [5 empty] | 0.0141 | 0.01‰ | 0.00018 | 0.00‰ | No | 1 | 0.99998 | 0.00098 | 0.00098 | 0.01466 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0141 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-05-27 | jenny dahl | 7:55 |  | Completed |  |  |  |  | 77013 | 2000000 |  |  | 77024 77025 77026 77027 77028 | 77024~77025~77026~77027~77028 | 3233 |  | 3233 3233 3233 3233 3233 | 2019-05-30 | 12:10 | Albert Zondervan |  | 0.67 mg |  | 0.67 | [5 empty] | 0.99938 | 0.00222 | 77024 |  | XCAMS | ST | 0.16% |  | TW3233 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0.16% system error * for excess run-to-run variability: x1.2 (cat#7), x1.3 (cat#32) primaries: * OxI cathodes: 7x, all via ST, all used for normalisation in CalAMS 5x large: no failures 2x small (subm = Rees): no failures d = 1.49 ugC extracted from the 4x small OxI in TW3185,3233 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result differs from consensus value by 4-sigma: ~450 years too old blanks: * Four groups have been identified to blank-correct: carbonates, large: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 4x blanks organics via AAA, large: 1x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks organics via AAA, small, via Excel: 2x unknowns + 1x secondaries + 3x blanks blanks used: Queets'A' only because it could not be resolved why the most recent lot of small KauriRRd appear worse than the few small Queets 'A': 2x Queets from this wheel + 1x from TW3185 m = 0.61 ugC unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-05-07 | Water | Complete |  | Take C13 split on water line and create new job for it. Seal both into breakseals | 67623 | Normal | 14047/11 |  |  | 2019-03-13 | 3890 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.0022 | 77025 | [25 empty] | 2019-03-13 | 2454 | Gas Splitting | Completed
[168 empty] | 0.99634 | 0.00221 | 77026 | [25 empty] | 2019-03-13 | 4279 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00218 | 77027 | [25 empty] | 2019-05-20 | 17023 | Graphite | Completed
[168 empty] | 1.0017 | 0.0022 | 77028 | [25 empty] | 2019-06-05 | 18350 | AMS Submission Results | Completed
[196 empty] | 2019-06-17 | 22048 | Close Job | Completed
[7 empty] | 111.7 | 88.3948 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45252 | 45251 | 45251 | 280 | 280 | 280 |  |  | Julia Collins | -996.42 | 0.12 ‰ | 2019-06-10 | 2018-12-12 | 11 | 180 |  |  |  |  | 4.39‰ | 4.39‰ | 2.00‰ | 2.00‰ | 4.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.42 ‰ | -996.42 ‰ | 0.12 | 0.12 | -996.38 | 0.13 ‰ | -996.38 | 0.13 ‰ |  |  | 2019-06-14 | Jocelyn Turnbull | 9:20 | 0.0035769 | 0.0035769 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00358 | 0.00012 | 18466 | 2019-06-11 | 16:34 | Albert Zondervan | 212582 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 27 | 5 | 5 | ... | 0.00344 | 0.00012 |  |  |  |  | 0.99767 | 0.99767 | Yes | Yes | [5 empty] | 0.00344 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00358 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00344 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-06-07 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 77113 | 2000000 |  |  | 76991 76994 77057 77058 77060 77061 | 76991~76994~77057~77058~77060~77061 | 3235 |  | 3235 3235 3235 3235 3235 3235 | 2019-06-11 | 16:47 | Albert Zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99769 | 0.00158 | 76991 |  | XCAMS | EA | 0% |  | TW3235 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C2 carbonate: result agrees with consensus value * 1x IAEA-C2 water: result agrees with consensus value blanks: * groups identified to blank-correct: carbonates: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 5x blanks (ground)waters: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-06-26 | Water |  |  |  | 67704 | Normal | 14047/11 |  |  | 2019-05-29 | 3994 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00158 | 76994 | [25 empty] | 2019-05-29 | 2531 | Gas Splitting | Completed
[168 empty] | 1.0003 | 0.00159 | 77057 | [25 empty] | 2019-05-29 | 4383 | Choose Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00161 | 77058 | [25 empty] | 2019-05-30 | 17123 | Graphite | Completed
[168 empty] | 1.0017 | 0.0016 | 77060 | [25 empty] | 2019-06-14 | 18466 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00159 | 77061 | [25 empty] | 2019-06-17 | 22074 | Close Job | Completed
[7 empty] | 94.3 | 87.9677 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37358 | 37358 | 37358 | 157 | 157 | 157 |  |  | jenny dahl | -990.44 | 0.19 ‰ | 2019-07-09 | 2019-01-10 | 11 | 180 |  |  |  |  | 5.11‰ | 5.11‰ | 2.00‰ | 2.00‰ | 5.11 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0002 | 0 | 0.8165 | -990.44 ‰ | -990.44 ‰ | 0.19 | 0.19 | -990.33 | 0.19 ‰ | -990.33 | 0.19 ‰ |  |  | 2019-07-15 | Cathy Ginnane | 14:20 | 0.0095557 | 0.0095557 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.6 |  |  | 0.00956 | 0.00019 | 18639 | 2019-07-10 | 16:45 | albert zondervan | 212714 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.96 | 0.02% | 10 | 5 | 5 | ... | 0.00919 | 0.00018 |  |  |  |  | 0.99356 | 0.99356 | Yes | Yes |  | AMS d13C 2.5 per mil heavy |  |  |  | 0.00919 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00956 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00919 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-05 | jenny dahl | 13:51 |  | Completed |  |  |  |  | 77273 | 2000000 |  |  | 77346 77347 77358 77359 77360 77361 | 77346~77347~77358~77359~77360~77361 | 3240 |  | 3240 3240 3240 3240 3240 3240 | 2019-07-10 | 17:02 | albert zondervan |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0017 | 0.00169 | 77346 |  | XCAMS | ST | 0% |  | TW3240 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: * excess run-to-run variability: zero * system error: zero primaries: * OxI cathodes: 6x, all via ST, no failures secondaries: * 1x IAEA-C2 carbonate, waters: result agrees with consensus value * 1x IAEA-C6 sucrose (changed to tuning not std as unknown; CG), lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organics via AAA: 6x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-07-19 | Water |  |  |  | 67911 | Normal | 14047/11 |  |  | 2019-06-27 | 4069 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99823 | 0.00177 | 77347 | [25 empty] | 2019-06-27 | 2552 | Gas Splitting | Completed
[168 empty] | 1.0008 | 0.00176 | 77358 | [25 empty] | 2019-06-27 | 4414 | Choose Gas Splitting | Completed
[168 empty] | 0.9997 | 0.0018 | 77359 | [25 empty] | 2019-06-28 | 17283 | Graphite | Completed
[168 empty] | 0.99837 | 0.00176 | 77360 | [25 empty] | 2019-07-15 | 18639 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00174 | 77361 | [25 empty] | 2019-07-16 | 22325 | Close Job | Completed
[7 empty] | 92.9 | 88.7338 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44452 | 44452 | 44452 | 275 | 274 | 274 | 44182.2 | 44732.1 | Julia Collins | -996.05 | 0.14 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.47‰ | 2.47‰ | 2.00‰ | 2.00‰ | 2.47 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.05 ‰ | -996.05 ‰ | 0.14 | 0.14 | -996 | 0.14 ‰ | -996 | 0.14 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0039512 | 0.0039512 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.6 |  |  | 0.00395 | 0.00014 | 18770 | 2019-07-23 | 15:56 | albert zondervan | 212849 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 30 | 5 | 5 | ... | 0.0038 | 0.00013 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.0038 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00395 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0038 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77319 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water | Complete | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 1 new material prepared in job 212871 | 68095 | Normal | 14047/11 |  |  | 2019-07-01 | 4075 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-01 | 2554 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-01 | 4418 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-02 | 17329 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18770 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22490 | Close Job | Completed
[196 empty] | 2019-09-19 | 9374 | Invoice | Completed
[7 empty] | 96.6 | 88.7946 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 45159 | 45159 | 45159 | 277 | 276 | 276 | 44886.8 | 45441 | Julia Collins | -996.38 | 0.12 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 1.79‰ | 1.79‰ | 2.00‰ | 2.00‰ | 1.79 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.38 ‰ | -996.38 ‰ | 0.12 | 0.12 | -996.34 | 0.13 ‰ | -996.34 | 0.13 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0036185 | 0.0036185 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00362 | 0.00012 | 18771 | 2019-07-23 | 15:56 | albert zondervan | 212851 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 29 | 5 | 5 | ... | 0.00348 | 0.00012 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00348 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00362 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00348 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77354 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 3 new material prepared in job 212871 | 68096 | Normal | 14047/11 |  |  | 2019-07-02 | 4086 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 2567 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-02 | 4432 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-04 | 17364 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18771 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22491 | Close Job | Completed
[196 empty] | 2019-09-19 | 9376 | Invoice | Completed
[7 empty] | 95.9 | 88.7357 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 45045 | 45044 | 45044 | 296 | 295 | 295 | 44754 | 45345.9 | Julia Collins | -996.33 | 0.14 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.33 ‰ | -996.33 ‰ | 0.14 | 0.14 | -996.28 | 0.14 ‰ | -996.28 | 0.14 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0036705 | 0.0036705 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00367 | 0.00014 | 18772 | 2019-07-23 | 15:56 | albert zondervan | 212852 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 31 | 5 | 3 | ... | 0.00353 | 0.00013 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00353 | 0.00‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00367 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00353 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77320 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.24 mg |  | 1.24 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | 2/7/19 This sample was too large on graphite. Gas was transferred back to bottle and split as a C13 to drop the size (CG). Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 1 new material prepared in job 212871 | 68097 | Normal | 14047/11 |  |  | 2019-07-01 | 4076 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 17330 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-08-07 | 18772 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-08-08 | 22492 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-09-19 | 9377 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 98.2 | 88.5721 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41028 | 41027 | 41027 | 193 | 193 | 193 | 40837.1 | 41223.6 | Julia Collins | -993.95 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 4.30‰ | 4.30‰ | 2.00‰ | 2.00‰ | 4.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -993.95 ‰ | -993.95 ‰ | 0.15 | 0.15 | -993.87 | 0.15 ‰ | -993.87 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0060516 | 0.0060516 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00605 | 0.00015 | 18773 | 2019-07-23 | 15:56 | albert zondervan | 212853 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 36 | 5 | 5 | ... | 0.00582 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00582 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00605 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00582 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77355 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 2 new material prepared in job 212871 | 68098 | Normal | 14047/11 |  |  | 2019-07-02 | 4081 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 2559 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-02 | 4424 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-04 | 17365 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18773 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22493 | Close Job | Completed
[196 empty] | 2019-09-19 | 9378 | Invoice | Completed
[7 empty] | 94.7 | 88.8006 | 0.0013 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 39252 | 39252 | 39252 | 177 | 177 | 177 | 39077 | 39431.1 | Julia Collins | -992.45 | 0.17 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 1.72‰ | 1.72‰ | 2.00‰ | 2.00‰ | 1.72 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -992.45 ‰ | -992.45 ‰ | 0.17 | 0.17 | -992.36 | 0.17 ‰ | -992.36 | 0.17 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0075489 | 0.0075489 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.6 |  |  | 0.00755 | 0.00017 | 18774 | 2019-07-23 | 15:56 | albert zondervan | 212854 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.75 | 0.02% | 32 | 5 | 3 | ... | 0.00726 | 0.00016 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00726 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00755 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00726 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77362 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.24 mg |  | 1.24 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water | Not Complete | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel flushed for 2 minutes without carbosorb run in set 3 new material prepared in job 212871 | 68099 | Normal | 14047/11 |  |  | 2019-07-03 | 4088 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-05 | 17372 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-08-07 | 18774 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-08-08 | 22494 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-09-19 | 9379 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 96.7 | 88.6529 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48110 | 48110 | 48110 | 300 | 299 | 299 |  |  | Julia Collins | -997.49 | 0.09 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 3.39‰ | 3.39‰ | 2.00‰ | 2.00‰ | 3.39 ‰ | 2.00 ‰ | AMS | 3.07‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -997.49 ‰ | -997.49 ‰ | 0.09 | 0.09 | -997.46 | 0.09 ‰ | -997.46 | 0.09 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0025059 | 0.0025059 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00251 | 9e-05 | 18775 | 2019-07-23 | 15:56 | albert zondervan | 212855 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.25 | 0.01% | 33 | 5 | 6 | ... | 0.00241 | 9e-05 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00241 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00251 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00241 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77321 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 3.1 | RRL | [5 empty] | Background Inorganic |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel as regular standard run in set 1 new material prepared in job 212871 | 68100 | Normal | 14047/11 |  |  | 2019-07-01 | 4077 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-01 | 2555 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-01 | 4419 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-02 | 17331 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-07-09 | 6694 | C13 Measurement | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-07 | 18775 | AMS Submission Results | Completed
[196 empty] | 2019-08-08 | 22495 | Close Job | Completed
[196 empty] | 2019-09-19 | 9380 | Invoice | Completed
[7 empty] | 91.8 | 88.6935 | 0.0005 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 46877 | 46876 | 46876 | 286 | 285 | 285 | 46596.1 | 47168 | Julia Collins | -997.08 | 0.10 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -997.08 ‰ | -997.08 ‰ | 0.1 | 0.1 | -997.04 | 0.11 ‰ | -997.04 | 0.11 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0029218 | 0.0029218 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00292 | 0.0001 | 18776 | 2019-07-23 | 15:56 | albert zondervan | 212856 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.29 | 0.01% | 34 | 5 | 5 | ... | 0.00281 | 0.0001 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00281 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00292 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00281 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77322 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.21 mg |  | 1.21 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel as regular standard run in set 1 new material prepared in job 212871 | 68101 | Normal | 14047/11 |  |  | 2019-07-01 | 4078 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-01 | 2556 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-01 | 4420 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-02 | 17332 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18776 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22496 | Close Job | Completed
[196 empty] | 2019-09-19 | 9381 | Invoice | Completed
[7 empty] | 99.6 | 88.5219 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41734 | 41734 | 41734 | 211 | 210 | 210 | 41526.2 | 41948.3 | Julia Collins | -994.46 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 4.87‰ | 4.87‰ | 2.00‰ | 2.00‰ | 4.87 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0001 | 0 | 0.8165 | -994.46 ‰ | -994.46 ‰ | 0.15 | 0.15 | -994.39 | 0.15 ‰ | -994.39 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.0055421 | 0.0055421 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00554 | 0.00015 | 18777 | 2019-07-23 | 15:56 | albert zondervan | 212857 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.55 | 0.01% | 37 | 5 | 6 | ... | 0.00533 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00533 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00554 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00533 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77356 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel as regular standard run in set 2 new material prepared in job 212871 | 68102 | Normal | 14047/11 |  |  | 2019-07-02 | 4082 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 2563 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-02 | 4428 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-04 | 17366 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-07-09 | 6704 | C13 Measurement | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-07 | 18777 | AMS Submission Results | Completed
[196 empty] | 2019-08-08 | 22497 | Close Job | Completed
[196 empty] | 2019-09-19 | 9382 | Invoice | Completed
[7 empty] | 100.2 | 88.7416 | 0.0026 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33332 | 33332 | 33332 | 132 | 132 | 132 | 33200.9 | 33465.7 | Julia Collins | -984.23 | 0.26 ‰ | 2019-07-22 | 2019-01-23 | 18 | 180 |  |  |  |  | 2.39‰ | 2.39‰ | 2.00‰ | 2.00‰ | 2.39 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -984.23 ‰ | -984.23 ‰ | 0.26 | 0.26 | -984.03 | 0.26 ‰ | -984.03 | 0.26 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.015774 | 0.015774 | 0.00026 | 0.00026 | 0.00026 | 0.00026 | 1.8 |  |  | 0.01577 | 0.00026 | 18778 | 2019-07-23 | 15:56 | albert zondervan | 212859 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.58 | 0.03% | 19 | 5 | 5 | ... | 0.01517 | 0.00025 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.01517 | 0.02‰ | 0.00025 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.01577 | 0 | 0.00026 | 0 | 0 |  |  | 0 | 0 | 0.01517 | 0 | 0.00025 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77357 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.13 mg |  | 1.13 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run in set 1 new material prepared in job 212871 | 68103 | Normal | 14047/11 |  |  | 2019-07-02 | 4083 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 2561 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-02 | 4426 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-04 | 17367 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18778 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22498 | Close Job | Completed
[196 empty] | 2019-09-19 | 9384 | Invoice | Completed
[7 empty] | 100.5 | 88.5317 | 0.0032 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31847 | 31847 | 31847 | 119 | 118 | 118 | 31729.4 | 31967.1 | Julia Collins | -981.02 | 0.28 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -981.02 ‰ | -981.02 ‰ | 0.28 | 0.28 | -980.79 | 0.28 ‰ | -980.79 | 0.28 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.018976 | 0.018976 | 0.00028 | 0.00028 | 0.00028 | 0.00028 | 1.5 |  |  | 0.01898 | 0.00028 | 18779 | 2019-07-23 | 15:56 | albert zondervan | 212860 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.9 | 0.03% | 20 | 5 | 3 | ... | 0.01825 | 0.00027 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.01825 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.01898 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01825 | 0 | 0.00027 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77363 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run in set 2 new material prepared in job 212871 | 68104 | Normal | 14047/11 |  |  | 2019-07-03 | 4084 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-05 | 17373 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-08-07 | 18779 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-08-08 | 22499 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-09-19 | 9385 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 100.8 | 88.6505 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42329 | 42328 | 42328 | 227 | 227 | 227 | 42104.6 | 42559.1 | Julia Collins | -994.85 | 0.15 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 3.42‰ | 3.42‰ | 2.00‰ | 2.00‰ | 3.42 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.85 ‰ | -994.85 ‰ | 0.15 | 0.15 | -994.79 | 0.15 ‰ | -994.79 | 0.15 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.005147 | 0.005147 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2 |  |  | 0.00515 | 0.00015 | 18780 | 2019-07-23 | 15:56 | albert zondervan | 212861 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 25 | 5 | 5 | ... | 0.00495 | 0.00014 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00495 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00515 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00495 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77364 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run small vessel no flushing run set 3 new material prepared in job 212871 | 68105 | Normal | 14047/11 |  |  | 2019-07-03 | 4091 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-03 | 2570 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-03 | 4443 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-05 | 17374 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18780 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22500 | Close Job | Completed
[196 empty] | 2019-09-19 | 9386 | Invoice | Completed
[7 empty] | 98.5 | 88.6483 | 0.0093 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23267 | 23266 | 23266 | 67 | 66 | 66 | 23200.2 | 23333.3 | Julia Collins | -944.78 | 0.46 ‰ | 2019-07-22 | 2019-01-23 | 20 | 180 |  |  |  |  | 3.44‰ | 3.44‰ | 2.00‰ | 2.00‰ | 3.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -944.78 ‰ | -944.78 ‰ | 0.46 | 0.46 | -944.09 | 0.46 ‰ | -944.09 | 0.46 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.055223 | 0.055223 | 0.00046 | 0.00046 | 0.00046 | 0.00046 | 1.7 |  |  | 0.05522 | 0.00046 | 18781 | 2019-07-23 | 15:56 | albert zondervan | 212862 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.52 | 0.05% | 22 | 5 | 3 | ... | 0.05311 | 0.00044 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.05311 | 0.05‰ | 0.00044 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05522 | 0 | 0.00046 | 0 | 0 |  |  | 0 | 0 | 0.05311 | 0 | 0.00044 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77323 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.20 mg |  | 1.2 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | 2/7/19 This sample was too large on graphite. Gas was transferred back to bottle and split as a C13 to drop the size (CG). Made to test water blanks Jeremy to run large vessel no flushing run in set 1 new material prepared in job 212871 | 68106 | Normal | 14047/11 |  |  | 2019-07-01 | 4079 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 17333 | Graphite | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-08-07 | 18781 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-08-08 | 22501 | Close Job | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-09-19 | 9387 | Invoice | Completed
[168 empty] | 1 | 0.00191 | 77444
[7 empty] | 102 | 88.7564 | 0.0087 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23762 | 23762 | 23762 | 66 | 65 | 65 | 23696.8 | 23828.7 | Julia Collins | -948.08 | 0.43 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.22‰ | 2.22‰ | 2.00‰ | 2.00‰ | 2.22 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -948.08 ‰ | -948.08 ‰ | 0.43 | 0.43 | -947.44 | 0.43 ‰ | -947.44 | 0.43 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.051917 | 0.051917 | 0.00043 | 0.00043 | 0.00043 | 0.00043 | 1.6 |  |  | 0.05192 | 0.00043 | 18782 | 2019-07-23 | 15:56 | albert zondervan | 212863 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.19 | 0.04% | 23 | 5 | 5 | ... | 0.04993 | 0.00041 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.04993 | 0.05‰ | 0.00041 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05192 | 0 | 0.00043 | 0 | 0 |  |  | 0 | 0 | 0.04993 | 0 | 0.00041 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77365 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel no flushing run in set 2 new material prepared in job 212871 | 68107 | Normal | 14047/11 |  |  | 2019-07-02 | 4085 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-02 | 2564 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-02 | 4429 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-05 | 17375 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18782 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22502 | Close Job | Completed
[196 empty] | 2019-09-19 | 9388 | Invoice | Completed
[7 empty] | 100 | 88.7555 | 0.0089 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 23613 | 23612 | 23612 | 65 | 64 | 64 | 23548.2 | 23677.7 | Julia Collins | -947.11 | 0.43 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.23‰ | 2.23‰ | 2.00‰ | 2.00‰ | 2.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -947.11 ‰ | -947.11 ‰ | 0.43 | 0.43 | -946.45 | 0.43 ‰ | -946.45 | 0.43 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.052894 | 0.052894 | 0.00043 | 0.00043 | 0.00043 | 0.00043 | 1.7 |  |  | 0.05289 | 0.00043 | 18783 | 2019-07-23 | 15:56 | albert zondervan | 212864 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 5.29 | 0.04% | 24 | 5 | 5 | ... | 0.05087 | 0.00041 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.05087 | 0.05‰ | 0.00041 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.05289 | 0 | 0.00043 | 0 | 0 |  |  | 0 | 0 | 0.05087 | 0 | 0.00041 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77366 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-14 | Water |  | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel no flushing run in set 3 new material prepared in job 212871 | 68108 | Normal | 14047/11 |  |  | 2019-07-03 | 4090 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-03 | 2571 | Gas Splitting | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-03 | 4444 | Choose Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-05 | 17376 | Graphite | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-08-07 | 18783 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-08 | 22503 | Close Job | Completed
[196 empty] | 2019-09-19 | 9389 | Invoice | Completed
[7 empty] | 95.7 | 88.7557 | 0.0012 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 39649 | 39648 | 39648 | 186 | 186 | 186 | 39465.1 | 39837.2 | Julia Collins | -992.82 | 0.17 ‰ | 2019-07-22 | 2019-01-23 | 17 | 180 |  |  |  |  | 2.23‰ | 2.23‰ | 2.00‰ | 2.00‰ | 2.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -992.82 ‰ | -992.82 ‰ | 0.17 | 0.17 | -992.73 | 0.17 ‰ | -992.73 | 0.17 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 14:42 | 0.007185 | 0.007185 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.00718 | 0.00017 | 18794 | 2019-07-23 | 15:56 | albert zondervan | 212895 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.72 | 0.02% | 26 | 5 | 6 | ... | 0.00691 | 0.00016 |  |  |  |  | 0.99959 | 0.99959 | Yes | Yes | [5 empty] | 0.00691 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00077 | 0.00077 | 0.00718 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00691 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-07-19 | Julia Collins | 13:27 |  | Completed |  |  |  |  | 77373 | 2000000 |  |  | 77428 77429 77430 77442 77443 77444 | 77428~77429~77430~77442~77443~77444 | 3244 |  | 3244 3244 3244 3244 3244 3244 |  |  |  |  | 1.17 mg |  | 1.17 | [5 empty] | 1.0009 | 0.00183 | 77428 |  | XCAMS | EA | 0% |  | TW3244 measurement: * no problems with XCAMS * all large samples * includes water-line blank tests CalAMS error adjustments: * excess run-to-run variability: global uncertainty multiplier = 1.05 * system error = zero primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result agrees with consensus value blanks: * groups identified to blank-correct: waters: NOTE: BLANK CORRECTION POSTPONED UNTIL virtual TW YAK RESULTS ARE AVAILABLE 2 waters blanks from tw3245 have been used as well 4x unknowns + 0x secondaries + 1 blanks 180 days: capturing 11x blanks 14x background tests R14047/11 organics via AAA: 10x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-08-28 | Water | Complete | Water blank test JCT 9/19 | Made to test water blanks Jeremy to run large vessel flushed for 2 minutes without carbosorb run in set 2 new material prepared in job 212871 Don't trust result due to vessel # being wrong looks to be large vessel from extract time and vessel number | 68109 | Normal | 14047/11 |  |  | 2019-07-02 | 4080 | Water CO2 Evolution | Cancelled |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99828 | 0.00188 | 77429 | [25 empty] | 2019-07-03 | 4087 | Water CO2 Evolution | Completed
[168 empty] | 0.99801 | 0.00187 | 77430 | [25 empty] | 2019-07-03 | 2573 | Gas Splitting | Completed
[168 empty] | 1.001 | 0.00189 | 77442 | [25 empty] | 2019-07-03 | 4445 | Choose Gas Splitting | Completed
[168 empty] | 1.0018 | 0.00187 | 77443 | [25 empty] | 2019-07-05 | 17383 | Graphite | Completed
[168 empty] | 1 | 0.00191 | 77444 | [25 empty] | 2019-08-07 | 18794 | AMS Submission Results | Completed
[196 empty] | 2019-08-08 | 22504 | Close Job | Completed
[196 empty] | 2019-09-19 | 9392 | Invoice | Completed
[7 empty] | 93.2 | 88.3795 | 0.0023 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 34329 | 34328 | 34328 | 107 | 107 | 107 |  |  | Julia Collins | -986.07 | 0.19 ‰ | 2019-08-03 | 2019-02-04 | 2 | 180 |  |  |  |  | 2.65‰ | 2.65‰ | 2.00‰ | 2.00‰ | 2.65 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -986.07 ‰ | -986.07 ‰ | 0.19 | 0.19 | -985.89 | 0.19 ‰ | -985.89 | 0.19 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.013933 | 0.013933 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 2 |  |  | 0.01393 | 0.00019 | 18809 | 2019-08-05 | 11:38 | Jeremy thompson | 212936 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.39 | 0.02% | 36 | 5 | 4 | ... | 0.0134 | 0.00018 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes |  | bad result. Jeremy to look into |  |  |  | 0.0134 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.01393 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.0134 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 77552 | 2000000 |  |  | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 |  | 3245 3245 3245 3245 3245 3245 | 2019-08-05 | 13:33 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 0.99964 | 0.00213 | 77527 |  | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2019-09-09 | Water | Complete | Water blank test JCT 9/19 | run as normal with new ss carbosorb trap use for blank corr of TW3244 waters | 68055 | Normal | 14047/11 |  |  | 2019-07-31 | 4116 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9995 | 0.0021 | 77528 | [25 empty] | 2019-08-01 | 17562 | Graphite | Completed
[168 empty] | 0.99635 | 0.0021 | 77529 | [25 empty] | 2019-08-05 | 6723 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.0021 | 77530 | [25 empty] | 2019-08-07 | 18809 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00208 | 77557 | [25 empty] | 2019-08-08 | 22505 | Close Job | Completed
[168 empty] | 1.0021 | 0.00211 | 77558 | [25 empty] | 2019-09-19 | 9393 | Invoice | Completed
[7 empty] | 93.3 | 88.4422 | 0.0026 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 33359 | 33358 | 33358 | 100 | 100 | 100 |  |  | Julia Collins | -984.28 | 0.20 ‰ | 2019-08-03 | 2019-02-04 | 2 | 180 |  |  |  |  | 1.94‰ | 1.94‰ | 2.00‰ | 2.00‰ | 1.94 ‰ | 2.00 ‰ | AMS | 2.54‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -984.28 ‰ | -984.28 ‰ | 0.2 | 0.2 | -984.08 | 0.20 ‰ | -984.08 | 0.20 ‰ |  |  | 2019-08-07 | Cathy Ginnane | 13:16 | 0.015722 | 0.015722 | 0.0002 | 0.0002 | 0.0002 | 0.0002 | 2 |  |  | 0.01572 | 0.0002 | 18827 | 2019-08-05 | 11:38 | Jeremy thompson | 212991 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 1.57 | 0.02% | 22 | 5 | 4 | ... | 0.01512 | 0.00019 |  |  |  |  | 0.99577 | 0.99577 | Yes | Yes |  | bad result. jeremy to look into |  |  |  | 0.01512 | 0.02‰ | 0.00019 | 0.00‰ | Yes | 1 | 0.99999 | 0.00086 | 0.00086 | 0.01572 | 0 | 0.0002 | 0 | 0 |  |  | 0 | 0 | 0.01512 | 0 | 0.00019 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-02 | Julia Collins | 13:31 |  | Completed |  |  |  |  | 77553 | 2000000 |  |  | 77527 77528 77529 77530 77557 77558 | 77527~77528~77529~77530~77557~77558 | 3245 |  | 3245 3245 3245 3245 3245 3245 | 2019-08-05 | 13:33 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 0.99964 | 0.00213 | 77527 |  | XCAMS | EA | 0.14% |  | TW3245 measurement: * no problems with XCAMS * all large samples CalAMS error adjustments: *system error: 0.14% primaries: * OxI cathodes: 6x, all via EA, no failures secondaries: * 1x IAEA-C6 sucrose, with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, with AAA group: result is 2 sigma younger than consensus value * 1x IAEA-C2/2 Freshwater Travertine, Result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine, result agrees with consensus value blanks: * groups identified to blank-correct: waters: 9x unknowns + 1x secondaries + 2x blanks 180 days: capturing 11x blanks organics via AAA: 11x unknowns + 2x secondaries + 0x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 4x blanks carbonate: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2019-09-17 | Water | Complete | Water blank test JCT 9/19 | run as normal with new ss carbosorb trap use for blank corr of TW3244 waters | 68056 | Normal | 14047/11 |  |  | 2019-07-31 | 4117 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9995 | 0.0021 | 77528 | [25 empty] | 2019-08-01 | 17563 | Graphite | Completed
[168 empty] | 0.99635 | 0.0021 | 77529 | [25 empty] | 2019-08-05 | 6726 | C13 Measurement | Completed
[168 empty] | 1.0019 | 0.0021 | 77530 | [25 empty] | 2019-08-07 | 18827 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00208 | 77557 | [25 empty] | 2019-08-08 | 22513 | Close Job | Completed
[168 empty] | 1.0021 | 0.00211 | 77558 | [25 empty] | 2019-09-19 | 9394 | Invoice | Completed
[7 empty] | 79.9 | 88.1653 | 0.0014 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38503 | 38502 | 38502 | 161 | 161 | 161 |  |  | jenny dahl | -991.71 | 0.17 ‰ | 2019-08-16 | 2019-02-17 | 1 | 180 |  |  |  |  | 6.14‰ | 6.14‰ | 2.00‰ | 2.00‰ | 6.14 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.89443 | -991.71 ‰ | -991.71 ‰ | 0.17 | 0.17 | -991.61 | 0.17 ‰ | -991.61 | 0.17 ‰ |  |  | 2019-08-20 | Cathy Ginnane | 11:05 | 0.0082872 | 0.0082872 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.6 |  |  | 0.00829 | 0.00017 | 18954 | 2019-08-19 | 14:01 | albert zondervan | 213051 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.83 | 0.02% | 37 | 5 | 5 | ... | 0.00797 | 0.00016 |  |  |  |  | 0.99682 | 0.99682 | Yes | Yes |  | d13C is more than 3.5 per mil heavy. odd c13 result |  |  |  | 0.00797 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00076 | 0.00076 | 0.00829 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00797 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-16 | jenny dahl | 12:31 |  | Completed |  |  |  |  | 77657 | 2000000 |  |  | 77589 77590 77633 77634 77635 | 77589~77590~77633~77634~77635 | 3247 |  | 3247 3247 3247 3247 3247 | 2019-08-19 | 14:23 | albert zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 0.99992 | 0.00165 | 77589 |  | XCAMS | EA | 0% |  | TW3247 measurement: * one problem with XCAMS: indexer error * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA 1x (TP77591) has failed because its rts = 0.990 and has been rejected as an outlier A complete retune mid-way did not improve its rts value. secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value * 1x FIRI-L whalebone: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 12x unknowns + 3x secondaries + 1x blanks 180 days: capturing 6x blanks organics (wood) via CE, Rafter method: 1x unknowns + 0x secondaries + 0x blanks 180 days: capturing 5x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-09-23 | Water |  |  |  | 68123 | Normal | 14047/11 |  |  | 2019-08-13 | 4133 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0017 | 0.00166 | 77590 | [25 empty] | 2019-08-13 | 2589 | Gas Splitting | Completed
[168 empty] | 0.99998 | 0.0017 | 77633 | [25 empty] | 2019-08-13 | 4465 | Choose Gas Splitting | Completed
[168 empty] | 1.0005 | 0.00174 | 77634 | [25 empty] | 2019-08-15 | 17667 | Graphite | Completed
[168 empty] | 0.99766 | 0.00176 | 77635 | [25 empty] | 2019-08-20 | 18954 | AMS Submission Results | Completed
[196 empty] | 2019-08-22 | 22576 | Close Job | Completed
[7 empty] | 105.9 | 88.4117 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46219 | 46218 | 46218 | 184 | 184 | 184 |  |  | jenny dahl | -996.83 | 0.07 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 1.14‰ | 1.14‰ | 2.00‰ | 2.00‰ | 1.14 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -996.83 ‰ | -996.83 ‰ | 0.07 | 0.07 | -996.79 | 0.07 ‰ | -996.79 | 0.07 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0031714 | 0.0031714 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.6 |  |  | 0.00317 | 7e-05 | 18979 | 2019-08-26 | 15:01 | Jeremy thompson | 213006 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 24 | 5 | 3 | ... | 0.00305 | 7e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00305 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00317 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.00305 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77653 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.27 mg |  | 1.27 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2019-09-17 | Water | Complete |  | use for blank correction of ZIP wheel | 68233 | Normal | 14047/11 |  |  | 2019-08-14 | 4132 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-15 | 17663 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18979 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22660 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 107.4 | 88.1079 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42443 | 42443 | 42443 | 148 | 148 | 148 | 42296.3 | 42592.6 | jenny dahl | -994.93 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.60‰ | 4.60‰ | 2.00‰ | 2.00‰ | 4.60 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.93 ‰ | -994.93 ‰ | 0.09 | 0.09 | -994.86 | 0.09 ‰ | -994.86 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0050742 | 0.0050742 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.7 |  |  | 0.00507 | 9e-05 | 18982 | 2019-08-26 | 15:01 | Jeremy thompson | 213053 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 23 | 5 | 3 | ... | 0.00488 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00488 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00507 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00488 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77662 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water | Not Complete | Water blank test JCT 9/19 | sample too large at graphite. had to transfer back to bottle and split. LARGE vessel 7 as normal blank TEST BLANK set 1 | 68234 | Normal | 14047/11 |  |  | 2019-08-14 | 4138 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17672 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18982 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22661 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9395 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 105.4 | 88.1144 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42643 | 42643 | 42643 | 152 | 151 | 151 | 42492.6 | 42796.4 | jenny dahl | -995.05 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 4.52‰ | 4.52‰ | 2.00‰ | 2.00‰ | 4.52 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.05 ‰ | -995.05 ‰ | 0.09 | 0.09 | -994.99 | 0.09 ‰ | -994.99 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0049494 | 0.0049494 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00495 | 9e-05 | 18983 | 2019-08-26 | 15:01 | Jeremy thompson | 213054 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 26 | 5 | 5 | ... | 0.00476 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00476 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00495 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00476 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77659 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | Large vessel 5 as normal blank TEST BLANK set 1 | 68235 | Normal | 14047/11 |  |  | 2019-08-14 | 4139 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-14 | 2590 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-14 | 4467 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-15 | 17669 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-08-29 | 18983 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [25 empty] | 2019-08-30 | 22662 | Close Job | Completed
[196 empty] | 2019-09-19 | 9396 | Invoice | Completed
[7 empty] | 104 | 88.3837 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43441 | 43440 | 43440 | 168 | 167 | 167 | 43274.8 | 43610.3 | jenny dahl | -995.52 | 0.09 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 1.46‰ | 1.46‰ | 2.00‰ | 2.00‰ | 1.46 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -995.52 ‰ | -995.52 ‰ | 0.09 | 0.09 | -995.46 | 0.09 ‰ | -995.46 | 0.09 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0044815 | 0.0044815 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.5 |  |  | 0.00448 | 9e-05 | 18984 | 2019-08-26 | 15:01 | Jeremy thompson | 213055 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.45 | 0.01% | 27 | 5 | 3 | ... | 0.00431 | 9e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00431 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00448 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00431 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77660 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.29 mg |  | 1.29 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | Large vessel 27 as normal blank TEST BLANK set 1 | 68236 | Normal | 14047/11 |  |  | 2019-08-14 | 4140 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-15 | 17670 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18984 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22663 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9397 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 104.7 | 88.4929 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41439 | 41438 | 41438 | 145 | 145 | 145 | 41294.6 | 41585.2 | jenny dahl | -994.25 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 8 | 180 |  |  |  |  | 0.23‰ | 0.23‰ | 2.00‰ | 2.00‰ | 0.23 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.25 ‰ | -994.25 ‰ | 0.1 | 0.1 | -994.18 | 0.11 ‰ | -994.18 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0057501 | 0.0057501 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00575 | 0.0001 | 18985 | 2019-08-26 | 15:01 | Jeremy thompson | 213056 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 29 | 5 | 3 | ... | 0.00553 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00553 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00575 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00553 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77661 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.28 mg |  | 1.28 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | Large vessel 4 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68237 | Normal | 14047/11 |  |  | 2019-08-14 | 4141 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-15 | 17671 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18985 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22664 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9398 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 105 | 88.3036 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41630 | 41629 | 41629 | 149 | 148 | 148 | 41482.3 | 41779.8 | jenny dahl | -994.39 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 2.37‰ | 2.37‰ | 2.00‰ | 2.00‰ | 2.37 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -994.39 ‰ | -994.39 ‰ | 0.1 | 0.1 | -994.32 | 0.11 ‰ | -994.32 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0056149 | 0.0056149 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.5 |  |  | 0.00561 | 0.0001 | 18986 | 2019-08-26 | 15:01 | Jeremy thompson | 213057 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 25 | 5 | 5 | ... | 0.0054 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.0054 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00561 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.0054 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77663 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.25 mg |  | 1.25 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | Large vessel 32 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68238 | Normal | 14047/11 |  |  | 2019-08-14 | 4142 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-14 | 2591 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-14 | 4468 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-16 | 17673 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-08-29 | 18986 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [25 empty] | 2019-08-30 | 22665 | Close Job | Completed
[196 empty] | 2019-09-19 | 9399 | Invoice | Completed
[7 empty] | 104.7 | 88.0268 | 0.0004 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 49771 | 49770 | 49770 | 246 | 245 | 245 | 49528.5 | 50020.5 | jenny dahl | -997.96 | 0.06 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 5.52‰ | 5.52‰ | 2.00‰ | 2.00‰ | 5.52 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -997.96 ‰ | -997.96 ‰ | 0.06 | 0.06 | -997.94 | 0.06 ‰ | -997.94 | 0.06 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.002038 | 0.002038 | 6e-05 | 6e-05 | 6e-05 | 6e-05 | 1.6 |  |  | 0.00204 | 6e-05 | 18987 | 2019-08-26 | 15:01 | Jeremy thompson | 213058 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.2 | 0.01% | 30 | 5 | 3 | ... | 0.00196 | 6e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00196 | 0.00‰ | 6e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00204 | 0 | 6e-05 | 0 | 0 |  |  | 0 | 0 | 0.00196 | 0 | 6e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77664 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.20 mg |  | 1.2 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | sample was oversized on graphite- had to return to bottle and split to reduce size. Large vessel 34 side arm flush 1 minute then regular 2 minute flush TEST BLANK set 1 | 68239 | Normal | 14047/11 |  |  | 2019-08-14 | 4143 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17674 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18987 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22666 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9400 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 103.4 | 88.271 | 0.001 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40837 | 40837 | 40837 | 135 | 134 | 134 | 40703.4 | 40973 | jenny dahl | -993.8 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 2.74‰ | 2.74‰ | 2.00‰ | 2.00‰ | 2.74 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.80 ‰ | -993.80 ‰ | 0.1 | 0.1 | -993.73 | 0.11 ‰ | -993.73 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0061972 | 0.0061972 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.8 |  |  | 0.0062 | 0.0001 | 18988 | 2019-08-26 | 15:01 | Jeremy thompson | 213059 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 31 | 5 | 5 | ... | 0.00596 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00596 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.0062 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00596 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77665 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | LARGE vessel 7 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68240 | Normal | 14047/11 |  |  | 2019-08-15 | 4144 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-15 | 2593 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-15 | 4470 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-16 | 17675 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-08-29 | 18988 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [25 empty] | 2019-08-30 | 22667 | Close Job | Completed
[196 empty] | 2019-09-19 | 9401 | Invoice | Completed
[7 empty] | 103 | 88.3448 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37589 | 37588 | 37588 | 117 | 116 | 116 | 37472.8 | 37706.7 | jenny dahl | -990.71 | 0.14 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 1.90‰ | 1.90‰ | 2.00‰ | 2.00‰ | 1.90 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -990.71 ‰ | -990.71 ‰ | 0.14 | 0.14 | -990.6 | 0.14 ‰ | -990.6 | 0.14 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0092854 | 0.0092854 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00929 | 0.00014 | 18989 | 2019-08-26 | 15:01 | Jeremy thompson | 213060 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 32 | 5 | 5 | ... | 0.00893 | 0.00013 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00893 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00929 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00893 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77666 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | LARGE vessel 27 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68241 | Normal | 14047/11 |  |  | 2019-08-15 | 4145 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-15 | 2594 | Gas Splitting | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-15 | 4471 | Choose Gas Splitting | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-16 | 17676 | Graphite | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-08-29 | 18989 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00137 | 77629 | [25 empty] | 2019-08-30 | 22668 | Close Job | Completed
[196 empty] | 2019-09-19 | 9402 | Invoice | Completed
[7 empty] | 104.3 | 87.9373 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44537 | 44537 | 44537 | 171 | 170 | 170 | 44368.3 | 44710.2 | jenny dahl | -996.09 | 0.08 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 6.54‰ | 6.54‰ | 2.00‰ | 2.00‰ | 6.54 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -996.09 ‰ | -996.09 ‰ | 0.08 | 0.08 | -996.04 | 0.08 ‰ | -996.04 | 0.08 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0039096 | 0.0039096 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 1.6 |  |  | 0.00391 | 8e-05 | 18990 | 2019-08-26 | 15:01 | Jeremy thompson | 213061 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.39 | 0.01% | 33 | 5 | 3 | ... | 0.00376 | 8e-05 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes |  | ~4 per mil difference in C13 |  |  |  | 0.00376 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00391 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00376 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77673 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.22 mg |  | 1.22 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | sample was oversized at graphite. Had to return to bottle and split to reduce size. LARGE vessel 5 with side arm flush 1 minute then regular 2 minute flush TEST BLANK set 2 | 68242 | Normal | 14047/11 |  |  | 2019-08-15 | 4146 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17683 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18990 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22669 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9403 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 103.3 | 88.1528 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37607 | 37606 | 37606 | 117 | 117 | 117 | 37490.6 | 37725 | jenny dahl | -990.74 | 0.14 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.08‰ | 4.08‰ | 2.00‰ | 2.00‰ | 4.08 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -990.74 ‰ | -990.74 ‰ | 0.14 | 0.14 | -990.62 | 0.14 ‰ | -990.62 | 0.14 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0092646 | 0.0092646 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.6 |  |  | 0.00926 | 0.00014 | 18991 | 2019-08-26 | 15:01 | Jeremy thompson | 213062 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 34 | 5 | 3 | ... | 0.00891 | 0.00013 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00891 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00926 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00891 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77674 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | LARGE vessel 32 as normal blank TEST BLANK set 3 | 68243 | Normal | 14047/11 |  |  | 2019-08-15 | 4147 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17684 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18991 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22670 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9404 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 104.9 | 88.0792 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40783 | 40783 | 40783 | 134 | 133 | 133 | 40650.5 | 40918.3 | jenny dahl | -993.76 | 0.10 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 4.92‰ | 4.92‰ | 2.00‰ | 2.00‰ | 4.92 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.76 ‰ | -993.76 ‰ | 0.1 | 0.1 | -993.68 | 0.11 ‰ | -993.68 | 0.11 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0062388 | 0.0062388 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00624 | 0.0001 | 18992 | 2019-08-26 | 15:01 | Jeremy thompson | 213063 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 36 | 5 | 3 | ... | 0.006 | 0.0001 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.006 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00624 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.006 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77675 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.14 mg |  | 1.14 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water |  | Water blank test JCT 9/19 | split on graphite line prior to transfer. don't trust measured pressure. LARGE vessel 4 as normal blank TEST BLANK set 3 | 68244 | Normal | 14047/11 |  |  | 2019-08-15 | 4148 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17685 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18992 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22671 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9405 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 102.7 | 88.341 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40468 | 40468 | 40468 | 142 | 141 | 141 | 40327.9 | 40611.1 | jenny dahl | -993.51 | 0.11 ‰ | 2019-08-23 | 2019-02-24 | 7 | 180 |  |  |  |  | 1.95‰ | 1.95‰ | 2.00‰ | 2.00‰ | 1.95 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0006 | 0 | 0.8165 | -993.51 ‰ | -993.51 ‰ | 0.11 | 0.11 | -993.43 | 0.12 ‰ | -993.43 | 0.12 ‰ |  |  | 2019-08-29 | Jocelyn Turnbull | 15:15 | 0.0064883 | 0.0064883 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.6 |  |  | 0.00649 | 0.00011 | 18993 | 2019-08-26 | 15:01 | Jeremy thompson | 213064 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.65 | 0.01% | 37 | 5 | 3 | ... | 0.00624 | 0.00011 |  |  |  |  | 0.99464 | 0.99464 | Yes | Yes | [5 empty] | 0.00624 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00056 | 0.00056 | 0.00649 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00624 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-08-22 | jenny dahl | 12:02 |  | Completed |  |  |  |  | 77676 | 2000000 |  |  | 77615 77616 77617 77624 77626 77629 | 77615~77616~77617~77624~77626~77629 | 3249 |  | 3249 3249 3249 3249 3249 3249 | 2019-08-26 | 15:23 | Jeremy thompson |  | 1.27 mg |  | 1.27 | [5 empty] | 0.99907 | 0.00136 | 77615 |  | XCAMS | ST | 0% |  | TW3249 measurement: * small and large samples blank-corrected via Excel: mass-threshold = 0.30 mgC CalAMS error adjustments: * 0% system error * for excess run-to-run variability: none primaries: * OxI cathodes: 8x, all via ST, all used for normalisation in CalAMS 6x large: no failures 2x small: no failures d = 2.5 ugC extracted from the 8x small OxI in TW3199,3202,3225,3233,3249 secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x FIRI-I, small: result is 2-sigma older than consensus value: ~250 years too old JCT - it is hard to explain this result from the MCC and DCC corrections - would need DCC to be 4x larger to get to the correct (younger) value, and this is clearly worse than the d values that we measure in the small OxIs in any of our wheels. Reducing MCC to zero leaves it still 200 years too old. So this suggests contamination in the sample material itself, rather than contamination coming from the reagents. Not satisfactory, but not much else to be done. blanks: * Three groups have been identified to blank-correct: waters, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA, large: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks organics via AAA, small, via Excel: 3x unknowns + 1x secondaries + 1x blanks blanks used: KauriRRd m = 1.34 ugC tests: * A set of 12x water blanks, prepared by Jeremy, were in this wheel. unknowns: * no failures |  | RRL | [5 empty] | Background Test |  | 2019-09-24 | Water | Complete | Water blank test JCT 9/19 | LARGE vessel 34 as normal blank TEST BLANK set 3 | 68245 | Normal | 14047/11 |  |  | 2019-08-15 | 4149 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99948 | 0.00137 | 77616 | [25 empty] | 2019-08-16 | 17686 | Graphite | Completed
[168 empty] | 1.002 | 0.00136 | 77617 | [25 empty] | 2019-08-29 | 18993 | AMS Submission Results | Completed
[168 empty] | 0.99799 | 0.00137 | 77624 | [25 empty] | 2019-08-30 | 22672 | Close Job | Completed
[168 empty] | 0.99994 | 0.00135 | 77626 | [25 empty] | 2019-09-19 | 9406 | Invoice | Completed
[168 empty] | 1.0015 | 0.00137 | 77629
[7 empty] | 89.4 | 88.3683 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42152 | 42152 | 42152 | 206 | 206 | 206 |  |  | Julia Collins | -994.74 | 0.14 ‰ | 2019-09-09 | 2019-03-13 | 13 | 180 |  |  |  |  | 2.92‰ | 2.92‰ | 2.00‰ | 2.00‰ | 2.92 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -994.74 ‰ | -994.74 ‰ | 0.14 | 0.14 | -994.67 | 0.14 ‰ | -994.67 | 0.14 ‰ |  |  | 2019-09-12 | Jocelyn Turnbull | 14:54 | 0.0052614 | 0.0052614 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 1.8 |  |  | 0.00526 | 0.00014 | 19075 | 2019-09-12 | 13:18 | Jeremy thompson | 213281 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.53 | 0.01% | 27 | 5 | 4 | ... | 0.00506 | 0.00013 |  |  |  |  | 0.99591 | 0.99591 | Yes | Yes | [5 empty] | 0.00506 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00071 | 0.00071 | 0.00526 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00506 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-06 | Julia Collins | 12:51 |  | Completed |  |  |  |  | 77728 | 2000000 |  |  | 77746 77747 77748 77763 77765 77768 | 77746~77747~77748~77763~77765~77768 | 3251 |  | 3251 3251 3251 3251 3251 3251 | 2019-09-12 | 13:41 | Jeremy thompson |  | 1.09 mg |  | 1.09 | [5 empty] | 1 | 0.00173 | 77746 |  | XCAMS | EA | 0% |  | TW3251 measurement: * no problem with XCAMS * regular precision reached * all large samples CalAMS error adjustments: * system error: 0% primaries: * OxI cathodes: 6x, all via EA no failures secondaries: * 1x IAEA-C6 sucrose, lumped in with AAA group: result agrees with consensus value * 1x FIRI-I cellulose, lumped in with AAA group: result agrees with consensus value * 1x IAEA-C2/12 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks organics via AAA: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Inorganic |  | 2019-10-17 | Water |  |  |  | 68313 | Normal | 14047/11 |  |  | 2019-08-26 | 4172 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99965 | 0.00175 | 77747 | [25 empty] | 2019-08-27 | 17738 | Graphite | Completed
[168 empty] | 1.0002 | 0.00176 | 77748 | [25 empty] | 2019-08-30 | 6775 | C13 Measurement | Completed
[168 empty] | 0.99975 | 0.00176 | 77763 | [25 empty] | 2019-09-12 | 19075 | AMS Submission Results | Completed
[168 empty] | 1.0003 | 0.00173 | 77765 | [25 empty] | 2019-09-13 | 22746 | Close Job | Completed
[168 empty] | 1.0001 | 0.00171 | 77768
[7 empty] | 93.9 | 88.9972 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40959 | 40959 | 40959 | 178 | 177 | 177 |  |  | jenny dahl | -993.9 | 0.14 ‰ | 2019-09-20 | 2019-03-24 | 7 | 180 |  |  |  |  | 0.92‰ | 0.92‰ | 2.00‰ | 2.00‰ | 0.92 ‰ | 2.00 ‰ | AMS | 2.68‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -993.90 ‰ | -993.90 ‰ | 0.14 | 0.14 | -993.82 | 0.14 ‰ | -993.82 | 0.14 ‰ |  |  | 2019-09-27 | Jocelyn Turnbull | 11:26 | 0.0061036 | 0.0061036 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.0061 | 0.00014 | 19147 | 2019-09-24 | 11:42 | albert zondervan | 213325 |  |  |  |  | 45% | 5 | 2/10/19 | 15:28:50 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 16 | 5 | 6 | ... | 0.00587 | 0.00013 |  |  |  |  | 1.001 | 1.001 | Yes | Yes | [5 empty] | 0.00587 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.0061 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00587 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-20 | jenny dahl | 12:34 |  | Completed |  |  |  |  | 77838 | 2000000 |  |  | 77778 77782 77784 77785 77786 77787 | 77778~77782~77784~77785~77786~77787 | 3253 |  | 3253 3253 3253 3253 3253 3253 | 2019-09-24 | 12:49 | albert zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99966 | 0.00171 | 77778 |  | XCAMS | ST | 0% |  | TW3253 measurement: * small & large samples * XCAMS: IML electrode 1 mm closer to ionizer, as part of UCI-cathode tests As observed in TW3234, a bias in delta13C is apparent in comparing the IRMS and AMS values. This effect appears sample size dependent and is probably generated by the modified different geometry in the ion source. CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 3x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x FIRI-I, small: result is NOT consistent with consensus value: 2.5-sigma older than consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 13x blanks large organics via AAA: 6x unknowns + 0x secondaries + 2x blanks 180 days: capturing 10x blanks: 1x Queets A from this wheel 1x Queets A from TW3228 1x Kauri Renton Rd from this wheel 7x Kauri Renton Rd from previous wheels small organics via AAA: 6x unknowns + 1x secondaries + 4x blanks Excel file 'to blank-correct small samples.xlsm': 2x other wheels selected also, to assist with determining d, m parameters: TW3249, 3233 d = 1.5 ug C m = 0.81 ug C, derived from 5x Queets A blanks + 3x Kauri Renton Road blanks (Queets A on its own would result in m = 0.72 ug C; KRR on its own: m = 1.02 ug C.) unknowns: * no failures * Gary Wilson submitted unknowns and blanks processed at Otago U. They fall into both our small and large sample ranges. Consistency of their Queets A blanks with our Kauri Renton Rd blanks was tested in both sample size ranges. | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-10-21 | Water |  |  |  | 68393 | Normal | 14047/11 |  |  | 2019-09-12 | 4191 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.00172 | 77782 | [25 empty] | 2019-09-12 | 2622 | Gas Splitting | Completed
[168 empty] | 0.99982 | 0.0017 | 77784 | [25 empty] | 2019-09-12 | 4508 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00163 | 77785 | [25 empty] | 2019-09-13 | 17848 | Graphite | Completed
[168 empty] | 1.0003 | 0.00171 | 77786 | [25 empty] | 2019-09-18 | 6792 | C13 Measurement | Completed
[168 empty] | 0.99819 | 0.00168 | 77787 | [25 empty] | 2019-09-27 | 19147 | AMS Submission Results | Completed
[196 empty] | 2019-10-01 | 22877 | Close Job | Completed
[7 empty] | 104.6 | 88.7704 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43862 | 43861 | 43861 | 216 | 216 | 216 |  |  | Julia Collins | -995.75 | 0.11 ‰ | 2019-09-27 | 2019-03-31 | 0 | 180 |  |  |  |  | 1.02‰ | 1.02‰ | 2.00‰ | 2.00‰ | 1.02 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0011 | 0 | 0.8165 | -995.75 ‰ | -995.75 ‰ | 0.11 | 0.11 | -995.69 | 0.12 ‰ | -995.69 | 0.12 ‰ |  |  | 2019-09-30 | Jocelyn Turnbull | 11:16 | 0.0042528 | 0.0042528 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.1 |  |  | 0.00425 | 0.00011 | 19174 | 2019-09-30 | 9:14 | jeremy thompson | 213282 |  |  |  |  | 45% | 5 | 12/11/20 | 09:31:58 | albert zondervan | 0 |  | 0.43 | 0.01% | 8 | 5 | 4 | ... | 0.00409 | 0.00011 |  |  |  |  | 0.99855 | 0.99855 | Yes | Yes | [5 empty] | 0.00409 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00425 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00409 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-09-27 | Julia Collins | 13:15 |  | Completed |  |  |  |  | 77973 | 2000000 |  |  | 77853 77854 77855 77856 77857 77861 | 77853~77854~77855~77856~77857~77861 | 3254 |  | 3254 3254 3254 3254 3254 3254 | 2019-09-30 | 9:30 | Jeremy thompson |  | 0.96 mg |  | 0.96 | [5 empty] | 0.99998 | 0.00164 | 77853 |  | XCAMS | ST | 0% |  | TW3254 measurement: * all large samples in NEC cathode holders * no problem with XCAMS * SNICS in modified geometry for UCI-cathode holder tests * regular precision reached CalAMS error adjustments: * system error: 0% possible problem: * sample-size dependent difference in del13C measured by IRMS and AMS * cause is unknown primaries: * OxI cathodes: 6x, all via ST no failures secondaries: * 1x IAEA-C6 sucrose, not used beyond tuning * 1x IAEA-C2 Freshwater Travertine: result agrees with consensus value blanks: * groups identified to blank-correct: waters: 31x unknowns + 1x secondaries + 1x blanks 180 days: capturing 14x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2019-10-17 | Water |  |  |  | 68419 | Normal | 14047/11 |  |  | 2019-09-26 | 4222 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00163 | 77854 | [25 empty] | 2019-09-27 | 17983 | Graphite | Completed
[168 empty] | 1 | 0.00166 | 77855 | [25 empty] | 2019-09-30 | 6837 | C13 Measurement | Completed
[168 empty] | 1 | 0.00167 | 77856 | [25 empty] | 2019-09-30 | 19174 | AMS Submission Results | Completed
[168 empty] | 0.99823 | 0.00167 | 77857 | [25 empty] | 2019-10-01 | 22873 | Close Job | Completed
[168 empty] | 1.0011 | 0.00171 | 77861
[7 empty] | 78.2 | 88.4498 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40455 | 40455 | 40455 | 231 | 231 | 231 | 40227.3 | 40690.1 | Julia Collins | -993.5 | 0.19 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 3.46‰ | 3.46‰ | 2.00‰ | 2.00‰ | 3.46 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -993.50 ‰ | -993.50 ‰ | 0.19 | 0.19 | -993.42 | 0.19 ‰ | -993.42 | 0.19 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0064987 | 0.0064987 | 0.00019 | 0.00019 | 0.00019 | 0.00019 | 1.9 |  |  | 0.0065 | 0.00019 | 19371 | 2019-10-25 | 8:47 | Jeremy thompson | 213679 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:31 | Jocelyn Turnbull | 0 |  | 0.65 | 0.02% | 9 | 5 | 3 | ... | 0.00625 | 0.00018 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00625 | 0.01‰ | 0.00018 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.0065 | 0 | 0.00019 | 0 | 0 |  |  | 0 | 0 | 0.00625 | 0 | 0.00018 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78153 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.08 mg |  | 1.08 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 3 minute extraction test water blanks to be run by Jeremy | 68593 | Normal | 14047/11 |  |  | 2019-10-21 | 4296 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-22 | 18163 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19371 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23119 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10102 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79.7 | 88.3921 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43385 | 43385 | 43385 | 278 | 277 | 277 | 43112.1 | 43667.6 | Julia Collins | -995.49 | 0.16 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 4.11‰ | 4.11‰ | 2.00‰ | 2.00‰ | 4.11 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -995.49 ‰ | -995.49 ‰ | 0.16 | 0.16 | -995.43 | 0.16 ‰ | -995.43 | 0.16 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0045127 | 0.0045127 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.8 |  |  | 0.00451 | 0.00016 | 19372 | 2019-10-25 | 8:47 | Jeremy thompson | 213680 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:32 | Jocelyn Turnbull | 0 |  | 0.45 | 0.02% | 8 | 5 | 3 | ... | 0.00434 | 0.00015 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00434 | 0.00‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00451 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00434 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78154 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.13 mg |  | 1.13 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 5 minute extraction test water blanks to be run by Jeremy | 68594 | Normal | 14047/11 |  |  | 2019-10-21 | 4297 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-22 | 18164 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19372 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23120 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10103 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 78.5 | 88.3622 | 0.0006 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44954 | 44954 | 44954 | 315 | 315 | 315 | 44645 | 45275.4 | Julia Collins | -996.29 | 0.15 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 4.45‰ | 4.45‰ | 2.00‰ | 2.00‰ | 4.45 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -996.29 ‰ | -996.29 ‰ | 0.15 | 0.15 | -996.24 | 0.15 ‰ | -996.24 | 0.15 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0037121 | 0.0037121 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00371 | 0.00015 | 19373 | 2019-10-25 | 8:47 | Jeremy thompson | 213681 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.37 | 0.01% | 10 | 5 | 3 | ... | 0.00357 | 0.00014 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00357 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00371 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00357 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78155 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.18 mg |  | 1.18 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 7 minute extraction test water blanks to be run by Jeremy | 68595 | Normal | 14047/11 |  |  | 2019-10-21 | 4298 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-22 | 18165 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19373 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23121 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10104 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 77 | 88.4959 | 0.0016 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 37367 | 37367 | 37367 | 184 | 183 | 183 | 37185.5 | 37553 | Julia Collins | -990.45 | 0.22 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 2.93‰ | 2.93‰ | 2.00‰ | 2.00‰ | 2.93 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -990.45 ‰ | -990.45 ‰ | 0.22 | 0.22 | -990.34 | 0.22 ‰ | -990.34 | 0.22 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0095453 | 0.0095453 | 0.00022 | 0.00022 | 0.00022 | 0.00022 | 1.7 |  |  | 0.00955 | 0.00022 | 19374 | 2019-10-25 | 8:47 | Jeremy thompson | 213682 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 11 | 5 | 3 | ... | 0.00918 | 0.00021 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00918 | 0.01‰ | 0.00021 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00955 | 0 | 0.00022 | 0 | 0 |  |  | 0 | 0 | 0.00918 | 0 | 0.00021 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78156 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.20 mg |  | 1.2 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 10 minute extraction test water blanks to be run by Jeremy | 68596 | Normal | 14047/11 |  |  | 2019-10-21 | 4299 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-22 | 18166 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19374 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23122 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10105 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79 | 88.451 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38373 | 38372 | 38372 | 198 | 198 | 198 | 38176.6 | 38573.4 | Julia Collins | -991.58 | 0.21 ‰ | 2019-10-23 | 2019-04-26 | 1 | 180 |  |  |  |  | 3.44‰ | 3.44‰ | 2.00‰ | 2.00‰ | 3.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -991.58 ‰ | -991.58 ‰ | 0.21 | 0.21 | -991.47 | 0.21 ‰ | -991.47 | 0.21 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0084223 | 0.0084223 | 0.00021 | 0.00021 | 0.00021 | 0.00021 | 1.9 |  |  | 0.00842 | 0.00021 | 19375 | 2019-10-25 | 8:47 | Jeremy thompson | 213683 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:33 | Jocelyn Turnbull | 0 |  | 0.84 | 0.02% | 12 | 5 | 3 | ... | 0.0081 | 0.0002 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.0081 | 0.01‰ | 0.0002 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00842 | 0 | 0.00021 | 0 | 0 |  |  | 0 | 0 | 0.0081 | 0 | 0.0002 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78157 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 |  |  |  |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 20 minute extraction test water blanks to be run by Jeremy | 68597 | Normal | 14047/11 |  |  | 2019-10-21 | 4300 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-22 | 18167 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19375 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23123 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10106 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 80.7 | 88.3629 | 0.001 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41482 | 41482 | 41482 | 234 | 233 | 233 | 41251.9 | 41719.4 | Julia Collins | -994.28 | 0.17 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.44‰ | 4.44‰ | 2.00‰ | 2.00‰ | 4.44 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.28 ‰ | -994.28 ‰ | 0.17 | 0.17 | -994.21 | 0.17 ‰ | -994.21 | 0.17 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0057189 | 0.0057189 | 0.00017 | 0.00017 | 0.00017 | 0.00017 |  |  |  | 0.00572 | 0.00017 | 19376 | 2019-10-25 | 8:47 | Jeremy thompson | 213684 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:34 | Jocelyn Turnbull | 0 |  | 0.57 | 0.02% | 13 | 5 | 3 | ... | 0.0055 | 0.00016 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.0055 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00572 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.0055 | 0 | 0.00016 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78171 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 3 minute extraction test water blanks to be run by Jeremy | 68598 | Normal | 14047/11 |  |  | 2019-10-22 | 4301 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-23 | 18181 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19376 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23124 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10107 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 77.2 | 88.3651 | 0.0009 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41964 | 41963 | 41963 | 248 | 248 | 248 | 41719.4 | 42215.8 | Julia Collins | -994.61 | 0.17 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.42‰ | 4.42‰ | 2.00‰ | 2.00‰ | 4.42 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -994.61 ‰ | -994.61 ‰ | 0.17 | 0.17 | -994.55 | 0.17 ‰ | -994.55 | 0.17 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.0053861 | 0.0053861 | 0.00017 | 0.00017 | 0.00017 | 0.00017 |  |  |  | 0.00539 | 0.00017 | 19377 | 2019-10-25 | 8:47 | Jeremy thompson | 213685 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:34 | Jocelyn Turnbull | 0 |  | 0.54 | 0.02% | 15 | 5 | 3 | ... | 0.00518 | 0.00016 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.00518 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.00539 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00518 | 0 | 0.00016 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78172 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 5 minute extraction test water blanks to be run by Jeremy | 68599 | Normal | 14047/11 |  |  | 2019-10-22 | 4302 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-23 | 18182 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19377 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23125 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10108 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79.6 | 88.3355 | 0.0018 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 36219 | 36219 | 36219 | 182 | 182 | 182 | 36039.3 | 36403.5 | Julia Collins | -988.99 | 0.25 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 4.76‰ | 4.76‰ | 2.00‰ | 2.00‰ | 4.76 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -988.99 ‰ | -988.99 ‰ | 0.25 | 0.25 | -988.85 | 0.25 ‰ | -988.85 | 0.25 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.011011 | 0.011011 | 0.00025 | 0.00025 | 0.00025 | 0.00025 |  |  |  | 0.01101 | 0.00025 | 19378 | 2019-10-25 | 8:47 | Jeremy thompson | 213686 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:35 | Jocelyn Turnbull | 0 |  | 1.1 | 0.02% | 16 | 5 | 3 | ... | 0.01059 | 0.00024 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.01059 | 0.01‰ | 0.00024 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.01101 | 0 | 0.00025 | 0 | 0 |  |  | 0 | 0 | 0.01059 | 0 | 0.00024 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78173 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 7 minute extraction test water blanks to be run by Jeremy | 68600 | Normal | 14047/11 |  |  | 2019-10-22 | 4303 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-23 | 18183 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19378 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23126 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10109 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 75.7 | 88.5405 | 0.0035 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 31104 | 31103 | 31103 | 124 | 124 | 124 | 30980.3 | 31229.1 | Julia Collins | -979.18 | 0.32 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 2.43‰ | 2.43‰ | 2.00‰ | 2.00‰ | 2.43 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -979.18 ‰ | -979.18 ‰ | 0.32 | 0.32 | -978.93 | 0.33 ‰ | -978.93 | 0.33 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.020817 | 0.020817 | 0.00032 | 0.00032 | 0.00032 | 0.00032 |  |  |  | 0.02082 | 0.00032 | 19379 | 2019-10-25 | 8:47 | Jeremy thompson | 213687 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:36 | Jocelyn Turnbull | 0 |  | 2.08 | 0.03% | 6 | 5 | 3 | ... | 0.02002 | 0.00031 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.02002 | 0.02‰ | 0.00031 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.02082 | 0 | 0.00032 | 0 | 0 |  |  | 0 | 0 | 0.02002 | 0 | 0.00031 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78174 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 10 minute extraction test water blanks to be run by Jeremy | 68601 | Normal | 14047/11 |  |  | 2019-10-22 | 4304 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-23 | 18184 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19379 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23127 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10110 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 79 | 88.5167 | 0.0026 | RRL-Test-NW | Background Test | 1950-01-01 | 1950-01-01 |  |  | 33407 | 33406 | 33406 | 144 | 144 | 144 | 33263.7 | 33552.3 | Julia Collins | -984.37 | 0.28 ‰ | 2019-10-23 | 2019-04-26 | 0 | 180 |  |  |  |  | 2.70‰ | 2.70‰ | 2.00‰ | 2.00‰ | 2.70 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0001 | 0 | 0.8165 | -984.37 ‰ | -984.37 ‰ | 0.28 | 0.28 | -984.18 | 0.28 ‰ | -984.18 | 0.28 ‰ |  |  | 2019-10-25 | Jocelyn Turnbull | 9:27 | 0.015628 | 0.015628 | 0.00028 | 0.00028 | 0.00028 | 0.00028 |  |  |  | 0.01563 | 0.00028 | 19380 | 2019-10-25 | 8:47 | Jeremy thompson | 213688 |  |  |  |  | 45% | 5 | 25/10/19 | 09:26:36 | Jocelyn Turnbull | 0 |  | 1.56 | 0.03% | 17 | 5 | 3 | ... | 0.01503 | 0.00027 |  |  |  |  | 0.99736 | 0.99736 | Yes | Yes | [5 empty] | 0.01503 | 0.02‰ | 0.00027 | 0.00‰ | Yes | 1 | 1 | 0.00093 | 0.00093 | 0.01563 | 0 | 0.00028 | 0 | 0 |  |  | 0 | 0 | 0.01503 | 0 | 0.00027 | 0 |  | Yes | 0.9878 | 0.95 | 1 | 2019-10-23 | Julia Collins | 14:17 |  | Completed |  |  |  |  | 78175 | 2000000 |  |  | 78059 78060 78095 78096 78097 78125 | 78059~78060~78095~78096~78097~78125 | 3259 |  | 3259 3259 3259 3259 3259 3259 | [12 empty] | 0.99874 | 0.00228 | 78059 |  | XCAMS | FLASK | 0% |  |  |  | RRL | [5 empty] | Background Test |  | 2019-11-29 | Water |  | Charge to GCT | 20 minute extraction test water blanks to be run by Jeremy | 68602 | Normal | 14047/11 |  |  | 2019-10-22 | 4305 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99616 | 0.00231 | 78060 | [25 empty] | 2019-10-23 | 18185 | Graphite | Completed
[168 empty] | 0.9984 | 0.00223 | 78095 | [25 empty] | 2019-10-25 | 19380 | AMS Submission Results | Completed
[168 empty] | 1.003 | 0.00229 | 78096 | [25 empty] | 2019-10-30 | 23128 | Close Job | Completed
[168 empty] | 1.0029 | 0.00231 | 78097 | [25 empty] | 2020-02-24 | 10111 | Invoice | Completed
[168 empty] | 1.0008 | 0.00221 | 78125
[7 empty] | 82.1 | 87.5748 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42443 | 42443 | 42443 | 165 | 164 | 164 |  |  | Julia Collins | -994.93 | 0.10 ‰ | 2019-11-27 | 2019-05-31 | 5 | 180 |  |  |  |  | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0004 | 0 | 0.8165 | -994.93 ‰ | -994.93 ‰ | 0.1 | 0.1 | -994.86 | 0.11 ‰ | -994.86 | 0.11 ‰ |  |  | 2019-11-29 | Jocelyn Turnbull | 14:47 | 0.0050742 | 0.0050742 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.6 |  |  | 0.00507 | 0.0001 | 19562 | 2019-11-29 | 13:44 | Jeremy thompson | 213929 |  |  |  |  | 45% | 5 | 29/11/19 | 14:46:38 | Jocelyn Turnbull | 0 |  | 0.51 | 0.01% | 1 | 5 | 5 | ... | 0.00488 | 0.0001 |  |  |  |  | 0.98696 | 0.98696 | Yes | Yes | [5 empty] | 0.00488 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00507 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00488 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-11-26 | Julia Collins | 13:11 |  | Completed |  |  |  |  | 78305 | 2000000 |  |  | 78212 78214 78216 78217 78218 78219 | 78212~78214~78216~78217~78218~78219 | 3264 |  | 3264 3264 3264 3264 3264 3264 | 2019-11-29 | 14:09 | Jeremy thompson |  | 1.25 mg |  | 1.25 | [5 empty] | 1.0003 | 0.0015 | 78212 |  | XCAMS | ST | 0% |  | TW3264 measurement: * large samples * XCAMS: - no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 water, large: result is consistent with consensus value * 1x IAEA-C2 carbonate, large: result is consistent with consensus value * 1x FIRI-I, large: result is consistent with consensus value * 1x Whalebone, large: result is consistent with consensus value blanks: * three groups identified to blank-correct: large Carbonates: 13x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large Waters: 2x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks large Organics via AAA: 10x unknowns + 2x secondaries + 1x blanks 180 days: capturing 7x blanks: unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-01-14 | Water |  |  |  | 68794 | Normal | 14047/11 |  |  | 2019-11-21 | 4328 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99964 | 0.0015 | 78214 | [25 empty] | 2019-11-21 | 2663 | Gas Splitting | Completed
[168 empty] | 0.99997 | 0.00152 | 78216 | [25 empty] | 2019-11-21 | 4595 | Choose Gas Splitting | Completed
[168 empty] | 0.99871 | 0.00152 | 78217 | [25 empty] | 2019-11-22 | 18315 | Graphite | Completed
[168 empty] | 1.0029 | 0.00151 | 78218 | [25 empty] | 2019-11-29 | 19562 | AMS Submission Results | Completed
[168 empty] | 0.99824 | 0.00161 | 78219 | [25 empty] | 2019-12-02 | 23334 | Close Job | Completed
[7 empty] | 106.9 | 87.2294 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40190 | 40189 | 40189 | 187 | 186 | 186 | 40005.5 | 40378.6 | Julia Collins | -993.28 | 0.16 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.94‰ | 2.94‰ | 2.00‰ | 2.00‰ | 2.94 ‰ | 2.00 ‰ | AMS | 2.63‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -993.28 ‰ | -993.28 ‰ | 0.16 | 0.16 | -993.2 | 0.16 ‰ | -993.2 | 0.16 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0067171 | 0.0067171 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00672 | 0.00016 | 19606 | 2019-12-09 | 11:49 | Jeremy thompson | 213869 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.67 | 0.02% | 1 | 5 | 4 | ... | 0.00646 | 0.00015 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00646 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00672 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00646 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78361 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.6 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water | Complete | Charge to GCT | test water blanks to be run by Jeremy | 68874 | Normal | 14047/11 |  |  | 2019-11-26 | 4332 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-28 | 18371 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6945 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19606 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23525 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10112 | Invoice | Completed
[7 empty] | 109.5 | 87.2901 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 43706 | 43706 | 43706 | 231 | 231 | 231 | 43478.2 | 43940.6 | Julia Collins | -995.66 | 0.12 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.24‰ | 2.24‰ | 2.00‰ | 2.00‰ | 2.24 ‰ | 2.00 ‰ | AMS | 2.92‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -995.66 ‰ | -995.66 ‰ | 0.12 | 0.12 | -995.61 | 0.13 ‰ | -995.61 | 0.13 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0043359 | 0.0043359 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00434 | 0.00012 | 19607 | 2019-12-09 | 11:49 | Jeremy thompson | 213870 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.43 | 0.01% | 2 | 5 | 4 | ... | 0.00417 | 0.00012 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00417 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00434 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00417 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78362 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.03 mg |  | 1.03 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68875 | Normal | 14047/11 |  |  | 2019-11-26 | 4333 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-28 | 18372 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6946 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19607 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23526 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10113 | Invoice | Completed
[7 empty] | 107.7 | 87.2256 | 0.0029 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 32517 | 32517 | 32517 | 115 | 114 | 114 | 32403.2 | 32632.8 | Julia Collins | -982.54 | 0.25 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.98‰ | 2.98‰ | 2.00‰ | 2.00‰ | 2.98 ‰ | 2.00 ‰ | AMS | 2.69‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -982.54 ‰ | -982.54 ‰ | 0.25 | 0.25 | -982.33 | 0.25 ‰ | -982.33 | 0.25 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.017458 | 0.017458 | 0.00025 | 0.00025 | 0.00025 | 0.00025 | 2.1 |  |  | 0.01746 | 0.00025 | 19608 | 2019-12-09 | 11:49 | Jeremy thompson | 213871 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 1.75 | 0.02% | 3 | 5 | 4 | ... | 0.01679 | 0.00024 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.01679 | 0.02‰ | 0.00024 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01746 | 0 | 0.00025 | 0 | 0 |  |  | 0 | 0 | 0.01679 | 0 | 0.00024 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78363 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 0.99 mg |  | 0.99 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.7 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68876 | Normal | 14047/11 |  |  | 2019-11-26 | 4334 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-28 | 18373 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6947 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19608 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23527 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10114 | Invoice | Completed
[7 empty] | 108.2 | 87.2638 | 0.0024 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 34186 | 34186 | 34186 | 130 | 129 | 129 | 34057.7 | 34316.9 | Julia Collins | -985.82 | 0.23 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.54‰ | 2.54‰ | 2.00‰ | 2.00‰ | 2.54 ‰ | 2.00 ‰ | AMS | 2.79‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -985.82 ‰ | -985.82 ‰ | 0.23 | 0.23 | -985.64 | 0.23 ‰ | -985.64 | 0.23 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.014183 | 0.014183 | 0.00023 | 0.00023 | 0.00023 | 0.00023 | 1.7 |  |  | 0.01418 | 0.00023 | 19609 | 2019-12-09 | 11:49 | Jeremy thompson | 213872 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 1.42 | 0.02% | 4 | 5 | 4 | ... | 0.01364 | 0.00022 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.01364 | 0.01‰ | 0.00022 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.01418 | 0 | 0.00023 | 0 | 0 |  |  | 0 | 0 | 0.01364 | 0 | 0.00022 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78364 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.06 mg |  | 1.06 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68877 | Normal | 14047/11 |  |  | 2019-11-26 | 4335 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-28 | 18374 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6948 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19609 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23528 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10115 | Invoice | Completed
[7 empty] | 109 | 87.2403 | 0.0013 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38853 | 38852 | 38852 | 168 | 168 | 168 | 38686.1 | 39023 | Julia Collins | -992.07 | 0.17 ‰ | 2019-12-05 | 2019-06-08 | 7 | 180 |  |  |  |  | 2.81‰ | 2.81‰ | 2.00‰ | 2.00‰ | 2.81 ‰ | 2.00 ‰ | AMS | 2.81‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -992.07 ‰ | -992.07 ‰ | 0.17 | 0.17 | -991.97 | 0.17 ‰ | -991.97 | 0.17 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0079336 | 0.0079336 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2 |  |  | 0.00793 | 0.00017 | 19610 | 2019-12-09 | 11:49 | Jeremy thompson | 213873 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.79 | 0.02% | 11 | 5 | 4 | ... | 0.00763 | 0.00016 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00763 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00793 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00763 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78365 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water | Complete | Charge to GCT | test water blanks to be run by Jeremy | 68878 | Normal | 14047/11 |  |  | 2019-11-26 | 4336 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-28 | 18375 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6949 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19610 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23529 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10116 | Invoice | Completed
[7 empty] | 107 | 87.1711 | 0.0012 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40079 | 40078 | 40078 | 184 | 183 | 183 | 39896.9 | 40264.9 | Julia Collins | -993.19 | 0.16 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 3.61‰ | 3.61‰ | 2.00‰ | 2.00‰ | 3.61 ‰ | 2.00 ‰ | AMS | 2.88‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -993.19 ‰ | -993.19 ‰ | 0.16 | 0.16 | -993.11 | 0.16 ‰ | -993.11 | 0.16 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0068107 | 0.0068107 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2 |  |  | 0.00681 | 0.00016 | 19611 | 2019-12-09 | 11:49 | Jeremy thompson | 213874 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.68 | 0.02% | 6 | 5 | 4 | ... | 0.00655 | 0.00015 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00655 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00681 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00655 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78366 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.02 mg |  | 1.02 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68879 | Normal | 14047/11 |  |  | 2019-11-27 | 4337 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-29 | 18376 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6950 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19611 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23530 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10117 | Invoice | Completed
[7 empty] | 108 | 87.2661 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 236 | 235 | 235 | 42393.8 | 42865.5 | Julia Collins | -995.04 | 0.15 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.51‰ | 2.51‰ | 2.00‰ | 2.00‰ | 2.51 ‰ | 2.00 ‰ | AMS | 2.83‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.15 | 0.15 | -994.98 | 0.15 ‰ | -994.98 | 0.15 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0049598 | 0.0049598 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00496 | 0.00015 | 19612 | 2019-12-09 | 11:49 | Jeremy thompson | 213875 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.5 | 0.01% | 7 | 5 | 4 | ... | 0.00477 | 0.00014 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00496 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78367 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68880 | Normal | 14047/11 |  |  | 2019-11-27 | 4338 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-29 | 18377 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6951 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19612 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23531 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10118 | Invoice | Completed
[7 empty] | 107 | 87.24 | 0.0009 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 41871 | 41871 | 41871 | 215 | 214 | 214 | 41659.5 | 42088.8 | Julia Collins | -994.55 | 0.15 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.81‰ | 2.81‰ | 2.00‰ | 2.00‰ | 2.81 ‰ | 2.00 ‰ | AMS | 2.87‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -994.55 ‰ | -994.55 ‰ | 0.15 | 0.15 | -994.48 | 0.15 ‰ | -994.48 | 0.15 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0054485 | 0.0054485 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00545 | 0.00015 | 19613 | 2019-12-09 | 11:49 | Jeremy thompson | 213876 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.54 | 0.01% | 8 | 5 | 4 | ... | 0.00524 | 0.00014 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00524 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00545 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00524 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78368 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.03 mg |  | 1.03 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68881 | Normal | 14047/11 |  |  | 2019-11-27 | 4339 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-29 | 18378 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6952 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19613 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23532 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10119 | Invoice | Completed
[7 empty] | 107.4 | 87.2268 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38645 | 38644 | 38644 | 174 | 174 | 174 | 38472.4 | 38821.2 | Julia Collins | -991.86 | 0.18 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.97‰ | 2.97‰ | 2.00‰ | 2.00‰ | 2.97 ‰ | 2.00 ‰ | AMS | 2.82‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -991.86 ‰ | -991.86 ‰ | 0.18 | 0.18 | -991.76 | 0.18 ‰ | -991.76 | 0.18 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0081416 | 0.0081416 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 1.9 |  |  | 0.00814 | 0.00018 | 19614 | 2019-12-09 | 11:49 | Jeremy thompson | 213877 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.81 | 0.02% | 9 | 5 | 4 | ... | 0.00783 | 0.00017 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00783 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00814 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00783 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78369 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.00 mg |  | 1 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.8 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68882 | Normal | 14047/11 |  |  | 2019-11-27 | 4340 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-29 | 18379 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6953 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19614 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23533 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10120 | Invoice | Completed
[7 empty] | 107.6 | 87.2266 | 0.0014 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 38759 | 38758 | 38758 | 166 | 166 | 166 | 38593.8 | 38926.8 | Julia Collins | -991.97 | 0.17 ‰ | 2019-12-05 | 2019-06-08 | 6 | 180 |  |  |  |  | 2.97‰ | 2.97‰ | 2.00‰ | 2.00‰ | 2.97 ‰ | 2.00 ‰ | AMS | 2.86‰ | 0.2‰ | C13 | No | -19 | -19.0008 | 0.82 | 0.8165 | -991.97 ‰ | -991.97 ‰ | 0.17 | 0.17 | -991.87 | 0.17 ‰ | -991.87 | 0.17 ‰ |  |  | 2019-12-09 | Jocelyn Turnbull | 11:51 | 0.0080272 | 0.0080272 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2 |  |  | 0.00803 | 0.00017 | 19615 | 2019-12-09 | 11:49 | Jeremy thompson | 213878 |  |  |  |  | 45% | 5 | 9/12/19 | 12:19:30 | Jeremy thompson | 0 |  | 0.8 | 0.02% | 12 | 5 | 4 | ... | 0.00772 | 0.00016 |  |  |  |  | 0.98309 | 0.98309 | Yes | Yes | [5 empty] | 0.00772 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00072 | 0.00072 | 0.00803 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00772 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2019-12-03 | Julia Collins | 9:23 |  | Completed |  |  |  |  | 78370 | 2000000 |  |  | 78372 78388 78389 78390 78391 78401 | 78372~78388~78389~78390~78391~78401 | 3266 |  | 3266 3266 3266 3266 3266 3266 | 2019-12-09 | 12:20 | Jeremy thompson |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0003 | 0.00181 | 78372 |  | XCAMS | ST | 0% |  | TW3266 measurement: * large samples * XCAMS: - no issues - short wheel 32 CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration secondaries: * 1x IAEA-C2 carbonate, large: result is consistent with consensus value blanks: * Two groups identified to blank-correct: large Carbonates: 12x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large Waters: 0x unknowns + 0x secondaries + 10x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.9 | RRL | [5 empty] | Background Test |  | 2019-12-24 | Water |  | Charge to GCT | test water blanks to be run by Jeremy | 68883 | Normal | 14047/11 |  |  | 2019-11-27 | 4341 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99799 | 0.00172 | 78388 | [25 empty] | 2019-11-29 | 18380 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 78389 | [25 empty] | 2019-12-10 | 6954 | C13 Measurement | Completed
[168 empty] | 0.99883 | 0.00179 | 78390 | [25 empty] | 2019-12-09 | 19615 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00181 | 78391 | [25 empty] | 2019-12-20 | 23534 | Close Job | Completed
[168 empty] | 0.99908 | 0.00174 | 78401 | [25 empty] | 2020-02-24 | 10121 | Invoice | Completed
[7 empty] | 73.6 | 88.1492 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43257 | 43256 | 43256 | 182 | 182 | 182 |  |  | jenny dahl | -995.41 | 0.10 ‰ | 2020-02-07 | 2019-08-11 | 4 | 180 |  |  |  |  | 5.43‰ | 5.43‰ | 2.00‰ | 2.00‰ | 5.43 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -995.41 ‰ | -995.41 ‰ | 0.1 | 0.1 | -995.36 | 0.11 ‰ | -995.36 | 0.11 ‰ |  |  | 2020-02-10 | Jocelyn Turnbull | 9:58 | 0.0045855 | 0.0045855 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.7 |  |  | 0.00459 | 0.0001 | 19858 | 2020-02-10 | 11:37 | albert zondervan | 214270 |  |  |  |  | 45% | 5 | 10/02/20 | 09:57:40 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 22 | 5 | 5 | ... | 0.00441 | 0.0001 |  |  |  |  | 0.99593 | 0.99593 | Yes | Yes | [5 empty] | 0.00441 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00459 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00441 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-05 | jenny dahl | 15:00 |  | Completed |  |  |  |  | 78621 | 2000000 |  |  | 78551 78552 78553 78583 78584 78586 | 78551~78552~78553~78583~78584~78586 | 3272 |  | 3272 3272 3272 3272 3272 3272 | 2020-02-10 | 11:55 | albert zondervan |  | 1.23 mg |  | 1.23 | [5 empty] | 0.99926 | 0.00153 | 78551 |  | XCAMS | EA | 0% |  | TW3272: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x FIRI-I: cellulose: result is consistent with consensus value * 1x IAEA-C2 carbonate: result is consistent with consensus value * 1x IAEA-C6 sucrose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-03-17 | Water |  |  |  | 69127 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99877 | 0.00154 | 78552 | [25 empty] | 2020-01-31 | 4450 | Water CO2 Evolution | Completed
[168 empty] | 1.0032 | 0.00157 | 78553 | [25 empty] | 2020-02-03 | 18631 | Graphite | Completed
[168 empty] | 1.0007 | 0.00153 | 78583 | [25 empty] | 2020-02-05 | 7065 | C13 Measurement | Completed
[168 empty] | 0.99885 | 0.00157 | 78584 | [25 empty] | 2020-02-10 | 19858 | AMS Submission Results | Completed
[168 empty] | 0.99931 | 0.00159 | 78586 | [25 empty] | 2020-02-12 | 23675 | Close Job | Completed
[7 empty] | 64.4 | 88.2168 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42216 | 42215 | 42215 | 224 | 224 | 224 |  |  | Julia Collins | -994.78 | 0.15 ‰ | 2020-02-26 | 2019-08-30 | 1 | 180 |  |  |  |  | 4.43‰ | 4.43‰ | 2.00‰ | 2.00‰ | 4.43 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19 | 0 | 0.8165 | -994.78 ‰ | -994.78 ‰ | 0.15 | 0.15 | -994.72 | 0.15 ‰ | -994.72 | 0.15 ‰ |  |  | 2020-02-28 | Jocelyn Turnbull | 14:14 | 0.0052198 | 0.0052198 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2 |  |  | 0.00522 | 0.00015 | 19916 | 2020-02-28 | 10:57 | albert zondervan | 214225 |  |  |  |  | 45% | 5 | 2/03/20 | 11:14:45 | Margaret Norris | 0 |  | 0.52 | 0.01% | 1 | 5 | 5 | ... | 0.00502 | 0.00014 |  |  |  |  | 0.9957 | 0.9957 | Yes | Yes | [5 empty] | 0.00502 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00081 | 0.00081 | 0.00522 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00502 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-02-25 | Julia Collins | 13:06 |  | Completed |  |  |  |  | 78739 | 2000000 |  |  | 78560 78561 78562 78590 78593 78697 | 78560~78561~78562~78590~78593~78697 | 3274 |  | 3274 3274 3274 3274 3274 3274 | 2020-02-28 | 11:01 | albert zondervan |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0002 | 0.00197 | 78560 |  | XCAMS | EA | 0% |  | TW3274: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: result is consistent with consensus value blanks: * One groups identified to blank-correct: waters: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-03-05 | Water | Complete |  |  | 69145 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99809 | 0.00197 | 78561 | [25 empty] | 2020-02-21 | 4496 | Water CO2 Evolution | Completed
[168 empty] | 0.99983 | 0.00199 | 78562 | [25 empty] | 2020-02-25 | 18749 | Graphite | Completed
[168 empty] | 1.0004 | 0.00196 | 78590 | [25 empty] | 2020-02-26 | 7119 | C13 Measurement | Completed
[168 empty] | 1.0024 | 0.00198 | 78593 | [25 empty] | 2020-02-28 | 19916 | AMS Submission Results | Completed
[168 empty] | 0.99902 | 0.00202 | 78697 | [25 empty] | 2020-03-03 | 23779 | Close Job | Completed
[7 empty] | 98.3 | 88.3622 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40677 | 40676 | 40676 | 185 | 184 | 184 |  |  | Julia Collins | -993.68 | 0.15 ‰ | 2020-03-03 | 2019-09-05 | 5 | 180 |  |  |  |  | 4.38‰ | 4.38‰ | 2.00‰ | 2.00‰ | 4.38 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0003 | 0 | 0.8165 | -993.68 ‰ | -993.68 ‰ | 0.15 | 0.15 | -993.6 | 0.15 ‰ | -993.6 | 0.15 ‰ |  |  | 2020-03-09 | Jocelyn Turnbull | 9:16 | 0.006322 | 0.006322 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.00632 | 0.00015 | 19993 | 2020-03-06 | 10:24 | albert zondervan | 214396 |  |  |  |  | 45% | 5 | 9/03/20 | 09:16:06 | Jocelyn Turnbull | 0 |  | 0.63 | 0.01% | 20 | 5 | 4 | ... | 0.00608 | 0.00014 |  |  |  |  | 0.99729 | 0.99729 | Yes | Yes | [5 empty] | 0.00608 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00088 | 0.00088 | 0.00632 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00608 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-02 | Julia Collins | 10:48 |  | Completed |  |  |  |  | 78782 | 2000000 |  |  | 78705 78706 78707 78724 78725 78726 | 78705~78706~78707~78724~78725~78726 | 3275 |  | 3275 3275 3275 3275 3275 3275 | 2020-03-06 | 10:35 | albert zondervan |  | 1.16 mg |  | 1.16 | [5 empty] | 0.99865 | 0.00214 | 78705 |  | XCAMS | EA | 0.14% |  | TW3275: * large samples * no issues with XCAMS CalAMS error adjustments: * 0.14% system error primaries: * OxI cathodes: 6x large no failures secondaries: * 1x IAEA-C2 carbonate: resulting CRA is higher than consensus value by 2.5 sigma * 1x FIRI-I cellulose: result is consistent with consensus value blanks: * Two groups identified to blank-correct: waters: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks organics: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-03-30 | Water | Complete |  |  | 69224 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9993 | 0.00214 | 78706 | [25 empty] | 2020-02-26 | 4510 | Water CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00215 | 78707 | [25 empty] | 2020-02-27 | 18792 | Graphite | Completed
[168 empty] | 0.99743 | 0.00215 | 78724 | [25 empty] | 2020-03-09 | 19993 | AMS Submission Results | Completed
[168 empty] | 1 | 0.00219 | 78725 | [25 empty] | 2020-03-10 | 23890 | Close Job | Completed
[168 empty] | 1.0037 | 0.00217 | 78726
[7 empty] | 93.7 | 87.4727 | 0.0015 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37771 | 37770 | 37770 | 147 | 147 | 147 |  |  | Julia Collins | -990.92 | 0.17 ‰ | 2020-03-12 | 2019-09-14 | 1 | 180 |  |  |  |  | 4.87‰ | 4.87‰ | 2.00‰ | 2.00‰ | 4.87 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19.0042 | 0 | 0.8165 | -990.92 ‰ | -990.92 ‰ | 0.17 | 0.17 | -990.81 | 0.17 ‰ | -990.81 | 0.17 ‰ |  |  | 2020-03-13 | Jocelyn Turnbull | 8:56 | 0.0090774 | 0.0090774 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 1.7 |  |  | 0.00908 | 0.00017 | 20076 | 2020-03-13 | 15:58 | albert zondervan | 214586 |  |  |  |  | 45% | 5 | 16/03/20 | 09:52:51 | Margaret Norris | 0 |  | 0.91 | 0.02% | 1 | 5 | 4 | ... | 0.00873 | 0.00016 |  |  |  |  | 0.98774 | 0.98774 | Yes | Yes | [5 empty] | 0.00873 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00908 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.00873 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-11 | Julia Collins | 14:51 |  | Completed |  |  |  |  | 78893 | 2000000 |  |  | 78833 78834 78835 78836 78854 78872 | 78833~78834~78835~78836~78854~78872 | 3278 |  | 3278 3278 3278 3278 3278 3278 | 2020-03-13 | 17:35 | albert zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 1.0005 | 0.00174 | 78833 |  | XCAMS | ST | 0% |  | TW3278 measurement: * small & large samples * XCAMS: no issues CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 1x small: no failures, used to determine 'small d' secondaries: * 1x IAEA-C2 water, large: result is 2-sigma younger than consensus * 1x FIRI-I, small: result is consistent with consensus value * 1x FIRI-I, small: result is consistent with consensus value blanks: * three groups identified to blank-correct: large waters: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks large organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks, all Kauri Renton Rd small organics via AAA: 5x unknowns + 1x secondaries + 3x blanks Excel file 'to blank-correct small samples.xlsm': 1x other wheels selected also, to assist with determining d, m parameters: TW3253 d = 1.8 ug C m = 0.77 ug C, derived from 4x Queets A blanks + 3x Kauri Renton Road blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-05-05 | Water |  |  |  | 69344 | Normal | 14047/11 |  |  | 2020-03-10 | 4535 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99952 | 0.00178 | 78834 | [25 empty] | 2020-03-11 | 18903 | Graphite | Completed
[168 empty] | 0.99963 | 0.00176 | 78835 | [25 empty] | 2020-03-12 | 7162 | C13 Measurement | Completed
[168 empty] | 1.0008 | 0.00175 | 78836 | [25 empty] | 2020-03-13 | 20076 | AMS Submission Results | Completed
[168 empty] | 0.99825 | 0.00184 | 78854 | [25 empty] | 2020-03-16 | 23963 | Close Job | Completed
[168 empty] | 1.0013 | 0.00183 | 78872
[7 empty] | 91.7 | 87.7883 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46088 | 46087 | 46087 | 207 | 207 | 207 |  |  | Margaret Norris | -996.78 | 0.08 ‰ | 2020-05-22 | 2019-11-24 | 64 | 180 |  |  |  |  | 2.94‰ | 2.94‰ | 2.00‰ | 2.00‰ | 2.94 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | Submitted | No | -19 | -19.0012 | 0 | 0.89443 | -996.78 ‰ | -996.78 ‰ | 0.08 | 0.08 | -996.74 | 0.08 ‰ | -996.74 | 0.08 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 12:02 | 0.0032234 | 0.0032234 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.7 |  |  | 0.00322 | 8e-05 | 20156 | 2020-05-26 | 15:28 | albert zondervan | 214580 |  |  |  |  | 45% | 5 | 31/05/20 | 12:01:38 | Jocelyn Turnbull | 0 |  | 0.32 | 0.01% | 1 | 5 | 7 | ... | 0.0031 | 8e-05 |  |  |  |  | 0.98939 | 0.98939 | Yes | Yes | [5 empty] | 0.0031 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00322 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.0031 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-03-24 | Margaret Norris | 13:07 |  | Completed |  |  |  |  | 78922 | 2000000 |  |  | 78934 78935 78943 78946 78949 | 78934~78935~78943~78946~78949 | 3280 |  | 3280 3280 3280 3280 3280 | 2020-05-26 | 18:11 | albert zondervan |  | 0.66 mg |  | 0.66 | [5 empty] | 0.99954 | 0.00144 | 78934 |  | XCAMS | EA | 0% |  | TW3280 measurement: * wheel stored under vacuum of source since before lockdown * first measurement after lockdown and subsequent Pelletron service * large samples * 3x ramped pyrolysis samples: results (to be) worked up separately by Cathy * quality of results was not affected by noteworthy deviations from regular behaviour: small, temporary drift in terminal voltage, which was resolved through more conditioning some variability amongst the OxIs in terms of 13C/12C ratio CalAMS error adjustments: * zero system error after rejection of one outlier OxI * global excess variability: factor 1.03 primaries: * OxI cathodes: 6x large: one outlier rejected (TP78954, 0.5 mgC) secondaries: * 1x FIRI-I: result is consistent with consensus value * 1x IAEA-C2: result is consistent with consensus value blanks: * two groups identified to blank-correct: organics, via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks waters: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures |  | RRL | [5 empty] | Background Inorganic |  | 2020-05-04 | Water |  |  | no CO2 in the bottle, no C13 result reported | 69425 | Normal | 14047/11 |  |  | 2020-01-10 | 6248 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99757 | 0.00149 | 78935 |  | XCAMS | [23 empty] | 2020-03-18 | 4540 | Water CO2 Evolution | Completed
[168 empty] | 1.0002 | 0.00148 | 78943 | [25 empty] | 2020-03-18 | 2756 | Gas Splitting | Completed
[168 empty] | 1.0016 | 0.00143 | 78946 | [25 empty] | 2020-03-18 | 4717 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00143 | 78949 | [25 empty] | 2020-03-19 | 18932 | Graphite | Completed
[196 empty] | 2020-03-23 | 7165 | C13 Measurement | Cancelled
[196 empty] | 2020-05-31 | 20156 | AMS Submission Results | Completed
[196 empty] | 2020-06-02 | 24117 | Close Job | Completed
[7 empty] | 89.2 | 87.6999 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37616 | 37615 | 37615 | 117 | 117 | 117 |  |  | Margaret Norris | -990.75 | 0.14 ‰ | 2020-05-27 | 2019-11-29 | 8 | 180 |  |  |  |  | 4.01‰ | 4.01‰ | 2.00‰ | 2.00‰ | 4.01 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -990.75 ‰ | -990.75 ‰ | 0.14 | 0.14 | -990.63 | 0.14 ‰ | -990.63 | 0.14 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:20 | 0.0092542 | 0.0092542 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2 |  |  | 0.00925 | 0.00014 | 20260 | 2020-05-29 | 9:36 | albert zondervan | 214869 |  |  |  |  | 45% | 5 | 31/05/20 | 13:20:27 | Jocelyn Turnbull | 0 |  | 0.93 | 0.01% | 36 | 5 | 4 | ... | 0.0089 | 0.00013 |  |  |  |  | 0.98946 | 0.98946 | Yes | Yes | [5 empty] | 0.0089 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00925 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0089 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-22 | Margaret Norris | 14:31 |  | Completed |  |  |  |  | 79054 | 2000000 |  |  | 79069 79071 79072 79091 79092 79098 | 79069~79071~79072~79091~79092~79098 | 3282 |  | 3282 3282 3282 3282 3282 3282 | 2020-05-29 | 9:46 | albert zondervan |  | 0.96 mg |  | 0.96 | [5 empty] | 1.0028 | 0.00143 | 79069 |  | XCAMS | ST | 0% |  | TW3282 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on carbonate-line: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * two groups identified to blank-correct: carbonate: 17x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks water: 11x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-05-14 | Water | Complete |  |  | 69529 | Normal | 14047/11 |  |  | 2020-05-18 | 4620 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99972 | 0.00142 | 79071 | [25 empty] | 2020-05-19 | 19064 | Graphite | Completed
[168 empty] | 0.99841 | 0.00143 | 79072 | [25 empty] | 2020-05-21 | 7245 | C13 Measurement | Completed
[168 empty] | 0.99967 | 0.00144 | 79091 | [25 empty] | 2020-05-31 | 20260 | AMS Submission Results | Completed
[168 empty] | 0.99857 | 0.00142 | 79092 | [25 empty] | 2020-06-02 | 24122 | Close Job | Completed
[168 empty] | 1.0008 | 0.00143 | 79098
[7 empty] | 89.9 | 87.7857 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40664 | 40663 | 40663 | 198 | 197 | 197 |  |  | Julia Collins | -993.67 | 0.16 ‰ | 2020-05-29 | 2019-12-01 | 7 | 180 |  |  |  |  | 3.74‰ | 3.74‰ | 2.00‰ | 2.00‰ | 3.74 ‰ | 2.00 ‰ | AMS | 2.27‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -993.67 ‰ | -993.67 ‰ | 0.16 | 0.16 | -993.59 | 0.16 ‰ | -993.59 | 0.16 ‰ |  |  | 2020-05-31 | Jocelyn Turnbull | 13:46 | 0.0063323 | 0.0063323 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 1.7 |  |  | 0.00633 | 0.00016 | 20279 | 2020-05-31 | 12:25 | albert zondervan | 213968 |  |  |  |  | 45% | 5 | 31/05/20 | 13:46:32 | Jocelyn Turnbull | 0 |  | 0.63 | 0.02% | 1 | 5 | 4 | ... | 0.00609 | 0.00015 |  |  |  |  | 0.99015 | 0.99015 | Yes | Yes | [5 empty] | 0.00609 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00067 | 0.00067 | 0.00633 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00609 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-05-28 | Julia Collins | 11:15 |  | Completed |  |  |  |  | 79099 | 2000000 |  |  | 79083 79085 79089 79139 79140 79141 | 79083~79085~79089~79139~79140~79141 | 3284 |  | 3284 3284 3284 3284 3284 3284 | 2020-05-31 | 12:57 | albert zondervan |  | 1.12 mg |  | 1.12 | [5 empty] | 0.9998 | 0.00162 | 79083 |  | XCAMS | EA | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks wood via CE, Rafter method: 3x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-01-16 | Water | Complete |  |  | 69548 | Normal | 14047/11 |  |  | 2020-05-20 | 4633 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99872 | 0.00162 | 79085 | [25 empty] | 2020-05-22 | 19109 | Graphite | Completed
[168 empty] | 1.0019 | 0.00161 | 79089 | [25 empty] | 2020-05-25 | 7260 | C13 Measurement | Completed
[168 empty] | 1.0005 | 0.00173 | 79139 | [25 empty] | 2020-05-31 | 20279 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00164 | 79140 | [25 empty] | 2020-06-02 | 24098 | Close Job | Completed
[168 empty] | 0.99777 | 0.00165 | 79141
[7 empty] | 97.3 | 87.7964 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 43184 | 43184 | 43184 | 199 | 198 | 198 |  |  | Julia Collins | -995.37 | 0.11 ‰ | 2020-06-05 | 2019-12-08 | 14 | 180 |  |  |  |  | 1.20‰ | 1.20‰ | 2.00‰ | 2.00‰ | 1.20 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -995.37 ‰ | -995.37 ‰ | 0.11 | 0.11 | -995.32 | 0.12 ‰ | -995.32 | 0.12 ‰ |  |  | 2020-06-06 | Jocelyn Turnbull | 9:29 | 0.0046271 | 0.0046271 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3.2 |  |  | 0.00463 | 0.00011 | 20380 | 2020-06-06 | 18:32 | albert zondervan | 214983 |  |  |  |  | 45% | 5 | 6/06/20 | 09:29:16 | Jocelyn Turnbull | 0 |  | 0.46 | 0.01% | 1 | 5 | 6 | ... | 0.00445 | 0.00011 |  |  |  |  | 0.98777 | 0.98777 | Yes | Yes | [5 empty] | 0.00445 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 0.99999 | 0.00069 | 0.00069 | 0.00463 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00445 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-04 | Julia Collins | 14:02 |  | Completed |  |  |  |  | 79103 | 2000000 |  |  | 79176 79177 79178 79179 79195 79196 | 79176~79177~79178~79179~79195~79196 | 3286 |  | 3286 3286 3286 3286 3286 3286 | 2020-06-06 | 18:43 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99861 | 0.00167 | 79176 |  | XCAMS | ST | 0% |  | TW3284 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 19x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-06-24 | Water |  |  |  | 69609 | Normal | 14047/11 |  |  | 2020-05-21 | 4640 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9978 | 0.00169 | 79177 | [25 empty] | 2020-05-21 | 2815 | Gas Splitting | Completed
[168 empty] | 0.99889 | 0.00168 | 79178 | [25 empty] | 2020-05-21 | 4787 | Choose Gas Splitting | Completed
[168 empty] | 1.0023 | 0.00174 | 79179 | [25 empty] | 2020-05-22 | 19113 | Graphite | Completed
[168 empty] | 1.0018 | 0.00169 | 79195 | [25 empty] | 2020-05-25 | 7268 | C13 Measurement | Completed
[168 empty] | 1.0007 | 0.00168 | 79196 | [25 empty] | 2020-06-06 | 20380 | AMS Submission Results | Completed
[196 empty] | 2020-06-08 | 24275 | Close Job | Completed
[7 empty] | 94.4 | 87.5996 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38991 | 38990 | 38990 | 150 | 149 | 149 |  |  | Margaret Norris | -992.2 | 0.15 ‰ | 2020-06-12 | 2019-12-15 | 9 | 180 |  |  |  |  | 3.19‰ | 3.19‰ | 2.00‰ | 2.00‰ | 3.19 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.20 ‰ | -992.20 ‰ | 0.15 | 0.15 | -992.11 | 0.15 ‰ | -992.11 | 0.15 ‰ |  |  | 2020-06-14 | Jocelyn Turnbull | 16:55 | 0.0077985 | 0.0077985 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.7 |  |  | 0.0078 | 0.00015 | 20443 | 2020-06-15 | 8:47 | albert zondervan | 215220 |  |  |  |  | 45% | 5 | 14/06/20 | 16:54:53 | Jocelyn Turnbull | 0 |  | 0.78 | 0.01% | 8 | 5 | 4 | ... | 0.0075 | 0.00014 |  |  |  |  | 0.98752 | 0.98752 | Yes | Yes | [5 empty] | 0.0075 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.0078 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.0075 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-11 | Margaret Norris | 13:41 |  | Completed |  |  |  |  | 79191 | 2000000 |  |  | 79258 79262 79265 79277 79279 79282 | 79258~79262~79265~79277~79279~79282 | 3288 |  | 3288 3288 3288 3288 3288 3288 | 2020-06-15 | 9:07 | albert zondervan |  | 1.11 mg |  | 1.11 | [5 empty] | 1.0011 | 0.00156 | 79258 |  | XCAMS | EA | 0% |  | TW3288 measurement: * large samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * 1x FIRI-I cellulose: result agrees with consensus value * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks wood via CE, Rafter method: 6x unknowns + 0x secondaries + 1x blanks 240 days: capturing 4x blanks organics via AAA: 6x unknowns + 1x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-07-22 | Water |  |  |  | 69672 | Normal | 14047/11 |  |  | 2020-06-02 | 4660 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9983 | 0.00156 | 79262 | [25 empty] | 2020-06-03 | 19201 | Graphite | Completed
[168 empty] | 1.0008 | 0.00157 | 79265 | [25 empty] | 2020-06-04 | 7301 | C13 Measurement | Completed
[168 empty] | 0.9983 | 0.00157 | 79277 | [25 empty] | 2020-06-14 | 20443 | AMS Submission Results | Completed
[168 empty] | 1.0015 | 0.00161 | 79279 | [25 empty] | 2020-06-15 | 24348 | Close Job | Completed
[168 empty] | 1 | 0.00155 | 79282
[7 empty] | 94.5 | 87.779 | 0.0005 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 46011 | 46010 | 46010 | 231 | 230 | 230 |  |  | Julia Collins | -996.75 | 0.09 ‰ | 2020-06-16 | 2019-12-19 | 6 | 180 |  |  |  |  | 3.88‰ | 3.88‰ | 2.00‰ | 2.00‰ | 3.88 ‰ | 2.00 ‰ | AMS | 2.24‰ | 0.2‰ | C13 | No | -19 | -19 | 0.82 | 0.8165 | -996.75 ‰ | -996.75 ‰ | 0.09 | 0.09 | -996.71 | 0.09 ‰ | -996.71 | 0.09 ‰ |  |  | 2020-06-18 | Jocelyn Turnbull | 17:58 | 0.0032546 | 0.0032546 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 3.4 |  |  | 0.00325 | 9e-05 | 20503 | 2020-06-18 | 14:06 | albert zondervan | 215221 |  |  |  |  | 45% | 5 | 22/06/20 | 10:34:38 | albert zondervan | 0 |  | 0.33 | 0.01% | 8 | 5 | 6 | ... | 0.00313 | 9e-05 |  |  |  |  | 0.99022 | 0.99022 | Yes | Yes | [5 empty] | 0.00313 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00325 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00313 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-15 | Julia Collins | 13:35 |  | Completed |  |  |  |  | 79274 | 2000000 |  |  | 79034 79035 79048 79063 79064 79165 | 79034~79035~79048~79063~79064~79165 | 3289 |  | 3289 3289 3289 3289 3289 3289 | 2020-06-22 | 10:30 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99805 | 0.00158 | 79034 |  | XCAMS | FLASK | 0% |  | TW3289 measurement: * large+small samples * sediments, waters CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 1x EA + 5x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * three groups identified to blank-correct: water: 16x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks organic via RP: 12x unknowns + 0x secondaries + 3x blanks + 3x OxI with blank-correction calculations outside RLIMS organic via AAA: 2x unknowns (replicates of RP unknowns) + 0x secondaries + 0x blanks 180 days: capturing 7x blanks unknowns: * no failures * Because a few cathodes gave noticably higher del13C_AMS than del13C_IRMS, these parameters were compared for all samples of which the del13C_IRMS was known. Their difference values correlate with sample mass, with ~6 permil/mgC proportionality in the 0.1-1.1 mgC range. The primary standard OxI graphites in this wheel are confined to a much narrower range: 0.40-0.63 mgC. The IAEA-C2 secondary standard (1.02 mgC) generated a result consistent with its consensus value. While this dependency needs more study (to attribute this to fractionation during graphitization or ion beam transport, it is expected that it is mostly (if not completely) accounted for through our use of 'full fractionation correction' mode 5 in CalAMS. | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-07-22 | Water | Complete |  |  | 69731 | Normal | 14047/11 |  |  | 2020-06-09 | 4693 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00168 | 79035 | [25 empty] | 2020-06-09 | 2849 | Gas Splitting | Completed
[168 empty] | 1.0021 | 0.00156 | 79048 | [25 empty] | 2020-06-09 | 4835 | Choose Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00158 | 79063 | [25 empty] | 2020-06-10 | 19284 | Graphite | Completed
[168 empty] | 1.0004 | 0.00157 | 79064 | [25 empty] | 2020-06-15 | 7332 | C13 Measurement | Completed
[168 empty] | 1.0002 | 0.00158 | 79165 | [25 empty] | 2020-06-18 | 20503 | AMS Submission Results | Completed
[196 empty] | 2020-06-22 | 24412 | Close Job | Completed
[7 empty] | 99.9 | 87.4138 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45113 | 45113 | 45113 | 207 | 206 | 206 |  |  | Julia Collins | -996.36 | 0.09 ‰ | 2020-06-19 | 2019-12-22 | 3 | 180 |  |  |  |  | 5.26‰ | 5.26‰ | 2.00‰ | 2.00‰ | 5.26 ‰ | 2.00 ‰ | AMS | 2.26‰ | 0.2‰ | C13 | No | -19 | -19.0031 | 0 | 0.8165 | -996.36 ‰ | -996.36 ‰ | 0.09 | 0.09 | -996.32 | 0.09 ‰ | -996.32 | 0.09 ‰ |  |  | 2020-06-22 | Jocelyn Turnbull | 12:46 | 0.0036393 | 0.0036393 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.8 |  |  | 0.00364 | 9e-05 | 20520 | 2020-06-20 | 18:56 | albert zondervan | 215380 |  |  |  |  | 45% | 5 | 22/06/20 | 12:46:17 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 1 | 5 | 5 | ... | 0.0035 | 9e-05 |  |  |  |  | 0.98745 | 0.98745 | Yes | Yes | [5 empty] | 0.0035 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00364 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0035 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-06-17 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 79317 | 2000000 |  |  | 79309 79310 79311 79312 79313 79319 | 79309~79310~79311~79312~79313~79319 | 3290 |  | 3290 3290 3290 3290 3290 3290 | 2020-06-20 | 19:29 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 1.0032 | 0.00165 | 79309 |  | XCAMS | ST | 0% |  | TW3290 measurement: * large samples * partial wheel CalAMS error adjustments: * none needed primaries: * OxI cathodes: no failures 6x FLASK secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * two groups identified to blank-correct: water: 15x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 10x unknowns + 1x secondaries + 1x blanks 180 days: capturing 8x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-08-04 | Water |  |  | new job for fresh water std prep | 69748 | Normal | 14047/11 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9998 | 0.00159 | 79310 | [25 empty] | 2020-06-15 | 4711 | Water CO2 Evolution | Completed
[168 empty] | 0.99809 | 0.00158 | 79311 | [25 empty] | 2020-06-16 | 19327 | Graphite | Completed
[168 empty] | 1.0007 | 0.00158 | 79312 | [25 empty] | 2020-06-22 | 7351 | C13 Measurement | Completed
[168 empty] | 0.9989 | 0.00161 | 79313 | [25 empty] | 2020-06-22 | 20520 | AMS Submission Results | Completed
[168 empty] | 0.99954 | 0.00159 | 79319 | [25 empty] | 2020-06-23 | 24422 | Close Job | Completed
[7 empty] | 127.2 | 87.3398 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45537 | 45537 | 45537 | 217 | 217 | 217 |  |  | Julia Collins | -996.55 | 0.09 ‰ | 2020-08-14 | 2020-02-16 | 14 | 180 |  |  |  |  | 2.72‰ | 2.72‰ | 2.00‰ | 2.00‰ | 2.72 ‰ | 2.00 ‰ | AMS | 2.43‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -996.55 ‰ | -996.55 ‰ | 0.09 | 0.09 | -996.51 | 0.09 ‰ | -996.51 | 0.09 ‰ |  |  | 2020-08-18 | Jocelyn Turnbull | 15:09 | 0.0034521 | 0.0034521 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00345 | 9e-05 | 20898 | 2020-08-16 | 13:00 | albert zondervan | 215602 |  |  |  |  | 45% | 5 | 18/08/20 | 15:08:37 | Jocelyn Turnbull | 0 |  | 0.35 | 0.01% | 15 | 5 | 5 | ... | 0.00332 | 9e-05 |  |  |  |  | 0.98413 | 0.98413 | Yes | Yes | [5 empty] | 0.00332 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 0.99999 | 0.00093 | 0.00093 | 0.00345 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00332 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-13 | Julia Collins | 13:49 |  | Completed |  |  |  |  | 79667 | 2000000 |  |  | 79630 79631 79632 79687 79688 79689 | 79630~79631~79632~79687~79688~79689 | 3299 |  | 3299 3299 3299 3299 3299 3299 | 2020-08-18 | 13:23 | albert zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 1.0004 | 0.00226 | 79630 |  | XCAMS | ST | 0.17% |  | TW3299 measurement: * organics, waters CalAMS error adjustments: * global factor 1.04 to account for excess run-to-run variability * system error 0.17% to account for excess OxI-to-OxI variability primaries: * OxI cathodes: no failures 6x ST * No clear cause was identified for the excess OxI-to-OxI variability, and no specific could be identified as the offending outlier. secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value * 1x FIRI-I: result agrees with consensus value blanks: * three groups identified to blank-correct: water: sample is combination of two extractions 1x unknowns + 0x secondaries + 2x blanks 180 days: capturing 2x blanks water: as per normal, i.e. through single extraction 20x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks organic via AAA: 5x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures * The AMS-IRMS difference in del13C is strongly associated with sample size, as has been seen in past measurements. | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-09-03 | Water |  |  |  | 70145 | Normal | 14047/11 |  |  | 2020-06-10 | 6603 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.004 | 0.00228 | 79631 | [25 empty] | 2020-07-30 | 4839 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00228 | 79632 | [25 empty] | 2020-07-31 | 19677 | Graphite | Completed
[168 empty] | 0.99929 | 0.00227 | 79687 | [25 empty] | 2020-08-03 | 7493 | C13 Measurement | Completed
[168 empty] | 0.99742 | 0.00226 | 79688 | [25 empty] | 2020-08-18 | 20898 | AMS Submission Results | Completed
[168 empty] | 0.99849 | 0.00228 | 79689 | [25 empty] | 2020-08-25 | 24869 | Close Job | Completed
[7 empty] | 118.8 | 87.643 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42779 | 42779 | 42779 | 222 | 222 | 222 |  |  | albert zondervan | -995.13 | 0.13 ‰ | 2020-08-25 | 2020-02-27 | 6 | 180 |  |  |  |  | 2.51‰ | 2.51‰ | 2.00‰ | 2.00‰ | 2.51 ‰ | 2.00 ‰ | AMS | 2.34‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -995.13 ‰ | -995.13 ‰ | 0.13 | 0.13 | -995.07 | 0.14 ‰ | -995.07 | 0.14 ‰ |  |  | 2020-08-27 | Jocelyn Turnbull | 11:19 | 0.0048662 | 0.0048662 | 0.00013 | 0.00013 | 0.00013 | 0.00013 | 2.4 |  |  | 0.00487 | 0.00013 | 20916 | 2020-08-26 | 14:42 | albert zondervan | 215278 |  |  |  |  | 45% | 5 | 27/08/20 | 11:19:06 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 37 | 5 | 6 | ... | 0.00468 | 0.00013 |  |  |  |  | 0.98733 | 0.98733 | Yes | Yes | [5 empty] | 0.00468 | 0.00‰ | 0.00013 | 0.00‰ | No | 1 | 0.99999 | 0.00125 | 0.00125 | 0.00487 | 0 | 0.00013 | 0 | 0 |  |  | 0 | 0 | 0.00468 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-08-25 | albert zondervan | 10:04 |  | Completed |  |  |  |  | 79745 | 2000000 |  |  | 79752 79753 79754 79755 79756 79757 | 79752~79753~79754~79755~79756~79757 | 3300 |  | 3300 3300 3300 3300 3300 3300 | 2020-08-26 | 14:48 | albert zondervan |  | 0.91 mg |  | 0.91 | [5 empty] | 1.0036 | 0.00307 | 79752 |  | XCAMS | EA | 0.25% |  | TW3300 measurement: * large water samples CalAMS error adjustments: * 0.25% primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 11x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2020-07-28 | Water | Complete |  |  | 70185 | Normal | 14047/11 |  |  | 2020-08-18 | 4866 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0012 | 0.00307 | 79753 | [25 empty] | 2020-08-18 | 2881 | Gas Splitting | Completed
[168 empty] | 0.99451 | 0.00304 | 79754 | [25 empty] | 2020-08-18 | 4884 | Choose Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00306 | 79755 | [25 empty] | 2020-08-19 | 19755 | Graphite | Completed
[168 empty] | 0.99897 | 0.00304 | 79756 | [25 empty] | 2020-08-20 | 7518 | C13 Measurement | Completed
[168 empty] | 1.001 | 0.00304 | 79757 | [25 empty] | 2020-08-27 | 20916 | AMS Submission Results | Completed
[196 empty] | 2020-08-27 | 24909 | Close Job | Completed
[7 empty] | 124.7 | 87.608 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39319 | 39318 | 39318 | 134 | 133 | 133 |  |  | jenny dahl | -992.51 | 0.12 ‰ | 2020-09-16 | 2020-03-20 | 6 | 180 |  |  |  |  | 1.03‰ | 1.03‰ | 2.00‰ | 2.00‰ | 1.03 ‰ | 2.00 ‰ | AMS | 2.23‰ | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -992.51 ‰ | -992.51 ‰ | 0.12 | 0.12 | -992.42 | 0.13 ‰ | -992.42 | 0.13 ‰ |  |  | 2020-09-17 | Jocelyn Turnbull | 11:32 | 0.0074865 | 0.0074865 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00749 | 0.00012 | 21134 | 2020-09-17 | 11:02 | albert zondervan | 216131 |  |  |  |  | 45% | 5 | 17/09/20 | 11:40:26 | jenny dahl | 0 |  | 0.75 | 0.01% | 1 | 5 | 4 | ... | 0.0072 | 0.00012 |  |  |  |  | 0.98548 | 0.98548 | Yes | Yes | [5 empty] | 0.0072 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00749 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0072 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-09-10 | jenny dahl | 12:45 |  | Completed |  |  |  |  | 79961 | 2000000 |  |  | 79924 79925 79926 79927 79928 79929 | 79924~79925~79926~79927~79928~79929 | 3305 |  | 3305 3305 3305 3305 3305 3305 | 2020-09-17 | 11:12 | albert zondervan |  | 1.05 mg |  | 1.05 | [5 empty] | 0.99796 | 0.00149 | 79924 |  | XCAMS | ST | 0% |  | TW3305 measurement: * large samples * partially filled wheel CalAMS error adjustments: * TP79942: error multiplier = 1.3 to account for excess run-to-run variability primaries: * OxI cathodes: 6x large: no failures secondaries: * 1x large FIRI-I cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value blanks: * two groups identified to blank-correct: large organics via AAA: 18x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks large water: 4x unknowns + 1x secondaries + 1x blanks 180 days: capturing 10x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2020-10-14 | Water |  |  |  | 70403 | Normal | 14047/11 |  |  | 2020-09-08 | 4943 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0027 | 0.00156 | 79925 | [25 empty] | 2020-09-10 | 19971 | Graphite | Completed
[168 empty] | 0.9993 | 0.00152 | 79926 | [25 empty] | 2020-09-11 | 7618 | C13 Measurement | Completed
[168 empty] | 0.99936 | 0.00151 | 79927 | [25 empty] | 2020-09-17 | 21134 | AMS Submission Results | Completed
[168 empty] | 1.0013 | 0.00151 | 79928 | [25 empty] | 2020-09-17 | 25106 | Close Job | Completed
[168 empty] | 0.9996 | 0.00157 | 79929
[7 empty] | 98.1 | 87.3683 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41338 | 41337 | 41337 | 186 | 186 | 186 |  |  | Margaret Norris | -994.18 | 0.14 ‰ | 2020-10-13 | 2020-04-16 | 12 | 180 |  |  |  |  | 6.75‰ | 6.75‰ | 2.00‰ | 2.00‰ | 6.75 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.75593 | -994.18 ‰ | -994.18 ‰ | 0.14 | 0.14 | -994.11 | 0.14 ‰ | -994.11 | 0.14 ‰ |  |  | 2020-10-15 | Jocelyn Turnbull | 14:18 | 0.0058229 | 0.0058229 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.00582 | 0.00014 | 21335 | 2020-10-15 | 12:50 | albert zondervan | 216380 |  |  |  |  | 45% | 5 | 15/10/20 | 14:17:37 | Jocelyn Turnbull | 0 |  | 0.58 | 0.01% | 1 | 5 | 6 | ... | 0.0056 | 0.00013 |  |  |  |  | 0.9884 | 0.9884 | Yes | Yes | [5 empty] | 0.0056 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00582 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.0056 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-10-12 | Margaret Norris | 15:40 |  | Completed |  |  |  |  | 80120 | 2000000 |  |  | 80140 80141 80142 80143 80161 80199 80205 | 80140~80141~80142~80143~80161~80199~80205 | 3310 |  | 3310 3310 3310 3310 3310 3310 3310 | 2020-10-15 | 13:07 | albert zondervan |  | 0.95 mg |  | 0.95 | [5 empty] | 1.0011 | 0.0018 | 80140 |  | XCAMS | FLASK | 0% |  | TW3310 measurement: * no issues * atm-14CH4 & water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x from flask + 1x from EA secondaries: * IAEA-C2: 1x result agrees with consensus value blanks: * Two groups identified to blank-correct: 14x air samples (atm-CH4 from NIWA): 14x unknowns, 0x secondaries, 0x blanks window = 360 days, capturing 7x blanks Extractions were done at NIWA - Hinrich Schaefer has determined that there is no quantifiable additional blank correction required related to the extraction procedure. Therefore we use our CO2 in air blank to estimate the graphite/measurement blank 14x water samples: 13x unknowns, 1x secondaries, 1x blanks window = 180 days, capturing 13x blanks unknowns: * no failures, however: * Some cathodes filled with graphite prepared from water samples generated a ca. 15% lower than average beam current, generating a consistent pattern throughout the measurement. Post-measurement the cathodes in question were tested for electrical and thermal conductivity, showing no clear pattern match with lower-current cathodes. The cause remains unclear at this stage, but the ratio measurements do not appear to have been affected. | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2020-11-20 | Water | Complete |  |  | 70579 | Normal | 14047/11 |  |  | 2020-09-30 | 4974 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00176 | 80141 | [25 empty] | 2020-09-30 | 2909 | Gas Splitting | Completed
[168 empty] | 0.99868 | 0.00174 | 80142 | [25 empty] | 2020-09-30 | 4924 | Choose Gas Splitting | Completed
[168 empty] | 1.0004 | 0.00172 | 80143 | [25 empty] | 2020-10-01 | 20130 | Graphite | Completed
[168 empty] | 1.0003 | 0.00169 | 80161 | [25 empty] | 2020-10-05 | 7646 | C13 Measurement | Completed
[168 empty] | 0.99829 | 0.00167 | 80199 | [25 empty] | 2020-10-15 | 21335 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00174 | 80205 | [25 empty] | 2020-10-19 | 25338 | Close Job | Completed
[7 empty] | 108.7 | 87.8247 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42865 | 42865 | 42865 | 191 | 190 | 190 |  |  | Julia Collins | -995.19 | 0.11 ‰ | 2020-12-10 | 2020-06-13 | 2 | 180 |  |  |  |  | 0.53‰ | 0.53‰ | 2.00‰ | 2.00‰ | 0.53 ‰ | 2.00 ‰ | AMS | 2.14‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.8165 | -995.19 ‰ | -995.19 ‰ | 0.11 | 0.11 | -995.13 | 0.12 ‰ | -995.13 | 0.12 ‰ |  |  | 2020-12-11 | Jocelyn Turnbull | 14:55 | 0.0048142 | 0.0048142 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 3.2 |  |  | 0.00481 | 0.00011 | 21784 | 2020-12-11 | 13:16 | albert zondervan | 216791 |  |  |  |  | 45% | 5 | 11/12/20 | 14:54:41 | Jocelyn Turnbull | 0 |  | 0.48 | 0.01% | 1 | 5 | 6 | ... | 0.00463 | 0.00011 |  |  |  |  | 0.98743 | 0.98743 | Yes | Yes | [5 empty] | 0.00463 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00481 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00463 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-08 | Julia Collins | 14:29 |  | Completed |  |  |  |  | 80600 | 2000000 |  |  | 80533 80602 80603 80604 80605 80606 | 80533~80602~80603~80604~80605~80606 | 3321 |  | 3321 3321 3321 3321 3321 3321 | 2020-12-11 | 13:22 | albert zondervan |  | 0.60 mg |  | 0.6 | [5 empty] | 0.99896 | 0.00161 | 80533 |  | XCAMS | ST | 0% |  | TW3321 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2021-01-19 | Water |  |  |  | 71051 | Normal | 14047/11 |  |  | 2020-12-07 | 5035 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99962 | 0.0016 | 80602 | [25 empty] | 2020-12-07 | 2973 | Gas Splitting | Completed
[168 empty] | 1.0007 | 0.0016 | 80603 | [25 empty] | 2020-12-07 | 5004 | Choose Gas Splitting | Completed
[168 empty] | 1.0014 | 0.00163 | 80604 | [25 empty] | 2020-12-08 | 20610 | Graphite | Completed
[168 empty] | 1 | 0.0016 | 80605 | [25 empty] | 2020-12-11 | 7743 | C13 Measurement | Completed
[168 empty] | 0.99935 | 0.0016 | 80606 | [25 empty] | 2020-12-11 | 21784 | AMS Submission Results | Completed
[196 empty] | 2020-12-14 | 25899 | Close Job | Completed
[7 empty] | 101.3 | 87.6827 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44431 | 44431 | 44431 | 210 | 209 | 209 |  |  | Julia Collins | -996.04 | 0.10 ‰ | 2020-12-16 | 2020-06-19 | 2 | 180 |  |  |  |  | 2.06‰ | 2.06‰ | 2.00‰ | 2.00‰ | 2.06 ‰ | 2.00 ‰ | AMS | 2.44‰ | 0.2‰ | C13 | No | -19 | -19.0013 | 0 | 0.8165 | -996.04 ‰ | -996.04 ‰ | 0.1 | 0.1 | -995.99 | 0.10 ‰ | -995.99 | 0.10 ‰ |  |  | 2020-12-22 | albert zondervan | 14:37 | 0.0039616 | 0.0039616 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 1.8 |  |  | 0.00396 | 0.0001 | 21856 | 2020-12-17 | 13:20 | albert zondervan | 216588 |  |  |  |  | 45% | 5 | 22/12/20 | 14:36:52 | albert zondervan | 0 |  | 0.4 | 0.01% | 11 | 5 | 4 | ... | 0.00381 | 0.0001 |  |  |  |  | 0.98734 | 0.98734 | Yes | Yes | [5 empty] | 0.00381 | 0.00‰ | 0.0001 | 0.00‰ | No | 1 | 1.0001 | 0.00124 | 0.00124 | 0.00396 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00381 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2020-12-15 | Julia Collins | 10:50 |  | Completed |  |  |  |  | 80648 | 2000000 |  |  | 80527 80529 80530 80531 80532 80607 | 80527~80529~80530~80531~80532~80607 | 3323 |  | 3323 3323 3323 3323 3323 3323 |  |  |  |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99717 | 0.003 | 80527 |  | XCAMS | ST | 0.25% |  | TW3323 measurement: * large & small samples CalAMS error adjustments: * 0.25% system error primaries: * OxI cathodes: 6x large: no failures, used in CalAMS for calibration 2x small: no failures, used to co-determine 'small d' small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 d = 0.79 ug C, see 'to blank-correct small samples.xlsm' secondaries: * 1x large FIRI-D cellulose: result agrees with consensus value * 1x large IAEA-C2 travertine: result agrees with consensus value * 2x small FIRI-D cellulose: result agrees with consensus value blanks: * three groups identified to blank-correct: large waters: 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks large organics via AAA: 6x unknowns + 0x secondaries + 0x blanks 180 days: capturing 17x blanks large organics via CE-Rafter: 3x unknowns + 1x secondaries + 1x blanks 360 days: capturing 4x blanks small organics via AAA & CE-Rafter: AAA: 4x unknowns + 0x secondaries + 1x blanks CE-Rafter: 0x unknowns + 2x secondaries + 2x blanks small OxI from 4x other wheels selected also: TW3302,3296,3253,3278 m = 0.74 ug C, see 'to blank-correct small samples.xlsm' unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2020-12-30 | Water | Complete |  |  | 71097 | Normal | 14047/11 |  |  | 2020-12-11 | 5042 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.004 | 0.0031 | 80529 | [25 empty] | 2020-12-14 | 20658 | Graphite | Completed
[168 empty] | 1.0027 | 0.00305 | 80530 | [25 empty] | 2020-12-17 | 7760 | C13 Measurement | Completed
[168 empty] | 0.99867 | 0.00302 | 80531 | [25 empty] | 2020-12-22 | 21856 | AMS Submission Results | Completed
[168 empty] | 1.0016 | 0.00301 | 80532 | [25 empty] | 2020-12-22 | 25961 | Close Job | Completed
[168 empty] | 0.99656 | 0.00302 | 80607
[7 empty] | 82.1 | 87.77 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 48843 | 48842 | 48842 | 256 | 255 | 255 |  |  | Julia Collins | -997.71 | 0.07 ‰ | 2021-02-09 | 2020-08-13 | 5 | 180 |  |  |  |  | 3.64‰ | 3.64‰ | 2.00‰ | 2.00‰ | 3.64 ‰ | 2.00 ‰ | AMS | 2.00‰ | 0.2‰ | C13 | No | -19 | -19.0002 | 0 | 0.8165 | -997.71 ‰ | -997.71 ‰ | 0.07 | 0.07 | -997.68 | 0.07 ‰ | -997.68 | 0.07 ‰ |  |  | 2021-02-12 | Cathy Ginnane | 8:15 | 0.0022875 | 0.0022875 | 7e-05 | 7e-05 | 7e-05 | 7e-05 | 1.9 |  |  | 0.00229 | 7e-05 | 21998 | 2021-02-11 | 11:57 | albert zondervan | 216938 |  |  |  |  | 45% | 5 | 12/02/21 | 08:14:31 | Cathy Ginnane | 0 |  | 0.23 | 0.01% | 1 | 5 | 4 | ... | 0.0022 | 7e-05 |  |  |  |  | 0.98988 | 0.98988 | Yes | Yes |  | The result for this blank is very low, much like a Kauri wood combustion blank. |  |  |  | 0.0022 | 0.00‰ | 7e-05 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00229 | 0 | 7e-05 | 0 | 0 |  |  | 0 | 0 | 0.0022 | 0 | 7e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-02-09 | Julia Collins | 8:29 |  | Completed |  |  |  |  | 80827 | 2000000 |  |  | 80685 80686 80696 80697 80698 80743 | 80685~80686~80696~80697~80698~80743 | 3327 |  | 3327 3327 3327 3327 3327 3327 | 2021-02-11 | 12:10 | albert zondervan |  | 1.10 mg |  | 1.1 | [5 empty] | 0.99874 | 0.00152 | 80685 |  | XCAMS | ST | 0% |  | TW3327 measurement: * large water samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * 1x IAEA-C2 on water-line: result agrees with consensus value blanks: * one groups identified to blank-correct: water: 29x unknowns + 1x secondaries + 1x blanks 180 days: capturing 9x blanks The result for the blank in this wheel is very low, much like a Kauri wood combustion blank. unknowns: * no failures | 2 | RRL | [5 empty] | Background Inorganic |  | 2021-02-12 | Water | Complete |  |  | 71224 | Normal | 14047/11 |  |  | 2021-02-03 | 5080 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1 | 0.00157 | 80686 | [25 empty] | 2021-02-04 | 20837 | Graphite | Completed
[168 empty] | 1.0009 | 0.00159 | 80696 | [25 empty] | 2021-02-10 | 7803 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00156 | 80697 | [25 empty] | 2021-02-12 | 21998 | AMS Submission Results | Completed
[168 empty] | 1.0004 | 0.00156 | 80698 | [25 empty] | 2021-02-12 | 26132 | Close Job | Completed
[168 empty] | 0.99949 | 0.00151 | 80743
[7 empty] | 91 | 87.9221 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45934 | 45933 | 45933 | 229 | 228 | 228 |  |  | Julia Collins | -996.71 | 0.09 ‰ | 2021-03-03 | 2020-09-04 | 0 | 180 |  |  |  |  | 4.30‰ | 4.30‰ | 2.00‰ | 2.00‰ | 4.30 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0028 | 0 | 0.8165 | -996.71 ‰ | -996.71 ‰ | 0.09 | 0.09 | -996.67 | 0.09 ‰ | -996.67 | 0.09 ‰ |  |  | 2021-03-05 | Jocelyn Turnbull | 11:26 | 0.0032858 | 0.0032858 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.00329 | 9e-05 | 22178 | 2021-03-05 | 9:41 | Jeremy Thompson | 217129 |  |  |  |  | 45% | 5 | 5/03/21 | 11:25:55 | Jocelyn Turnbull | 0 |  | 0.33 | 0.01% | 22 | 5 | 4 | ... | 0.00316 | 9e-05 |  |  |  |  | 0.99225 | 0.99225 | Yes | Yes | [5 empty] | 0.00316 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.0006 | 0.0006 | 0.00329 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00316 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-03 | Julia Collins | 14:13 |  | Completed |  |  |  |  | 81063 | 2000000 |  |  | 80926 80933 80935 80965 80977 80981 | 80926~80933~80935~80965~80977~80981 | 3331 |  | 3331 3331 3331 3331 3331 3331 | 2021-03-05 | 10:43 | Jeremy Thompson |  | 1.05 mg |  | 1.05 | [5 empty] | 0.9993 | 0.00144 | 80926 |  | XCAMS | EA | 0% |  | TW3331 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2: Freshwater Travertine (single extraction) 1x result agree with expectation value IAEA-C2: Freshwater Travertine (combined extractions) 1x result 2 sigma younger than Expected value FIRI-I: cellulose 1x result is 5-sigma lower than expected value blanks: * 4 groups identified to blank-correct: Water (Single Extraction): 14x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Water (combined Extraction): 2x unknowns + 1x secondaries + 1x blanks Selected blanks TP#;81047, 81049, 79669, 79711 AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks Cellulose: 1x unknowns + 0x secondaries + 1x blanks 180 days: capturing 2x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-03-23 | Water | Complete |  | reg sized blank | 71445 | Normal | 14047/11 |  |  | 2021-03-02 | 5128 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99767 | 0.0015 | 80933 | [25 empty] | 2021-03-03 | 21073 | Graphite | Completed
[168 empty] | 1.0018 | 0.00143 | 80935 | [25 empty] | 2021-03-22 | 7846 | C13 Measurement | Completed
[168 empty] | 1.0012 | 0.00164 | 80965 | [25 empty] | 2021-03-05 | 22178 | AMS Submission Results | Completed
[168 empty] | 0.99925 | 0.00145 | 80977 | [25 empty] | 2021-03-22 | 26533 | Close Job | Completed
[168 empty] | 1.0009 | 0.00146 | 80981
[7 empty] | 89.2 | 87.7381 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40341 | 40340 | 40340 | 152 | 152 | 152 |  |  | Julia Collins | -993.41 | 0.12 ‰ | 2021-03-08 | 2020-09-09 | 3 | 180 |  |  |  |  | 4.62‰ | 4.62‰ | 2.00‰ | 2.00‰ | 4.62 ‰ | 2.00 ‰ | AMS | 2.53‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.41 ‰ | -993.41 ‰ | 0.12 | 0.12 | -993.33 | 0.13 ‰ | -993.33 | 0.13 ‰ |  |  | 2021-03-10 | Jocelyn Turnbull | 9:55 | 0.0065923 | 0.0065923 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00659 | 0.00012 | 22231 | 2021-03-10 | 9:04 | Jeremy Thompson | 216996 |  |  |  |  | 45% | 5 | 10/03/21 | 09:54:54 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 1 | 5 | 4 | ... | 0.00634 | 0.00012 |  |  |  |  | 0.99048 | 0.99048 | Yes | Yes | [5 empty] | 0.00634 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00659 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00634 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-08 | Julia Collins | 8:44 |  | Completed |  |  |  |  | 81078 | 2000000 |  |  | 81022 81061 81062 81090 81091 81092 | 81022~81061~81062~81090~81091~81092 | 3333 |  | 3333 3333 3333 3333 3333 3333 | 2021-03-10 | 9:38 | Jeremy Thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99767 | 0.00162 | 81022 |  | XCAMS | EA | 0% |  | TW3333 measurement: * All Large samples CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * FIRI-D: cellulose 1x result agrees with consensus value blanks: * 2 groups identified to blank-correct: Water: 13x unknowns + 0x secondaries + 1x blanks(terrible blank) 180 days: capturing 7x blanks AAA : 8x unknowns + 1x secondaries + 1x blanks 180 days: capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-03-08 | Water |  |  |  | 71497 | Normal | 14047/11 |  |  | 2021-03-04 | 5142 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00158 | 81061 | [25 empty] | 2021-03-05 | 21088 | Graphite | Completed
[168 empty] | 0.99992 | 0.00157 | 81062 | [25 empty] | 2021-03-22 | 7863 | C13 Measurement | Completed
[168 empty] | 0.99984 | 0.00159 | 81090 | [25 empty] | 2021-03-10 | 22231 | AMS Submission Results | Completed
[168 empty] | 1.0002 | 0.00156 | 81091 | [25 empty] | 2021-03-30 | 26548 | Close Job | Completed
[168 empty] | 1.0016 | 0.00158 | 81092
[7 empty] | 102.3 | 87.6554 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40391 | 40391 | 40391 | 153 | 153 | 153 |  |  | Julia Collins | -993.45 | 0.12 ‰ | 2021-03-16 | 2020-09-17 | 4 | 180 |  |  |  |  | 4.67‰ | 4.67‰ | 2.00‰ | 2.00‰ | 4.67 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.45 ‰ | -993.45 ‰ | 0.12 | 0.12 | -993.37 | 0.13 ‰ | -993.37 | 0.13 ‰ |  |  | 2021-03-18 | Jocelyn Turnbull | 14:12 | 0.0065507 | 0.0065507 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.1 |  |  | 0.00655 | 0.00012 | 22300 | 2021-03-18 | 13:44 | albert zondervan | 217217 |  |  |  |  | 45% | 5 | 22/03/21 | 10:14:27 | jenny dahl | 0 |  | 0.66 | 0.01% | 27 | 5 | 6 | ... | 0.0063 | 0.00012 |  |  |  |  | 0.9896 | 0.9896 | Yes | Yes | [5 empty] | 0.0063 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00655 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.0063 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-03-15 | Julia Collins | 15:10 |  | Completed |  |  |  |  | 81152 | 2000000 |  |  | 81013 81017 81020 81053 81057 81059 | 81013~81017~81020~81053~81057~81059 | 3334 |  | 3334 3334 3334 3334 3334 3334 | 2021-03-18 | 13:52 | albert zondervan |  | 0.97 mg |  | 0.97 | [5 empty] | 0.99968 | 0.00164 | 81013 |  | XCAMS | EA | 0% |  | TW3334 measurement: * large waters + large carbonates CalAMS error adjustments: * none primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on carbonate line: 1x result agree with consensus value * IAEA-C2 on water line: 1x result agree with consensus value blanks: * 2 groups identified to blank-correct: water: 21x unknowns + 1x secondaries + 1x blanks 180 days: capturing 7x blanks carbonate: 7x unknowns + 1x secondaries + 1x blanks 180 days: capturing 4x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-04-13 | Water | Complete |  |  | 71567 | Normal | 14047/11 |  |  | 2021-03-11 | 5166 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0007 | 0.00166 | 81017 | [25 empty] | 2021-03-11 | 3038 | Gas Splitting | Completed
[168 empty] | 0.99929 | 0.00161 | 81020 | [25 empty] | 2021-03-11 | 5080 | Choose Gas Splitting | Completed
[168 empty] | 0.99856 | 0.00161 | 81053 | [25 empty] | 2021-03-12 | 21162 | Graphite | Completed
[168 empty] | 0.99921 | 0.0016 | 81057 | [25 empty] | 2021-07-16 | 7874 | C13 Measurement | Completed
[168 empty] | 1.0026 | 0.00163 | 81059 | [25 empty] | 2021-03-18 | 22300 | AMS Submission Results | Completed
[196 empty] | 2021-07-28 | 27521 | Close Job | Completed
[7 empty] | 92.3 | 87.9469 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42296 | 42296 | 42296 | 129 | 129 | 129 |  |  | Julia Collins | -994.83 | 0.08 ‰ | 2021-04-04 | 2020-10-06 | 3 | 180 |  |  |  |  | 4.83‰ | 4.83‰ | 2.00‰ | 2.00‰ | 4.83 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.89443 | -994.83 ‰ | -994.83 ‰ | 0.08 | 0.08 | -994.77 | 0.08 ‰ | -994.77 | 0.08 ‰ |  |  | 2021-04-19 | Jocelyn Turnbull | 15:08 | 0.0051678 | 0.0051678 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00517 | 8e-05 | 22467 | 2021-04-08 | 9:20 | albert zondervan | 217370 |  |  |  |  | 45% | 5 | 19/04/21 | 15:07:30 | Jocelyn Turnbull | 0 |  | 0.52 | 0.01% | 1 | 5 | 4 | ... | 0.00497 | 8e-05 |  |  |  |  | 0.99305 | 0.99305 | Yes | [6 empty] | 0.00497 | 0.00‰ | 8e-05 | 0.00‰ | No | 1 | 1 | 0.00128 | 0.00128 | 0.00517 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00497 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-01 | Julia Collins | 14:19 |  | Completed |  |  |  |  | 81316 | 2000000 |  |  | 81099 81100 81111 81290 81291 | 81099~81100~81111~81290~81291 | 3338 |  | 3338 3338 3338 3338 3338 | 2021-04-08 | 9:40 | albert zondervan |  | 1.04 mg |  | 1.04 | [5 empty] | 0.99969 | 0.00286 | 81099 |  | XCAMS | FLASK | 0.26% |  | *** JCT/AZ results of further tests show that the flask Ox is behaving normally again. Our best guess (but only a guess) is room air contamination on the graphite for hte flask 1 Ox in this wheel. Also not sure about the EA Ox in this wheel, likely a separate issue. System error is 0.26% using all 5 flask Oxs, which is the same as the wheel to wheel error we currently assign to carbonates. Therefore have reported with wheel to wheel error. TW3338 measurement: * large waters * complete system retune during measurement did not resolve OxI repeatability issues primaries: * OxI cathodes: 5x accepted, from flask 1x rejected, via EA (J216823), due to being 4-sigma outlier CalAMS error adjustments: * 1.05 global error multiplier due to excess run-to-run variability * 0.26% system error: determined from excess variability in 5x accepted flask OxI-CO2 cause of this exceptionally large system error is not identifiable in XCAMS raw data variability appears associated with OxI-CO2 flask number: taking means, flask 2 has ca. 0.5% higher rts than flask 1 system error not removed in RLIMS because of its size secondaries: * IAEA-C2 on water line: 1x result differs from consensus value by ca. 2-sigma blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days capturing 8x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-05-03 | Water | Complete |  | D bank stuck at ~0.0028 for prebake | 71696 | Normal | 14047/11 |  |  | 2021-03-31 | 5196 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99725 | 0.00286 | 81100 | [25 empty] | 2021-04-01 | 21326 | Graphite | Completed
[168 empty] | 0.99743 | 0.00286 | 81111 | [25 empty] | 2021-04-06 | 7920 | C13 Measurement | Completed
[168 empty] | 1.0021 | 0.00288 | 81290 | [25 empty] | 2021-04-19 | 22467 | AMS Submission Results | Completed
[168 empty] | 1.0038 | 0.00289 | 81291 | [25 empty] | 2021-04-22 | 26672 | Close Job | Completed
[7 empty] | 104.2 | 87.8537 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41887 | 41886 | 41886 | 184 | 184 | 184 |  |  | Julia Collins | -994.56 | 0.12 ‰ | 2021-04-21 | 2020-10-23 | 7 | 180 |  |  |  |  | 2.30‰ | 2.30‰ | 2.00‰ | 2.00‰ | 2.30 ‰ | 2.00 ‰ | AMS | 2.50‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -994.56 ‰ | -994.56 ‰ | 0.12 | 0.12 | -994.49 | 0.13 ‰ | -994.49 | 0.13 ‰ |  |  | 2021-04-22 | Cathy Ginnane | 22:53 | 0.0054381 | 0.0054381 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00544 | 0.00012 | 22530 | 2021-04-22 | 12:35 | albert zondervan | 217316 |  |  |  |  | 45% | 5 | 22/04/21 | 22:52:34 | Cathy Ginnane | 0 |  | 0.54 | 0.01% | 15 | 5 | 6 | ... | 0.00523 | 0.00012 |  |  |  |  | 0.9895 | 0.9895 | Yes | Yes | [5 empty] | 0.00523 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 0.99999 | 0.00084 | 0.00084 | 0.00544 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00523 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-04-20 | Julia Collins | 12:58 |  | Completed |  |  |  |  | 81343 | 2000000 |  |  | 81348 81349 81350 81351 81407 81408 | 81348~81349~81350~81351~81407~81408 | 3341 |  | 3341 3341 3341 3341 3341 3341 | 2021-04-22 | 13:26 | albert zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 1.0029 | 0.00202 | 81348 |  | XCAMS | ST | 0.12% |  | TW3341 measurement: * large samples CalAMS error adjustments: * 1.03 global error multiplier to account for excess run-to-run variability * 0.12% system error to account for excess Ox-I cathode-to-cathode variability primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value * FIRI-I cellulose: 2x result agrees with consensus value: R41572/4 prepared by Otlet, assuming this is a FIRI-I R24889/9 prepared by RRL blanks: * 3 groups identified to blank-correct: water: 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks RRL org: 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Otlet org: 6x unknowns + 1x secondaries + 2x blanks also picking up TW3324, capturing 5x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-04-23 | Water | Not Complete |  | D bank only went down to 0.0030 mbar | 71763 | Normal | 14047/11 |  |  | 2021-04-13 | 5199 | Water CO2 Evolution | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00203 | 81349 | [25 empty] | 2021-04-13 | 3075 | Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00207 | 81350 | [25 empty] | 2021-04-13 | 5121 | Choose Gas Splitting | Completed
[168 empty] | 0.99811 | 0.00202 | 81351 | [25 empty] | 2021-04-14 | 21353 | Graphite | Completed
[168 empty] | 0.99895 | 0.00212 | 81407 | [25 empty] | 2021-04-16 | 7929 | C13 Measurement | Completed
[168 empty] | 0.99781 | 0.00203 | 81408 | [25 empty] | 2021-04-22 | 22530 | AMS Submission Results | Completed
[196 empty] | 2021-04-23 | 26696 | Close Job | Completed
[7 empty] | 130.5 | 87.5549 | 0.0007 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 44182 | 44182 | 44182 | 225 | 224 | 224 |  |  | Julia Collins | -995.91 | 0.11 ‰ | 2021-06-10 | 2020-12-12 | 6 | 180 |  |  |  |  | 3.00‰ | 3.00‰ | 2.00‰ | 2.00‰ | 3.00 ‰ | 2.00 ‰ | AMS | 2.41‰ | 0.2‰ | C13 | No | -19 | -19.0009 | 0 | 0.8165 | -995.91 ‰ | -995.91 ‰ | 0.11 | 0.11 | -995.86 | 0.12 ‰ | -995.86 | 0.12 ‰ |  |  | 2021-06-11 | Jocelyn Turnbull | 12:02 | 0.0040864 | 0.0040864 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2.5 |  |  | 0.00409 | 0.00011 | 22969 | 2021-06-11 | 11:47 | Albert Zondervan | 217605 |  |  |  |  | 45% | 5 | 11/06/21 | 12:02:27 | Jocelyn Turnbull | 0 |  | 0.41 | 0.01% | 1 | 5 | 5 | ... | 0.00393 | 0.00011 |  |  |  |  | 0.98682 | 0.98682 | Yes | Yes | [5 empty] | 0.00393 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00068 | 0.00068 | 0.00409 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00393 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-09 | Julia Collins | 11:50 |  | Completed |  |  |  |  | 81829 | 2000000 |  |  | 81782 81783 81808 81809 81828 81864 | 81782~81783~81808~81809~81828~81864 | 3352 |  | 3352 3352 3352 3352 3352 3352 | 2021-06-11 | 11:51 | Albert Zondervan |  | 0.88 mg |  | 0.88 | [5 empty] | 0.99999 | 0.00172 | 81782 |  | XCAMS | ST | 0% |  | TW3352 measurement: * large water samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x accepted secondaries: * IAEA-C2 on waterline: 1x result agrees with consensus value blanks: * 1 groups identified to blank-correct: water: 30x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-06-07 | Water | Complete |  |  | 72154 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99939 | 0.00164 | 81783 | [25 empty] | 2021-06-03 | 5298 | Water CO2 Evolution | Completed
[168 empty] | 1.0021 | 0.00166 | 81808 | [25 empty] | 2021-06-04 | 21839 | Graphite | Completed
[168 empty] | 0.99924 | 0.0017 | 81809 | [25 empty] | 2021-06-08 | 8102 | C13 Measurement | Completed
[168 empty] | 1.0003 | 0.00168 | 81828 | [25 empty] | 2021-06-11 | 22969 | AMS Submission Results | Completed
[168 empty] | 0.99904 | 0.00163 | 81864 | [25 empty] | 2021-06-11 | 27256 | Close Job | Completed
[7 empty] | 136 | 87.72 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41238 | 41237 | 41237 | 169 | 169 | 169 |  |  | Julia Collins | -994.1 | 0.12 ‰ | 2021-06-22 | 2020-12-24 | 4 | 180 |  |  |  |  | 2.07‰ | 2.07‰ | 2.00‰ | 2.00‰ | 2.07 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.001 | 0 | 0.8165 | -994.10 ‰ | -994.10 ‰ | 0.12 | 0.12 | -994.03 | 0.13 ‰ | -994.03 | 0.13 ‰ |  |  | 2021-06-24 | Cathy Ginnane | 23:58 | 0.0058956 | 0.0058956 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.2 |  |  | 0.0059 | 0.00012 | 23100 | 2021-06-23 | 12:12 | Albert Zondervan | 217909 |  |  |  |  | 45% | 5 | 24/06/21 | 23:58:21 | Cathy Ginnane | 0 |  | 0.59 | 0.01% | 37 | 5 | 7 | ... | 0.00567 | 0.00012 |  |  |  |  | 0.98776 | 0.98776 | Yes | Yes | [5 empty] | 0.00567 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00113 | 0.00113 | 0.0059 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00567 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-06-21 | Julia Collins | 10:44 |  | Completed |  |  |  |  | 81922 | 2000000 |  |  | 81865 81893 81894 81895 81896 81937 | 81865~81893~81894~81895~81896~81937 | 3355 |  | 3355 3355 3355 3355 3355 3355 | 2021-06-23 | 12:29 | Albert Zondervan |  | 0.94 mg |  | 0.94 | [5 empty] | 1.0025 | 0.00275 | 81865 |  | XCAMS | ST | 0.22% |  | TW3355 measurement CalAMS error adjustments: * factor 1.07 due to excess global run-to-run variability * 0.22% system error due to excess OxI cat-to-cat variability primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 4-sigma blanks: * 2 groups identified to blank-correct: organics via AAA: 11x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks water: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks unknowns: * no failures TO BE SORTED: * The IRMS del13C values of J218004 and J217995 appear swapped, while the XCAMS data do not suggest a problem. *refer to e-log 223 on possible swap concerns | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2021-07-07 | Water | Complete |  | C13 split initially sealed then transfered to MS bottle. | 72353 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0024 | 0.0028 | 81893 | [25 empty] | 2021-06-17 | 5350 | Water CO2 Evolution | Completed
[168 empty] | 0.99658 | 0.00273 | 81894 | [25 empty] | 2021-06-17 | 3119 | Gas Splitting | Completed
[168 empty] | 0.99815 | 0.00279 | 81895 | [25 empty] | 2021-06-17 | 5215 | Choose Gas Splitting | Completed
[168 empty] | 0.99796 | 0.00277 | 81896 | [25 empty] | 2021-06-18 | 21932 | Graphite | Completed
[168 empty] | 1.0025 | 0.00279 | 81937 | [25 empty] | 2021-06-22 | 8167 | C13 Measurement | Completed
[196 empty] | 2021-06-24 | 23100 | AMS Submission Results | Completed
[196 empty] | 2021-06-25 | 27479 | Close Job | Completed
[7 empty] | 106.5 | 87.9491 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 185 | 185 | 185 |  |  | Taylor Ferrick | -995.04 | 0.11 ‰ | 2021-07-22 | 2021-01-23 | 6 | 180 |  |  |  |  | 2.91‰ | 2.91‰ | 2.00‰ | 2.00‰ | 2.91 ‰ | 2.00 ‰ | AMS | 2.48‰ | 0.2‰ | C13 | No | -19 | -19.0059 | 0 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.11 | 0.11 | -994.98 | 0.12 ‰ | -994.98 | 0.12 ‰ |  |  | 2021-07-26 | Jocelyn Turnbull | 10:23 | 0.0049598 | 0.0049598 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00496 | 0.00011 | 23228 | 2021-07-23 | 15:27 | Jeremy Thompson | 218145 |  |  |  |  | 45% | 5 | 26/07/21 | 10:22:52 | Jocelyn Turnbull | 0 |  | 0.5 | 0.01% | 1 | 5 | 5 | ... | 0.00477 | 0.00011 |  |  |  |  | 0.99118 | 0.99118 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00069 | 0.00069 | 0.00496 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-07-21 | Taylor Ferrick | 13:09 |  | Completed |  |  |  |  | 82086 | 2000000 |  |  | 82082 82083 82084 82085 82096 82101 | 82082~82083~82084~82085~82096~82101 | 3358 |  | 3358 3358 3358 3358 3358 3358 | 2021-07-23 | 15:46 | Jeremy Thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0032 | 0.00169 | 82082 |  | XCAMS | ST | 0% |  | TW3358 measurement CalAMS error adjustments: * no error adjustment needed primaries: * large OxI cathodes: 6x accepted secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result comes out too young by 2-sigma blanks: * 3 groups identified to blank-correct: Organics via AAA: 13x unknowns + 1x secondaries + 1x blanks 180 days, capturing 8x blanks Water: 14x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Carbonate 1x unknowns + 0x secondaries + 0x blanks 180 days, capturing 8x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-08-04 | Water |  |  | Using liquid nitrogen and ethanol slushies. | 72473 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00166 | 82083 | [25 empty] | 2021-07-15 | 5402 | Water CO2 Evolution | Completed
[168 empty] | 0.99757 | 0.00165 | 82084 | [25 empty] | 2021-07-16 | 22096 | Graphite | Completed
[168 empty] | 0.99887 | 0.00166 | 82085 | [25 empty] | 2021-07-19 | 8223 | C13 Measurement | Completed
[168 empty] | 1.0014 | 0.00174 | 82096 | [25 empty] | 2021-07-26 | 23228 | AMS Submission Results | Completed
[168 empty] | 0.99909 | 0.00168 | 82101 | [25 empty] | 2021-07-28 | 27572 | Close Job | Completed
[7 empty] | 93.4 | 88.0575 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40252 | 40252 | 40252 | 163 | 162 | 162 |  |  | Taylor Ferrick | -993.33 | 0.14 ‰ | 2021-09-29 | 2021-04-02 | 7 | 180 |  |  |  |  | 1.15‰ | 1.15‰ | 2.00‰ | 2.00‰ | 1.15 ‰ | 2.00 ‰ | AMS | 2.47‰ | 0.2‰ | C13 | No | -19 | -19.0051 | 0 | 0.8165 | -993.33 ‰ | -993.33 ‰ | 0.14 | 0.14 | -993.25 | 0.14 ‰ | -993.25 | 0.14 ‰ |  |  | 2021-10-01 | Jocelyn Turnbull | 11:31 | 0.0066651 | 0.0066651 | 0.00014 | 0.00014 | 0.00014 | 0.00014 | 2.1 |  |  | 0.00667 | 0.00014 | 23629 | 2021-10-01 | 8:51 | Jeremy Thompson | 218818 |  |  |  |  | 45% | 5 | 1/10/21 | 11:30:52 | Jocelyn Turnbull | 0 |  | 0.67 | 0.01% | 22 | 5 | 5 | ... | 0.00641 | 0.00013 |  |  |  |  | 0.99065 | 0.99065 | Yes | Yes | [5 empty] | 0.00641 | 0.01‰ | 0.00013 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00667 | 0 | 0.00014 | 0 | 0 |  |  | 0 | 0 | 0.00641 | 0 | 0.00013 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-09-24 | Taylor Ferrick | 8:49 |  | Completed |  |  |  |  | 82456 | 2000000 |  |  | 82438 82439 82440 82441 82458 82459 | 82438~82439~82440~82441~82458~82459 | 3368 |  | 3368 3368 3368 3368 3368 3368 | 2021-10-01 | 11:00 | Jeremy Thompson |  | 1.01 mg |  | 1.01 | [5 empty] | 1.0017 | 0.00148 | 82438 |  | XCAMS | ST | 0% |  | TW3368 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 19x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 0x blanks 180 days, capturing 1x blanks + 9 AAA blanks Water 7x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-10-20 | Water |  |  |  | 72865 | Normal | 14047/11 |  |  | 2021-04-13 | 7242 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0001 | 0.00145 | 82439 | [25 empty] | 2021-09-20 | 5464 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.00146 | 82440 | [25 empty] | 2021-09-22 | 22466 | Graphite | Completed
[168 empty] | 0.99716 | 0.00142 | 82441 | [25 empty] | 2021-09-23 | 8372 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00159 | 82458 | [25 empty] | 2021-10-01 | 23629 | AMS Submission Results | Completed
[168 empty] | 0.99945 | 0.00161 | 82459 | [25 empty] | 2021-10-01 | 28074 | Close Job | Completed
[7 empty] | 96.8 | 87.7945 | 0.0008 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42728 | 42727 | 42727 | 188 | 187 | 187 |  |  | Taylor Ferrick | -995.1 | 0.11 ‰ | 2021-10-04 | 2021-04-07 | 5 | 180 |  |  |  |  | 3.13‰ | 3.13‰ | 2.00‰ | 2.00‰ | 3.13 ‰ | 2.00 ‰ | AMS | 2.46‰ | 0.2‰ | C13 | No | -19 | -19.0024 | 0 | 0.8165 | -995.10 ‰ | -995.10 ‰ | 0.11 | 0.11 | -995.04 | 0.12 ‰ | -995.04 | 0.12 ‰ |  |  | 2021-10-13 | Jocelyn Turnbull | 8:51 | 0.0048974 | 0.0048974 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 1.8 |  |  | 0.0049 | 0.00011 | 23638 | 2021-10-06 | 8:27 | Jeremy Thompson | 218996 |  |  |  |  | 45% | 5 | 13/10/21 | 08:50:43 | Jocelyn Turnbull | 0 |  | 0.49 | 0.01% | 1 | 5 | 5 | ... | 0.00471 | 0.00011 |  |  |  |  | 0.98965 | 0.98965 | Yes | Yes | [5 empty] | 0.00471 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.0049 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00471 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-01 | Taylor Ferrick | 11:09 |  | Completed |  |  |  |  | 82500 | 2000000 |  |  | 82472 82473 82506 82518 82519 82522 | 82472~82473~82506~82518~82519~82522 | 3369 |  | 3369 3369 3369 3369 3369 3369 | 2021-10-06 | 8:52 | Jeremy Thompson |  | 1.06 mg |  | 1.06 | [5 empty] | 0.99683 | 0.00157 | 82472 |  | XCAMS | EA | 0% |  | TW3369 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C2 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 16x unknowns + 1x secondaries + 0x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 2x unknowns + 0x secondaries + 1x blanks 180 days, capturing 2x blanks + 9 AAA blanks Water 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 6x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-11-17 | Water | Complete |  | For new treatment of marble material. Please make jobs for KIT, LAG, MEH, etc | 72905 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99911 | 0.00162 | 82473 | [25 empty] | 2021-09-28 | 5483 | Water CO2 Evolution | Completed
[168 empty] | 1.0014 | 0.00159 | 82506 | [25 empty] | 2021-09-29 | 22510 | Graphite | Completed
[168 empty] | 1.0005 | 0.00154 | 82518 | [25 empty] | 2021-09-30 | 8394 | C13 Measurement | Completed
[168 empty] | 0.99896 | 0.00158 | 82519 | [25 empty] | 2021-10-13 | 23638 | AMS Submission Results | Completed
[168 empty] | 1.0034 | 0.0016 | 82522 | [25 empty] | 2021-10-13 | 28161 | Close Job | Completed
[7 empty] | 107.4 | 87.7447 | 0.0018 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 36404 | 36403 | 36403 | 124 | 124 | 124 |  |  | Taylor Ferrick | -989.24 | 0.17 ‰ | 2021-10-10 | 2021-04-13 | 3 | 180 |  |  |  |  | 5.04‰ | 5.04‰ | 2.00‰ | 2.00‰ | 5.04 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0022 | 0 | 0.8165 | -989.24 ‰ | -989.24 ‰ | 0.17 | 0.17 | -989.11 | 0.17 ‰ | -989.11 | 0.17 ‰ |  |  | 2021-10-13 | Jocelyn Turnbull | 9:00 | 0.010762 | 0.010762 | 0.00017 | 0.00017 | 0.00017 | 0.00017 | 2.2 |  |  | 0.01076 | 0.00017 | 23678 | 2021-10-12 | 11:06 | Albert Zondervan | 219248 |  |  |  |  | 45% | 5 | 13/10/21 | 09:41:48 | jenny dahl | 0 |  | 1.08 | 0.02% | 1 | 5 | 6 | ... | 0.01035 | 0.00016 |  |  |  |  | 0.99097 | 0.99097 | Yes | Yes |  | poor blank |  |  |  | 0.01035 | 0.01‰ | 0.00016 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.01076 | 0 | 0.00017 | 0 | 0 |  |  | 0 | 0 | 0.01035 | 0 | 0.00016 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-08 | Taylor Ferrick | 11:44 |  | Completed |  |  |  |  | 82573 | 2000000 |  |  | 82523 82559 82560 82577 82578 82600 | 82523~82559~82560~82577~82578~82600 | 3370 |  | 3370 3370 3370 3370 3370 3370 | 2021-10-12 | 11:27 | Albert Zondervan |  | 0.84 mg |  | 0.84 | [5 empty] | 1.0001 | 0.00161 | 82523 |  | XCAMS | EA | 0% |  | TW3370 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x result agrees with consensus value * IAEA-C6 1x result agrees with consensus value blanks: * 3 groups identified to blank-correct: Organics via AAA: 17x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Organics via Cellulose Extraction: 1x unknowns + 0x secondaries + 0x blanks 360 days, capturing 3x blanks Water 11x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks blank TP82573 is high unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Complete |  | refreshed material under J 218996 | 72945 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.0016 | 82559 | [25 empty] | 2021-10-06 | 5494 | Water CO2 Evolution | Cancelled
[168 empty] | 0.99776 | 0.00162 | 82560 | [25 empty] | 2021-10-06 | 5498 | Water CO2 Evolution | Completed
[168 empty] | 1.0012 | 0.00161 | 82577 | [25 empty] | 2021-10-07 | 22583 | Graphite | Completed
[168 empty] | 0.99969 | 0.00156 | 82578 | [25 empty] | 2021-10-11 | 8405 | C13 Measurement | Completed
[168 empty] | 1.001 | 0.00164 | 82600 | [25 empty] | 2021-10-13 | 23678 | AMS Submission Results | Completed
[196 empty] | 2021-10-13 | 28163 | Close Job | Completed
[7 empty] | 98.1 | 88.1074 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40878 | 40877 | 40877 | 163 | 162 | 162 |  |  | Taylor Ferrick | -993.83 | 0.12 ‰ | 2021-10-22 | 2021-04-25 | 3 | 180 |  |  |  |  | 1.82‰ | 1.82‰ | 2.00‰ | 2.00‰ | 1.82 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0006 | 0 | 0.8165 | -993.83 ‰ | -993.83 ‰ | 0.12 | 0.12 | -993.76 | 0.13 ‰ | -993.76 | 0.13 ‰ |  |  | 2021-10-26 | Jocelyn Turnbull | 9:40 | 0.006166 | 0.006166 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00617 | 0.00012 | 23798 | 2021-10-23 | 14:38 | Albert Zondervan | 219249 |  |  |  |  | 45% | 5 | 26/10/21 | 09:40:01 | Jocelyn Turnbull | 0 |  | 0.62 | 0.01% | 1 | 5 | 7 | ... | 0.00593 | 0.00012 |  |  |  |  | 0.99188 | 0.99188 | Yes | Yes | [5 empty] | 0.00593 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00617 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00593 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-10-19 | Taylor Ferrick | 11:39 |  | Completed |  |  |  |  | 82685 | 2000000 |  |  | 82609 82610 82611 82612 82613 82614 | 82609~82610~82611~82612~82613~82614 | 3373 |  | 3373 3373 3373 3373 3373 3373 | 2021-10-23 | 14:42 | Albert Zondervan |  | 0.83 mg |  | 0.83 | [5 empty] | 1.0011 | 0.00163 | 82609 |  | XCAMS | ST | 0% |  | TW3373 measurement * A retune was necessary as measurement _2 had poor OxI repeatability. * Only _3 measurement data have been imported into RLIMS. CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * IAEA-C2 travertine 1x result 2-sigma younger than consensus age * LAA1 coral 1x result seems to be in keeping with previously measured values * LAC1 coral 1x result seems to be in keeping with previously measured values blanks: * 1 groups identified to blank-correct: Water 28x unknowns + 3x secondaries + 1x blanks 180 days, capturing 7x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water |  |  | refreshed material under J 218996 | 73066 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0002 | 0.00158 | 82610 | [25 empty] | 2021-10-18 | 5527 | Water CO2 Evolution | Completed
[168 empty] | 1.0003 | 0.00159 | 82611 | [25 empty] | 2021-10-18 | 3229 | Gas Splitting | Completed
[168 empty] | 0.99922 | 0.00161 | 82612 | [25 empty] | 2021-10-18 | 5385 | Choose Gas Splitting | Completed
[168 empty] | 1.0001 | 0.0016 | 82613 | [25 empty] | 2021-10-19 | 22695 | Graphite | Completed
[168 empty] | 0.99915 | 0.00159 | 82614 | [25 empty] | 2021-10-19 | 8453 | C13 Measurement | Completed
[196 empty] | 2021-10-26 | 23798 | AMS Submission Results | Completed
[196 empty] | 2021-11-01 | 28360 | Close Job | Completed
[7 empty] | 141.4 | 87.8688 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37499 | 37499 | 37499 | 107 | 106 | 106 |  |  | Taylor Ferrick | -990.61 | 0.12 ‰ | 2021-11-05 | 2021-05-09 | 1 | 180 |  |  |  |  | -0.35‰ | -0.35‰ | 2.00‰ | 2.00‰ | -0.35 ‰ | 2.00 ‰ | AMS | 2.22‰ | 0.2‰ | C13 | No | -19 | -19.0025 | 0 | 0.8165 | -990.61 ‰ | -990.61 ‰ | 0.12 | 0.12 | -990.49 | 0.13 ‰ | -990.49 | 0.13 ‰ |  |  | 2021-11-08 | Jocelyn Turnbull | 9:36 | 0.0093893 | 0.0093893 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.8 |  |  | 0.00939 | 0.00012 | 23918 | 2021-11-08 | 8:56 | Jeremy Thompson | 219250 |  |  |  |  | 45% | 5 | 9/11/21 | 12:11:33 | jenny dahl | 0 |  | 0.94 | 0.01% | 1 | 5 | 5 | ... | 0.00903 | 0.00012 |  |  |  |  | 0.98706 | 0.98706 | Yes | Yes | [5 empty] | 0.00903 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00054 | 0.00054 | 0.00939 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00903 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-04 | Taylor Ferrick | 12:16 |  | Completed |  |  |  |  | 82800 | 2000000 |  |  | 82752 82753 82758 82759 82760 82761 | 82752~82753~82758~82759~82760~82761 | 3376 |  | 3376 3376 3376 3376 3376 3376 | 2021-11-08 | 9:23 | Jeremy Thompson |  | 0.74 mg |  | 0.74 | [5 empty] | 1.0019 | 0.00135 | 82752 |  | XCAMS | ST | 0% |  | TW3376 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consenous IAEA-C2: Freshwater Travertine 1x Result 2 sigma out LAC: LAC1 coral 1x Result agrees with consenous LAA: LAA1 coral 1x Result agrees with consenous blanks: * 3 groups identified to blank-correct: Organics via AAA: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 17x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 9x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Not Complete |  | refreshed material under J 218996 | 73185 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9996 | 0.00132 | 82753 | [25 empty] | 2021-11-03 | 5569 | Water CO2 Evolution | Completed
[168 empty] | 0.99998 | 0.0013 | 82758 | [25 empty] | 2021-11-04 | 22810 | Graphite | Completed
[168 empty] | 0.99929 | 0.00128 | 82759 | [25 empty] | 2021-11-09 | 8508 | C13 Measurement | Completed
[168 empty] | 0.99934 | 0.00132 | 82760 | [25 empty] | 2021-11-08 | 23918 | AMS Submission Results | Completed
[168 empty] | 1 | 0.0013 | 82761 | [25 empty] | 2021-11-12 | 28423 | Close Job | Completed
[7 empty] | 115.2 | 87.8949 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37411 | 37411 | 37411 | 150 | 149 | 149 |  |  | Taylor Ferrick | -990.51 | 0.18 ‰ | 2021-11-11 | 2021-05-15 | 1 | 180 |  |  |  |  | 2.77‰ | 2.77‰ | 2.00‰ | 2.00‰ | 2.77 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0036 | 0 | 0.8165 | -990.51 ‰ | -990.51 ‰ | 0.18 | 0.18 | -990.39 | 0.18 ‰ | -990.39 | 0.18 ‰ |  |  | 2021-11-12 | Jocelyn Turnbull | 14:27 | 0.0094933 | 0.0094933 | 0.00018 | 0.00018 | 0.00018 | 0.00018 | 2.4 |  |  | 0.00949 | 0.00018 | 23989 | 2021-11-12 | 13:02 | Albert Zondervan | 219251 |  |  |  |  | 45% | 5 | 12/11/21 | 14:27:10 | Jocelyn Turnbull | 0 |  | 0.95 | 0.02% | 1 | 5 | 7 | ... | 0.00913 | 0.00017 |  |  |  |  | 0.99043 | 0.99043 | Yes | Yes | [5 empty] | 0.00913 | 0.01‰ | 0.00017 | 0.00‰ | Yes | 1 | 1 | 0.00064 | 0.00064 | 0.00949 | 0 | 0.00018 | 0 | 0 |  |  | 0 | 0 | 0.00913 | 0 | 0.00017 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-10 | Taylor Ferrick | 12:36 |  | Completed |  |  |  |  | 82861 | 2000000 |  |  | 82790 82791 82792 82850 82853 82856 | 82790~82791~82792~82850~82853~82856 | 3377 |  | 3377 3377 3377 3377 3377 3377 | 2021-11-12 | 13:27 | Albert Zondervan |  | 0.78 mg |  | 0.78 | [5 empty] | 1.0009 | 0.00145 | 82790 |  | XCAMS | EA | 0% |  | TW3377 measurement CalAMS error adjustments: * none needed primaries: * 6 OxI cathodes: no failures secondaries: * water IAEA-C2: 1x result differs from consensus value by 2.5 sigma LAC1: 1x result agrees with expectation value LAA1: 1x Result agrees with expectation value * organic FIRI-I 1x result agrees with consensus value blanks: * water 14x unknowns + 3x secondaries + 1x blanks 180 days, capturing 9x blanks * organic via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | 2.1 | RRL | [5 empty] | Background Inorganic |  | 2021-11-24 | Water | Not Complete |  | refreshed material under J 218996 | 73225 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0005 | 0.00171 | 82791 | [25 empty] | 2021-11-08 | 5587 | Water CO2 Evolution | Completed
[168 empty] | 0.99918 | 0.00157 | 82792 | [25 empty] | 2021-11-08 | 3243 | Gas Splitting | Completed
[168 empty] | 1.0009 | 0.00163 | 82850 | [25 empty] | 2021-11-08 | 5405 | Choose Gas Splitting | Completed
[168 empty] | 0.99657 | 0.00154 | 82853 | [25 empty] | 2021-11-10 | 22871 | Graphite | Completed
[168 empty] | 1.002 | 0.00155 | 82856 | [25 empty] | 2021-11-16 | 8524 | C13 Measurement | Completed
[196 empty] | 2021-11-12 | 23989 | AMS Submission Results | Completed
[196 empty] | 2021-11-29 | 28556 | Close Job | Completed
[7 empty] | 79.8 | 88.3545 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41979 | 41979 | 41979 | 186 | 186 | 186 |  |  | Taylor Ferrick | -994.62 | 0.12 ‰ | 2021-11-22 | 2021-05-26 | 4 | 180 |  |  |  |  | 0.31‰ | 0.31‰ | 2.00‰ | 2.00‰ | 0.31 ‰ | 2.00 ‰ | AMS | 2.19‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -994.62 ‰ | -994.62 ‰ | 0.12 | 0.12 | -994.56 | 0.13 ‰ | -994.56 | 0.13 ‰ |  |  | 2021-11-24 | Jocelyn Turnbull | 10:35 | 0.0053757 | 0.0053757 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.3 |  |  | 0.00538 | 0.00012 | 24038 | 2021-11-24 | 9:50 | Jeremy Thompson | 219337 |  |  |  |  | 45% | 5 | 24/11/21 | 10:35:18 | Jocelyn Turnbull | 0 |  | 0.54 | 0.01% | 1 | 5 | 7 | ... | 0.00517 | 0.00012 |  |  |  |  | 0.99316 | 0.99316 | Yes | Yes | [5 empty] | 0.00517 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00065 | 0.00065 | 0.00538 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00517 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-19 | Taylor Ferrick | 11:32 |  | Completed |  |  |  |  | 82938 | 2000000 |  |  | 82859 82891 82892 82893 82894 82895 | 82859~82891~82892~82893~82894~82895 | 3379 |  | 3379 3379 3379 3379 3379 3379 | 2021-11-24 | 10:01 | Jeremy Thompson |  | 0.80 mg |  | 0.8 | [5 empty] | 1.0013 | 0.00155 | 82859 |  | XCAMS | EA | 0% |  | TW3379 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with consensus blanks: * 1 groups identified to blank-correct: Waters 29x unknowns + 2x secondaries (1 coral) + 1x blanks 180 days, capturing 10x blanks unknowns: * no failures | 2.2 | RRL | [5 empty] | Background Inorganic |  | 2021-12-06 | Water | Complete |  |  | 73265 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99928 | 0.00161 | 82891 | [25 empty] | 2021-11-17 | 5623 | Water CO2 Evolution | Completed
[168 empty] | 0.9979 | 0.00163 | 82892 | [25 empty] | 2021-11-17 | 3253 | Gas Splitting | Completed
[168 empty] | 0.99989 | 0.00159 | 82893 | [25 empty] | 2021-11-17 | 5415 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00158 | 82894 | [25 empty] | 2021-11-18 | 22948 | Graphite | Completed
[168 empty] | 1.0001 | 0.00162 | 82895 | [25 empty] | 2021-11-22 | 8569 | C13 Measurement | Completed
[196 empty] | 2021-11-24 | 24038 | AMS Submission Results | Completed
[196 empty] | 2021-11-24 | 28490 | Close Job | Completed
[7 empty] | 151.1 | 87.8361 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 41719 | 41719 | 41719 | 150 | 150 | 150 |  |  | jacob leath | -994.45 | 0.10 ‰ | 2021-11-25 | 2021-05-29 | 1 | 180 |  |  |  |  | 2.40‰ | 2.40‰ | 2.00‰ | 2.00‰ | 2.40 ‰ | 2.00 ‰ | AMS |  | 0.2‰ | C13 | No | -19 | -19.0008 | 0 | 0.8165 | -994.45 ‰ | -994.45 ‰ | 0.1 | 0.1 | -994.38 | 0.11 ‰ | -994.38 | 0.11 ‰ |  |  | 2021-11-26 | Jocelyn Turnbull | 9:22 | 0.0055525 | 0.0055525 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2.3 |  |  | 0.00555 | 0.0001 | 24111 | 2021-11-26 | 7:59 | Jeremy Thompson | 219492 |  |  |  |  | 45% | 5 | 26/11/21 | 09:22:07 | Jocelyn Turnbull | 0 |  | 0.56 | 0.01% | 20 | 5 | 7 | ... | 0.00534 | 0.0001 |  |  |  |  | 0.9894 | 0.9894 | Yes | Yes | [5 empty] | 0.00534 | 0.01‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00555 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00534 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-11-24 | jacob leath | 11:48 |  | Completed |  |  |  |  | 82996 | 2000000 |  |  | 82960 82961 82970 82971 82972 82973 | 82960~82961~82970~82971~82972~82973 | 3380 |  | 3380 3380 3380 3380 3380 3380 | 2021-11-26 | 8:16 | Jeremy Thompson |  | 0.81 mg |  | 0.81 | [5 empty] | 1.0008 | 0.00151 | 82960 |  | XCAMS | ST | 0% |  | TW3380 measurement CalAMS error adjustments: * no error adjustment needed primaries: * 6 large OxI cathodes, no failures secondaries: * FIRI-I 1x Result agrees with consensus IAEA-C2: Freshwater Travertine 1x Result agrees with consensus LAC: LAC1 coral 1x Result agrees with expected value LAA: LAA1 coral 1x Result agrees with expected value blanks: * 2 groups identified to blank-correct: Organics via AAA: 12x unknowns + 1x secondaries + 1x blanks 180 days, capturing 11x blanks Waters 14x unknowns + 3x secondaries (2 coral) + 1x blanks 180 days, capturing 11x blanks unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2021-12-29 | Water | Complete |  |  | 73305 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0009 | 0.00151 | 82961 | [25 empty] | 2021-11-23 | 5643 | Water CO2 Evolution | Completed
[168 empty] | 0.99808 | 0.00153 | 82970 | [25 empty] | 2021-11-23 | 3261 | Gas Splitting | Completed
[168 empty] | 1.0004 | 0.0015 | 82971 | [25 empty] | 2021-11-23 | 5426 | Choose Gas Splitting | Completed
[168 empty] | 1.0013 | 0.00144 | 82972 | [25 empty] | 2021-11-24 | 23006 | Graphite | Completed
[168 empty] | 0.99849 | 0.00146 | 82973 | [25 empty] | 2021-11-26 | 8595 | C13 Measurement | Completed
[196 empty] | 2021-11-26 | 24111 | AMS Submission Results | Completed
[196 empty] | 2021-11-29 | 28519 | Close Job | Completed
[7 empty] | 89.5 | 88.2171 | 0.0016 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 37634 | 37634 | 37634 | 136 | 135 | 135 |  |  | Taylor Ferrick | -990.77 | 0.16 ‰ | 2021-12-03 | 2021-06-06 | 3 | 180 |  |  |  |  | -0.03‰ | -0.03‰ | 2.00‰ | 2.00‰ | -0.03 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -990.77 ‰ | -990.77 ‰ | 0.16 | 0.16 | -990.65 | 0.16 ‰ | -990.65 | 0.16 ‰ |  |  | 2021-12-06 | Jocelyn Turnbull | 8:50 | 0.0092334 | 0.0092334 | 0.00016 | 0.00016 | 0.00016 | 0.00016 | 2.2 |  |  | 0.00923 | 0.00016 | 24158 | 2021-12-04 | 19:50 | Albert Zondervan | 219534 |  |  |  |  | 45% | 5 | 6/12/21 | 08:49:35 | Jocelyn Turnbull | 0 |  | 0.92 | 0.02% | 1 | 5 | 5 | ... | 0.00888 | 0.00015 |  |  |  |  | 0.99128 | 0.99128 | Yes | Yes | [5 empty] | 0.00888 | 0.01‰ | 0.00015 | 0.00‰ | Yes | 1 | 1 | 0.00063 | 0.00063 | 0.00923 | 0 | 0.00016 | 0 | 0 |  |  | 0 | 0 | 0.00888 | 0 | 0.00015 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2021-12-02 | Taylor Ferrick | 12:52 |  | Completed |  |  |  |  | 83048 | 2000000 |  |  | 83028 83029 83030 83055 83056 83057 | 83028~83029~83030~83055~83056~83057 | 3382 |  | 3382 3382 3382 3382 3382 3382 | 2021-12-04 | 20:09 | Albert Zondervan |  | 0.93 mg |  | 0.93 | [5 empty] | 1.001 | 0.00161 | 83028 |  | XCAMS | EA | 0% |  | TW3382 measurement * partially filled wheel CalAMS error adjustments: * error increased for TP83036,83031 to account for excess run-to-run variability primaries: * 6 large OxI cathodes, no failures secondaries: organics via AAA: FIRI-I result agrees with consensus value IAEA-C6 result agrees with consensus value waters: IAEA-C2 result deviates from consensus value by 2.5-sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses organics via CE: none in this wheel blanks: * 2 groups identified to blank-correct: organics via AAA: 13x unknowns + 2x secondaries + 1x blanks 180 days, capturing 13x blanks waters: 6x unknowns + 3x secondaries + 1x blanks 180 days, capturing 12x blanks organics via CE: 2x unknowns + 0x secondaries + 0x blanks 360 days, capturing 4x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-01-04 | Water |  |  |  | 73385 | Normal | 14047/11 |  |  | 2021-09-28 | 7531 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0013 | 0.00155 | 83029 | [25 empty] | 2021-11-29 | 5663 | Water CO2 Evolution | Completed
[168 empty] | 1.0004 | 0.00155 | 83030 | [25 empty] | 2021-11-30 | 23058 | Graphite | Completed
[168 empty] | 0.9985 | 0.00152 | 83055 | [25 empty] | 2021-12-01 | 8625 | C13 Measurement | Completed
[168 empty] | 1.0006 | 0.00152 | 83056 | [25 empty] | 2021-12-06 | 24158 | AMS Submission Results | Completed
[168 empty] | 0.99841 | 0.00154 | 83057 | [25 empty] | 2021-12-08 | 28614 | Close Job | Completed
[7 empty] | 110.1 | 87.9711 | 0.0012 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 39909 | 39908 | 39908 | 144 | 144 | 144 |  |  | Taylor Ferrick | -993.04 | 0.12 ‰ | 2022-02-04 | 2021-08-08 | 2 | 180 |  |  |  |  | 4.11‰ | 4.11‰ | 2.00‰ | 2.00‰ | 4.11 ‰ | 2.00 ‰ | AMS | 2.49‰ | 0.2‰ | C13 | No | -19 | -19.0012 | 0 | 0.8165 | -993.04 ‰ | -993.04 ‰ | 0.12 | 0.12 | -992.96 | 0.13 ‰ | -992.96 | 0.13 ‰ |  |  | 2022-02-08 | Jocelyn Turnbull | 10:18 | 0.0069562 | 0.0069562 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 1.9 |  |  | 0.00696 | 0.00012 | 24295 | 2022-02-08 | 8:24 | Jeremy Thompson | 219613 |  |  |  |  | 45% | 5 | 8/02/22 | 10:17:31 | Jocelyn Turnbull | 0 |  | 0.7 | 0.01% | 1 | 5 | 5 | ... | 0.00669 | 0.00012 |  |  |  |  | 0.99262 | 0.99262 | Yes | Yes | [5 empty] | 0.00669 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00696 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00669 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-02-02 | Taylor Ferrick | 12:27 |  | Completed |  |  |  |  | 83189 | 2000000 |  |  | 83113 83114 83115 83116 83158 83169 | 83113~83114~83115~83116~83158~83169 | 3385 |  | 3385 3385 3385 3385 3385 3385 | 2022-02-08 | 8:41 | Jeremy Thompson |  | 1.12 mg |  | 1.12 | [5 empty] | 0.99812 | 0.00158 | 83113 |  | XCAMS | FLASK | 0% |  | TW3385 measurement *all large targets CalAMS error adjustments: * primaries: * 6 large OxI cathodes, no failures secondaries: Waters: IAEA-C2 result deviates from consensus value by 3 sigma LAA1 result agrees with mean of previous analyses LAC1 result agrees with mean of previous analyses blanks: * 2 groups identified to blank-correct: Waters: 8x unknowns + 3x secondaries + 1x blanks 180 days, capturing 10x blanks C02 from CO: 20x unknowns + 0x secondaries + 4x blanks Speicfic blanks chosen to repeat past blank correction approach(TP: 68062, 68063, 68064, 68065 from TW3007) unknowns: * no failures | 2.5 | RRL | [5 empty] | Background Inorganic |  | 2022-01-13 | Water | Not Complete |  | New batch made for Marble water standard IAEA-C1. Has been crushed into powder instead of smaller chunks. | 73498 | Normal | 14047/11 |  |  | 2021-11-22 | 7706 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99944 | 0.00152 | 83114 | [25 empty] | 2022-02-01 | 5696 | Water CO2 Evolution | Completed
[168 empty] | 1.0013 | 0.0015 | 83115 | [25 empty] | 2022-02-02 | 23199 | Graphite | Completed
[168 empty] | 1.0005 | 0.00155 | 83116 | [25 empty] | 2022-02-04 | 8667 | C13 Measurement | Completed
[168 empty] | 0.9995 | 0.00155 | 83158 | [25 empty] | 2022-02-08 | 24295 | AMS Submission Results | Completed
[168 empty] | 1.001 | 0.00148 | 83169 | [25 empty] | 2022-02-11 | 28819 | Close Job | Completed
[7 empty] | 111.9 | 87.5215 | 0.0009 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 42152 | 42152 | 42152 | 175 | 174 | 174 |  |  | Taylor Ferrick | -994.74 | 0.11 ‰ | 2022-05-02 | 2021-11-03 | 4 | 180 |  |  |  |  | 4.36‰ | 4.36‰ | 2.00‰ | 2.00‰ | 4.36 ‰ | 2.00 ‰ | AMS | 2.31‰ | 0.2‰ | C13 | No | -19 | -19.0005 | 0 | 0.75593 | -994.74 ‰ | -994.74 ‰ | 0.11 | 0.11 | -994.67 | 0.12 ‰ | -994.67 | 0.12 ‰ |  |  | 2022-05-04 | Cathy Ginnane | 23:03 | 0.0052614 | 0.0052614 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00526 | 0.00011 | 24991 | 2022-05-03 | 16:31 | Albert Zondervan | 219488 |  |  |  |  | 45% | 5 | 4/05/22 | 23:02:43 | Cathy Ginnane | 0 |  | 0.53 | 0.01% | 15 | 5 | 5 | ... | 0.00506 | 0.00011 |  |  |  |  | 0.98779 | 0.98779 | Yes | Yes | [5 empty] | 0.00506 | 0.01‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00058 | 0.00058 | 0.00526 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00506 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-04-29 | Taylor Ferrick | 12:03 |  | Completed |  |  |  |  | 83859 | 2000000 |  |  | 83786 83831 83832 83840 83843 83844 83845 | 83786~83831~83832~83840~83843~83844~83845 | 3403 |  | 3403 3403 3403 3403 3403 3403 3403 | 2022-05-04 | 9:10 | Albert Zondervan |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99936 | 0.00154 | 83786 |  | XCAMS | ST | 0% |  | TW3403 measurement: * regular-size samples only * EA tests, GIRI-D, and usual suite of unknowns CalAMS error adjustments: * none needed primaries: * OxI cathodes: 1x EA + 6x ST all 7x accepted secondaries: * carbonates: 2x corals agree with expectation value * waters: 2x corals agree with expectation value * organics via CE: 1x FIRI-D agrees with AMS consensus value * IAEA-C6 for EA test: 2x IAEA-C6 agree with each other blanks: * 4 groups identified to blank-correct: carbonates: 12x unknowns + 2x secondaries + 1x blanks 180 days, capturing 6x blanks waters: 9x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks organics via CE: 0x unknowns + 1x secondaries + 1x blanks 180 days, capturing 5x blanks organics via AAA: 1x unknowns + 2x secondaries + 0x blanks 180 days, capturing 12x blanks unknowns: * no failures | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2021-12-29 | Water | Not Complete |  | job for new treatment of marble std for water line | 74109 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99956 | 0.0015 | 83831 | [25 empty] | 2022-04-27 | 5836 | Water CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00157 | 83832 | [25 empty] | 2022-04-28 | 23869 | Graphite | Completed
[168 empty] | 0.99867 | 0.00152 | 83840 | [25 empty] | 2022-04-29 | 8881 | C13 Measurement | Completed
[168 empty] | 0.9983 | 0.00154 | 83843 | [25 empty] | 2022-05-04 | 24991 | AMS Submission Results | Completed
[168 empty] | 1.0009 | 0.00157 | 83844 | [25 empty] | 2022-05-06 | 29427 | Close Job | Completed
[168 empty] | 1.0024 | 0.00155 | 83845
[7 empty] | 133.3 | 87.6495 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45883 | 45883 | 45883 | 227 | 227 | 227 |  |  | jacob leath | -996.69 | 0.09 ‰ | 2022-05-09 | 2021-11-10 | 3 | 180 |  |  |  |  | 2.89‰ | 2.89‰ | 2.00‰ | 2.00‰ | 2.89 ‰ | 2.00 ‰ | AMS | 2.39‰ | 0.2‰ | C13 | No | -19 | -19.0015 | 0 | 0.8165 | -996.69 ‰ | -996.69 ‰ | 0.09 | 0.09 | -996.65 | 0.09 ‰ | -996.65 | 0.09 ‰ |  |  | 2022-05-12 | Cathy Ginnane | 10:02 | 0.0033065 | 0.0033065 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 1.9 |  |  | 0.00331 | 9e-05 | 25102 | 2022-05-11 | 9:35 | Albert Zondervan | 221031 |  |  |  |  | 45% | 5 | 12/05/22 | 10:02:00 | Cathy Ginnane | 0 |  | 0.33 | 0.01% | 27 | 5 | 5 | ... | 0.00318 | 9e-05 |  |  |  |  | 0.98778 | 0.98778 | Yes | Yes | [5 empty] | 0.00318 | 0.00‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00057 | 0.00057 | 0.00331 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00318 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-05-06 | jacob leath | 13:06 |  | Completed |  |  |  |  | 83954 | 2000000 |  |  | 83921 83930 83931 83932 83933 83934 | 83921~83930~83931~83932~83933~83934 | 3405 |  | 3405 3405 3405 3405 3405 3405 | 2022-05-11 | 11:05 | Albert Zondervan |  | 1.03 mg |  | 1.03 | [5 empty] | 0.99951 | 0.00134 | 83921 |  | XCAMS | ST | 0% |  | TW3405 measurement: * large waters + few RP samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value blanks: * 2 groups identified to blank-correct: waters: 26x unknowns + 1x secondaries + 1x blanks 180 days, capturing 7x blanks sediments via RP: to be blank-corrected still unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-06-23 | Water |  |  |  | 74265 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99765 | 0.00144 | 83930 | [25 empty] | 2022-05-05 | 5871 | Water CO2 Evolution | Completed
[168 empty] | 1.0011 | 0.00144 | 83931 | [25 empty] | 2022-05-06 | 23964 | Graphite | Completed
[168 empty] | 1.0008 | 0.00139 | 83932 | [25 empty] | 2022-05-09 | 8967 | C13 Measurement | Completed
[168 empty] | 1.0009 | 0.00136 | 83933 | [25 empty] | 2022-05-12 | 25102 | AMS Submission Results | Completed
[168 empty] | 0.99993 | 0.00142 | 83934 | [25 empty] | 2022-05-18 | 29581 | Close Job | Completed
[7 empty] | 146.7 | 87.3751 | 0.001 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40918 | 40918 | 40918 | 123 | 122 | 122 |  |  | jacob leath | -993.87 | 0.09 ‰ | 2022-06-09 | 2021-12-11 | 13 | 180 |  |  |  |  | 6.06‰ | 6.06‰ | 2.00‰ | 2.00‰ | 6.06 ‰ | 2.00 ‰ | AMS | 2.33‰ | 0.2‰ | C13 | No | -19 | -19.0021 | 0.82 | 0.8165 | -993.87 ‰ | -993.87 ‰ | 0.09 | 0.09 | -993.79 | 0.09 ‰ | -993.79 | 0.09 ‰ |  |  | 2022-06-16 | Jocelyn Turnbull | 9:45 | 0.0061348 | 0.0061348 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2.3 |  |  | 0.00613 | 9e-05 | 25270 | 2022-06-15 | 19:42 | Albert Zondervan | 221035 |  |  |  |  | 45% | 5 | 16/06/22 | 09:45:01 | Jocelyn Turnbull | 0 |  | 0.61 | 0.01% | 1 | 5 | 7 | ... | 0.0059 | 9e-05 |  |  |  |  | 0.9878 | 0.9878 | Yes | Yes | [5 empty] | 0.0059 | 0.01‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00055 | 0.00055 | 0.00613 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.0059 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-06-08 | jacob leath | 11:59 |  | Completed |  |  |  |  | 84143 | 2000000 |  |  | 84044 84045 84195 84196 84197 84198 | 84044~84045~84195~84196~84197~84198 | 3410 |  | 3410 3410 3410 3410 3410 3410 | 2022-06-15 | 19:45 | Albert Zondervan |  | 0.89 mg |  | 0.89 | [5 empty] | 0.99934 | 0.00126 | 84044 |  | XCAMS | ST | 0% |  | TW3410 measurement: * large waters * unexpected inter- and intra-cathode variability: only in beam current for unclear reason, probably related to SNICS condition no additional variability in ratio-to-standard * (del13C_AMS - del13C_IRMS) difference correlates with beam current * SNICS is being tested before next batch of sample graphites is measured CalAMS error adjustments: * none needed: good inter-OxI repeatability, reasonable intra-cathode (run-to-run) repeatability * data from first two rotations ignored, to cull most runs with low current primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value 1x LAC1 coral agrees with expectation value blanks: * 1 groups identified to blank-correct: waters: 28x unknowns + 2x secondaries + 1x blanks 180 days, capturing 4x blanks unknowns: * cat#18 (TP84174) 1st Q-flag set: failed in loading, not measured * cat#16 (TP84188) 2nd Q-flag set: foreign object near surface | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-06-23 | Water |  |  |  | 74471 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.9992 | 0.00127 | 84045 | [25 empty] | 2022-05-26 | 5933 | Water CO2 Evolution | Completed
[168 empty] | 1.001 | 0.00145 | 84195 | [25 empty] | 2022-05-26 | 3343 | Gas Splitting | Completed
[168 empty] | 1.0001 | 0.00146 | 84196 | [25 empty] | 2022-05-26 | 5584 | Choose Gas Splitting | Completed
[168 empty] | 0.99937 | 0.00128 | 84197 | [25 empty] | 2022-05-27 | 24153 | Graphite | Completed
[168 empty] | 1.0016 | 0.00146 | 84198 | [25 empty] | 2022-05-27 | 9043 | C13 Measurement | Completed
[196 empty] | 2022-06-16 | 25270 | AMS Submission Results | Completed
[196 empty] | 2022-06-23 | 29822 | Close Job | Completed
[7 empty] | 113.2 | 88.0196 | 0.0013 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 38842 | 38842 | 38842 | 147 | 147 | 147 |  |  | Taylor Ferrick | -992.06 | 0.15 ‰ | 2022-06-20 | 2021-12-22 | 11 | 180 |  |  |  |  | 4.19‰ | 4.19‰ | 2.00‰ | 2.00‰ | 4.19 ‰ | 2.00 ‰ | AMS | 2.29‰ | 0.2‰ | C13 | No | -19 | -19.0016 | 0.82 | 0.8165 | -992.06 ‰ | -992.06 ‰ | 0.15 | 0.15 | -991.96 | 0.15 ‰ | -991.96 | 0.15 ‰ |  |  | 2022-06-21 | Jocelyn Turnbull | 17:01 | 0.007944 | 0.007944 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 1.9 |  |  | 0.00794 | 0.00015 | 25304 | 2022-06-21 | 16:08 | Albert Zondervan | 221105 |  |  |  |  | 45% | 5 | 21/06/22 | 17:01:07 | Jocelyn Turnbull | 0 |  | 0.79 | 0.01% | 1 | 5 | 5 | ... | 0.00764 | 0.00014 |  |  |  |  | 0.99324 | 0.99324 | Yes | Yes | [5 empty] | 0.00764 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00066 | 0.00066 | 0.00794 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00764 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-06-17 | Taylor Ferrick | 13:55 |  | Completed |  |  |  |  | 84216 | 2000000 |  |  | 84225 84226 84227 84228 84229 84239 | 84225~84226~84227~84228~84229~84239 | 3411 |  | 3411 3411 3411 3411 3411 3411 | 2022-06-21 | 16:17 | Albert Zondervan |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99919 | 0.00162 | 84225 |  | XCAMS | ST | 0% |  | TW3411 measurement: * large waters + large organics via AAA * Cs delivery tube of SNICS has been replaced and position of bandheater on Cs reservoir has been raised * although SNICS better under control than during TW3410, still some undesirable inter-cathode variability: only in beam current no additional variability in ratio-to-standard * (del13C_AMS - del13C_IRMS) difference correlates with beam current * MAN (high-prec) will be run after REP1 (reg-prec) CalAMS error adjustments: * error multiplier of 1.05 to account for (small) excess run-to-run variability primaries: * OxI cathodes: 6x ST, accepted secondaries: * waters: 1x LAA1 coral agrees with expectation value 1x LAC1 coral agrees with expectation value * organics 1x IAEA-C6 sucrose agrees with expectation value blanks: * 1 groups identified to blank-correct: waters: 15x unknowns + 2x secondaries + 1x blanks 180 days, capturing 4x blanks organics: 13x unknowns + 1x secondaries + 1x blanks 180 days, capturing 4x blanks unknowns: * all accepted | 2.3 | RRL | [5 empty] | Background Inorganic |  | 2022-07-01 | Water | Not Complete |  | Untested (new) water vessel used. | 74510 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0011 | 0.00163 | 84226 | [25 empty] | 2022-06-08 | 5967 | Water CO2 Evolution | Completed
[168 empty] | 0.99811 | 0.00163 | 84227 | [25 empty] | 2022-06-09 | 24226 | Graphite | Completed
[168 empty] | 0.99967 | 0.00165 | 84228 | [25 empty] | 2022-06-10 | 9082 | C13 Measurement | Combined
[168 empty] | 1.0006 | 0.00166 | 84229 | [25 empty] | 2022-06-21 | 25304 | AMS Submission Results | Completed
[168 empty] | 1.0012 | 0.00156 | 84239 | [25 empty] | 2022-06-23 | 29826 | Close Job | Completed
[7 empty] | 57.9 | 88.2774 | 0.0004 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 47946 | 47945 | 47945 | 261 | 261 | 261 |  |  | Taylor Ferrick | -997.44 | 0.08 ‰ | 2022-07-15 | 2022-01-16 | 2 | 180 |  |  |  |  | 4.35‰ | 4.35‰ | 2.00‰ | 2.00‰ | 4.35 ‰ | 2.00 ‰ | AMS | 2.65‰ | 0.2‰ | C13 | No | -19 | -19.0018 | 0 | 0.8165 | -997.44 ‰ | -997.44 ‰ | 0.08 | 0.08 | -997.41 | 0.08 ‰ | -997.41 | 0.08 ‰ |  |  | 2022-07-20 | Jocelyn Turnbull | 10:02 | 0.0025579 | 0.0025579 | 8e-05 | 8e-05 | 8e-05 | 8e-05 | 2.1 |  |  | 0.00256 | 8e-05 | 25489 | 2022-07-18 | 14:42 | Christian Lewis | 221628 |  |  |  |  | 45% | 5 | 20/07/22 | 10:02:10 | Jocelyn Turnbull | 0 |  | 0.26 | 0.01% | 22 | 5 | 5 | ... | 0.00246 | 8e-05 |  |  |  |  | 0.99631 | 0.99631 | Yes | Yes | [5 empty] | 0.00246 | 0.00‰ | 8e-05 | 0.00‰ | Yes | 1 | 1 | 0.00062 | 0.00062 | 0.00256 | 0 | 8e-05 | 0 | 0 |  |  | 0 | 0 | 0.00246 | 0 | 8e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-14 | Taylor Ferrick | 12:42 |  | Completed |  |  |  |  | 84472 | 2000000 |  |  | 84462 84463 84487 84488 84489 84490 | 84462~84463~84487~84488~84489~84490 | 3416 |  | 3416 3416 3416 3416 3416 3416 | 2022-07-20 | 9:40 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 1.0001 | 0.00148 | 84462 |  | XCAMS | ST | 0% |  | TW3416 measurement: * large CalAMS error adjustments: * none needed Primaries: * OxI cathodes: 6x ST, accepted Secondaries: * organics via CE 1x FIRI-D agrees with expectation value * Waters 1x LAC1 coral agrees with expectation value 1x LAA1 coral agrees with expectation value Blanks: * 3 groups identified to blank-correct: organics via AAA: 2x unknowns + 0x secondaries + 0x blanks 180 days, capturing 12x blanks organics via CE: 10x unknowns + 1x secondaries + 1x blanks 180 days, capturing 9x blanks Waters CO2 extraction 17x unknowns + 2x secondaries + 1x blanks 180 days, capturing 6x blanks Unknowns: * all accepted | 2.7 | RRL | [5 empty] | Background Inorganic |  | 2022-08-22 | Water |  |  |  | 74735 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 1.0008 | 0.00144 | 84463 | [25 empty] | 2022-07-12 | 6000 | Water CO2 Evolution | Completed
[168 empty] | 1.0014 | 0.00161 | 84487 | [25 empty] | 2022-07-13 | 24482 | Graphite | Completed
[168 empty] | 0.99928 | 0.00158 | 84488 | [25 empty] | 2022-07-15 | 9244 | C13 Measurement | Completed
[168 empty] | 1.0001 | 0.00154 | 84489 | [25 empty] | 2022-07-20 | 25489 | AMS Submission Results | Completed
[168 empty] | 0.99829 | 0.00154 | 84490 | [25 empty] | 2022-07-21 | 30268 | Close Job | Completed
[7 empty] | 94 | 88.107 | 0.0006 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 45229 | 45228 | 45228 | 233 | 232 | 232 |  |  | Taylor Ferrick | -996.41 | 0.10 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.06‰ | 2.06‰ | 2.00‰ | 2.00‰ | 2.06 ‰ | 2.00 ‰ | AMS | 2.84‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -996.41 ‰ | -996.41 ‰ | 0.1 | 0.1 | -996.37 | 0.11 ‰ | -996.37 | 0.11 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0035873 | 0.0035873 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 2 |  |  | 0.00359 | 0.0001 | 25689 | 2022-08-05 | 10:56 | Christian Lewis | 221951 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:44 | Jocelyn Turnbull | 0 |  | 0.36 | 0.01% | 8 | 5 | 5 | ... | 0.00345 | 0.0001 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00345 | 0.00‰ | 0.0001 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00359 | 0 | 0.0001 | 0 | 0 |  |  | 0 | 0 | 0.00345 | 0 | 0.0001 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84630 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 1.02 mg |  | 1.02 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Inorganic |  | 2022-09-14 | Water |  |  | Normal water blank | 74938 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [25 empty] | 2022-07-27 | 6046 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [25 empty] | 2022-07-28 | 24640 | Graphite | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [25 empty] | 2022-07-29 | 9352 | C13 Measurement | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [25 empty] | 2022-08-05 | 25689 | AMS Submission Results | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [25 empty] | 2022-08-05 | 30460 | Close Job | Completed
[7 empty] | 94.6 | 88.0729 | 0.0011 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 40190 | 40189 | 40189 | 174 | 174 | 174 | 40017.7 | 40365.9 | Taylor Ferrick | -993.28 | 0.15 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.45‰ | 2.45‰ | 2.00‰ | 2.00‰ | 2.45 ‰ | 2.00 ‰ | AMS | 2.76‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -993.28 ‰ | -993.28 ‰ | 0.15 | 0.15 | -993.2 | 0.15 ‰ | -993.2 | 0.15 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0067171 | 0.0067171 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 2.1 |  |  | 0.00672 | 0.00015 | 25693 | 2022-08-05 | 10:56 | Christian Lewis | 221623 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:45 | Jocelyn Turnbull | 0 |  | 0.67 | 0.01% | 15 | 5 | 6 | ... | 0.00646 | 0.00014 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00646 | 0.01‰ | 0.00014 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00672 | 0 | 0.00015 | 0 | 0 |  |  | 0 | 0 | 0.00646 | 0 | 0.00014 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84625 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.92 mg |  | 0.92 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water |  |  | RECOMBUST!- Test of recombusted water blank MN not completed ST process to make additional jobs for future wheels | 74939 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [25 empty] | 2022-07-12 | 6003 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [25 empty] | 2022-07-27 | 5297 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [25 empty] | 2022-07-28 | 24635 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [25 empty] | 2022-07-29 | 9347 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [25 empty] | 2022-08-05 | 25693 | AMS Submission Results | Completed
[196 empty] | 2022-08-05 | 30445 | Close Job | Completed
[197 empty] | 13978 | Invoice | Started
[7 empty] | 96.5 | 88.0947 | 0.0007 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 44327 | 44326 | 44326 | 229 | 228 | 228 | 44100.9 | 44558.9 | Taylor Ferrick | -995.99 | 0.11 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.20‰ | 2.20‰ | 2.00‰ | 2.00‰ | 2.20 ‰ | 2.00 ‰ | AMS | 2.79‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -995.99 ‰ | -995.99 ‰ | 0.11 | 0.11 | -995.94 | 0.12 ‰ | -995.94 | 0.12 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0040136 | 0.0040136 | 0.00011 | 0.00011 | 0.00011 | 0.00011 | 2 |  |  | 0.00401 | 0.00011 | 25694 | 2022-08-05 | 10:56 | Christian Lewis | 221625 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:48 | Jocelyn Turnbull | 0 |  | 0.4 | 0.01% | 25 | 5 | 6 | ... | 0.00386 | 0.00011 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00386 | 0.00‰ | 0.00011 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00401 | 0 | 0.00011 | 0 | 0 |  |  | 0 | 0 | 0.00386 | 0 | 0.00011 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84627 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water |  |  | RECOMBUST!- Test of recombusted water blank | 74940 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [25 empty] | 2022-07-12 | 6005 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [25 empty] | 2022-07-27 | 5299 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [25 empty] | 2022-07-28 | 24637 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [25 empty] | 2022-07-29 | 9349 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [25 empty] | 2022-08-05 | 25694 | AMS Submission Results | Completed
[196 empty] | 2022-08-05 | 30447 | Close Job | Completed
[197 empty] | 13980 | Invoice | Started
[7 empty] | 97.8 | 88.0621 | 0.0008 | RRL-Test-LG | Background Test | 1950-01-01 | 1950-01-01 |  |  | 42626 | 42626 | 42626 | 202 | 202 | 202 | 42426.6 | 42830.9 | Taylor Ferrick | -995.04 | 0.12 ‰ | 2022-08-04 | 2022-02-05 | 7 | 180 |  |  |  |  | 2.57‰ | 2.57‰ | 2.00‰ | 2.00‰ | 2.57 ‰ | 2.00 ‰ | AMS | 2.81‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.8165 | -995.04 ‰ | -995.04 ‰ | 0.12 | 0.12 | -994.98 | 0.13 ‰ | -994.98 | 0.13 ‰ |  |  | 2022-08-05 | Jocelyn Turnbull | 12:55 | 0.0049598 | 0.0049598 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2 |  |  | 0.00496 | 0.00012 | 25695 | 2022-08-05 | 10:56 | Christian Lewis | 221624 |  |  |  |  | 45% | 5 | 5/08/22 | 12:54:49 | Jocelyn Turnbull | 0 |  | 0.5 | 0.01% | 34 | 5 | 6 | ... | 0.00477 | 0.00012 |  |  |  |  | 0.99211 | 0.99211 | Yes | Yes | [5 empty] | 0.00477 | 0.00‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00073 | 0.00073 | 0.00496 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00477 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-07-29 | Taylor Ferrick | 12:01 |  | Completed |  |  |  |  | 84626 | 2000000 |  |  | 84537 84538 84632 84633 84634 84635 | 84537~84538~84632~84633~84634~84635 | 3421 |  | 3421 3421 3421 3421 3421 3421 | 2022-08-05 | 11:18 | Christian Lewis |  | 0.98 mg |  | 0.98 | [5 empty] | 0.99904 | 0.00178 | 84537 |  | XCAMS | ST | 0% |  | TW3421 measurement: * regular-size samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 6x ST, all accepted secondaries: * LAA1 coral. Within error of expected age. blanks: * 1 groups identified to blank-correct: Waters corrected via Carrara Marble WATER LINE 14047/11 24x unknowns + 1x secondaries + 4x blanks 180 days, capturing 6x blanks This wheel also contains test of recombusted waters and carbonate blanks. Combusted marble blanks look higher than the normal marble blanks. The combusted water blanks test look within the range of typical water blanks. These recombusted water blanks/tests are NOT currently used in this wheel's blank correction. | 2.8 | RRL | [5 empty] | Background Test |  | 2022-08-22 | Water | Complete |  | RECOMBUST!- Test of recombusted water blank | 74941 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99884 | 0.00182 | 84538 | [25 empty] | 2022-07-12 | 6004 | Water CO2 Evolution | Completed
[168 empty] | 1.0016 | 0.0018 | 84632 | [25 empty] | 2022-07-27 | 5298 | Sealed Tube Combustion | Completed
[168 empty] | 1.0006 | 0.00169 | 84633 | [25 empty] | 2022-07-28 | 24636 | Graphite | Completed
[168 empty] | 1.0031 | 0.00177 | 84634 | [25 empty] | 2022-07-29 | 9348 | C13 Measurement | Completed
[168 empty] | 0.99659 | 0.00182 | 84635 | [25 empty] | 2022-08-05 | 25695 | AMS Submission Results | Completed
[196 empty] | 2022-08-05 | 30446 | Close Job | Completed
[197 empty] | 13979 | Invoice | Started
[7 empty] | 138.2 | 87.8261 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40341 | 40340 | 40340 | 152 | 152 | 152 |  |  | Taylor Ferrick | -993.41 | 0.12 ‰ | 2022-08-16 | 2022-02-17 | 13 | 180 |  |  |  |  | 2.35‰ | 2.35‰ | 2.00‰ | 2.00‰ | 2.35 ‰ | 2.00 ‰ | AMS | 3.20‰ | 0.2‰ | C13 | No | -19 | -19.0004 | 0 | 0.8165 | -993.41 ‰ | -993.41 ‰ | 0.12 | 0.12 | -993.33 | 0.13 ‰ | -993.33 | 0.13 ‰ |  |  | 2022-08-18 | Cathy Ginnane | 13:32 | 0.0065923 | 0.0065923 | 0.00012 | 0.00012 | 0.00012 | 0.00012 | 2.4 |  |  | 0.00659 | 0.00012 | 25803 | 2022-08-17 | 10:13 | Christian Lewis | 221896 |  |  |  |  | 45% | 5 | 18/08/22 | 13:31:53 | Cathy Ginnane | 0 |  | 0.66 | 0.01% | 1 | 5 | 5 | ... | 0.00634 | 0.00012 |  |  |  |  | 0.98924 | 0.98924 | Yes | Yes | [5 empty] | 0.00634 | 0.01‰ | 0.00012 | 0.00‰ | Yes | 1 | 1 | 0.00061 | 0.00061 | 0.00659 | 0 | 0.00012 | 0 | 0 |  |  | 0 | 0 | 0.00634 | 0 | 0.00012 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-08-15 | Taylor Ferrick | 8:17 |  | Completed |  |  |  |  | 84676 | 2000000 |  |  | 84735 84736 84748 84752 84753 84754 | 84735~84736~84748~84752~84753~84754 | 3424 |  | 3424 3424 3424 3424 3424 3424 | 2022-08-17 | 10:49 | Christian Lewis |  | 0.81 mg |  | 0.81 | [5 empty] | 1.0014 | 0.00149 | 84735 |  | XCAMS | ST | 0% |  | TW3424 measurement: * large CalAMS error adjustments: * none needed Primaries: * OxI cathodes: 6x ST, accepted Secondaries: * organics via CE 1x FIRI-D agrees with expectation value * Waters 1x LAC1 coral agrees with expectation value 1x LAA1 coral agrees with expectation value Blanks: * 3 groups identified to blank-correct: organics via AAA: 7x unknowns + 0x secondaries + 1x blanks 180 days, capturing 16x blanks organics via CE: 9x unknowns + 1x secondaries + 1x blanks 180 days, capturing 10x blanks (this correction has been updated from 11 to 10 blanks, to eliminate TP83997 (TW3408) which had a quality flag to mark it as an outlier; CG; 23/8/22) Waters CO2 extraction 10x unknowns + 2x secondaries + 1x blanks 180 days, capturing 7x blanks Unknowns: * all accepted | 3.2 | RRL | [5 empty] | Background Inorganic |  | 2022-09-13 | Water | Complete |  | C13 0.8 per mil out | 75023 | Normal | 14047/11 |  |  | 2022-04-27 | 8200 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99971 | 0.00151 | 84736 | [25 empty] | 2022-08-02 | 6063 | Water CO2 Evolution | Completed
[168 empty] | 0.9999 | 0.00149 | 84748 | [25 empty] | 2022-08-03 | 24686 | Graphite | Completed
[168 empty] | 1.002 | 0.00149 | 84752 | [25 empty] | 2022-08-05 | 9393 | C13 Measurement | Completed
[168 empty] | 0.99879 | 0.00151 | 84753 | [25 empty] | 2022-08-18 | 25803 | AMS Submission Results | Completed
[168 empty] | 0.99821 | 0.00148 | 84754 | [27 empty] | Close Job
[7 empty] | 118.5 | 87.7108 | 0.0011 | RRL-CBIn-LG | Background Inorganic | 1950-01-01 | 1950-01-01 |  |  | 40328 | 40327 | 40327 | 114 | 113 | 113 |  |  | Taylor Ferrick | -993.4 | 0.09 ‰ | 2022-08-26 | 2022-02-27 | 3 | 180 |  |  |  |  | 5.07‰ | 5.07‰ | 2.00‰ | 2.00‰ | 5.07 ‰ | 2.00 ‰ | AMS | 2.36‰ | 0.2‰ | C13 | No | -19 | -19.0007 | 0 | 0.75593 | -993.40 ‰ | -993.40 ‰ | 0.09 | 0.09 | -993.32 | 0.09 ‰ | -993.32 | 0.09 ‰ |  |  | 2022-08-30 | Jocelyn Turnbull | 17:07 | 0.0066027 | 0.0066027 | 9e-05 | 9e-05 | 9e-05 | 9e-05 | 2 |  |  | 0.0066 | 9e-05 | 25963 | 2022-08-29 | 11:03 | Christian Lewis | 221989 |  |  |  |  | 45% | 5 | 30/08/22 | 17:07:09 | Jocelyn Turnbull | 0 |  | 0.66 | 0.01% | 31 | 5 | 5 | ... | 0.00635 | 9e-05 |  |  |  |  | 0.99062 | 0.99062 | Yes | Yes | [5 empty] | 0.00635 | 0.01‰ | 9e-05 | 0.00‰ | Yes | 1 | 1 | 0.00048 | 0.00048 | 0.0066 | 0 | 9e-05 | 0 | 0 |  |  | 0 | 0 | 0.00635 | 0 | 9e-05 | 0 | Regular | Yes | 0.9878 | 0.95 | 1 | 2022-08-25 | Taylor Ferrick | 12:48 |  | Completed |  |  |  |  | 84897 | 2000000 |  |  | 84694 84695 84709 84717 84718 84908 84920 | 84694~84695~84709~84717~84718~84908~84920 | 3428 |  | 3428 3428 3428 3428 3428 3428 3428 | 2022-08-29 | 12:00 | Christian Lewis |  | 1.09 mg |  | 1.09 | [5 empty] | 1.0015 | 0.0013 | 84694 |  | XCAMS | FLASK | 0% |  | TW3428 measurement: Water samples Air samples CalAMS error adjustments: * none needed primaries: * OxI cathodes: 7x ST All accepted; however, one is very low. TP 84709 has RTS of 0.99591, which is on the cusp of 3-sigma for this set of standards. secondaries: * Inorganics via water evolution: Both agree with expectation values blanks: * 1 groups identified to blank-correct: Inorganics via CO2 evolution on waterline unknowns: * no failures | 2.4 | RRL | [5 empty] | Background Inorganic |  | 2022-09-22 | Water |  |  | New job with refreshed IAEA-C1 material for waterline blanks. | 75173 | Normal | 14047/11 |  |  | 2022-08-03 | 8437 | Physical Pretreatment | Completed |  |  | Slab of freshly cut Carrara Marble, IMEG, Vareggio, Italy | Use this R 14047/11 for CO2 produced on water line, to match with water samples. Use R14047/1 for CO2 produced on carbonate line for carbonate samples. |  | 2.5 | 0.02 ± 0.02 pMC |  | IAEA-C1: Carrara Marble WATER LINE | IAEA-C1 | Labelled as water since this R is for processing on water line, even though it is a carbonate
[168 empty] | 0.99988 | 0.00128 | 84695 | [25 empty] | 2022-08-22 | 6078 | Water CO2 Evolution | Completed
[168 empty] | 0.99591 | 0.00123 | 84709 | [25 empty] | 2022-08-23 | 24907 | Graphite | Completed
[168 empty] | 1.0014 | 0.00129 | 84717 | [25 empty] | 2022-08-26 | 9489 | C13 Measurement | Completed
[168 empty] | 0.99936 | 0.00127 | 84718 | [25 empty] | 2022-08-30 | 25963 | AMS Submission Results | Completed
[168 empty] | 1.0028 | 0.00126 | 84908 | [27 empty] | Close Job
[168 empty] | 0.99947 | 0.00124 | 84920

## Sheet1
Biobased content
Biobased content error
Biobased content multiplier
Blank rts assigned
Blank rts err assigned
Blanks fm av err Calc
Blanks fm average Calc
C12
C12C13 ratio
C14C13 ratio
Category Field
Category In Calculation
Collection Date
Collection Date Saved
Colour Description
Comment
CRA
CRA Report
CRA Report Saved
CRA_Error
CRA_Error Report
CRA_Error Report Saved
CRA_OneSigmaMinus
CRA_OneSigmaPlus
Creation User
D14C
D14C_Error
Date Run
DateLastRun
DaysAsGraphite
DaysFromCurrentRun
DCC
DCC_error
DCCstd
DCCstd_error
delta13C_AMS
delta13C_AMS Saved
delta13C_AMS_Error
delta13C_AMS_Error Stored
delta13C_In_Calculation
delta13C_In_Calculation_Error
delta13C_In_Calculation_Source
delta13C_IRMS
delta13C_IRMS_Error
delta13C_IRMS_Source
delta13C_std_av_Match
delta13C_stds_av
delta13C_stds_av_Calculation
delta13C_stds_av_error
delta13C_stds_av_error_Calculation
DELTA14C
DELTA14C Saved
DELTA14C_Error
DELTA14C_Error Saved
deltalower14C
deltalower14C_Error
dlower14C
dlower14C_Error
Email Body
Email sender
End Date
End Operator
End Time
F_corrected_normed
F_corrected_normed Saved
F_corrected_normed_error
F_corrected_normed_error Saved
F_corrected_normed_error_no_WTW_Error
F_corrected_normed_error_no_WTW_noSYS_Error
FeC
Find_TW Global
Flake Description
Fm_normalized
Fm_normalized_error
ID
Import XCAMS Results Date
Import XCAMS Results Time
Import XCAMS Results User
Job
MCC
MCC Saved
MCC_error
MCC_error Saved
MCC_fractional_error
Mode
Modification Date
Modification Time
Modification User
Normalization_exp_factor
Packing Comments
pM
pM_Error
Position
Process Group
Processes Order
Quality Flag
Ratio to standard
Ratio to standard error
Ratio_13C_standard
Ratio13C_denominator
Ratio13C_numerator
Ratio13C_standard_error
Ratio13C_std_av
Ratio13C_std_av Calc
Ratio13C_std_av_Match
Remove System Error
Result Comment for Submitter
Result Comment Internal
rts_bl_av_Calculation
rts_bl_av_Calculation Old
rts_bl_av_error_Calculation
RTS_corrected
RTS_corrected Saved
RTS_corrected_error
RTS_corrected_error Saved
rts_std_av_Match
rts_stds_av
rts_stds_av_Calculation
rts_stds_av_error
rts_stds_av_error_Calculation
Silvia_F_corrected_normed
Silvia_F_corrected_normed_Check
Silvia_F_corrected_normed_error
Silvia_F_corrected_normed_error Check
Silvia_F_corrected_normed_error_no_WTW_Error Check
Silvia_MCC
Silvia_MCC Check
Silvia_MCC_error
Silvia_MCC_error Check
Silvia_RTS_corrected
Silvia_RTS_corrected Check
Silvia_RTS_corrected_error
Silvia_RTS_corrected_error Check
Size Category
SmallSamples values calculated via Excel
Standard 13C Value Constant
Standard Specific Activity Constant
Standard Type Number
Start Date
Start Operator
Start Time
Static Description
Status
tmp_jct_F_new
tmp_jct_F_old
tmp_jct_rts_corr_new
tmp_jct_rts_corr_old
TP
TP Graphite Submitted Minimum Value
TPs Blanks
TPs Blanks Report
TPs Standards
TPs Standards Report
TW
TWs Blanks
TWs Standards
Update Blanks Average Date
Update Blanks Average Time
Update Blanks Average User
Use Assigned Blank
Weight Initial
Wheel To Wheel Error
wtgraph
Date Run
Ratio to standard
TP
TW
wtgraph
Ratio to standard
Ratio to standard error
TP
Standard Prep Type
AMS System
Standard Prep Type
System Error
Wheel Result Comment For Submitter
Wheel Result Comment Internal
delta13C
Surname
delta13C
Position Number
Run Number
delta13C
Run Number
AMS Category
delta13C Estimate
External Due Date
Fraction to date
GlobalStatus
Job
Job Invoicing Notes
Job notes
NZA
Precision
R
Virtual TW
Contaminant observed suspected
Process End Date
Process ID
Process Name
Process Status
R Composed
Sample Description
Comments From Submitter
Comments On Sample Lab Internal
Contaminants
delta13C Submitted
Expected age
Pretreatment by Submitter
Sample Description
Sample ID
Submitter Description
Submitter Requested Pretreatment